## IR BT target parameter
Method |  | Differentiation of L. monocytogenes serogroupy by Bruker Lis mono serogrouper classifier, v.4.0 | [7 empty] | target parameter | Listeria monocytogenes | ST 1/2
Differentiation |  | External Validation of the a.m. classifier; incubation on COL-SB 37°C 24h | Isolates | 40 | [5 empty] | non-target parameter | Listeria monocytogenes | ST 3; ST 4; ST 7
Date of validation
|  |  | Spectra | 269
[5 empty] | classified | Number | Percentage |  |  | outlier score
[5 empty] | correct | 269 | 100% |  |  | Number in green | 263
[5 empty] | questionable | 0 | 0% |  |  | Number in yellow | 6
[5 empty] | wrong | 0 | 0% |  |  | Number in red | 0
[11 empty] | 269 | 0 | 0
Isolate |  | Name (incl. spot no.) | Isolate description | Isolate ID | QT result: true=pass; false=fail | Factual classification | IR BT Classification | Classifi-cation correct? => 1 | classification score: if <0.6 => invalid | outlier score: <1: green, 1-2: yellow, >2: red | if classification score >/=0.6 AND outlier score </=2 AND classification = factual classification => 1 (correct) | if (classification score <0.6 OR outlier score >2) AND classification = factual classification => 1 (questionable) | if classification <> factual classification => 1 (wrong)
# | 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_D10.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.408 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_D11.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.499 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_D12.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.451 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_E10.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.416 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_E11.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.325 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_E12.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.372 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_F10.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.578 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_F11.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.656 | 1 | 0 | 0
| 1 | 220308-1421-Lis_mono_IR-ST_Vali_09L 30642 A1_F12.0 | MB 5340 | 09L 30642 A1 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 0.46 | 1 | 0 | 0
| 1 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_09L 30642 A1_A4.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.751 | 1 | 0 | 0
| 1 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_09L 30642 A1_A5.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.677 | 1 | 0 | 0
| 1 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_09L 30642 A1_A6.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.697 | 1 | 0 | 0
| 1 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_09L 30642 A1_A7.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.403 | 1 | 0 | 0
| 1 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_09L 30642 A1_A8.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.51 | 1 | 0 | 0
| 1 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_09L 30642 A1_A9.0 | MB 5340 | 09L 30642 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.506 | 1 | 0 | 0
# | 2 | 220301-1142-Lis_mono_IR-ST_Vali_09L 31072 006 DIR A1_E7.0 | MB 5342 | 09L 31072 006 DIR A1 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 0.856 | 1 | 0 | 0
| 2 | 220301-1142-Lis_mono_IR-ST_Vali_09L 31072 006 DIR A1_E8.0 | MB 5342 | 09L 31072 006 DIR A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.855 | 1 | 0 | 0
| 2 | 220301-1142-Lis_mono_IR-ST_Vali_09L 31072 006 DIR A1_E9.0 | MB 5342 | 09L 31072 006 DIR A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.787 | 1 | 0 | 0
| 2 | 220304-1020-Lis_mono_IR-ST_Vali_09L 31072 006 DIR A1_D7.0 | MB 5342 | 09L 31072 006 DIR A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.562 | 1 | 0 | 0
| 2 | 220304-1020-Lis_mono_IR-ST_Vali_09L 31072 006 DIR A1_D8.0 | MB 5342 | 09L 31072 006 DIR A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.559 | 1 | 0 | 0
| 2 | 220304-1020-Lis_mono_IR-ST_Vali_09L 31072 006 DIR A1_D9.0 | MB 5342 | 09L 31072 006 DIR A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.533 | 1 | 0 | 0
# | 3 | 220301-1142-Lis_mono_IR-ST_Vali_09L 32701 001 A1_F7.0 | MB 5424 | 09L 32701 001 A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 1.102 | 1 | 0 | 0
| 3 | 220301-1142-Lis_mono_IR-ST_Vali_09L 32701 001 A1_F8.0 | MB 5424 | 09L 32701 001 A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.749 | 1 | 0 | 0
| 3 | 220301-1142-Lis_mono_IR-ST_Vali_09L 32701 001 A1_F9.0 | MB 5424 | 09L 32701 001 A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.7 | 1 | 0 | 0
| 3 | 220304-1020-Lis_mono_IR-ST_Vali_09L 32701 001 A1_E7.0 | MB 5424 | 09L 32701 001 A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.674 | 1 | 0 | 0
| 3 | 220304-1020-Lis_mono_IR-ST_Vali_09L 32701 001 A1_E8.0 | MB 5424 | 09L 32701 001 A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.301 | 1 | 0 | 0
| 3 | 220304-1020-Lis_mono_IR-ST_Vali_09L 32701 001 A1_E9.0 | MB 5424 | 09L 32701 001 A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.386 | 1 | 0 | 0
| 3 | 220323-1410-_09L 32701 001 A1_C4.0 | MB 5424 | 09L 32701 001 A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.641 | 1 | 0 | 0
| 3 | 220323-1410-_09L 32701 001 A1_C5.0 | MB 5424 | 09L 32701 001 A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.605 | 1 | 0 | 0
| 3 | 220323-1410-_09L 32701 001 A1_C6.0 | MB 5424 | 09L 32701 001 A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.395 | 1 | 0 | 0
# | 4 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0330_E7.0 | MB 6557 | 10BfR L0330 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.66 | 1 | 0 | 0
| 4 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0330_E8.0 | MB 6557 | 10BfR L0330 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.755 | 1 | 0 | 0
| 4 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0330_E9.0 | MB 6557 | 10BfR L0330 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.563 | 1 | 0 | 0
| 4 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0330_E7.0 | MB 6557 | 10BfR L0330 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.478 | 1 | 0 | 0
| 4 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0330_E8.0 | MB 6557 | 10BfR L0330 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.548 | 1 | 0 | 0
| 4 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0330_E9.0 | MB 6557 | 10BfR L0330 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.512 | 1 | 0 | 0
# | 5 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0360_F7.0 | MB 6558 | 10BfR L0360 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.521 | 1 | 0 | 0
| 5 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0360_F8.0 | MB 6558 | 10BfR L0360 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.562 | 1 | 0 | 0
| 5 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0360_F9.0 | MB 6558 | 10BfR L0360 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.509 | 1 | 0 | 0
| 5 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0360_F7.0 | MB 6558 | 10BfR L0360 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.798 | 1 | 0 | 0
| 5 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0360_F8.0 | MB 6558 | 10BfR L0360 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.621 | 1 | 0 | 0
| 5 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0360_F9.0 | MB 6558 | 10BfR L0360 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.599 | 1 | 0 | 0
# | 6 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0363_H7.0 | MB 6560 | 10BfR L0363 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.52 | 1 | 0 | 0
| 6 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0363_H8.0 | MB 6560 | 10BfR L0363 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.474 | 1 | 0 | 0
| 6 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0363_H9.0 | MB 6560 | 10BfR L0363 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.811 | 1 | 0 | 0
| 6 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0363_H7.0 | MB 6560 | 10BfR L0363 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.571 | 1 | 0 | 0
| 6 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0363_H8.0 | MB 6560 | 10BfR L0363 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.594 | 1 | 0 | 0
| 6 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0363_H9.0 | MB 6560 | 10BfR L0363 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.526 | 1 | 0 | 0
# | 7 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0365_A10.0 | MB 6561 | 10BfR L0365 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.727 | 1 | 0 | 0
| 7 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0365_A11.0 | MB 6561 | 10BfR L0365 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.799 | 1 | 0 | 0
| 7 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0365_A12.0 | MB 6561 | 10BfR L0365 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.657 | 1 | 0 | 0
| 7 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0365_A10.0 | MB 6561 | 10BfR L0365 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.448 | 1 | 0 | 0
| 7 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0365_A11.0 | MB 6561 | 10BfR L0365 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.365 | 1 | 0 | 0
| 7 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0365_A12.0 | MB 6561 | 10BfR L0365 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.347 | 1 | 0 | 0
# | 8 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0444_C10.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.988 | 1 | 0 | 0
| 8 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0444_C11.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.647 | 1 | 0 | 0
| 8 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0444_C10.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.778 | 1 | 0 | 0
| 8 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0444_C11.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.754 | 1 | 0 | 0
| 8 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0444_C12.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.765 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0444_E7.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 0.999 | 0.605 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0444_E8.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.395 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0444_E9.0 | MB 6563 | 10BfR L0444 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.364 | 1 | 0 | 0
# | 9 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0519_D10.0 | MB 6564 | 10BfR L0519 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.433 | 1 | 0 | 0
| 9 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0519_D11.0 | MB 6564 | 10BfR L0519 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.644 | 1 | 0 | 0
| 9 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0519_D12.0 | MB 6564 | 10BfR L0519 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.639 | 1 | 0 | 0
| 9 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0519_D10.0 | MB 6564 | 10BfR L0519 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.584 | 1 | 0 | 0
| 9 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0519_D11.0 | MB 6564 | 10BfR L0519 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.508 | 1 | 0 | 0
| 9 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0519_D12.0 | MB 6564 | 10BfR L0519 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.586 | 1 | 0 | 0
# | 10 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0528_F10.0 | MB 6567 | 10BfR L0528 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.7 | 1 | 0 | 0
| 10 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0528_F11.0 | MB 6567 | 10BfR L0528 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.713 | 1 | 0 | 0
| 10 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0528_F12.0 | MB 6567 | 10BfR L0528 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.865 | 1 | 0 | 0
| 10 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0528_F10.0 | MB 6567 | 10BfR L0528 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.781 | 1 | 0 | 0
| 10 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0528_F11.0 | MB 6567 | 10BfR L0528 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.679 | 1 | 0 | 0
| 10 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0528_F12.0 | MB 6567 | 10BfR L0528 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.765 | 1 | 0 | 0
# | 11 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0712_B4.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.321 | 1 | 0 | 0
| 11 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0712_B5.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.659 | 1 | 0 | 0
| 11 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0712_B6.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.331 | 1 | 0 | 0
| 11 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0712_B4.0 | MB 6574 | 10BfR L0712 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 1.252 | 1 | 0 | 0
| 11 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0712_B5.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.913 | 1 | 0 | 0
| 11 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0712_B6.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.717 | 1 | 0 | 0
| 11 | 220323-1410-_10BfR L0712_D7.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.398 | 1 | 0 | 0
| 11 | 220323-1410-_10BfR L0712_D8.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.401 | 1 | 0 | 0
| 11 | 220323-1410-_10BfR L0712_D9.0 | MB 6574 | 10BfR L0712 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.44 | 1 | 0 | 0
# | 12 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0713_C4.0 | MB 6575 | 10BfR L0713 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.37 | 1 | 0 | 0
| 12 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0713_C5.0 | MB 6575 | 10BfR L0713 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.927 | 1 | 0 | 0
| 12 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0713_C6.0 | MB 6575 | 10BfR L0713 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.369 | 1 | 0 | 0
| 12 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0713_C4.0 | MB 6575 | 10BfR L0713 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.554 | 1 | 0 | 0
| 12 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0713_C5.0 | MB 6575 | 10BfR L0713 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.849 | 1 | 0 | 0
| 12 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0713_C6.0 | MB 6575 | 10BfR L0713 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.58 | 1 | 0 | 0
# | 13 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0714_D4.0 | MB 6576 | 10BfR L0714 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.521 | 1 | 0 | 0
| 13 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0714_D5.0 | MB 6576 | 10BfR L0714 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.303 | 1 | 0 | 0
| 13 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0714_D6.0 | MB 6576 | 10BfR L0714 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.408 | 1 | 0 | 0
| 13 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0714_D4.0 | MB 6576 | 10BfR L0714 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.772 | 1 | 0 | 0
| 13 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0714_D5.0 | MB 6576 | 10BfR L0714 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.809 | 1 | 0 | 0
| 13 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0714_D6.0 | MB 6576 | 10BfR L0714 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.615 | 1 | 0 | 0
# | 14 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0989_G4.0 | MB 6589 | 10BfR L0989 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.593 | 1 | 0 | 0
| 14 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0989_G5.0 | MB 6589 | 10BfR L0989 | TRUE | 1/2a | IR-ST 1/2 | 1 | 0.999 | 0.692 | 1 | 0 | 0
| 14 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0989_G6.0 | MB 6589 | 10BfR L0989 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.573 | 1 | 0 | 0
| 14 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0989_G4.0 | MB 6589 | 10BfR L0989 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.471 | 1 | 0 | 0
| 14 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0989_G5.0 | MB 6589 | 10BfR L0989 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.647 | 1 | 0 | 0
| 14 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0989_G6.0 | MB 6589 | 10BfR L0989 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.62 | 1 | 0 | 0
# | 15 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1276_B7.0 | MB 6603 | 10BfR L1276 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.286 | 1 | 0 | 0
| 15 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1276_B8.0 | MB 6603 | 10BfR L1276 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.534 | 1 | 0 | 0
| 15 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1276_B9.0 | MB 6603 | 10BfR L1276 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.259 | 1 | 0 | 0
| 15 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1276_B7.0 | MB 6603 | 10BfR L1276 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.649 | 1 | 0 | 0
| 15 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1276_B8.0 | MB 6603 | 10BfR L1276 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.541 | 1 | 0 | 0
| 15 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1276_B9.0 | MB 6603 | 10BfR L1276 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.375 | 1 | 0 | 0
# | 16 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 25552 001_H7.0 | MB 6055 | 10KA 25552 001 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.534 | 1 | 0 | 0
| 16 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 25552 001_H8.0 | MB 6055 | 10KA 25552 001 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.555 | 1 | 0 | 0
| 16 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 25552 001_H9.0 | MB 6055 | 10KA 25552 001 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.358 | 1 | 0 | 0
| 16 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 25552 001_G7.0 | MB 6055 | 10KA 25552 001 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.515 | 1 | 0 | 0
| 16 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 25552 001_G8.0 | MB 6055 | 10KA 25552 001 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.433 | 1 | 0 | 0
| 16 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 25552 001_G9.0 | MB 6055 | 10KA 25552 001 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.419 | 1 | 0 | 0
# | 17 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29328 A1_A10.0 | MB 6144 | 10L 29328 A1 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.891 | 1 | 0 | 0
| 17 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29328 A1_A11.0 | MB 6144 | 10L 29328 A1 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.872 | 1 | 0 | 0
| 17 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29328 A1_A12.0 | MB 6144 | 10L 29328 A1 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 1.242 | 1 | 0 | 0
| 17 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29328 A1_H7.0 | MB 6144 | 10L 29328 A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.514 | 1 | 0 | 0
| 17 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29328 A1_H8.0 | MB 6144 | 10L 29328 A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.337 | 1 | 0 | 0
| 17 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29328 A1_H9.0 | MB 6144 | 10L 29328 A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.29 | 1 | 0 | 0
| 17 | 220323-1410-_10L 29328 A1_E4.0 | MB 6144 | 10L 29328 A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.401 | 1 | 0 | 0
| 17 | 220323-1410-_10L 29328 A1_E5.0 | MB 6144 | 10L 29328 A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.558 | 1 | 0 | 0
| 17 | 220323-1410-_10L 29328 A1_E6.0 | MB 6144 | 10L 29328 A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.43 | 1 | 0 | 0
# | 18 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29333 002 P3_B10.0 | MB 6151 | 10L 29333 002 P3 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.696 | 1 | 0 | 0
| 18 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29333 002 P3_B11.0 | MB 6151 | 10L 29333 002 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.823 | 1 | 0 | 0
| 18 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29333 002 P3_B12.0 | MB 6151 | 10L 29333 002 P3 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.799 | 1 | 0 | 0
| 18 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29333 002 P3_A10.0 | MB 6151 | 10L 29333 002 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.256 | 1 | 0 | 0
| 18 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29333 002 P3_A11.0 | MB 6151 | 10L 29333 002 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.419 | 1 | 0 | 0
| 18 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29333 002 P3_A12.0 | MB 6151 | 10L 29333 002 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.281 | 1 | 0 | 0
| 18 | 220323-1410-_10L 29333 002 P3_F4.0 | MB 6151 | 10L 29333 002 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.571 | 1 | 0 | 0
| 18 | 220323-1410-_10L 29333 002 P3_F5.0 | MB 6151 | 10L 29333 002 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.465 | 1 | 0 | 0
| 18 | 220323-1410-_10L 29333 002 P3_F6.0 | MB 6151 | 10L 29333 002 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.339 | 1 | 0 | 0
# | 19 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29340 P3_C10.0 | MB 6153 | 10L 29340 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 1.061 | 1 | 0 | 0
| 19 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29340 P3_C11.0 | MB 6153 | 10L 29340 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.567 | 1 | 0 | 0
| 19 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29340 P3_C12.0 | MB 6153 | 10L 29340 P3 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.713 | 1 | 0 | 0
| 19 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29340 P3_B10.0 | MB 6153 | 10L 29340 P3 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.298 | 1 | 0 | 0
| 19 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29340 P3_B11.0 | MB 6153 | 10L 29340 P3 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.264 | 1 | 0 | 0
| 19 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29340 P3_B12.0 | MB 6153 | 10L 29340 P3 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.248 | 1 | 0 | 0
# | 20 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29439 A5_D10.0 | MB 6163 | 10L 29439 A5 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.797 | 1 | 0 | 0
| 20 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29439 A5_D11.0 | MB 6163 | 10L 29439 A5 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.691 | 1 | 0 | 0
| 20 | 220301-1142-Lis_mono_IR-ST_Vali_10L 29439 A5_D12.0 | MB 6163 | 10L 29439 A5 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.681 | 1 | 0 | 0
| 20 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29439 A5_C10.0 | MB 6163 | 10L 29439 A5 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.243 | 1 | 0 | 0
| 20 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29439 A5_C11.0 | MB 6163 | 10L 29439 A5 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.293 | 1 | 0 | 0
| 20 | 220304-1020-Lis_mono_IR-ST_Vali_10L 29439 A5_C12.0 | MB 6163 | 10L 29439 A5 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.296 | 1 | 0 | 0
| 20 | 220323-1410-_10L 29439 A5_G4.0 | MB 6163 | 10L 29439 A5 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.661 | 1 | 0 | 0
| 20 | 220323-1410-_10L 29439 A5_G5.0 | MB 6163 | 10L 29439 A5 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.624 | 1 | 0 | 0
| 20 | 220323-1410-_10L 29439 A5_G6.0 | MB 6163 | 10L 29439 A5 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.644 | 1 | 0 | 0
# | 21 | 220302-1444-Lis_mono_IR-ST_Vali_10L 63950 P1_E7.0 | MB 6903 | 10L 63950 P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.295 | 1 | 0 | 0
| 21 | 220302-1444-Lis_mono_IR-ST_Vali_10L 63950 P1_E8.0 | MB 6903 | 10L 63950 P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.258 | 1 | 0 | 0
| 21 | 220302-1444-Lis_mono_IR-ST_Vali_10L 63950 P1_E9.0 | MB 6903 | 10L 63950 P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.305 | 1 | 0 | 0
| 21 | 220303-1155-Lis_mono_IR-ST_Vali_10L 63950 P1_E7.0 | MB 6903 | 10L 63950 P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.708 | 1 | 0 | 0
| 21 | 220303-1155-Lis_mono_IR-ST_Vali_10L 63950 P1_E8.0 | MB 6903 | 10L 63950 P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.719 | 1 | 0 | 0
| 21 | 220303-1155-Lis_mono_IR-ST_Vali_10L 63950 P1_E9.0 | MB 6903 | 10L 63950 P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.642 | 1 | 0 | 0
# | 22 | 220302-1444-Lis_mono_IR-ST_Vali_10L 85241 A1_G7.0 | MB 7137 | 10L 85241 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.462 | 1 | 0 | 0
| 22 | 220302-1444-Lis_mono_IR-ST_Vali_10L 85241 A1_G8.0 | MB 7137 | 10L 85241 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.351 | 1 | 0 | 0
| 22 | 220302-1444-Lis_mono_IR-ST_Vali_10L 85241 A1_G9.0 | MB 7137 | 10L 85241 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.551 | 1 | 0 | 0
| 22 | 220303-1155-Lis_mono_IR-ST_Vali_10L 85241 A1_G7.0 | MB 7137 | 10L 85241 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.578 | 1 | 0 | 0
| 22 | 220303-1155-Lis_mono_IR-ST_Vali_10L 85241 A1_G8.0 | MB 7137 | 10L 85241 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.483 | 1 | 0 | 0
| 22 | 220303-1155-Lis_mono_IR-ST_Vali_10L 85241 A1_G9.0 | MB 7137 | 10L 85241 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.646 | 1 | 0 | 0
# | 23 | 220303-1403-Lis_mono_IR-ST_Vali_10L 88661 A1_A4.0 | MB 7238 | 10L 88661 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.637 | 1 | 0 | 0
| 23 | 220303-1403-Lis_mono_IR-ST_Vali_10L 88661 A1_A5.0 | MB 7238 | 10L 88661 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.459 | 1 | 0 | 0
| 23 | 220303-1403-Lis_mono_IR-ST_Vali_10L 88661 A1_A6.0 | MB 7238 | 10L 88661 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.528 | 1 | 0 | 0
| 23 | 220316-1359-Lis_mono_IR-ST_Vali_10L 88661 A1_A4.0 | MB 7238 | 10L 88661 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.544 | 1 | 0 | 0
| 23 | 220316-1359-Lis_mono_IR-ST_Vali_10L 88661 A1_A5.0 | MB 7238 | 10L 88661 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.629 | 1 | 0 | 0
| 23 | 220316-1359-Lis_mono_IR-ST_Vali_10L 88661 A1_A6.0 | MB 7238 | 10L 88661 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.563 | 1 | 0 | 0
# | 24 | 220302-1444-Lis_mono_IR-ST_Vali_10L 92343 A1_D10.0 | MB 7235 | 10L 92343 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.413 | 1 | 0 | 0
| 24 | 220302-1444-Lis_mono_IR-ST_Vali_10L 92343 A1_D11.0 | MB 7235 | 10L 92343 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.292 | 1 | 0 | 0
| 24 | 220302-1444-Lis_mono_IR-ST_Vali_10L 92343 A1_D12.0 | MB 7235 | 10L 92343 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.288 | 1 | 0 | 0
| 24 | 220303-1155-Lis_mono_IR-ST_Vali_10L 92343 A1_D10.0 | MB 7235 | 10L 92343 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.819 | 1 | 0 | 0
| 24 | 220303-1155-Lis_mono_IR-ST_Vali_10L 92343 A1_D11.0 | MB 7235 | 10L 92343 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.617 | 1 | 0 | 0
| 24 | 220303-1155-Lis_mono_IR-ST_Vali_10L 92343 A1_D12.0 | MB 7235 | 10L 92343 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.696 | 1 | 0 | 0
# | 25 | 220303-1403-Lis_mono_IR-ST_Vali_11BAV 63047 001_A10.0 | MB 8284 | 11BAV 63047 001 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.433 | 1 | 0 | 0
| 25 | 220303-1403-Lis_mono_IR-ST_Vali_11BAV 63047 001_A11.0 | MB 8284 | 11BAV 63047 001 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.474 | 1 | 0 | 0
| 25 | 220303-1403-Lis_mono_IR-ST_Vali_11BAV 63047 001_A12.0 | MB 8284 | 11BAV 63047 001 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.481 | 1 | 0 | 0
| 25 | 220316-1359-Lis_mono_IR-ST_Vali_11BAV 63047 001_A10.0 | MB 8284 | 11BAV 63047 001 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.997 | 1 | 0 | 0
| 25 | 220316-1359-Lis_mono_IR-ST_Vali_11BAV 63047 001_A11.0 | MB 8284 | 11BAV 63047 001 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.672 | 1 | 0 | 0
| 25 | 220316-1359-Lis_mono_IR-ST_Vali_11BAV 63047 001_A12.0 | MB 8284 | 11BAV 63047 001 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.695 | 1 | 0 | 0
# | 26 | 220303-1403-Lis_mono_IR-ST_Vali_11L 36996 A4_A7.0 | MB 8070 | 11L 36996 A4 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.322 | 1 | 0 | 0
| 26 | 220303-1403-Lis_mono_IR-ST_Vali_11L 36996 A4_A8.0 | MB 8070 | 11L 36996 A4 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.753 | 1 | 0 | 0
| 26 | 220303-1403-Lis_mono_IR-ST_Vali_11L 36996 A4_A9.0 | MB 8070 | 11L 36996 A4 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.696 | 1 | 0 | 0
| 26 | 220316-1359-Lis_mono_IR-ST_Vali_11L 36996 A4_A7.0 | MB 8070 | 11L 36996 A4 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.663 | 1 | 0 | 0
| 26 | 220316-1359-Lis_mono_IR-ST_Vali_11L 36996 A4_A8.0 | MB 8070 | 11L 36996 A4 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.696 | 1 | 0 | 0
| 26 | 220316-1359-Lis_mono_IR-ST_Vali_11L 36996 A4_A9.0 | MB 8070 | 11L 36996 A4 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.831 | 1 | 0 | 0
# | 27 | 220303-1403-Lis_mono_IR-ST_Vali_11L 38607 003 A1_D7.0 | MB 8074 | 11L 38607 003 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.408 | 1 | 0 | 0
| 27 | 220303-1403-Lis_mono_IR-ST_Vali_11L 38607 003 A1_D8.0 | MB 8074 | 11L 38607 003 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.386 | 1 | 0 | 0
| 27 | 220303-1403-Lis_mono_IR-ST_Vali_11L 38607 003 A1_D9.0 | MB 8074 | 11L 38607 003 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.705 | 1 | 0 | 0
| 27 | 220316-1359-Lis_mono_IR-ST_Vali_11L 38607 003 A1_D7.0 | MB 8074 | 11L 38607 003 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.671 | 1 | 0 | 0
| 27 | 220316-1359-Lis_mono_IR-ST_Vali_11L 38607 003 A1_D8.0 | MB 8074 | 11L 38607 003 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.828 | 1 | 0 | 0
| 27 | 220316-1359-Lis_mono_IR-ST_Vali_11L 38607 003 A1_D9.0 | MB 8074 | 11L 38607 003 A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.822 | 1 | 0 | 0
# | 28 | 220303-1403-Lis_mono_IR-ST_Vali_11L 46488 KEZ P1_F7.0 | MB 8152 | 11L 46488 KEZ P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.497 | 1 | 0 | 0
| 28 | 220303-1403-Lis_mono_IR-ST_Vali_11L 46488 KEZ P1_F8.0 | MB 8152 | 11L 46488 KEZ P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.316 | 1 | 0 | 0
| 28 | 220303-1403-Lis_mono_IR-ST_Vali_11L 46488 KEZ P1_F9.0 | MB 8152 | 11L 46488 KEZ P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.35 | 1 | 0 | 0
| 28 | 220316-1359-Lis_mono_IR-ST_Vali_11L 46488 KEZ P1_F7.0 | MB 8152 | 11L 46488 KEZ P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.817 | 1 | 0 | 0
| 28 | 220316-1359-Lis_mono_IR-ST_Vali_11L 46488 KEZ P1_F8.0 | MB 8152 | 11L 46488 KEZ P1 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 1.021 | 1 | 0 | 0
| 28 | 220316-1359-Lis_mono_IR-ST_Vali_11L 46488 KEZ P1_F9.0 | MB 8152 | 11L 46488 KEZ P1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.859 | 1 | 0 | 0
# | 29 | 220303-1403-Lis_mono_IR-ST_Vali_11L 51132 A5_G7.0 | MB 8155 | 11L 51132 A5 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.351 | 1 | 0 | 0
| 29 | 220303-1403-Lis_mono_IR-ST_Vali_11L 51132 A5_G8.0 | MB 8155 | 11L 51132 A5 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.324 | 1 | 0 | 0
| 29 | 220303-1403-Lis_mono_IR-ST_Vali_11L 51132 A5_G9.0 | MB 8155 | 11L 51132 A5 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.274 | 1 | 0 | 0
| 29 | 220316-1359-Lis_mono_IR-ST_Vali_11L 51132 A5_G7.0 | MB 8155 | 11L 51132 A5 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.897 | 1 | 0 | 0
| 29 | 220316-1359-Lis_mono_IR-ST_Vali_11L 51132 A5_G8.0 | MB 8155 | 11L 51132 A5 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 0.881 | 1 | 0 | 0
| 29 | 220316-1359-Lis_mono_IR-ST_Vali_11L 51132 A5_G9.0 | MB 8155 | 11L 51132 A5 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.861 | 1 | 0 | 0
# | 30 | 220303-1403-Lis_mono_IR-ST_Vali_11L 51330 P7_H7.0 | MB 8157 | 11L 51330 P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.469 | 1 | 0 | 0
| 30 | 220303-1403-Lis_mono_IR-ST_Vali_11L 51330 P7_H8.0 | MB 8157 | 11L 51330 P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.35 | 1 | 0 | 0
| 30 | 220303-1403-Lis_mono_IR-ST_Vali_11L 51330 P7_H9.0 | MB 8157 | 11L 51330 P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.325 | 1 | 0 | 0
| 30 | 220316-1359-Lis_mono_IR-ST_Vali_11L 51330 P7_H7.0 | MB 8157 | 11L 51330 P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.865 | 1 | 0 | 0
| 30 | 220316-1359-Lis_mono_IR-ST_Vali_11L 51330 P7_H8.0 | MB 8157 | 11L 51330 P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.59 | 1 | 0 | 0
| 30 | 220316-1359-Lis_mono_IR-ST_Vali_11L 51330 P7_H9.0 | MB 8157 | 11L 51330 P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.572 | 1 | 0 | 0
# | 31 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77765 001 A2_B10.0 | MB 8459 | 11L 77765 001 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.381 | 1 | 0 | 0
| 31 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77765 001 A2_B11.0 | MB 8459 | 11L 77765 001 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.308 | 1 | 0 | 0
| 31 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77765 001 A2_B12.0 | MB 8459 | 11L 77765 001 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.318 | 1 | 0 | 0
| 31 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77765 001 A2_B10.0 | MB 8459 | 11L 77765 001 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.63 | 1 | 0 | 0
| 31 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77765 001 A2_B11.0 | MB 8459 | 11L 77765 001 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.741 | 1 | 0 | 0
| 31 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77765 001 A2_B12.0 | MB 8459 | 11L 77765 001 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.841 | 1 | 0 | 0
# | 32 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77922 A2_D10.0 | MB 8544 | 11L 77922 A2 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.487 | 1 | 0 | 0
| 32 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77922 A2_D11.0 | MB 8544 | 11L 77922 A2 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.441 | 1 | 0 | 0
| 32 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77922 A2_D12.0 | MB 8544 | 11L 77922 A2 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.436 | 1 | 0 | 0
| 32 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77922 A2_D10.0 | MB 8544 | 11L 77922 A2 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.875 | 1 | 0 | 0
| 32 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77922 A2_D11.0 | MB 8544 | 11L 77922 A2 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.833 | 1 | 0 | 0
| 32 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77922 A2_D12.0 | MB 8544 | 11L 77922 A2 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.823 | 1 | 0 | 0
# | 33 | 220301-1142-Lis_mono_IR-ST_Vali_13L 01724 001 HF A2_B7.0 | MB 3266,2 | 13L 01724 001 HF A2 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.737 | 1 | 0 | 0
| 33 | 220301-1142-Lis_mono_IR-ST_Vali_13L 01724 001 HF A2_B8.0 | MB 3266,2 | 13L 01724 001 HF A2 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 1.096 | 1 | 0 | 0
| 33 | 220301-1142-Lis_mono_IR-ST_Vali_13L 01724 001 HF A2_B9.0 | MB 3266,2 | 13L 01724 001 HF A2 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 0.799 | 1 | 0 | 0
| 33 | 220304-1020-Lis_mono_IR-ST_Vali_13L 01724 001 HF A2_B7.0 | MB 3266,2 | 13L 01724 001 HF A2 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.34 | 1 | 0 | 0
| 33 | 220304-1020-Lis_mono_IR-ST_Vali_13L 01724 001 HF A2_B8.0 | MB 3266,2 | 13L 01724 001 HF A2 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.327 | 1 | 0 | 0
| 33 | 220304-1020-Lis_mono_IR-ST_Vali_13L 01724 001 HF A2_B9.0 | MB 3266,2 | 13L 01724 001 HF A2 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.293 | 1 | 0 | 0
| 33 | 220323-1410-_13L 01724 001 HF A2_B4.0 | MB 3266,2 | 13L 01724 001 HF A2 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.39 | 1 | 0 | 0
| 33 | 220323-1410-_13L 01724 001 HF A2_B5.0 | MB 3266,2 | 13L 01724 001 HF A2 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.434 | 1 | 0 | 0
| 33 | 220323-1410-_13L 01724 001 HF A2_B6.0 | MB 3266,2 | 13L 01724 001 HF A2 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.557 | 1 | 0 | 0
# | 34 | 220301-1142-Lis_mono_IR-ST_Vali_13L 06711 A2_B4.0 | MB 240,2 | 13L 06711 A2 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.623 | 1 | 0 | 0
| 34 | 220301-1142-Lis_mono_IR-ST_Vali_13L 06711 A2_B5.0 | MB 240,2 | 13L 06711 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.73 | 1 | 0 | 0
| 34 | 220301-1142-Lis_mono_IR-ST_Vali_13L 06711 A2_B6.0 | MB 240,2 | 13L 06711 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.72 | 1 | 0 | 0
| 34 | 220304-1020-Lis_mono_IR-ST_Vali_13L 06711 A2_B4.0 | MB 240,2 | 13L 06711 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.298 | 1 | 0 | 0
| 34 | 220304-1020-Lis_mono_IR-ST_Vali_13L 06711 A2_B5.0 | MB 240,2 | 13L 06711 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.416 | 1 | 0 | 0
| 34 | 220304-1020-Lis_mono_IR-ST_Vali_13L 06711 A2_B6.0 | MB 240,2 | 13L 06711 A2 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.355 | 1 | 0 | 0
# | 35 | 220301-1142-Lis_mono_IR-ST_Vali_13L 08512 VF A1_C7.0 | MB 3280,2 | 13L 08512 VF A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.56 | 1 | 0 | 0
| 35 | 220301-1142-Lis_mono_IR-ST_Vali_13L 08512 VF A1_C8.0 | MB 3280,2 | 13L 08512 VF A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.577 | 1 | 0 | 0
| 35 | 220301-1142-Lis_mono_IR-ST_Vali_13L 08512 VF A1_C9.0 | MB 3280,2 | 13L 08512 VF A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.527 | 1 | 0 | 0
| 35 | 220304-1020-Lis_mono_IR-ST_Vali_13L 08512 VF A1_C7.0 | MB 3280,2 | 13L 08512 VF A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.397 | 1 | 0 | 0
| 35 | 220304-1020-Lis_mono_IR-ST_Vali_13L 08512 VF A1_C8.0 | MB 3280,2 | 13L 08512 VF A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.353 | 1 | 0 | 0
| 35 | 220304-1020-Lis_mono_IR-ST_Vali_13L 08512 VF A1_C9.0 | MB 3280,2 | 13L 08512 VF A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.352 | 1 | 0 | 0
# | 36 | 220301-1142-Lis_mono_IR-ST_Vali_13L 14056 001 HF P7_C4.0 | MB 777,2 | 13L 14056 001 HF P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.515 | 1 | 0 | 0
| 36 | 220301-1142-Lis_mono_IR-ST_Vali_13L 14056 001 HF P7_C5.0 | MB 777,2 | 13L 14056 001 HF P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.409 | 1 | 0 | 0
| 36 | 220301-1142-Lis_mono_IR-ST_Vali_13L 14056 001 HF P7_C6.0 | MB 777,2 | 13L 14056 001 HF P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.607 | 1 | 0 | 0
| 36 | 220304-1020-Lis_mono_IR-ST_Vali_13L 14056 001 HF P7_C4.0 | MB 777,2 | 13L 14056 001 HF P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.367 | 1 | 0 | 0
| 36 | 220304-1020-Lis_mono_IR-ST_Vali_13L 14056 001 HF P7_C5.0 | MB 777,2 | 13L 14056 001 HF P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.306 | 1 | 0 | 0
| 36 | 220304-1020-Lis_mono_IR-ST_Vali_13L 14056 001 HF P7_C6.0 | MB 777,2 | 13L 14056 001 HF P7 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.274 | 1 | 0 | 0
# | 37 | 220301-1142-Lis_mono_IR-ST_Vali_13L 14056 003 HF A1_D4.0 | MB 784,2 | 13L 14056 003 HF A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.526 | 1 | 0 | 0
| 37 | 220301-1142-Lis_mono_IR-ST_Vali_13L 14056 003 HF A1_D5.0 | MB 784,2 | 13L 14056 003 HF A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.782 | 1 | 0 | 0
| 37 | 220301-1142-Lis_mono_IR-ST_Vali_13L 14056 003 HF A1_D6.0 | MB 784,2 | 13L 14056 003 HF A1 | FAIL | 1/2c | IR-ST 1/2 | 1 | 1 | 0.808 | 1 | 0 | 0
| 37 | 220304-1020-Lis_mono_IR-ST_Vali_13L 14056 003 HF A1_D4.0 | MB 784,2 | 13L 14056 003 HF A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.367 | 1 | 0 | 0
| 37 | 220304-1020-Lis_mono_IR-ST_Vali_13L 14056 003 HF A1_D5.0 | MB 784,2 | 13L 14056 003 HF A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.32 | 1 | 0 | 0
| 37 | 220304-1020-Lis_mono_IR-ST_Vali_13L 14056 003 HF A1_D6.0 | MB 784,2 | 13L 14056 003 HF A1 | TRUE | 1/2c | IR-ST 1/2 | 1 | 1 | 0.341 | 1 | 0 | 0
# | 38 | 220301-1142-Lis_mono_IR-ST_Vali_13L 17778 002 HF A1_F4.0 | MB 1061,2 | 13L 17778 002 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.63 | 1 | 0 | 0
| 38 | 220301-1142-Lis_mono_IR-ST_Vali_13L 17778 002 HF A1_F5.0 | MB 1061,2 | 13L 17778 002 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.37 | 1 | 0 | 0
| 38 | 220301-1142-Lis_mono_IR-ST_Vali_13L 17778 002 HF A1_F6.0 | MB 1061,2 | 13L 17778 002 HF A1 | FAIL | 1/2a | IR-ST 1/2 | 1 | 1 | 0.565 | 1 | 0 | 0
| 38 | 220304-1020-Lis_mono_IR-ST_Vali_13L 17778 002 HF A1_F4.0 | MB 1061,2 | 13L 17778 002 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.325 | 1 | 0 | 0
| 38 | 220304-1020-Lis_mono_IR-ST_Vali_13L 17778 002 HF A1_F5.0 | MB 1061,2 | 13L 17778 002 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.402 | 1 | 0 | 0
| 38 | 220304-1020-Lis_mono_IR-ST_Vali_13L 17778 002 HF A1_F6.0 | MB 1061,2 | 13L 17778 002 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.562 | 1 | 0 | 0
# | 39 | 220301-1142-Lis_mono_IR-ST_Vali_13L 17778 003 HF A1_G4.0 | MB 1065,2 | 13L 17778 003 HF A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.647 | 1 | 0 | 0
| 39 | 220301-1142-Lis_mono_IR-ST_Vali_13L 17778 003 HF A1_G5.0 | MB 1065,2 | 13L 17778 003 HF A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.692 | 1 | 0 | 0
| 39 | 220301-1142-Lis_mono_IR-ST_Vali_13L 17778 003 HF A1_G6.0 | MB 1065,2 | 13L 17778 003 HF A1 | FAIL | 1/2b | IR-ST 1/2 | 1 | 1 | 0.646 | 1 | 0 | 0
| 39 | 220304-1020-Lis_mono_IR-ST_Vali_13L 17778 003 HF A1_G4.0 | MB 1065,2 | 13L 17778 003 HF A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.383 | 1 | 0 | 0
| 39 | 220304-1020-Lis_mono_IR-ST_Vali_13L 17778 003 HF A1_G5.0 | MB 1065,2 | 13L 17778 003 HF A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.337 | 1 | 0 | 0
| 39 | 220304-1020-Lis_mono_IR-ST_Vali_13L 17778 003 HF A1_G6.0 | MB 1065,2 | 13L 17778 003 HF A1 | TRUE | 1/2b | IR-ST 1/2 | 1 | 1 | 0.328 | 1 | 0 | 0
# | 40 | 220301-1142-Lis_mono_IR-ST_Vali_13L 19334 HF A2_E4.0 | MB 927,2 | 13L 19334 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.668 | 1 | 0 | 0
| 40 | 220301-1142-Lis_mono_IR-ST_Vali_13L 19334 HF A2_E5.0 | MB 927,2 | 13L 19334 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.795 | 1 | 0 | 0
| 40 | 220301-1142-Lis_mono_IR-ST_Vali_13L 19334 HF A2_E6.0 | MB 927,2 | 13L 19334 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.739 | 1 | 0 | 0
| 40 | 220304-1020-Lis_mono_IR-ST_Vali_13L 19334 HF A1_E4.0 | MB 927,2 | 13L 19334 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.537 | 1 | 0 | 0
| 40 | 220304-1020-Lis_mono_IR-ST_Vali_13L 19334 HF A1_E5.0 | MB 927,2 | 13L 19334 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.323 | 1 | 0 | 0
| 40 | 220304-1020-Lis_mono_IR-ST_Vali_13L 19334 HF A1_E6.0 | MB 927,2 | 13L 19334 HF A1 | TRUE | 1/2a | IR-ST 1/2 | 1 | 1 | 0.625 | 1 | 0 | 0

## IR BT non-target parameter
Method |  | Differentiation of L. monocytogenes serogroupy by Bruker Lis mono serogrouper classifier, v.4.0 | [7 empty] | target parameter | ST 1/2
Differentiation |  | External Validation of the a.m. classifier; incubation on COL-SB 37°C 24h | Isolates | 54 | [5 empty] | non-target parameter | ST 3; ST 4; ST 7
Date of validation
|  |  | Spectra | 361
[5 empty] | classified | Number | Percentage |  |  | outlier score
[5 empty] | correct | 361 | 100% |  |  | Number in green | 299
[5 empty] | questionable | 0 | 0% |  |  | Number in yellow | 62
[5 empty] | wrong | 0 | 0% |  |  | Number in red | 0
[11 empty] | 361 | 0 | 0
Isolate |  | Spectrum name (incl. spot no.) | Isolate description | Isolate ID | QT result: true=pass; false=fail | Factual classification | IR BT Classification | Classifi-cation correct? => 1 | classifi-cation score: if <0.6 => invalid | outlier score: <1: green, 1-2: yellow, >2: red | if classification score >/=0.6 AND outlier score </=2 AND classification = factual classification => 1 (correct) | if (classification score <0.6 OR outlier score >2) AND classification = factual classification => 1 (questionable) | if classification <> factual classification => 1 (wrong)
# | 1 | 220301-1142-Lis_mono_IR-ST_Vali_03L 34842_A4.0 | MB 18 | 03L 34842 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.436 | 1 | 0 | 0
| 1 | 220301-1142-Lis_mono_IR-ST_Vali_03L 34842_A5.0 | MB 18 | 03L 34842 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.795 | 1 | 0 | 0
| 1 | 220301-1142-Lis_mono_IR-ST_Vali_03L 34842_A6.0 | MB 18 | 03L 34842 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.617 | 1 | 0 | 0
| 1 | 220304-1020-Lis_mono_IR-ST_Vali_03L 34842_A4.0 | MB 18 | 03L 34842 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.256 | 1 | 0 | 0
| 1 | 220304-1020-Lis_mono_IR-ST_Vali_03L 34842_A5.0 | MB 18 | 03L 34842 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.365 | 1 | 0 | 0
| 1 | 220304-1020-Lis_mono_IR-ST_Vali_03L 34842_A6.0 | MB 18 | 03L 34842 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.308 | 1 | 0 | 0
# | 2 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0064_A4.0 | MB 6544 | 10BfR L0064 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.656 | 1 | 0 | 0
| 2 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0064_A5.0 | MB 6544 | 10BfR L0064 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.647 | 1 | 0 | 0
| 2 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0064_A6.0 | MB 6544 | 10BfR L0064 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.767 | 1 | 0 | 0
| 2 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0064_A4.0 | MB 6544 | 10BfR L0064 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.869 | 1 | 0 | 0
| 2 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0064_A5.0 | MB 6544 | 10BfR L0064 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.717 | 1 | 0 | 0
| 2 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0064_A6.0 | MB 6544 | 10BfR L0064 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.573 | 1 | 0 | 0
# | 3 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0066_B4.0 | MB 6545 | 10BfR L0066 | FAIL | 3c | IR-ST 3 | 1 | 1 | 0.747 | 1 | 0 | 0
| 3 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0066_B5.0 | MB 6545 | 10BfR L0066 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.954 | 1 | 0 | 0
| 3 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0066_B6.0 | MB 6545 | 10BfR L0066 | FAIL | 3c | IR-ST 3 | 1 | 1 | 1.112 | 1 | 0 | 0
| 3 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0066_B4.0 | MB 6545 | 10BfR L0066 | TRUE | 3c | IR-ST 3 | 1 | 1 | 1.256 | 1 | 0 | 0
| 3 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0066_B5.0 | MB 6545 | 10BfR L0066 | TRUE | 3c | IR-ST 3 | 1 | 1 | 1.276 | 1 | 0 | 0
| 3 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0066_B6.0 | MB 6545 | 10BfR L0066 | FAIL | 3c | IR-ST 3 | 1 | 1 | 1.402 | 1 | 0 | 0
| 3 | 220323-1410-_10BfR L0066_B7.0 | MB 6545 | 10BfR L0066 | TRUE | 3c | IR-ST 3 | 1 | 1 | 1.367 | 1 | 0 | 0
| 3 | 220323-1410-_10BfR L0066_B8.0 | MB 6545 | 10BfR L0066 | FAIL | 3c | IR-ST 3 | 1 | 1 | 1.431 | 1 | 0 | 0
| 3 | 220323-1410-_10BfR L0066_B9.0 | MB 6545 | 10BfR L0066 | FAIL | 3c | IR-ST 3 | 1 | 1 | 1.513 | 1 | 0 | 0
# | 4 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0067_C4.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 1 | 0.687 | 1 | 0 | 0
| 4 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0067_C5.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 0.903 | 1 | 0 | 0
| 4 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0067_C6.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 0.994 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_A10.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 1.123 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_A11.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 1.117 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_A12.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 1.129 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_B10.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 1.317 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_B11.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.998 | 1.424 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_B12.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 1.122 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_C10.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 0.981 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_C11.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 1.07 | 1 | 0 | 0
| 4 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0067_C12.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 1.393 | 1 | 0 | 0
| 4 | 220323-1410-_10BfR L0067_E7.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 1 | 0.969 | 1 | 0 | 0
| 4 | 220323-1410-_10BfR L0067_E8.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.997 | 1.454 | 1 | 0 | 0
| 4 | 220323-1410-_10BfR L0067_E9.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.998 | 1.421 | 1 | 0 | 0
| 4 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_10BfR L0067_B4.0 | MB 6546 | 10BfR L0067 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 1.008 | 1 | 0 | 0
| 4 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_10BfR L0067_B7.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 1.079 | 1 | 0 | 0
| 4 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_10BfR L0067_B8.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.998 | 1.25 | 1 | 0 | 0
| 4 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_10BfR L0067_B9.0 | MB 6546 | 10BfR L0067 | TRUE | 4a | IR-ST 4x | 1 | 0.998 | 1.228 | 1 | 0 | 0
# | 5 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0206_D4.0 | MB 6547 | 10BfR L0206 | FAIL | 3b | IR-ST 3 | 1 | 1 | 1.036 | 1 | 0 | 0
| 5 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0206_D5.0 | MB 6547 | 10BfR L0206 | TRUE | 3b | IR-ST 3 | 1 | 1 | 1.014 | 1 | 0 | 0
| 5 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0206_D6.0 | MB 6547 | 10BfR L0206 | TRUE | 3b | IR-ST 3 | 1 | 1 | 1.022 | 1 | 0 | 0
| 5 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0206_D4.0 | MB 6547 | 10BfR L0206 | TRUE | 3b | IR-ST 3 | 1 | 1 | 1.08 | 1 | 0 | 0
| 5 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0206_D5.0 | MB 6547 | 10BfR L0206 | TRUE | 3b | IR-ST 3 | 1 | 1 | 1.033 | 1 | 0 | 0
| 5 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0206_D6.0 | MB 6547 | 10BfR L0206 | TRUE | 3b | IR-ST 3 | 1 | 1 | 0.907 | 1 | 0 | 0
# | 6 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0208_E4.0 | MB 6548 | 10BfR L0208 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.51 | 1 | 0 | 0
| 6 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0208_E5.0 | MB 6548 | 10BfR L0208 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.763 | 1 | 0 | 0
| 6 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0208_E6.0 | MB 6548 | 10BfR L0208 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.57 | 1 | 0 | 0
| 6 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0208_E4.0 | MB 6548 | 10BfR L0208 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.847 | 1 | 0 | 0
| 6 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0208_E5.0 | MB 6548 | 10BfR L0208 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.717 | 1 | 0 | 0
| 6 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0208_E6.0 | MB 6548 | 10BfR L0208 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.802 | 1 | 0 | 0
# | 7 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0210_F4.0 | MB 6549 | 10BfR L0210 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.815 | 1 | 0 | 0
| 7 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0210_F5.0 | MB 6549 | 10BfR L0210 | TRUE | 3c | IR-ST 3 | 1 | 0.999 | 0.591 | 1 | 0 | 0
| 7 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0210_F6.0 | MB 6549 | 10BfR L0210 | TRUE | 3c | IR-ST 3 | 1 | 1 | 1.085 | 1 | 0 | 0
| 7 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0210_F4.0 | MB 6549 | 10BfR L0210 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.697 | 1 | 0 | 0
| 7 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0210_F5.0 | MB 6549 | 10BfR L0210 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.561 | 1 | 0 | 0
| 7 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0210_F6.0 | MB 6549 | 10BfR L0210 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.565 | 1 | 0 | 0
# | 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_F7.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 1 | 1.3 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_F8.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 1 | 1.063 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_F9.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 1.434 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_G7.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 1 | 1.174 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_G8.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 1.418 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_G9.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 0.999 | 1.12 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_H7.0 | MB 6550 | 10BfR L0211 | TRUE | 4a | IR-ST 4x | 1 | 1 | 0.719 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_H8.0 | MB 6550 | 10BfR L0211 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 0.946 | 1 | 0 | 0
| 8 | 220308-1421-Lis_mono_IR-ST_Vali_10BfR L0211_H9.0 | MB 6550 | 10BfR L0211 | TRUE | 4a | IR-ST 4x | 1 | 1 | 0.79 | 1 | 0 | 0
| 8 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_10BfR L0211_C5.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 0.997 | 1.108 | 1 | 0 | 0
| 8 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_10BfR L0211_C7.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 0.996 | 1.447 | 1 | 0 | 0
| 8 | 220329-1041-Lis_mono_IR-ST_Vali(v.2)_10BfR L0211_C8.0 | MB 6550 | 10BfR L0211 | FAIL | 4a | IR-ST 4x | 1 | 0.995 | 1.598 | 1 | 0 | 0
# | 9 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0217_G4.0 | MB 6551 | 10BfR L0217 | TRUE | 4d | IR-ST 4x | 1 | 0.999 | 0.519 | 1 | 0 | 0
| 9 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0217_G5.0 | MB 6551 | 10BfR L0217 | TRUE | 4d | IR-ST 4x | 1 | 0.999 | 0.583 | 1 | 0 | 0
| 9 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0217_G6.0 | MB 6551 | 10BfR L0217 | TRUE | 4d | IR-ST 4x | 1 | 0.999 | 0.754 | 1 | 0 | 0
| 9 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0217_G4.0 | MB 6551 | 10BfR L0217 | TRUE | 4d | IR-ST 4x | 1 | 0.999 | 0.711 | 1 | 0 | 0
| 9 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0217_G5.0 | MB 6551 | 10BfR L0217 | TRUE | 4d | IR-ST 4x | 1 | 0.999 | 0.508 | 1 | 0 | 0
| 9 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0217_G6.0 | MB 6551 | 10BfR L0217 | TRUE | 4d | IR-ST 4x | 1 | 0.999 | 0.567 | 1 | 0 | 0
# | 10 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0219_H4.0 | MB 6552 | 10BfR L0219 | TRUE | 4d | IR-ST 4x | 1 | 1 | 0.647 | 1 | 0 | 0
| 10 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0219_H5.0 | MB 6552 | 10BfR L0219 | TRUE | 4d | IR-ST 4x | 1 | 1 | 0.632 | 1 | 0 | 0
| 10 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0219_H6.0 | MB 6552 | 10BfR L0219 | TRUE | 4d | IR-ST 4x | 1 | 1 | 0.64 | 1 | 0 | 0
| 10 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0219_H4.0 | MB 6552 | 10BfR L0219 | TRUE | 4d | IR-ST 4x | 1 | 1 | 0.562 | 1 | 0 | 0
| 10 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0219_H5.0 | MB 6552 | 10BfR L0219 | TRUE | 4d | IR-ST 4x | 1 | 1 | 0.569 | 1 | 0 | 0
| 10 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0219_H6.0 | MB 6552 | 10BfR L0219 | TRUE | 4d | IR-ST 4x | 1 | 1 | 0.612 | 1 | 0 | 0
# | 11 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0223_A7.0 | MB 6553 | 10BfR L0223 | TRUE | 4e | IR-ST 4x | 1 | 1 | 0.521 | 1 | 0 | 0
| 11 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0223_A8.0 | MB 6553 | 10BfR L0223 | TRUE | 4e | IR-ST 4x | 1 | 1 | 0.52 | 1 | 0 | 0
| 11 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0223_A9.0 | MB 6553 | 10BfR L0223 | TRUE | 4e | IR-ST 4x | 1 | 1 | 0.713 | 1 | 0 | 0
| 11 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0223_A7.0 | MB 6553 | 10BfR L0223 | FAIL | 4e | IR-ST 4x | 1 | 0.999 | 1.365 | 1 | 0 | 0
| 11 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0223_A8.0 | MB 6553 | 10BfR L0223 | TRUE | 4e | IR-ST 4x | 1 | 1 | 1.073 | 1 | 0 | 0
| 11 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0223_A9.0 | MB 6553 | 10BfR L0223 | FAIL | 4e | IR-ST 4x | 1 | 1 | 0.939 | 1 | 0 | 0
| 11 | 220323-1410-_10BfR L0223_C7.0 | MB 6553 | 10BfR L0223 | FAIL | 4e | IR-ST 4x | 1 | 1 | 1.07 | 1 | 0 | 0
| 11 | 220323-1410-_10BfR L0223_C8.0 | MB 6553 | 10BfR L0223 | FAIL | 4e | IR-ST 4x | 1 | 0.999 | 1.587 | 1 | 0 | 0
| 11 | 220323-1410-_10BfR L0223_C9.0 | MB 6553 | 10BfR L0223 | FAIL | 4e | IR-ST 4x | 1 | 0.999 | 1.914 | 1 | 0 | 0
# | 12 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0278_B7.0 | MB 6554 | 10BfR L0278 | TRUE | 3a | IR-ST 7 | 1 | 1 | 0.708 | 1 | 0 | 0
| 12 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0278_B8.0 | MB 6554 | 10BfR L0278 | TRUE | 3a | IR-ST 7 | 1 | 1 | 0.631 | 1 | 0 | 0
| 12 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0278_B9.0 | MB 6554 | 10BfR L0278 | TRUE | 3a | IR-ST 7 | 1 | 1 | 1.116 | 1 | 0 | 0
| 12 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0278_B7.0 | MB 6554 | 10BfR L0278 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.589 | 1 | 0 | 0
| 12 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0278_B8.0 | MB 6554 | 10BfR L0278 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.588 | 1 | 0 | 0
| 12 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0278_B9.0 | MB 6554 | 10BfR L0278 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.503 | 1 | 0 | 0
# | 13 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0286_C7.0 | MB 6555 | 10BfR L0286 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.573 | 1 | 0 | 0
| 13 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0286_C8.0 | MB 6555 | 10BfR L0286 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.643 | 1 | 0 | 0
| 13 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0286_C9.0 | MB 6555 | 10BfR L0286 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.619 | 1 | 0 | 0
| 13 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0286_C7.0 | MB 6555 | 10BfR L0286 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.415 | 1 | 0 | 0
| 13 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0286_C8.0 | MB 6555 | 10BfR L0286 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.596 | 1 | 0 | 0
| 13 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0286_C9.0 | MB 6555 | 10BfR L0286 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.592 | 1 | 0 | 0
# | 14 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0287_D7.0 | MB 6556 | 10BfR L0287 | TRUE | 3a | IR-ST 7 | 1 | 1 | 0.627 | 1 | 0 | 0
| 14 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0287_D8.0 | MB 6556 | 10BfR L0287 | TRUE | 3a | IR-ST 7 | 1 | 1 | 0.694 | 1 | 0 | 0
| 14 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0287_D9.0 | MB 6556 | 10BfR L0287 | TRUE | 3a | IR-ST 7 | 1 | 1 | 0.657 | 1 | 0 | 0
| 14 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0287_D7.0 | MB 6556 | 10BfR L0287 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.647 | 1 | 0 | 0
| 14 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0287_D8.0 | MB 6556 | 10BfR L0287 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.834 | 1 | 0 | 0
| 14 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0287_D9.0 | MB 6556 | 10BfR L0287 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.699 | 1 | 0 | 0
# | 15 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0362_G7.0 | MB 6559 | 10BfR L0362 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.592 | 1 | 0 | 0
| 15 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0362_G8.0 | MB 6559 | 10BfR L0362 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.456 | 1 | 0 | 0
| 15 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0362_G9.0 | MB 6559 | 10BfR L0362 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.513 | 1 | 0 | 0
| 15 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0362_G7.0 | MB 6559 | 10BfR L0362 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.508 | 1 | 0 | 0
| 15 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0362_G8.0 | MB 6559 | 10BfR L0362 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.369 | 1 | 0 | 0
| 15 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0362_G9.0 | MB 6559 | 10BfR L0362 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.31 | 1 | 0 | 0
# | 16 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0435_B10.0 | MB 6562 | 10BfR L0435 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 0.746 | 1 | 0 | 0
| 16 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0435_B11.0 | MB 6562 | 10BfR L0435 | TRUE | 4a | IR-ST 4x | 1 | 0.996 | 1.102 | 1 | 0 | 0
| 16 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0435_B12.0 | MB 6562 | 10BfR L0435 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 0.497 | 1 | 0 | 0
| 16 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0435_B10.0 | MB 6562 | 10BfR L0435 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 0.492 | 1 | 0 | 0
| 16 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0435_B11.0 | MB 6562 | 10BfR L0435 | TRUE | 4a | IR-ST 4x | 1 | 0.999 | 0.671 | 1 | 0 | 0
| 16 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0435_B12.0 | MB 6562 | 10BfR L0435 | TRUE | 4a | IR-ST 4x | 1 | 1 | 0.455 | 1 | 0 | 0
# | 17 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0524_E10.0 | MB 6566 | 10BfR L0524 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.337 | 1 | 0 | 0
| 17 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0524_E11.0 | MB 6566 | 10BfR L0524 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.419 | 1 | 0 | 0
| 17 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0524_E12.0 | MB 6566 | 10BfR L0524 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.378 | 1 | 0 | 0
| 17 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0524_E10.0 | MB 6566 | 10BfR L0524 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.597 | 1 | 0 | 0
| 17 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0524_E11.0 | MB 6566 | 10BfR L0524 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.455 | 1 | 0 | 0
| 17 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0524_E12.0 | MB 6566 | 10BfR L0524 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.434 | 1 | 0 | 0
# | 18 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0614_G10.0 | MB 6571 | 10BfR L0614 | TRUE | 4ab | IR-ST 4x | 1 | 0.999 | 1.084 | 1 | 0 | 0
| 18 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0614_G11.0 | MB 6571 | 10BfR L0614 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.828 | 1 | 0 | 0
| 18 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0614_G12.0 | MB 6571 | 10BfR L0614 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.819 | 1 | 0 | 0
| 18 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0614_G10.0 | MB 6571 | 10BfR L0614 | TRUE | 4ab | IR-ST 4x | 1 | 1 | 0.859 | 1 | 0 | 0
| 18 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0614_G11.0 | MB 6571 | 10BfR L0614 | TRUE | 4ab | IR-ST 4x | 1 | 0.971 | 0.839 | 1 | 0 | 0
| 18 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0614_G12.0 | MB 6571 | 10BfR L0614 | TRUE | 4ab | IR-ST 4x | 1 | 0.998 | 0.752 | 1 | 0 | 0
# | 19 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0709_H10.0 | MB 6572 | 10BfR L0709 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.88 | 1 | 0 | 0
| 19 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0709_H11.0 | MB 6572 | 10BfR L0709 | TRUE | 3a | IR-ST 3 | 1 | 1 | 1.095 | 1 | 0 | 0
| 19 | 220301-1507-Lis_mono_IR-ST_Vali_10BfR L0709_H12.0 | MB 6572 | 10BfR L0709 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.372 | 1 | 0 | 0
| 19 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0709_H10.0 | MB 6572 | 10BfR L0709 | TRUE | 3a | IR-ST 3 | 1 | 1 | 1.174 | 1 | 0 | 0
| 19 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0709_H11.0 | MB 6572 | 10BfR L0709 | FAIL | 3a | IR-ST 3 | 1 | 1 | 1.984 | 1 | 0 | 0
| 19 | 220302-1210-Lis_mono_IR-ST_Vali_10BfR L0709_H12.0 | MB 6572 | 10BfR L0709 | TRUE | 3a | IR-ST 3 | 1 | 0.997 | 0.656 | 1 | 0 | 0
# | 20 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0711_A4.0 | MB 6573 | 10BfR L0711 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.517 | 1 | 0 | 0
| 20 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0711_A5.0 | MB 6573 | 10BfR L0711 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.317 | 1 | 0 | 0
| 20 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0711_A6.0 | MB 6573 | 10BfR L0711 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.3 | 1 | 0 | 0
| 20 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0711_F10.0 | MB 6573 | 10BfR L0711 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.277 | 1 | 0 | 0
| 20 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0711_A4.0 | MB 6573 | 10BfR L0711 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.293 | 1 | 0 | 0
| 20 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0711_A5.0 | MB 6573 | 10BfR L0711 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.639 | 1 | 0 | 0
| 20 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0711_A6.0 | MB 6573 | 10BfR L0711 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.603 | 1 | 0 | 0
# | 21 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0873_E4.0 | MB 6582 | 10BfR L0873 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.274 | 1 | 0 | 0
| 21 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0873_E5.0 | MB 6582 | 10BfR L0873 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.275 | 1 | 0 | 0
| 21 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0873_E6.0 | MB 6582 | 10BfR L0873 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.358 | 1 | 0 | 0
| 21 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0873_E4.0 | MB 6582 | 10BfR L0873 | TRUE | 3a | IR-ST 3 | 1 | 1 | 1.299 | 1 | 0 | 0
| 21 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0873_E5.0 | MB 6582 | 10BfR L0873 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.52 | 1 | 0 | 0
| 21 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0873_E6.0 | MB 6582 | 10BfR L0873 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.54 | 1 | 0 | 0
# | 22 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0935_F4.0 | MB 6587 | 10BfR L0935 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.421 | 1 | 0 | 0
| 22 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0935_F5.0 | MB 6587 | 10BfR L0935 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.797 | 1 | 0 | 0
| 22 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L0935_F6.0 | MB 6587 | 10BfR L0935 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.355 | 1 | 0 | 0
| 22 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0935_F4.0 | MB 6587 | 10BfR L0935 | FAIL | 3a | IR-ST 3 | 1 | 1 | 0.422 | 1 | 0 | 0
| 22 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0935_F5.0 | MB 6587 | 10BfR L0935 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.523 | 1 | 0 | 0
| 22 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L0935_F6.0 | MB 6587 | 10BfR L0935 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.501 | 1 | 0 | 0
# | 23 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1269_H4.0 | MB 6601 | 10BfR L1269 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.417 | 1 | 0 | 0
| 23 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1269_H5.0 | MB 6601 | 10BfR L1269 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.474 | 1 | 0 | 0
| 23 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1269_H6.0 | MB 6601 | 10BfR L1269 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.459 | 1 | 0 | 0
| 23 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1269_H4.0 | MB 6601 | 10BfR L1269 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.545 | 1 | 0 | 0
| 23 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1269_H5.0 | MB 6601 | 10BfR L1269 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.611 | 1 | 0 | 0
| 23 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1269_H6.0 | MB 6601 | 10BfR L1269 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.519 | 1 | 0 | 0
# | 24 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1270_A7.0 | MB 6602 | 10BfR L1270 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.642 | 1 | 0 | 0
| 24 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1270_A8.0 | MB 6602 | 10BfR L1270 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.623 | 1 | 0 | 0
| 24 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1270_A9.0 | MB 6602 | 10BfR L1270 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.664 | 1 | 0 | 0
| 24 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1270_A7.0 | MB 6602 | 10BfR L1270 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.641 | 1 | 0 | 0
| 24 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1270_A8.0 | MB 6602 | 10BfR L1270 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.628 | 1 | 0 | 0
| 24 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1270_A9.0 | MB 6602 | 10BfR L1270 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.6 | 1 | 0 | 0
# | 25 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1297_C7.0 | MB 6607 | 10BfR L1297 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.325 | 1 | 0 | 0
| 25 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1297_C8.0 | MB 6607 | 10BfR L1297 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.419 | 1 | 0 | 0
| 25 | 220302-1444-Lis_mono_IR-ST_Vali_10BfR L1297_C9.0 | MB 6607 | 10BfR L1297 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.35 | 1 | 0 | 0
| 25 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1297_C7.0 | MB 6607 | 10BfR L1297 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.777 | 1 | 0 | 0
| 25 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1297_C8.0 | MB 6607 | 10BfR L1297 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.51 | 1 | 0 | 0
| 25 | 220303-1155-Lis_mono_IR-ST_Vali_10BfR L1297_C9.0 | MB 6607 | 10BfR L1297 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.588 | 1 | 0 | 0
# | 26 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 29381 002_E10.0 | MB 6235 | 10KA 29381 002 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 0.724 | 1 | 0 | 0
| 26 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 29381 002_E11.0 | MB 6235 | 10KA 29381 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.666 | 1 | 0 | 0
| 26 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 29381 002_E12.0 | MB 6235 | 10KA 29381 002 | TRUE | 4b | IR-ST 4x | 1 | 0.998 | 1.169 | 1 | 0 | 0
| 26 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 29381 002_D10.0 | MB 6235 | 10KA 29381 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.485 | 1 | 0 | 0
| 26 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 29381 002_D11.0 | MB 6235 | 10KA 29381 002 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.512 | 1 | 0 | 0
| 26 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 29381 002_D12.0 | MB 6235 | 10KA 29381 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.574 | 1 | 0 | 0
| 26 | 220323-1410-_10KA 29381 002_H4.0 | MB 6235 | 10KA 29381 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.69 | 1 | 0 | 0
| 26 | 220323-1410-_10KA 29381 002_H5.0 | MB 6235 | 10KA 29381 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.517 | 1 | 0 | 0
| 26 | 220323-1410-_10KA 29381 002_H6.0 | MB 6235 | 10KA 29381 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.668 | 1 | 0 | 0
# | 27 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 29853 002_F10.0 | MB 6236 | 10KA 29853 002 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 0.979 | 1 | 0 | 0
| 27 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 29853 002_F11.0 | MB 6236 | 10KA 29853 002 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 0.815 | 1 | 0 | 0
| 27 | 220301-1142-Lis_mono_IR-ST_Vali_10KA 29853 002_F12.0 | MB 6236 | 10KA 29853 002 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.787 | 1 | 0 | 0
| 27 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 29853 002_E10.0 | MB 6236 | 10KA 29853 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.634 | 1 | 0 | 0
| 27 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 29853 002_E11.0 | MB 6236 | 10KA 29853 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.687 | 1 | 0 | 0
| 27 | 220304-1020-Lis_mono_IR-ST_Vali_10KA 29853 002_E12.0 | MB 6236 | 10KA 29853 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.639 | 1 | 0 | 0
| 27 | 220323-1410-_10KA 29853 002_A7.0 | MB 6236 | 10KA 29853 002 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.734 | 1 | 0 | 0
| 27 | 220323-1410-_10KA 29853 002_A8.0 | MB 6236 | 10KA 29853 002 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.557 | 1 | 0 | 0
| 27 | 220323-1410-_10KA 29853 002_A9.0 | MB 6236 | 10KA 29853 002 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.635 | 1 | 0 | 0
# | 28 | 220301-1142-Lis_mono_IR-ST_Vali_10L 06243 008 A1_G7.0 | MB 5798 | 10L 06243 008 A1 | FAIL | 3a | IR-ST 3 | 1 | 1 | 1.068 | 1 | 0 | 0
| 28 | 220301-1142-Lis_mono_IR-ST_Vali_10L 06243 008 A1_G8.0 | MB 5798 | 10L 06243 008 A1 | FAIL | 3a | IR-ST 3 | 1 | 1 | 0.782 | 1 | 0 | 0
| 28 | 220301-1142-Lis_mono_IR-ST_Vali_10L 06243 008 A1_G9.0 | MB 5798 | 10L 06243 008 A1 | FAIL | 3a | IR-ST 3 | 1 | 1 | 0.951 | 1 | 0 | 0
| 28 | 220304-1020-Lis_mono_IR-ST_Vali_10L 06243 008 A1_F7.0 | MB 5798 | 10L 06243 008 A1 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.465 | 1 | 0 | 0
| 28 | 220304-1020-Lis_mono_IR-ST_Vali_10L 06243 008 A1_F8.0 | MB 5798 | 10L 06243 008 A1 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.507 | 1 | 0 | 0
| 28 | 220304-1020-Lis_mono_IR-ST_Vali_10L 06243 008 A1_F9.0 | MB 5798 | 10L 06243 008 A1 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.509 | 1 | 0 | 0
| 28 | 220323-1410-_10L 06243 008 A1_D4.0 | MB 5798 | 10L 06243 008 A1 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.643 | 1 | 0 | 0
| 28 | 220323-1410-_10L 06243 008 A1_D5.0 | MB 5798 | 10L 06243 008 A1 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.794 | 1 | 0 | 0
| 28 | 220323-1410-_10L 06243 008 A1_D6.0 | MB 5798 | 10L 06243 008 A1 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.977 | 1 | 0 | 0
# | 29 | 220302-1444-Lis_mono_IR-ST_Vali_10L 67834 A1_D7.0 | MB 6902 | 10L 67834 A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.512 | 1 | 0 | 0
| 29 | 220302-1444-Lis_mono_IR-ST_Vali_10L 67834 A1_D8.0 | MB 6902 | 10L 67834 A1 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.537 | 1 | 0 | 0
| 29 | 220302-1444-Lis_mono_IR-ST_Vali_10L 67834 A1_D9.0 | MB 6902 | 10L 67834 A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.566 | 1 | 0 | 0
| 29 | 220303-1155-Lis_mono_IR-ST_Vali_10L 67834 A1_D7.0 | MB 6902 | 10L 67834 A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.685 | 1 | 0 | 0
| 29 | 220303-1155-Lis_mono_IR-ST_Vali_10L 67834 A1_D8.0 | MB 6902 | 10L 67834 A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.936 | 1 | 0 | 0
| 29 | 220303-1155-Lis_mono_IR-ST_Vali_10L 67834 A1_D9.0 | MB 6902 | 10L 67834 A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.966 | 1 | 0 | 0
# | 30 | 220302-1444-Lis_mono_IR-ST_Vali_10L 88874 KELI A1_F7.0 | MB 7133 | 10L 88874 KELI A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.456 | 1 | 0 | 0
| 30 | 220302-1444-Lis_mono_IR-ST_Vali_10L 88874 KELI A1_F8.0 | MB 7133 | 10L 88874 KELI A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.455 | 1 | 0 | 0
| 30 | 220302-1444-Lis_mono_IR-ST_Vali_10L 88874 KELI A1_F9.0 | MB 7133 | 10L 88874 KELI A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.468 | 1 | 0 | 0
| 30 | 220303-1155-Lis_mono_IR-ST_Vali_10L 88874 KELI A1_F7.0 | MB 7133 | 10L 88874 KELI A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.951 | 1 | 0 | 0
| 30 | 220303-1155-Lis_mono_IR-ST_Vali_10L 88874 KELI A1_F8.0 | MB 7133 | 10L 88874 KELI A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.537 | 1 | 0 | 0
| 30 | 220303-1155-Lis_mono_IR-ST_Vali_10L 88874 KELI A1_F9.0 | MB 7133 | 10L 88874 KELI A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.468 | 1 | 0 | 0
# | 31 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90650 002 SK2_H7.0 | MB 7198 | 10L 90650 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.483 | 1 | 0 | 0
| 31 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90650 002 SK2_H8.0 | MB 7198 | 10L 90650 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.512 | 1 | 0 | 0
| 31 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90650 002 SK2_H9.0 | MB 7198 | 10L 90650 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.472 | 1 | 0 | 0
| 31 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90650 002 SK2_H7.0 | MB 7198 | 10L 90650 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.389 | 1 | 0 | 0
| 31 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90650 002 SK2_H8.0 | MB 7198 | 10L 90650 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.543 | 1 | 0 | 0
| 31 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90650 002 SK2_H9.0 | MB 7198 | 10L 90650 002 SK2 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.653 | 1 | 0 | 0
# | 32 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90650 003 SK4_A10.0 | MB 7200 | 10L 90650 003 SK4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.571 | 1 | 0 | 0
| 32 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90650 003 SK4_A11.0 | MB 7200 | 10L 90650 003 SK4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.54 | 1 | 0 | 0
| 32 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90650 003 SK4_A12.0 | MB 7200 | 10L 90650 003 SK4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.497 | 1 | 0 | 0
| 32 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90650 003 SK4_A10.0 | MB 7200 | 10L 90650 003 SK4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.534 | 1 | 0 | 0
| 32 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90650 003 SK4_A11.0 | MB 7200 | 10L 90650 003 SK4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.525 | 1 | 0 | 0
| 32 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90650 003 SK4_A12.0 | MB 7200 | 10L 90650 003 SK4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.534 | 1 | 0 | 0
# | 33 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90654 002 SK1_B10.0 | MB 7202 | 10L 90654 002 SK1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.496 | 1 | 0 | 0
| 33 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90654 002 SK1_B11.0 | MB 7202 | 10L 90654 002 SK1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.477 | 1 | 0 | 0
| 33 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90654 002 SK1_B12.0 | MB 7202 | 10L 90654 002 SK1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.472 | 1 | 0 | 0
| 33 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90654 002 SK1_B10.0 | MB 7202 | 10L 90654 002 SK1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.66 | 1 | 0 | 0
| 33 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90654 002 SK1_B11.0 | MB 7202 | 10L 90654 002 SK1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.475 | 1 | 0 | 0
| 33 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90654 002 SK1_B12.0 | MB 7202 | 10L 90654 002 SK1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.488 | 1 | 0 | 0
# | 34 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90654 002 SK2_C10.0 | MB 7203 | 10L 90654 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.469 | 1 | 0 | 0
| 34 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90654 002 SK2_C11.0 | MB 7203 | 10L 90654 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.514 | 1 | 0 | 0
| 34 | 220302-1444-Lis_mono_IR-ST_Vali_10L 90654 002 SK2_C12.0 | MB 7203 | 10L 90654 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.506 | 1 | 0 | 0
| 34 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90654 002 SK2_C10.0 | MB 7203 | 10L 90654 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.753 | 1 | 0 | 0
| 34 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90654 002 SK2_C11.0 | MB 7203 | 10L 90654 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.725 | 1 | 0 | 0
| 34 | 220303-1155-Lis_mono_IR-ST_Vali_10L 90654 002 SK2_C12.0 | MB 7203 | 10L 90654 002 SK2 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.791 | 1 | 0 | 0
# | 35 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94279 K A1_E4.0 | MB 7245 | 10L 94279 K A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.432 | 1 | 0 | 0
| 35 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94279 K A1_E5.0 | MB 7245 | 10L 94279 K A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.48 | 1 | 0 | 0
| 35 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94279 K A1_E6.0 | MB 7245 | 10L 94279 K A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.452 | 1 | 0 | 0
| 35 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94279 K A1_E4.0 | MB 7245 | 10L 94279 K A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.526 | 1 | 0 | 0
| 35 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94279 K A1_E5.0 | MB 7245 | 10L 94279 K A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.537 | 1 | 0 | 0
| 35 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94279 K A1_E6.0 | MB 7245 | 10L 94279 K A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.632 | 1 | 0 | 0
# | 36 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94279 K P2_D4.0 | MB 7244 | 10L 94279 K P2 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.362 | 1 | 0 | 0
| 36 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94279 K P2_D5.0 | MB 7244 | 10L 94279 K P2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.412 | 1 | 0 | 0
| 36 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94279 K P2_D6.0 | MB 7244 | 10L 94279 K P2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.438 | 1 | 0 | 0
| 36 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94279 K P2_D4.0 | MB 7244 | 10L 94279 K P2 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.813 | 1 | 0 | 0
| 36 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94279 K P2_D5.0 | MB 7244 | 10L 94279 K P2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.656 | 1 | 0 | 0
| 36 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94279 K P2_D6.0 | MB 7244 | 10L 94279 K P2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.648 | 1 | 0 | 0
# | 37 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94644 A A3_B4.0 | MB 7241 | 10L 94644 A A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.407 | 1 | 0 | 0
| 37 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94644 A A3_B5.0 | MB 7241 | 10L 94644 A A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.489 | 1 | 0 | 0
| 37 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94644 A A3_B6.0 | MB 7241 | 10L 94644 A A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.503 | 1 | 0 | 0
| 37 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94644 A A3_B4.0 | MB 7241 | 10L 94644 A A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.683 | 1 | 0 | 0
| 37 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94644 A A3_B5.0 | MB 7241 | 10L 94644 A A3 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.931 | 1 | 0 | 0
| 37 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94644 A A3_B6.0 | MB 7241 | 10L 94644 A A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.801 | 1 | 0 | 0
# | 38 | 220302-1444-Lis_mono_IR-ST_Vali_10L 94644 K P5_E10.0 | MB 7237 | 10L 94644 K P5 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.45 | 1 | 0 | 0
| 38 | 220302-1444-Lis_mono_IR-ST_Vali_10L 94644 K P5_E11.0 | MB 7237 | 10L 94644 K P5 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.432 | 1 | 0 | 0
| 38 | 220302-1444-Lis_mono_IR-ST_Vali_10L 94644 K P5_E12.0 | MB 7237 | 10L 94644 K P5 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.452 | 1 | 0 | 0
| 38 | 220303-1155-Lis_mono_IR-ST_Vali_10L 94644 K P5_E10.0 | MB 7237 | 10L 94644 K P5 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.434 | 1 | 0 | 0
| 38 | 220303-1155-Lis_mono_IR-ST_Vali_10L 94644 K P5_E11.0 | MB 7237 | 10L 94644 K P5 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.488 | 1 | 0 | 0
| 38 | 220303-1155-Lis_mono_IR-ST_Vali_10L 94644 K P5_E12.0 | MB 7237 | 10L 94644 K P5 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.564 | 1 | 0 | 0
# | 39 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94645 A P4_C4.0 | MB 7242 | 10L 94645 A P4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.432 | 1 | 0 | 0
| 39 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94645 A P4_C5.0 | MB 7242 | 10L 94645 A P4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.475 | 1 | 0 | 0
| 39 | 220303-1403-Lis_mono_IR-ST_Vali_10L 94645 A P4_C6.0 | MB 7242 | 10L 94645 A P4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.439 | 1 | 0 | 0
| 39 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94645 A P4_C4.0 | MB 7242 | 10L 94645 A P4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.478 | 1 | 0 | 0
| 39 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94645 A P4_C5.0 | MB 7242 | 10L 94645 A P4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.514 | 1 | 0 | 0
| 39 | 220316-1359-Lis_mono_IR-ST_Vali_10L 94645 A P4_C6.0 | MB 7242 | 10L 94645 A P4 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.545 | 1 | 0 | 0
# | 40 | 220303-1403-Lis_mono_IR-ST_Vali_11L 15492 001 KEZ A2_G4.0 | MB 7702 | 11L 15492 001 KEZ A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.42 | 1 | 0 | 0
| 40 | 220303-1403-Lis_mono_IR-ST_Vali_11L 15492 001 KEZ A2_G5.0 | MB 7702 | 11L 15492 001 KEZ A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.51 | 1 | 0 | 0
| 40 | 220303-1403-Lis_mono_IR-ST_Vali_11L 15492 001 KEZ A2_G6.0 | MB 7702 | 11L 15492 001 KEZ A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.506 | 1 | 0 | 0
| 40 | 220316-1359-Lis_mono_IR-ST_Vali_11L 15492 001 KEZ A2_G4.0 | MB 7702 | 11L 15492 001 KEZ A2 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.682 | 1 | 0 | 0
| 40 | 220316-1359-Lis_mono_IR-ST_Vali_11L 15492 001 KEZ A2_G5.0 | MB 7702 | 11L 15492 001 KEZ A2 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.67 | 1 | 0 | 0
| 40 | 220316-1359-Lis_mono_IR-ST_Vali_11L 15492 001 KEZ A2_G6.0 | MB 7702 | 11L 15492 001 KEZ A2 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.945 | 1 | 0 | 0
# | 41 | 220303-1403-Lis_mono_IR-ST_Vali_11L 15492 005 KEZ A10_F4.0 | MB 7701 | 11L 15492 005 KEZ A10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.629 | 1 | 0 | 0
| 41 | 220303-1403-Lis_mono_IR-ST_Vali_11L 15492 005 KEZ A10_F5.0 | MB 7701 | 11L 15492 005 KEZ A10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.45 | 1 | 0 | 0
| 41 | 220303-1403-Lis_mono_IR-ST_Vali_11L 15492 005 KEZ A10_F6.0 | MB 7701 | 11L 15492 005 KEZ A10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.454 | 1 | 0 | 0
| 41 | 220316-1359-Lis_mono_IR-ST_Vali_11L 15492 005 KEZ A10_F4.0 | MB 7701 | 11L 15492 005 KEZ A10 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.695 | 1 | 0 | 0
| 41 | 220316-1359-Lis_mono_IR-ST_Vali_11L 15492 005 KEZ A10_F5.0 | MB 7701 | 11L 15492 005 KEZ A10 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.499 | 1 | 0 | 0
| 41 | 220316-1359-Lis_mono_IR-ST_Vali_11L 15492 005 KEZ A10_F6.0 | MB 7701 | 11L 15492 005 KEZ A10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.606 | 1 | 0 | 0
# | 42 | 220303-1403-Lis_mono_IR-ST_Vali_11L 25437 004 Anli A2_H4.0 | MB 7845 | 11L 25437 004 Anli A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.552 | 1 | 0 | 0
| 42 | 220303-1403-Lis_mono_IR-ST_Vali_11L 25437 004 Anli A2_H5.0 | MB 7845 | 11L 25437 004 Anli A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.659 | 1 | 0 | 0
| 42 | 220303-1403-Lis_mono_IR-ST_Vali_11L 25437 004 Anli A2_H6.0 | MB 7845 | 11L 25437 004 Anli A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.553 | 1 | 0 | 0
| 42 | 220316-1359-Lis_mono_IR-ST_Vali_11L 25437 004 Anli A2_H4.0 | MB 7845 | 11L 25437 004 Anli A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.55 | 1 | 0 | 0
| 42 | 220316-1359-Lis_mono_IR-ST_Vali_11L 25437 004 Anli A2_H5.0 | MB 7845 | 11L 25437 004 Anli A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.691 | 1 | 0 | 0
| 42 | 220316-1359-Lis_mono_IR-ST_Vali_11L 25437 004 Anli A2_H6.0 | MB 7845 | 11L 25437 004 Anli A2 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.702 | 1 | 0 | 0
# | 43 | 220303-1403-Lis_mono_IR-ST_Vali_11L 37481 A P2_E7.0 | MB 8080 | 11L 37481 A P2 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.351 | 1 | 0 | 0
| 43 | 220303-1403-Lis_mono_IR-ST_Vali_11L 37481 A P2_E8.0 | MB 8080 | 11L 37481 A P2 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.59 | 1 | 0 | 0
| 43 | 220303-1403-Lis_mono_IR-ST_Vali_11L 37481 A P2_E9.0 | MB 8080 | 11L 37481 A P2 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.351 | 1 | 0 | 0
| 43 | 220316-1359-Lis_mono_IR-ST_Vali_11L 37481 A P2_E7.0 | MB 8080 | 11L 37481 A P2 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.955 | 1 | 0 | 0
| 43 | 220316-1359-Lis_mono_IR-ST_Vali_11L 37481 A P2_E8.0 | MB 8080 | 11L 37481 A P2 | TRUE | 3a | IR-ST 3 | 1 | 1 | 1.015 | 1 | 0 | 0
| 43 | 220316-1359-Lis_mono_IR-ST_Vali_11L 37481 A P2_E9.0 | MB 8080 | 11L 37481 A P2 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.508 | 1 | 0 | 0
# | 44 | 220303-1403-Lis_mono_IR-ST_Vali_11L 39018 A3_B7.0 | MB 8072 | 11L 39018 A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.531 | 1 | 0 | 0
| 44 | 220303-1403-Lis_mono_IR-ST_Vali_11L 39018 A3_B8.0 | MB 8072 | 11L 39018 A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.562 | 1 | 0 | 0
| 44 | 220303-1403-Lis_mono_IR-ST_Vali_11L 39018 A3_B9.0 | MB 8072 | 11L 39018 A3 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.56 | 1 | 0 | 0
| 44 | 220316-1359-Lis_mono_IR-ST_Vali_11L 39018 A3_B7.0 | MB 8072 | 11L 39018 A3 | TRUE | 4b | IR-ST 4x | 1 | 0.998 | 1.003 | 1 | 0 | 0
| 44 | 220316-1359-Lis_mono_IR-ST_Vali_11L 39018 A3_B8.0 | MB 8072 | 11L 39018 A3 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.809 | 1 | 0 | 0
| 44 | 220316-1359-Lis_mono_IR-ST_Vali_11L 39018 A3_B9.0 | MB 8072 | 11L 39018 A3 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.818 | 1 | 0 | 0
# | 45 | 220303-1403-Lis_mono_IR-ST_Vali_11L 39018 P2_C7.0 | MB 8073 | 11L 39018 P2 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.485 | 1 | 0 | 0
| 45 | 220303-1403-Lis_mono_IR-ST_Vali_11L 39018 P2_C8.0 | MB 8073 | 11L 39018 P2 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.476 | 1 | 0 | 0
| 45 | 220303-1403-Lis_mono_IR-ST_Vali_11L 39018 P2_C9.0 | MB 8073 | 11L 39018 P2 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.468 | 1 | 0 | 0
| 45 | 220316-1359-Lis_mono_IR-ST_Vali_11L 39018 P2_C7.0 | MB 8073 | 11L 39018 P2 | FAIL | 4b | IR-ST 4x | 1 | 0.998 | 0.91 | 1 | 0 | 0
| 45 | 220316-1359-Lis_mono_IR-ST_Vali_11L 39018 P2_C8.0 | MB 8073 | 11L 39018 P2 | TRUE | 4b | IR-ST 4x | 1 | 0.998 | 0.942 | 1 | 0 | 0
| 45 | 220316-1359-Lis_mono_IR-ST_Vali_11L 39018 P2_C9.0 | MB 8073 | 11L 39018 P2 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 1.06 | 1 | 0 | 0
# | 46 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77922 A1_C10.0 | MB 8543 | 11L 77922 A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.613 | 1 | 0 | 0
| 46 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77922 A1_C11.0 | MB 8543 | 11L 77922 A1 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.63 | 1 | 0 | 0
| 46 | 220303-1403-Lis_mono_IR-ST_Vali_11L 77922 A1_C12.0 | MB 8543 | 11L 77922 A1 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.629 | 1 | 0 | 0
| 46 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77922 A1_C10.0 | MB 8543 | 11L 77922 A1 | FAIL | 4b | IR-ST 4x | 1 | 0.987 | 1.446 | 1 | 0 | 0
| 46 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77922 A1_C11.0 | MB 8543 | 11L 77922 A1 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.912 | 1 | 0 | 0
| 46 | 220316-1359-Lis_mono_IR-ST_Vali_11L 77922 A1_C12.0 | MB 8543 | 11L 77922 A1 | TRUE | 4b | IR-ST 4x | 1 | 0.999 | 0.957 | 1 | 0 | 0
# | 47 | 220303-1403-Lis_mono_IR-ST_Vali_11L 84950 002 A3_E10.0 | MB 8646 | 11L 84950 002 A3 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.284 | 1 | 0 | 0
| 47 | 220303-1403-Lis_mono_IR-ST_Vali_11L 84950 002 A3_E11.0 | MB 8646 | 11L 84950 002 A3 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.321 | 1 | 0 | 0
| 47 | 220303-1403-Lis_mono_IR-ST_Vali_11L 84950 002 A3_E12.0 | MB 8646 | 11L 84950 002 A3 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.313 | 1 | 0 | 0
| 47 | 220316-1359-Lis_mono_IR-ST_Vali_11L 84950 002 A3_E10.0 | MB 8646 | 11L 84950 002 A3 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.534 | 1 | 0 | 0
| 47 | 220316-1359-Lis_mono_IR-ST_Vali_11L 84950 002 A3_E11.0 | MB 8646 | 11L 84950 002 A3 | TRUE | 3a | IR-ST 3 | 1 | 1 | 1.126 | 1 | 0 | 0
| 47 | 220316-1359-Lis_mono_IR-ST_Vali_11L 84950 002 A3_E12.0 | MB 8646 | 11L 84950 002 A3 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.483 | 1 | 0 | 0
# | 48 | 220301-1142-Lis_mono_IR-ST_Vali_13L 04475 001 HF P10_A7.0 | MB 1073,2 | 13L 04475 001 HF P10 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 0.838 | 1 | 0 | 0
| 48 | 220301-1142-Lis_mono_IR-ST_Vali_13L 04475 001 HF P10_A8.0 | MB 1073,2 | 13L 04475 001 HF P10 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 0.829 | 1 | 0 | 0
| 48 | 220301-1142-Lis_mono_IR-ST_Vali_13L 04475 001 HF P10_A9.0 | MB 1073,2 | 13L 04475 001 HF P10 | TRUE | 4b | IR-ST 4x | 1 | 0.998 | 1.192 | 1 | 0 | 0
| 48 | 220304-1020-Lis_mono_IR-ST_Vali_13L 04475 001 HF P10_A7.0 | MB 1073,2 | 13L 04475 001 HF P10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.526 | 1 | 0 | 0
| 48 | 220304-1020-Lis_mono_IR-ST_Vali_13L 04475 001 HF P10_A8.0 | MB 1073,2 | 13L 04475 001 HF P10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.508 | 1 | 0 | 0
| 48 | 220304-1020-Lis_mono_IR-ST_Vali_13L 04475 001 HF P10_A9.0 | MB 1073,2 | 13L 04475 001 HF P10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.598 | 1 | 0 | 0
| 48 | 220323-1410-_13L 04475 001 HF P10_A4.0 | MB 1073,2 | 13L 04475 001 HF P10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.543 | 1 | 0 | 0
| 48 | 220323-1410-_13L 04475 001 HF P10_A5.0 | MB 1073,2 | 13L 04475 001 HF P10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.402 | 1 | 0 | 0
| 48 | 220323-1410-_13L 04475 001 HF P10_A6.0 | MB 1073,2 | 13L 04475 001 HF P10 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.689 | 1 | 0 | 0
# | 49 | 220301-1142-Lis_mono_IR-ST_Vali_13L 04475 001 HF P7_H4.0 | MB 1072,2 | 13L 04475 001 HF P7 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 0.814 | 1 | 0 | 0
| 49 | 220301-1142-Lis_mono_IR-ST_Vali_13L 04475 001 HF P7_H5.0 | MB 1072,2 | 13L 04475 001 HF P7 | FAIL | 4b | IR-ST 4x | 1 | 0.999 | 0.751 | 1 | 0 | 0
| 49 | 220301-1142-Lis_mono_IR-ST_Vali_13L 04475 001 HF P7_H6.0 | MB 1072,2 | 13L 04475 001 HF P7 | FAIL | 4b | IR-ST 4x | 1 | 1 | 0.792 | 1 | 0 | 0
| 49 | 220304-1020-Lis_mono_IR-ST_Vali_13L 04475 001 HF P7_H4.0 | MB 1072,2 | 13L 04475 001 HF P7 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.491 | 1 | 0 | 0
| 49 | 220304-1020-Lis_mono_IR-ST_Vali_13L 04475 001 HF P7_H5.0 | MB 1072,2 | 13L 04475 001 HF P7 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.484 | 1 | 0 | 0
| 49 | 220304-1020-Lis_mono_IR-ST_Vali_13L 04475 001 HF P7_H6.0 | MB 1072,2 | 13L 04475 001 HF P7 | TRUE | 4b | IR-ST 4x | 1 | 1 | 0.499 | 1 | 0 | 0
# | 50 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L0065_C7.0 | MB 33647 | BfR L0065 | TRUE | 3b | IR-ST 3 | 1 | 1 | 1.197 | 1 | 0 | 0
| 50 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L0065_C8.0 | MB 33647 | BfR L0065 | TRUE | 3b | IR-ST 3 | 1 | 1 | 0.857 | 1 | 0 | 0
| 50 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L0065_C9.0 | MB 33647 | BfR L0065 | TRUE | 3b | IR-ST 3 | 1 | 1 | 1.066 | 1 | 0 | 0
| 50 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L0065_E4.0 | MB 33647 | BfR L0065 | TRUE | 3b | IR-ST 3 | 1 | 1 | 1.788 | 1 | 0 | 0
| 50 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L0065_E6.0 | MB 33647 | BfR L0065 | FAIL | 3b | IR-ST 3 | 1 | 1 | 1.808 | 1 | 0 | 0
# | 51 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1416_A4.0 | MB 33637 | BfR L1416 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.356 | 1 | 0 | 0
| 51 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1416_A5.0 | MB 33637 | BfR L1416 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.334 | 1 | 0 | 0
| 51 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1416_A6.0 | MB 33637 | BfR L1416 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.275 | 1 | 0 | 0
| 51 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1416_A4.0 | MB 33637 | BfR L1416 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.693 | 1 | 0 | 0
| 51 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1416_A5.0 | MB 33637 | BfR L1416 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.861 | 1 | 0 | 0
| 51 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1416_A6.0 | MB 33637 | BfR L1416 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.46 | 1 | 0 | 0
# | 52 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1437_B4.0 | MB 33638 | BfR L1437 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.431 | 1 | 0 | 0
| 52 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1437_B5.0 | MB 33638 | BfR L1437 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.421 | 1 | 0 | 0
| 52 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1437_B6.0 | MB 33638 | BfR L1437 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.569 | 1 | 0 | 0
| 52 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1437_B4.0 | MB 33638 | BfR L1437 | TRUE | 3a | IR-ST 3 | 1 | 1 | 1.063 | 1 | 0 | 0
| 52 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1437_B5.0 | MB 33638 | BfR L1437 | TRUE | 3a | IR-ST 3 | 1 | 1 | 0.953 | 1 | 0 | 0
| 52 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1437_B6.0 | MB 33638 | BfR L1437 | TRUE | 3a | IR-ST 3 | 1 | 1 | 1.006 | 1 | 0 | 0
# | 53 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1463_C4.0 | MB 33639 | BfR L1463 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.59 | 1 | 0 | 0
| 53 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1463_C5.0 | MB 33639 | BfR L1463 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.349 | 1 | 0 | 0
| 53 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1463_C6.0 | MB 33639 | BfR L1463 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.393 | 1 | 0 | 0
| 53 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1463_C4.0 | MB 33639 | BfR L1463 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.466 | 1 | 0 | 0
| 53 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1463_C5.0 | MB 33639 | BfR L1463 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.515 | 1 | 0 | 0
| 53 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1463_C6.0 | MB 33639 | BfR L1463 | TRUE | 3c | IR-ST 3 | 1 | 1 | 0.422 | 1 | 0 | 0
# | 54 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1836_E4.0 | MB 33641 | BfR L1836 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.49 | 1 | 0 | 0
| 54 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1836_E5.0 | MB 33641 | BfR L1836 | TRUE | 3a | IR-ST 7 | 1 | 0.998 | 0.542 | 1 | 0 | 0
| 54 | 220308-1056-Lis_mono_IR-ST_Vali_BfR L1836_E6.0 | MB 33641 | BfR L1836 | TRUE | 3a | IR-ST 7 | 1 | 0.998 | 0.56 | 1 | 0 | 0
| 54 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1836_D4.0 | MB 33641 | BfR L1836 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.595 | 1 | 0 | 0
| 54 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1836_D5.0 | MB 33641 | BfR L1836 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.563 | 1 | 0 | 0
| 54 | 220317-1314-Lis_mono_IR-ST_Vali_BfR L1836_D6.0 | MB 33641 | BfR L1836 | TRUE | 3a | IR-ST 7 | 1 | 0.999 | 0.571 | 1 | 0 | 0

## Report
Targeted Identification
following the Guidelines for Validating Species Identifications using MALDI-TOF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022
Listeria monocytogenes ST 1/2
Version: | 2
generated by | Dyk | checked by | ObH
generated on | 2023-01-31 | checked on | 2023-02-27
Application | Identification of Listeria (L.) monocytogenes serogroups
Environment | L. monocytogenes serogroups 1/2, 3, 4x, 7
Applied database / software version |  | Bruker classifier "Lmono serogrouper" |  | IR Biotyper v4.0 01/2023
Validation set
Number of target parameter isolates: 40 |  | number of non-target parameter isolates: 54
number of target parameter spectra: 269 |  | number of non-target parameter spectra: 361
cultivation conditions: Columbia Sheep Blood Agar, produced by Becton Dickinson (COL-SB BD), 37°C 24h aerobic
prep method: preparation acc. to producer instructions (Instruction for use IR Biotyper Kit, Revision B (03/2021))
Only identified samples are used for determination of the inclusivity and exclusivity.
True positive: classification score >/=0.6 AND outlier score </=2 AND target parameter classification = label |  | True negative: classification score >/=0.6 AND outlier score </=2 AND non-target parameter classification <> target parameter
False negative: true-positive criteria are not fullfilled |  | False positive: true-negative criteria are not fullfilled
Validation result | Target parameter: | Listeria monocytogenes ST 1/2 |  | target | fulfilled
Total number of spectra | 269
of which with ID result | 269 | rate of identified spectra | 100% | 90% | yes | 1
Total number of isolates | 40 |  |  | 20 | yes | 1
With respect to identified spectra: |  | Inklusivity
positive accordance (PA) | 269 | True positive rate | 100% | 95% | yes | 1
negative deviation (ND) | 0 | False negative rate | 0% | 1% | yes | 1
Non-target parameter
Total number of spectra | 361
of which with ID result | 361 | rate of identified spectra | 100% | 90% | yes | 1
Total number of isolates | 54 |  |  | 30 | yes | 1
With respect to identified spectra: |  | Exklusivity
negative accordance (NA) | 361 | True negative rate | 100% | 99% | yes | 1
positive deviation (PD) | 0 | False-Positive Rate | 0% | 1% | yes | 1
Out of 269 identified target parameter spectra, 269 (=100%) were classified correctly (inclusivity). 0 (=0%) of identified target parameter spectra were allocated incorrectly to a different serogroup.
Out of 361 identified non-target parameter spectra, 361 (=100%) were counted correctly as negative (exclusivity). 0 (=0%) of all spectra were allocated incorrectly as the target parameter.
Criteria for the | targeted identification | are thus | fullfilled
Comment: | Serogroup differentiation to be applied only onto L. monocytogenes positively identified on the species level.
Released for site | for department | on (date) | by (signature)
CVUA S | T, MT, D | 2023-02-28 | RAU
The procedure is suitable for the intended purpose and is thereby released for usage at site.
Analytical parameter added to LIMS:
A signed version of this validation report is kept within the generating lab.
yes
no