## Settings
Settings für die Validierung
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Institution / Labor | CVUA S
Parameter | Gattung (Genus) |  | Art (Spezies)
|  | Allium | ursinum
Unterscheidungsgrenzen
| Score zur Unterscheidung der Gattung | 1.7 | Score zur Unterscheidung der Spezies | 2

## Parameter (Spezies)
Parameter: | Allium | ursinum | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | nicht zugeordnet | richtig-positiv | falsch-negativ | falsch zugeordnet
|  |  | 1 | 1 | TX 017.1 | 3; determination by expert | Plant extr. | 2021-07-19 | Allium | ursinum | 2.32 | Allium | ursinum | 2.22 | A | 1 | 0 | 1 | 0 | 0
Zahl aller Spektren | 23 |  | 1 | 1 | TX 019 | 3; determination by expert | Plant extr. | 2021-07-19 | Allium | ursinum | 2.46 | Allium | ursinum | 2.25 | A | 1 | 0 | 1 | 0 | 0
davon Identifiziert | 23 |  | 1 | 1 | TX 051 | 3; determination by expert | Plant extr. | 2021-07-15 | Allium | ursinum | 2.28 | Allium | ursinum | 2.07 | A | 1 | 0 | 1 | 0 | 0
nicht identifiziert: | 0 |  | 1 | 1 | TX 198.1 | 3; determination by expert | Plant extr. | 2022-09-09 | Allium | ursinum | 2.38 | Allium | ursinum | 2.23 | A | 1 | 0 | 1 | 0 | 0
von allen identifizierten Isolaten |  |  | 1 | 1 | TX 497 | 1; label | Plant extr. | 2023-11-21 | Allium | ursinum | 2.29 | Allium | ursinum | 1.76 | A | 1 | 0 | 1 | 0 | 0
richtig-positiv: | 23 |  | 1 | 1 | TX 498 | 3; label, PCR | Plant extr.-R | 2022-11-23 | Allium | ursinum | 2.35 | Allium | ursinum | 1.68 | A | 1 | 0 | 1 | 0 | 0
falsch zugeordnet | 0 |  | 1 | 1 | TX 693 | 3; determination by expert | Plant extr. | 2023-11-20 | Allium | ursinum | 2.22 | Allium | ursinum | 2.18 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 696 | 3; determination by expert | Plant extr. | 2023-11-15 | Allium | ursinum | 2.15 | Allium | ursinum | 2.11 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 702 | 3; determination by expert | Plant extr. | 2023-11-15 | Allium | ursinum | 2.18 | Allium | ursinum | 2.07 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 722.2 | 3; determination by expert | Plant extr. | 2023-10-26 | Allium | ursinum | 2.42 | Allium | ursinum | 2.31 | A | 1 | 0 | 1 | 0 | 0
Scorewertgrenzen: | 2 | Species | 1 | 1 | TX 763.2 | 3; determination by expert | Plant extr. | 2023-10-27 | Allium | ursinum | 2.41 | Allium | ursinum | 2.4 | A | 1 | 0 | 1 | 0 | 0
| 1.7 | Genus | 1 | 1 | TX 766.2 | 3; determination by expert | Plant extr. | 2023-11-08 | Allium | ursinum | 2.12 | Allium | ursinum | 2.12 | A | 1 | 0 | 1 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 1 | 1 | TX 767.2 | 3; determination by expert | Plant extr. | 2023-10-27 | Allium | ursinum | 2.38 | Allium | ursinum | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 772 | 3; determination by expert | Plant extr. | 2023-11-08 | Allium | ursinum | 2.18 | Allium | ursinum | 2.07 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 773 | 3; determination by expert | Plant extr. | 2023-10-31 | Allium | ursinum | 2.14 | Allium | ursinum | 2.08 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 780 | 3; determination by expert | Plant extr. | 2023-11-06 | Allium | ursinum | 2.3 | Allium | ursinum | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 786 | 3; determination by expert | Plant extr. | 2023-11-09 | Allium | ursinum | 2.14 | Allium | ursinum | 2.08 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 787 | 3; determination by expert | Plant extr. | 2023-11-14 | Allium | ursinum | 2.19 | Allium | ursinum | 2.17 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 788 | 3; determination by expert | Plant extr. | 2023-11-20 | Allium | ursinum | 2.41 | Allium | ursinum | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 794 | 3; determination by expert | Plant extr. | 2023-11-20 | Allium | ursinum | 2.11 | Allium | ursinum | 1.97 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 805 | 3; determination by expert | Plant extr. | 2023-11-14 | Allium | ursinum | 2.24 | Allium | ursinum | 2.08 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 806 | 3; determination by expert | Plant extr. | 2023-11-20 | Allium | ursinum | 2.25 | Allium | ursinum | 2.18 | A | 1 | 0 | 1 | 0 | 0
davon mit ID-Ergebnis | 23 |  | 1 | 1 | TX 853 | 3; label; PCR | Plant extr. | 2023-11-22 | Allium | ursinum | 2.21 | Allium | ursinum | 1.66 | A | 1 | 0 | 1 | 0 | 0
Positive Übereinstimmungen (PA) | 23
Negative Abweichungen (ND) | 0
nicht identifiziert in % | 0%
falsch zugeordnert in % | 0%

## #Parameter (Spezies) DB
Parameter: | Allium ursinum |  | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | verwendete Gattung | verwendete Art | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | richtige ID | keine richtige ID | 1. Hit | 2. Hit | 1. Hit Treffer Parameter | 2. Hit Treffer Parameter | falsch als Parameter identifiziert | PM unter den ersten 2 Hits
|  |  | 1 | 1 | TX 003 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus | precatorius | Abrus | precatorius | 2.26 | Abrus | precatorius | 2.24 | A | 1 | 1 | 0 | Abrus | Abrus | 0 | 0 | 0 | 0
Anzahl Isolate: | 578 |  | 1 | 1 | TX 006 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus | precatorius | Abrus | precatorius | 2.39 | Abrus | precatorius | 2.18 | A | 1 | 1 | 0 | Abrus | Abrus | 0 | 0 | 0 | 0
davon Identifiziert | 261 |  | 1 | 1 | TX 379.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Aconitum | carmichaelii | Aconitum | carmichaelii | 2.53 | Camellia | sinensis | 1.21 | A | 1 | 1 | 0 | Aconitum | Camellia | 0 | 0 | 0 | 0
von allen identifizierten Isolaten |  |  | 1 | 1 | TX 038.1 | 3; determination by expert | Plant extr. | 2022-01-14 | Aconitum | napellus | Aconitum | napellus | 2.29 | Aconitum | carmichaelii | 1.53 | A | 1 | 1 | 0 | Aconitum | Aconitum | 0 | 0 | 0 | 0
richtige ID | 261 |  | 1 | 0 | TX 038.2 | 3; determination by expert | Plant extr. | 2022-01-26 | Aconitum | napellus | Camellia | sinensis | 1.17 | Aconitum | carmichaelii | 0.95 | B | 0 | 0 | 1 | Camellia | Aconitum | 0 | 0 | 0 | 0
falsch als Parameter identifiziert | 0 |  | 1 | 0 | TX 055.1 | 2; determination by expert | Plant extr. | 2022-01-20 | Aconitum | napellus | Aconitum | carmichaelii | 1.35 | Asarum | europaeum | 1.22 | B | 0 | 0 | 1 | Aconitum | Asarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 076.1 | 3; determination by expert | Plant extr. | 2022-01-14 | Aconitum | napellus | Aconitum | napellus | 1.88 | Aconitum | carmichaelii | 1.45 | B | 0 | 0 | 1 | Aconitum | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 089 | 2; determination by expert | Plant extr. | 2022-01-14 | Aconitum | napellus | Petroselinum | crispum | 1.36 | Aconitum | carmichaelii | 1.35 | B | 0 | 0 | 1 | Petroselinum | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 122.1 | 2; determination by expert | Plant extr. | 2022-01-14 | Aconitum | napellus | Aconitum | carmichaelii | 1.65 | Aconitum | napellus | 1.53 | B | 0 | 0 | 1 | Aconitum | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 037.1 | 3; determination by expert | Plant extr. | 2022-10-07 | Aethusa | cynapium | Aethusa | cynapium | 2.23 | Conium | maculatum | 1.88 | C | 0 | 0 | 1 | Aethusa | Conium | 0 | 0 | 0 | 0
Parameter wird getroffen mit: |  |  | 1 | 0 | TX 037.2 | 3; determination by expert | Plant extr. | 2022-01-25 | Aethusa | cynapium | Aethusa | cynapium | 2.53 | Conium | maculatum | 2.4 | C | 0 | 0 | 1 | Aethusa | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 854.1 | 3; label, sensory analysis | Plant extr. | 2023-11-23 | Allium | fistulosum | Allium | fistulosum | 2.36 | Allium | schoenoprasum | 1.98 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
1. Hit Treffer Parameter | 0 |  | 1 | 1 | TX 851 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium | porrum | Allium | porrum | 2.46 | Allium | tuberosum | 1.15 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
2. Hit Treffer Parameter | 0 |  | 1 | 1 | TX 082.1 | 3; label; sensorical analysis | Plant extr. | 2021-07-30 | Allium | sativum | Allium | sativum | 2.21 | Allium | schoenoprasum | 1.32 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
PM unter den ersten 2 Hits | 0 |  | 1 | 0 | TX 082.2 | 3; label; sensorical analysis | Plant extr. | 2022-01-26 | Allium | sativum | Allium | sativum | 1.29 | Allium | tuberosum | 1.12 | B | 0 | 0 | 1 | Allium | Allium | 0 | 0 | 0 | 0
in % | 0% |  | 1 | 1 | TX 084 | 3; determination by expert | Plant extr. | 2021-07-27 | Allium | schoenoprasum | Allium | schoenoprasum | 2.54 | Allium | schoenoprasum | 2 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 171.1 | 3; determination by expert | Plant extr. | 2023-11-15 | Allium | schoenoprasum | Allium | schoenoprasum | 2.25 | Allium | schoenoprasum | 1.83 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 848 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium | schoenoprasum | Allium | schoenoprasum | 2.21 | Allium | schoenoprasum | 1.77 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
Scorewertgrenzen: | 2 | Species | 1 | 1 | TX 849 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium | schoenoprasum | Allium | schoenoprasum | 2.52 | Allium | vineale | 1.63 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
| 1.7 | Genus | 1 | 1 | TX 856 | 3; label; sensory analysis | Plant extr. | 2023-11-24 | Allium | schoenoprasum | Allium | schoenoprasum | 2.12 | Allium | schoenoprasum | 2.04 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 1 | 1 | TX 847 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium | tuberosum | Allium | tuberosum | 2.3 | Allium | schoenoprasum | 1.35 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 178 | 3; determination by expert | Plant extr. | 2023-11-15 | Allium | vineale | Allium | vineale | 2.28 | Allium | tuberosum | 1.5 | A | 1 | 1 | 0 | Allium | Allium | 0 | 0 | 0 | 0
davon mit ID-Ergebnis | 261 |  | 1 | 0 | TX 601.1 | 2; expert opinion | Plant extr.-R | 2022-12-21 | Aloe | vera | Vaccinium | myrtillus | 0.94 | Vaccinium | myrtillus | 0.92 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
Negative Übereinstimmungen (NA) | 261 |  | 1 | 1 | TX 413 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Alpinia | officinarum | Alpinia | officinarum | 2.3 | Aethusa | cynapium | 1.25 | A | 1 | 1 | 0 | Alpinia | Aethusa | 0 | 0 | 0 | 0
Positive Abweichungen (PD) | 0 |  | 1 | 0 | TX 402.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Amaranthus | sp | Lycium | chinense | 1.29 | Arum | italicum | 1.23 | B | 0 | 0 | 1 | Lycium | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 457.1 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Ampelopsis | sinica | Camellia | sinensis | 1.22 | Leucojum | vernum | 1.17 | B | 0 | 0 | 1 | Camellia | Leucojum | 0 | 0 | 0 | 0
nicht identifiziert in % | 54.84% |  | 1 | 0 | TX 457.2 | 3; botanical collection | Plant extr.-R | 2022-11-02 | Ampelopsis | sinica | Vaccinium | myrtillus | 1.08 | Vaccinium | sect. | 1.02 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
falsch zugeordnert in % | 0% |  | 1 | 1 | TX 132.1 | 2; label | Plant extr. | 2022-02-02 | Anethum | graveolens | Anethum | graveolens | 2.02 | Petroselinum | crispum | 1.37 | A | 1 | 1 | 0 | Anethum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 132.2 | 2; label | Plant extr. | 2022-01-14 | Anethum | graveolens | Anethum | graveolens | 2.32 | Petroselinum | crispum | 1.42 | A | 1 | 1 | 0 | Anethum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 132.3 | 2; label | Plant extr. | 2022-01-14 | Anethum | graveolens | Anethum | graveolens | 2.02 | Chamaecyparis | lawsoniana | 1.27 | A | 1 | 1 | 0 | Anethum | Chamaecyparis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 136 | 2; label | Plant extr. | 2022-01-14 | Anethum | graveolens | Datura | inoxia | 0.95 | Vaccinum | vitis-idaea | 0.9 | B | 0 | 0 | 1 | Datura | Vaccinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 363 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Apocynum | cannabinum | Apocynum | cannabinum | 2.24 | Coffea | arabica | 1.21 | A | 1 | 1 | 0 | Apocynum | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 488.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Arctostaphylos | uva-ursi | Vaccinium | sect. | 1.09 | Vaccinium | sect. | 1.07 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 371 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Aristolochia | clematitis | Convallaria | majalis | 1.24 | Galanthus | nivalis | 1.2 | B | 0 | 0 | 1 | Convallaria | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 372 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Aristolochia | rotunda | Digitalis | lanata | 1.07 | Aconitum | carmichaelii | 1.04 | B | 0 | 0 | 1 | Digitalis | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 028 | 3; determination by expert | Plant extr. | 2021-07-20 | Artemisia | absinthium | Artemisia | absinthium | 2.27 | Artemisia | vulgaris | 1.6 | A | 1 | 1 | 0 | Artemisia | Artemisia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 073.2 | 3; determination by expert | Plant extr. | 2022-01-14 | Artemisia | absinthum | Vaccinium | sect. | 1.01 | Vaccinium | myrtillus | 1 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 152.1 | 2; label | Plant extr. | 2022-02-04 | Artemisia | dracunculus | Convallaria | majalis | 1.1 | Colchicum | autumnale | 1.06 | B | 0 | 0 | 1 | Convallaria | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 152.2 | 2; label | Plant extr. | 2022-02-25 | Artemisia | dracunculus | Viburnum | lantana | 1.15 | Allium | ursinum | 1.08 | B | 0 | 0 | 1 | Viburnum | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 152.3 | 2; label | Plant extr. | 2022-02-21 | Artemisia | dracunculus | Digitalis | purpurea | 1.36 | Artemisia | vulgaris | 1.27 | B | 0 | 0 | 1 | Digitalis | Artemisia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 071 | 2; Label; determination by expert | Plant extr. | 2021-07-27 | Artemisia | vulgaris | Artemisia | vulgaris | 2.37 | Camellia | sinensis | 1.3 | A | 1 | 1 | 0 | Artemisia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 072.2 | 2; Label; determination by expert | Plant extr.-R. | 2022-01-26 | Artemisia | vulgaris | Digitalis | lanata | 1.15 | Galanthus | nivalis | 1.07 | B | 0 | 0 | 1 | Digitalis | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 367 | 3; botanical collection | Plant extr.-R | 2022-10-07 | Arum | italicum | Arum | italicum | 2.6 | Arum | maculatum | 2.4 | B | 0 | 0 | 1 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 015 | 2; determination by expert | Plant extr. | 2021-07-20 | Arum | maculatum | Arum | maculatum | 2.41 | Arum | maculatum | 2.4 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 368 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Arum | maculatum | Arum | maculatum | 2.55 | Arum | italicum | 2.18 | B | 0 | 0 | 1 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 695 | 3; determination by expert | Plant extr. | 2023-11-15 | Arum | maculatum | Arum | maculatum | 2.41 | Arum | maculatum | 2.38 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 699 | 3; determination by expert | Plant extr. | 2023-11-15 | Arum | maculatum | Arum | maculatum | 2.29 | Arum | maculatum | 2.11 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 723 | 3; determination by expert | Plant extr. | 2023-11-08 | Arum | maculatum | Arum | maculatum | 2.23 | Arum | maculatum | 1.98 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 764 | 3; determination by expert | Plant extr. | 2023-11-09 | Arum | maculatum | Arum | maculatum | 2.48 | Arum | maculatum | 2.23 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 765 | 3; determination by expert | Plant extr. | 2023-11-09 | Arum | maculatum | Arum | maculatum | 2.43 | Arum | maculatum | 2.38 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 770 | 3; determination by expert | Plant extr. | 2023-10-31 | Arum | maculatum | Arum | maculatum | 2.32 | Arum | maculatum | 2.23 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 771 | 3; determination by expert | Plant extr. | 2023-10-31 | Arum | maculatum | Arum | maculatum | 2.35 | Arum | maculatum | 2.31 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 779 | 3; determination by expert | Plant extr. | 2023-11-06 | Arum | maculatum | Arum | maculatum | 2.13 | Arum | maculatum | 2.01 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 785 | 3; determination by expert | Plant extr. | 2023-11-09 | Arum | maculatum | Arum | maculatum | 2.36 | Arum | maculatum | 2.34 | A | 1 | 1 | 0 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 241.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Arum | sp | Arum | maculatum | 1.24 | Arum | maculatum | 1.1 | B | 0 | 0 | 1 | Arum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 056.1 | 2; label | Plant extr. | 2022-02-21 | Asarum | europaeum | Asarum | europaeum | 2.5 | Conium | maculatum | 1.26 | A | 1 | 1 | 0 | Asarum | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 056.2 | 2; label | Plant extr. | 2022-01-26 | Asarum | europaeum | Vaccinium | sect. | 0.98 | Ligustrum | vulgare | 0.93 | B | 0 | 0 | 1 | Vaccinium | Ligustrum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 357.1 | 3; botanical collection | Plant extr.-R | 2022-10-07 | Asimina | triloba | Leucojum | vernum | 1.31 | Viburnum | opulus | 1.25 | B | 0 | 0 | 1 | Leucojum | Viburnum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 011 | 3; botanical collection | Plant extr. | 2022-01-19 | Atropa | belladonna | Vaccinium | angustifolium | 0.77 | Vaccinium | myrtillus | 0.68 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 125.2 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Atropa | belladonna | Atropa | belladonna | 1.93 | Atropa | belladonna | 1.13 | B | 0 | 0 | 1 | Atropa | Atropa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 375.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Atropa | belladonna | Atropa | belladonna | 2.52 | Atropa | belladonna | 1.44 | A | 1 | 1 | 0 | Atropa | Atropa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 375.2 | 3; botanical collection | Plant extr.-R | 2022-10-19 | Atropa | belladonna | Atropa | belladonna | 2.45 | Atropa | belladonna | 1.36 | A | 1 | 1 | 0 | Atropa | Atropa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 448 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Backhousia | citriodora | Backhousia | citriodora | 2.5 | Symphoricarpos | albus | 1.13 | A | 1 | 1 | 0 | Backhousia | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 500 | 2; label; MALDI | Plant extr.-R | 2022-11-23 | Backhousia | citriodora | Backhousia | citriodora | 1.78 | Viburnum | lantana | 1.3 | B | 0 | 0 | 1 | Backhousia | Viburnum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 475.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Berberis | thunbergii | Allium | sativum | 1.38 | Allium | vineale | 1.18 | B | 0 | 0 | 1 | Allium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 026 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Brugmansia | pittieri | Allium | ursinum | 1.22 | Allium | ursinum | 1.14 | B | 0 | 0 | 1 | Allium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 027 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Brugmansia | sanguinea | Ligustrum | vulgare | 1.17 | Jacobaea | maritima | 0.98 | B | 0 | 0 | 1 | Ligustrum | Jacobaea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 120 | 3; botanical collection | Plant extr.-R | 2022-01-27 | Bryonia | cretica | Cinnamomum | camphora | 1.08 | Coffea | arabica | 0.92 | B | 0 | 0 | 1 | Cinnamomum | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 356.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Bryonia | cretica | Allium | schoenoprasum | 1.16 | Leucojum | vernum | 1.1 | B | 0 | 0 | 1 | Allium | Leucojum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 376 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Caltha | palustris | Colchicum | autumnale | 1.48 | Ligustrum | vulgare | 1.28 | B | 0 | 0 | 1 | Colchicum | Ligustrum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 290 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Camellia | sinensis | Camellia | sinensis | 2.3 | Camellia | sinensis | 1.73 | A | 1 | 1 | 0 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 456.1 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Camellia | sinensis | Camellia | sinensis | 2.56 | Camellia | sinensis | 2.36 | A | 1 | 1 | 0 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 456.2 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Camellia | sinensis | Camellia | sinensis | 2.52 | Camellia | sinensis | 1.89 | A | 1 | 1 | 0 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 456.3 | 3; botanical collection | Plant extr.-R | 2022-10-26 | Camellia | sinensis | Camellia | sinensis | 2.5 | Camellia | sinensis | 2.22 | A | 1 | 1 | 0 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 471 | 1; label | Plant extr.-R | 2022-10-27 | Camellia | sinensis | Camellia | sinensis | 1.77 | Camellia | sinensis | 1.34 | B | 0 | 0 | 1 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 472 | 1; label | Plant extr.-R | 2022-11-15 | Camellia | sinensis | Camellia | sinensis | 1.57 | Alpinia | officinarum | 1.39 | B | 0 | 0 | 1 | Camellia | Alpinia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 473 | 1; label | Plant extr.-R | 2022-11-15 | Camellia | sinensis | Camellia | sinensis | 1.43 | Camellia | sinensis | 1.39 | B | 0 | 0 | 1 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 499 | 1; label | Plant extr.-R | 2022-11-23 | Camellia | sinensis | Camellia | sinensis | 1.71 | Camellia | sinensis | 1.47 | B | 0 | 0 | 1 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 552 | 1; label | Plant extr.-R | 2022-12-02 | Camellia | sinensis | Camellia | sinensis | 1.42 | Camellia | sinensis | 1.37 | B | 0 | 0 | 1 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 567 | 1; label | Plant extr.-AceRL42S90 | 2022-12-12 | Camellia | sinensis | Camellia | sinensis | 1.6 | Camellia | sinensis | 1.46 | B | 0 | 0 | 1 | Camellia | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 364 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Capsicum | chinense | Lycium | chinense | 1.27 | Solanum | dulcamara | 1.11 | B | 0 | 0 | 1 | Lycium | Solanum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 355 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Chamaecyparis | lawsoniana | Chamaecyparis | lawsoniana | 2.19 | Ligustrum | vulgare | 1.33 | A | 1 | 1 | 0 | Chamaecyparis | Ligustrum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 115.2 | 3; botanical collection | Plant extr.-R | 2022-01-27 | Chelidonium | majus | Polygonatum | multiflorum | 0.94 | Allium | vineale | 0.93 | B | 0 | 0 | 1 | Polygonatum | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 046 | 3; botanical collection | Plant extr.-R | 2022-01-28 | Cicuta | virosa | Polygonatum | multiflorum | 1.06 | Symphoricarpos | albus | 0.91 | B | 0 | 0 | 1 | Polygonatum | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 412 | 3; botanical collection | Plant extr.-R | 2022-11-04 | Cinnamomum | camphora | Cinnamomum | camphora | 2.21 | Clinopodium | nepeta | 1.17 | A | 1 | 1 | 0 | Cinnamomum | Clinopodium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 410 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Cinnamomum | pittosporoides | Cinnamomum | pittosporoides | 2.2 | Cinnamomum | verum | 2.12 | B | 0 | 0 | 1 | Cinnamomum | Cinnamomum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 408 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Cinnamomum | verum | Cinnamomum | verum | 2.45 | Cinnamomum | pittosporoides | 2.41 | B | 0 | 0 | 1 | Cinnamomum | Cinnamomum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 138 | 2; expert opinion | Plant extr.-R | 2022-02-07 | Cistus | sp | Allium | schoenoprasum | 1.06 | Asarum | europaeum | 0.97 | B | 0 | 0 | 1 | Allium | Asarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 263 | 3; botanical collection | Plant extr. | 2022-09-16 | Citrus | cavaleriei | Petroselinum | crispum | 1.1 | Petroselinum | crispum | 1.08 | B | 0 | 0 | 1 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 078 | 2; label; determination by expert | Plant extr. | 2021-07-30 | Clinopodium | nepeta | Clinopodium | nepeta | 2.11 | Convallaria | majalis | 1.41 | A | 1 | 1 | 0 | Clinopodium | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 251 | 3; botanical collection | Plant extr. | 2022-09-15 | Coffea | arabica | Coffea | arabica | 2.5 | Coffea | arabica | 2.43 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 254 | 3; botanical collection | Plant extr. | 2022-09-15 | Coffea | arabica | Coffea | arabica | 2.6 | Coffea | arabica | 2.32 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 256.1 | 3; botanical collection | Plant extr. | 2022-09-14 | Coffea | arabica | Coffea | arabica | 2.21 | Coffea | arabica | 2.06 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 256.2 | 3; botanical collection | Plant extr. | 2022-10-04 | Coffea | arabica | Coffea | arabica | 2.57 | Coffea | arabica | 2.5 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 260.1 | 3; botanical collection | Plant extr. | 2022-09-14 | Coffea | arabica | Coffea | arabica | 2.48 | Coffea | arabica | 2.03 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 260.2 | 3; botanical collection | Plant extr. | 2022-09-29 | Coffea | arabica | Coffea | arabica | 2.23 | Coffea | arabica | 2.22 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 328.1 | 3; botanical collection | Plant extr. | 2022-09-27 | Coffea | arabica | Coffea | arabica | 2.22 | Coffea | arabica | 2.22 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 440 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Coffea | arabica | Coffea | arabica | 2.54 | Coffea | arabica | 2.52 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 453 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Coffea | arabica | Coffea | arabica | 2.66 | Coffea | arabica | 2.5 | A | 1 | 1 | 0 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 351 | 3; botanical collection | Plant extr. | 2022-09-28 | Coffea | canephora | Coffea | canephora | 2.35 | Coffea | arabica | 2.31 | B | 0 | 0 | 1 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 018 | 3; determination by expert | Plant extr. | 2021-07-20 | Colchicum | autumnale | Colchicum | autumnale | 2.54 | Colchicum | autumnale | 2.26 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 110.1 | 3; PCR | Plant extr. | 2023-11-17 | Colchicum | autumnale | Colchicum | autumnale | 2.1 | Colchicum | autumnale | 2.01 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 110.3 | 3; PCR | Plant extr. | 2023-11-30 | Colchicum | autumnale | Colchicum | autumnale | 2.34 | Colchicum | autumnale | 2.19 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 226.4 | 3; PCR | Plant extr. | 2023-11-09 | Colchicum | autumnale | Colchicum | autumnale | 2.21 | Colchicum | autumnale | 2.06 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 226.5 | 3; PCR | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.36 | Colchicum | autumnale | 2.07 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 227.1 | 3; PCR | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.49 | Colchicum | autumnale | 2.36 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 227 | 3; PCR | Plant extr. | 2023-11-09 | Colchicum | autumnale | Colchicum | autumnale | 2.2 | Colchicum | autumnale | 1.85 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 244.1 | 3; determination by expert | Plant extr. | 2023-11-30 | Colchicum | autumnale | Colchicum | autumnale | 2.16 | Colchicum | autumnale | 1.96 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 244 | 3; determination by expert | Plant extr. | 2022-09-13 | Colchicum | autumnale | Colchicum | autumnale | 2.35 | Colchicum | autumnale | 2.27 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 701.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.21 | Colchicum | autumnale | 2.03 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 701 | 3; determination by expert | Plant extr. | 2023-11-17 | Colchicum | autumnale | Colchicum | autumnale | 2.21 | Colchicum | autumnale | 2.05 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 724.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.18 | Colchicum | autumnale | 2.17 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 724 | 3; determination by expert | Plant extr. | 2023-11-08 | Colchicum | autumnale | Colchicum | autumnale | 2.16 | Colchicum | autumnale | 2.02 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 768.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.36 | Colchicum | autumnale | 2.23 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 768 | 3; determination by expert | Plant extr. | 2023-11-09 | Colchicum | autumnale | Colchicum | autumnale | 2.33 | Colchicum | autumnale | 2.18 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 769.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.36 | Colchicum | autumnale | 2.26 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 769 | 3; determination by expert | Plant extr. | 2023-10-31 | Colchicum | autumnale | Colchicum | autumnale | 2.32 | Colchicum | autumnale | 2.25 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 777.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.17 | Colchicum | autumnale | 2.05 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 777 | 3; determination by expert | Plant extr. | 2023-11-06 | Colchicum | autumnale | Colchicum | autumnale | 2.25 | Colchicum | autumnale | 2.1 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 778.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum | autumnale | Colchicum | autumnale | 2.43 | Colchicum | autumnale | 2.41 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 778 | 3; determination by expert | Plant extr. | 2023-11-06 | Colchicum | autumnale | Colchicum | autumnale | 2.35 | Colchicum | autumnale | 2.05 | A | 1 | 1 | 0 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 041.1 | 3; determination by expert | Plant extr. | 2022-02-21 | Conium | maculatum | Conium | maculatum | 2.58 | Aethusa | cynapium | 2.03 | C | 0 | 0 | 1 | Conium | Aethusa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 041.2 | 3; determination by expert | Plant extr. | 2022-01-27 | Conium | maculatum | Conium | maculatum | 2.44 | Aethusa | cynapium | 2.3 | C | 0 | 0 | 1 | Conium | Aethusa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 021.1 | 3; determination by expert | Plant extr. | 2021-07-20 | Convallaria | majalis | Convallaria | majalis | 2.18 | Convallaria | majalis | 2.14 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 059 | 3; label; sensory analysis | Plant extr. | 2022-01-21 | Convallaria | majalis | Convallaria | majalis | 2.21 | Convallaria | majalis | 2.03 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 063 | 3; label; sensory analysis | Plant extr. | 2023-11-16 | Convallaria | majalis | Convallaria | majalis | 2.36 | Convallaria | majalis | 2.33 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 128 | 3; botanical collection | Plant extr. | 2023-11-27 | Convallaria | majalis | Convallaria | majalis | 2.16 | Convallaria | majalis | 2.05 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 197 | 3; determination by expert | Plant extr. | 2022-09-09 | Convallaria | majalis | Convallaria | majalis | 2.27 | Convallaria | majalis | 1.99 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 220 | 3; determination by expert | Plant extr. | 2022-09-09 | Convallaria | majalis | Convallaria | majalis | 2.32 | Convallaria | majalis | 2.18 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 242.1 | 3; determination by expert | Plant extr. | 2022-09-09 | Convallaria | majalis | Convallaria | majalis | 2.2 | Convallaria | majalis | 2.2 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 559.1 | 3; determination by expert | Plant extr. | 2023-11-17 | Convallaria | majalis | Convallaria | majalis | 2.03 | Convallaria | majalis | 1.99 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 725.1 | 3; label, sensory analysis | Plant extr. | 2023-11-08 | Convallaria | majalis | Convallaria | majalis | 2.32 | Convallaria | majalis | 2.27 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 752.1 | 3; determination by expert | Plant extr. | 2023-11-17 | Convallaria | majalis | Convallaria | majalis | 2.38 | Convallaria | majalis | 2.23 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 756.1 | 3; determination by expert | Plant extr. | 2023-11-17 | Convallaria | majalis | Convallaria | majalis | 2.32 | Convallaria | majalis | 2.1 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 759.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria | majalis | Convallaria | majalis | 2.36 | Convallaria | majalis | 2.26 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 760.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria | majalis | Convallaria | majalis | 2.1 | Convallaria | majalis | 2 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 761.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria | majalis | Convallaria | majalis | 2.01 | Convallaria | majalis | 2.01 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 762.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria | majalis | Convallaria | majalis | 2.2 | Convallaria | majalis | 2.19 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 774 | 3; determination by expert | Plant extr. | 2023-10-31 | Convallaria | majalis | Convallaria | majalis | 2.42 | Convallaria | majalis | 2.28 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 775 | 3; determination by expert | Plant extr. | 2023-10-31 | Convallaria | majalis | Convallaria | majalis | 2.29 | Convallaria | majalis | 2.11 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 776 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria | majalis | Convallaria | majalis | 2.04 | Convallaria | majalis | 1.99 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 781 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria | majalis | Convallaria | majalis | 2.25 | Convallaria | majalis | 2.2 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 782 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria | majalis | Convallaria | majalis | 2.32 | Convallaria | majalis | 2.21 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 783 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria | majalis | Convallaria | majalis | 2.32 | Convallaria | majalis | 2.26 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 784 | 3; determination by expert | Plant extr. | 2023-11-09 | Convallaria | majalis | Convallaria | majalis | 2.35 | Convallaria | majalis | 2.25 | A | 1 | 1 | 0 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 153.1 | 3; botanical collection | Plant extr. | 2022-02-04 | Coriandrum | sativum | Anethum | graveolens | 1.27 | Arum | maculatum | 1.24 | B | 0 | 0 | 1 | Anethum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 153.2 | 3; botanical collection | Plant extr. | 2022-02-04 | Coriandrum | sativum | Mansoa | alliacea | 1.2 | Jacobaea | maritima | 1.11 | B | 0 | 0 | 1 | Mansoa | Jacobaea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 153.3 | 3; botanical collection | Plant extr. | 2022-02-25 | Coriandrum | sativum | Galanthus | nivalis | 1.28 | Colchicum | autumnale | 1.28 | B | 0 | 0 | 1 | Galanthus | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 153.4 | 3; botanical collection | Plant extr. | 2022-02-21 | Coriandrum | sativum | Convallaria | majalis | 1.37 | Camellia | sinensis | 1.27 | B | 0 | 0 | 1 | Convallaria | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 154 | 2; label | Plant extr. | 2022-01-21 | Coriandrum | sativum | Allium | tuberosum | 1.03 | Symphoricarpos | albus | 0.97 | B | 0 | 0 | 1 | Allium | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 422 | 3; botanical collection | Plant extr.-R | 2022-11-02 | Costus | malortieanus | Symphoricarpos | albus | 1.41 | Allium | schoenoprasum | 1.35 | B | 0 | 0 | 1 | Symphoricarpos | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 382.1 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Cotoneaster | ballatus | Chamaecyparis | lawsoniana | 1.21 | Symphoricarpos | albus | 1.2 | B | 0 | 0 | 1 | Chamaecyparis | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 397.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Crataegus | monogyna | Cinnamomum | camphora | 1.31 | Symphoricarpos | albus | 1.22 | B | 0 | 0 | 1 | Cinnamomum | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 397.2 | 2; expert opinion | Plant extr.-R-A | 2022-10-18 | Crataegus | monogyna | Vaccinium | sect. | 1.06 | Vaccinium | sect. | 1.03 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 204.1 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Crataegus | sp | Lycium | europaeum | 0.94 | Allium | porrum | 0.92 | B | 0 | 0 | 1 | Lycium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 212 | 2; determination by expert | Plant extr. | 2022-09-09 | Cunninghamia | konishii | Chamaecyparis | lawsoniana | 1.35 | Alpinia | officinarum | 1.24 | B | 0 | 0 | 1 | Chamaecyparis | Alpinia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 432 | 3; botanical collection | Plant extr.-R | 2022-11-04 | Curcuma | longa | Curcuma | longa | 2.31 | Camellia | sinensis | 1.27 | A | 1 | 1 | 0 | Curcuma | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 394.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Datura | innoxia | Datura | inoxia | 2.26 | Lycium | chinense | 1.05 | A | 0 | 0 | 1 | Datura | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 034 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Datura | stramonium | Datura | stramonium | 1.72 | Datura | inoxia | 1.33 | B | 0 | 0 | 1 | Datura | Datura | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 378 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Datura | stramonium | Datura | stramonium | 2.04 | Allium | ursinum | 1.12 | A | 1 | 1 | 0 | Datura | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 393.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Datura | stramonium | Datura | stramonium | 2.23 | Datura | inoxia | 1.4 | A | 1 | 1 | 0 | Datura | Datura | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 377 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Digitalis | grandiflora | Jacobaea | maritima | 1.4 | Symphoricarpos | albus | 1.26 | B | 0 | 0 | 1 | Jacobaea | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 380 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Digitalis | lanata | Digitalis | lanata | 2.52 | Digitalis | purpurea | 1.48 | A | 1 | 1 | 0 | Digitalis | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 039.1 | 3; determination by expert | Plant extr. | 2022-01-20 | Digitalis | purpurea | Digitalis | purpurea | 2.43 | Digitalis | purpurea | 1.31 | A | 1 | 1 | 0 | Digitalis | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 039.2 | 3; determination by expert | Plant extr. | 2022-01-20 | Digitalis | purpurea | Digitalis | purpurea | 2.25 | Conium | maculatum | 1.26 | A | 1 | 1 | 0 | Digitalis | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 054.1 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Digitalis | purpurea | Digitalis | purpurea | 1.84 | Digitalis | lanata | 1.24 | B | 0 | 0 | 1 | Digitalis | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 054.2 | 2; determination by expert | Plant extr. | 2022-01-14 | Digitalis | purpurea | Viscum | album | 0.73 | Vaccinium | sect. | 0.68 | B | 0 | 0 | 1 | Viscum | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 060 | 2; label | Plant extr. | 2022-01-14 | Digitalis | purpurea | Digitalis | purpurea | 2.1 | Digitalis | lanata | 1.24 | A | 1 | 1 | 0 | Digitalis | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 061.1 | 2; label | Plant extr. | 2022-01-14 | Digitalis | purpurea | Digitalis | purpurea | 1.95 | Digitalis | purpurea | 1.59 | B | 0 | 0 | 1 | Digitalis | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 061.2 | 2; label | Plant extr. | 2022-01-27 | Digitalis | purpurea | Chamaecyparis | lawsoniana | 1.04 | Vaccinium | sect. | 0.81 | B | 0 | 0 | 1 | Chamaecyparis | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 062 | 2; label | Plant extr. | 2022-01-14 | Digitalis | purpurea | Digitalis | purpurea | 1.92 | Digitalis | lanata | 1.59 | B | 0 | 0 | 1 | Digitalis | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 217 | 2; determination by expert | Plant extr. | 2022-09-09 | Digitalis | purpurea | Digitalis | purpurea | 1.49 | Allium | sativum | 1.25 | B | 0 | 0 | 1 | Digitalis | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 103 | 2; expert opinion | Plant extr.-R | 2022-10-10 | Dryopteris | filix-mas | Arum | maculatum | 1.19 | Viburnum | opulus | 1.03 | B | 0 | 0 | 1 | Arum | Viburnum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 209 | 2; determination by expert | Plant extr. | 2022-09-09 | Echium | candicans | Ligustrum | vulgare | 1.28 | Aconitum | carmichaelii | 1.26 | B | 0 | 0 | 1 | Ligustrum | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 208 | 2; determination by expert | Plant extr. | 2022-09-09 | Echium | sp | Prunus | lusitanica | 1.25 | Convallaria | majalis | 1.15 | B | 0 | 0 | 1 | Prunus | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 210 | 2; determination by expert | Plant extr. | 2022-09-09 | Echium | sp | Convallaria | majalis | 1.24 | Anethum | graveolens | 1.15 | B | 0 | 0 | 1 | Convallaria | Anethum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 022 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Erythroxylum | coca | Allium | ursinum | 1.19 | Galium | odoratum | 1.08 | B | 0 | 0 | 1 | Allium | Galium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 111 | 2; expert opinion | Plant extr. | 2022-01-28 | Euonymus | europaeus | Vaccinium | fuscatum | 0.76 | Atropa | belladonna | 0.74 | B | 0 | 0 | 1 | Vaccinium | Atropa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 113.2 | 2; expert opinion | Plant extr. | 2022-01-21 | Euonymus | europaeus | Cinnamomum | camphora | 0.96 | Colchicum | autumnale | 0.88 | B | 0 | 0 | 1 | Cinnamomum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 459.1 | 2; expert opinion | Plant extr.-R | 2022-10-28 | Euonymus | europaeus | Mansoa | alliacea | 1.3 | Digitalis | lanata | 1.14 | B | 0 | 0 | 1 | Mansoa | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 481.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Euonymus | europaeus | Arum | maculatum | 1.13 | Mansoa | alliacea | 1.07 | B | 0 | 0 | 1 | Arum | Mansoa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 477.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Euonymus | fortunei | Convallaria | majalis | 1.19 | Allium | tuberosum | 1.17 | B | 0 | 0 | 1 | Convallaria | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 166.1 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus | nivalis | Galanthus | nivalis | 2.28 | Galanthus | nivalis | 1.99 | A | 1 | 1 | 0 | Galanthus | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 690 | 3; determination by expert | Plant extr. | 2023-11-27 | Galanthus | nivalis | Galanthus | nivalis | 2.52 | Galanthus | nivalis | 2.13 | A | 1 | 1 | 0 | Galanthus | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 184.1 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus | sp | Galanthus | nivalis | 1.87 | Galanthus | nivalis | 1.82 | B | 0 | 0 | 1 | Galanthus | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 236.1 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus | sp | Galanthus | nivalis | 1.67 | Galanthus | nivalis | 1.53 | B | 0 | 0 | 1 | Galanthus | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 236.2 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus | sp | Galanthus | nivalis | 1.19 | Leucojum | vernum | 1.02 | B | 0 | 0 | 1 | Galanthus | Leucojum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 086 | 1; determination by expert | Plant extr. | 2021-07-27 | Galium | mollugo | Galium | mollugo | 2.23 | Galium | odoratum | 1.37 | A | 1 | 1 | 0 | Galium | Galium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 085 | 2; determination by expert | Plant extr. | 2021-07-27 | Galium | odoratum | Galium | odoratum | 2.12 | Jacobaea | maritima | 1.2 | A | 1 | 1 | 0 | Galium | Jacobaea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 620 | 1; label | Plant extr.-AceR | 2023-01-10 | Ganoderma | lucidum | Colchicum | autumnale | 0.65 | Coffea | arabica | 0.61 | B | 0 | 0 | 1 | Colchicum | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 087.2 | 2; expert opinion | Plant extr.-R. | 2022-01-26 | Gentiana | lutea | Galanthus | nivalis | 1.03 | Allium | schoenoprasum | 0.99 | B | 0 | 0 | 1 | Galanthus | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 214 | 2; determination by expert | Plant extr. | 2022-09-09 | Gunnera | sp | Viburnum | lantana | 1.13 | Allium | ursinum | 1.06 | B | 0 | 0 | 1 | Viburnum | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 391.1 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Heliotropium | arborescens | Vaccinium | sect. | 0.89 | Vaccinium | sect. | 0.88 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 066.2 | 2; label | Plant extr.-R. | 2022-01-26 | Helleborus | niger | Ligustrum | vulgare | 1.12 | Zingiber | officinale | 1.01 | B | 0 | 0 | 1 | Ligustrum | Zingiber | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 067.1 | 2; label | Plant extr. | 2022-01-14 | Helleborus | niger | Ocimum | basilicum | 1.05 | Apocynum | cannabinum | 1.03 | B | 0 | 0 | 1 | Ocimum | Apocynum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 068.1 | 2; label | OS extr.-Pl | 2018-01-23 | Helleborus | niger | Digitalis | purpurea | 1.29 | Aconitum | napellus | 1.24 | B | 0 | 0 | 1 | Digitalis | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 118 | 3; botanical collection | OS extr.-Pl | 2018-01-23 | Helleborus | niger | Digitalis | purpurea | 1.35 | Digitalis | purpurea | 1.24 | B | 0 | 0 | 1 | Digitalis | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 119 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Humulus | lupulus | Camellia | sinensis | 1.26 | Curcuma | longa | 1.23 | B | 0 | 0 | 1 | Camellia | Curcuma | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 043 | 3; botanical collection | Plant extr. | 2022-01-19 | Hyoscyamus | niger | Coffea | arabica | 1.01 | Arum | italicum | 0.95 | B | 0 | 0 | 1 | Coffea | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 206 | 1; determination by expert | Plant extr. | 2022-09-09 | Jacobaea | maritima | Jacobaea | maritima | 1.89 | Jacobaea | vulgaris | 1.35 | B | 0 | 0 | 1 | Jacobaea | Jacobaea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 245.1 | 2; determination by expert | Plant extr. | 2022-09-13 | Jacobaea | maritima | Jacobaea | maritima | 2.31 | Jacobaea | vulgaris | 1.65 | A | 1 | 1 | 0 | Jacobaea | Jacobaea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | PL 0047 | 2; label; sensorical analysis | OS extr.-Pl | 2018-01-23 | Juglans | regia | Juglans | regia | 2.74 | Juglans | regia | 2.34 | A | 1 | 1 | 0 | Juglans | Juglans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | PL 0048 | 2; label; sensorical analysis | OS extr.-Pl | 2018-01-23 | Juglans | regia | Juglans | regia | 2.52 | Juglans | regia | 2.37 | A | 1 | 1 | 0 | Juglans | Juglans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 237 | 2; determination by expert | Plant extr. | 2022-09-09 | Juglans | regia | Camellia | sinensis | 1.22 | Ligustrum | vulgare | 1.2 | B | 0 | 0 | 1 | Camellia | Ligustrum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 042.1 | 3; botanical collection | Plant extr.-R | 2022-02-21 | Juniperus | sabina | Chamaecyparis | lawsoniana | 1.29 | Camellia | sinensis | 1.1 | B | 0 | 0 | 1 | Chamaecyparis | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 042.2 | 3; botanical collection | Plant extr. | 2022-01-18 | Juniperus | sabina | Viscum | album | 1.3 | Convallaria | majalis | 0.9 | B | 0 | 0 | 1 | Viscum | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 261 | 3; botanical collection | Plant extr. | 2022-09-16 | Justicia | adhotoda | Curcuma | longa | 1.19 | Galanthus | nivalis | 1.06 | B | 0 | 0 | 1 | Curcuma | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 047 | 3; botanical collection | Plant extr. | 2022-01-19 | Laburnum | anagyroides | Vaccinium | sect. | 0.94 | Vaccinium | sect. | 0.87 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 231 | 2; determination by expert | Plant extr. | 2022-09-06 | Laburnum | sp | Zingiber | officinale | 0.86 | Allium | ursinum | 0.86 | B | 0 | 0 | 1 | Zingiber | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 145 | 2; determination by expert | Plant extr.-R | 2022-02-03 | Laurus | nobilis | Artemisia | vulgaris | 1.21 | Cinnamomum | camphora | 1.21 | B | 0 | 0 | 1 | Artemisia | Cinnamomum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 269 | 3; botanical collection | Plant extr. | 2022-09-16 | Laurus | nobilis | Jacobaea | maritima | 1.25 | Zingiber | officinale | 1.22 | B | 0 | 0 | 1 | Jacobaea | Zingiber | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 312 | 3; botanical collection | Plant extr. | 2022-09-15 | Lemna | minor | Allium | schoenoprasum | 1.19 | Colchicum | autumnale | 1.16 | B | 0 | 0 | 1 | Allium | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 438 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Leptospermum | citratum | Leptospermum | citratum | 2.48 | Backhousia | citriodora | 1.2 | A | 1 | 1 | 0 | Leptospermum | Backhousia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 698 | 3; determination by expert | Plant extr. | 2023-11-27 | Leucojum | vernum | Leucojum | vernum | 2.52 | Lycium | chinense | 1.17 | A | 1 | 1 | 0 | Leucojum | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 404.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Ligustrum | vulgare | Ligustrum | vulgare | 2.05 | Arum | maculatum | 1.41 | A | 1 | 1 | 0 | Ligustrum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 404.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Ligustrum | vulgare | Ligustrum | vulgare | 1.57 | Solanum | betaceum | 1.17 | B | 0 | 0 | 1 | Ligustrum | Solanum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 460.1 | 2; expert opinion | Plant extr.-R | 2022-10-28 | Ligustrum | vulgare | Ligustrum | vulgare | 1.67 | Digitalis | lanata | 1.09 | B | 0 | 0 | 1 | Ligustrum | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 491.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Ligustrum | vulgare | Ligustrum | vulgare | 1.73 | Petroselinum | crispum | 1.23 | B | 0 | 0 | 1 | Ligustrum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 284.2 | 2; botanical collection | Plant extr. | 2022-10-04 | Lonicera | sp | Vaccinium | myrtillus | 1.07 | Vaccinium | vitis-idaea | 0.99 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 023 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Lophophora | williamsi | Allium | ursinum | 1.09 | Convallaria | majalis | 1 | B | 0 | 0 | 1 | Allium | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | PL 0001 | 2; label; sensorical analysis | Osextr-Pl | 2017-11-03 | Lupinus | sp | Vaccinium | sect. | 1.08 | Lycium | europaeum | 0.85 | B | 0 | 0 | 1 | Vaccinium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 443.1 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | barbarum | Lycium | barbarum | 2.44 | Symphoricarpos | albus | 1.22 | A | 1 | 1 | 0 | Lycium | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 443.2 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | barbarum | Lycium | barbarum | 2.61 | Lycium | chinense | 2.25 | B | 0 | 0 | 1 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 443.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | barbarum | Lycium | barbarum | 2.34 | Lycium | europaeum | 2.11 | B | 0 | 0 | 1 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 447.1 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | chinense | Lycium | chinense | 2.35 | Lycium | europaeum | 2.21 | B | 0 | 0 | 1 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 447.2 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | chinense | Lycium | chinense | 2.47 | Lycium | europaeum | 2.06 | B | 0 | 0 | 1 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 447.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | chinense | Lycium | chinense | 2.48 | Lycium | chinense | 2.25 | A | 1 | 1 | 0 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 449.1 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | europaeum | Lycium | europaeum | 2.45 | Lycium | chinense | 2.29 | B | 0 | 0 | 1 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 449.2 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | europaeum | Lycium | europaeum | 2.6 | Lycium | barbarum | 1.98 | A | 1 | 1 | 0 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 449.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | europaeum | Lycium | europaeum | 2.13 | Lycium | chinense | 1.41 | A | 1 | 1 | 0 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 474.2 | 1; label | Plant extr.-R | 2022-11-21 | Lycium | sp | Lycium | barbarum | 1.64 | Ricinus | communis | 1.06 | B | 0 | 0 | 1 | Lycium | Ricinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 496.1 | 1; label | Plant extr.-R | 2022-11-23 | Lycium | sp | Vaccinium | sect. | 0.85 | Vaccinium | sect. | 0.77 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 496.2 | 1; label | Plant extr.-R | 2022-11-23 | Lycium | sp | Lycium | barbarum | 1.45 | Lycium | europaeum | 1.33 | B | 0 | 0 | 1 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 390.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Malus | sylvestris | Vaccinium | vitis-idaea | 0.95 | Allium | vineale | 0.92 | B | 0 | 0 | 1 | Vaccinium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 124 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Mandragora | officinarum | Vaccinium | myrtillus | 1.06 | Vaccinium | myrtillus | 1.05 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 277 | 3; botanical collection | Plant extr. | 2023-11-27 | Mansoa | alliacea | Mansoa | alliacea | 2.22 | Camellia | sinensis | 1.25 | A | 1 | 1 | 0 | Mansoa | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 058.1 | 2; determination by expert | Plant extr. | 2022-02-21 | Mentha | pulegium | Digitalis | purpurea | 1.48 | Clinopodium | nepeta | 1.37 | B | 0 | 0 | 1 | Digitalis | Clinopodium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 058.2 | 2; determination by expert | Plant extr. | 2022-01-18 | Mentha | pulegium | Vaccinium | corymbosum | 0.95 | Lycium | barbarum | 0.86 | B | 0 | 0 | 1 | Vaccinium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 213 | 2; determination by expert | Plant extr. | 2022-09-09 | Metasequoia | glyptostroboides | Chamaecyparis | lawsoniana | 1.17 | Artemisia | absinthium | 1.09 | B | 0 | 0 | 1 | Chamaecyparis | Artemisia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 257.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Microcitrus | australasia | Allium | sativum | 1.04 | Galanthus | nivalis | 1.04 | B | 0 | 0 | 1 | Allium | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 455 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Moringa | stenopetala | Moringa | stenopetala | 2.21 | Allium | ursinum | 1.06 | A | 1 | 1 | 0 | Moringa | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 012 | 2; determination by expert | Plant extr. | 2021-07-15 | Nerium | oleander | Nerium | oleander | 2.45 | Digitalis | purpurea | 1.1 | A | 1 | 1 | 0 | Nerium | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 024 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Nerium | oleander | Nerium | oleander | 1.96 | Solanum | dulcamara | 1.19 | B | 0 | 0 | 1 | Nerium | Solanum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 360 | 3; botanical collection | Plant extr.-R | 2022-10-07 | Nicotiana | sylvestris | Vaccinium | vitis-idaea | 1.19 | Solanum | betaceum | 1.15 | B | 0 | 0 | 1 | Vaccinium | Solanum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 216 | 2; determination by expert | Plant extr. | 2022-09-09 | Nothofagus | allessandrii | Clinopodium | nepeta | 1.11 | Vaccinium | sect. | 1.05 | B | 0 | 0 | 1 | Clinopodium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 131.1 | 2; label | Plant extr. | 2022-02-01 | Ocimum | basilicum | Ocimum | basilicum | 2.65 | Anethum | graveolens | 1.1 | A | 1 | 1 | 0 | Ocimum | Anethum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 131.2 | 2; label | Plant extr. | 2022-02-03 | Ocimum | basilicum | Ocimum | basilicum | 2.11 | Aconitum | carmichaelii | 1.45 | A | 1 | 1 | 0 | Ocimum | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 131.3 | 2; label | Plant extr. | 2022-02-21 | Ocimum | basilicum | Ocimum | basilicum | 2.15 | Solanum | dulcamara | 1.35 | A | 1 | 1 | 0 | Ocimum | Solanum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 134 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Ocimum | basilicum | Conium | maculatum | 1.14 | Allium | ursinum | 1.13 | B | 0 | 0 | 1 | Conium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 139 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Olea | europaea | Viburnum | lantana | 0.91 | Vaccinium | sect. | 0.58 | B | 0 | 0 | 1 | Viburnum | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 135 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Origanum | vulgare | Digitalis | purpurea | 1.04 | Allium | tuberosum | 1.01 | B | 0 | 0 | 1 | Digitalis | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 147.1 | 2; determination by expert | Plant extr. | 2022-02-03 | Origanum | vulgare | Apocynum | cannabinum | 1.23 | Mansoa | alliacea | 1.16 | B | 0 | 0 | 1 | Apocynum | Mansoa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 147.2 | 2; determination by expert | Plant extr. | 2022-02-03 | Origanum | vulgare | Ocimum | basilicum | 1.3 | Digitalis | purpurea | 1.29 | B | 0 | 0 | 1 | Ocimum | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 147.3 | 2; determination by expert | Plant extr. | 2022-02-21 | Origanum | vulgare | Camellia | sinensis | 1.16 | Digitalis | purpurea | 1.16 | B | 0 | 0 | 1 | Camellia | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 148 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Origanum | vulgare | Spinacia | oleracea | 0.92 | Vaccinium | myrtillus | 0.92 | B | 0 | 0 | 1 | Spinacia | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 036.3 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Papaver | somniferum | Allium | schoenoprasum | 1.04 | Leucojum | vernum | 0.94 | B | 0 | 0 | 1 | Allium | Leucojum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 270.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Passiflora | sp | Curcuma | longa | 1.16 | Jacobaea | vulgaris | 1.1 | B | 0 | 0 | 1 | Curcuma | Jacobaea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 255 | 3; botanical collection | Plant extr. | 2022-09-16 | Pelargonium | sp | Digitalis | lanata | 1.21 | Conium | maculatum | 1.01 | B | 0 | 0 | 1 | Digitalis | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 020.1 | 2; determination by expert | Plant extr. | 2022-01-24 | Petasites | hybridus | Colchicum | autumnale | 1.19 | Colchicum | autumnale | 1.17 | B | 0 | 0 | 1 | Colchicum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 020.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Petasites | hybridus | Conium | maculatum | 1.08 | Vaccinium | sect. | 0.93 | B | 0 | 0 | 1 | Conium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 081 | 3; label; determination by expert | Plant extr. | 2021-07-27 | Petroselinum | crispum | Petroselinum | crispum | 2.32 | Petroselinum | crispum | 1.82 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 130.1 | 2; label | Plant extr. | 2022-02-01 | Petroselinum | crispum | Petroselinum | crispum | 2.62 | Petroselinum | crispum | 2.07 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 130.2 | 2; label | Plant extr. | 2022-02-01 | Petroselinum | crispum | Petroselinum | crispum | 2.5 | Petroselinum | crispum | 2.09 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 130.3 | 2; label | Plant extr. | 2022-02-21 | Petroselinum | crispum | Petroselinum | crispum | 2 | Petroselinum | crispum | 1.81 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 130.4 | 2; label | Plant extr. | 2022-02-21 | Petroselinum | crispum | Petroselinum | crispum | 2.17 | Petroselinum | crispum | 2.03 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 130.5 | 2; label | Plant extr. | 2022-02-21 | Petroselinum | crispum | Petroselinum | crispum | 2.16 | Petroselinum | crispum | 2.12 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 137 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Petroselinum | crispum | Petroselinum | crispum | 1.8 | Petroselinum | crispum | 1.77 | B | 0 | 0 | 1 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 150.1 | 2; label | Plant extr. | 2022-02-04 | Petroselinum | crispum | Petroselinum | crispum | 2.28 | Petroselinum | crispum | 1.76 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 150.2 | 2; label | Plant extr. | 2022-02-04 | Petroselinum | crispum | Petroselinum | crispum | 2.01 | Petroselinum | crispum | 1.74 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 151.1 | 2; label | Plant extr. | 2022-02-04 | Petroselinum | crispum | Petroselinum | crispum | 2.09 | Petroselinum | crispum | 1.62 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 151.2 | 2; label | Plant extr. | 2022-02-04 | Petroselinum | crispum | Petroselinum | crispum | 2.11 | Petroselinum | crispum | 1.81 | A | 1 | 1 | 0 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 162 | 2; label | Plant extr.-R | 2022-02-23 | Petroselinum | crispum | Jacobaea | vulgaris | 1.12 | Curcuma | longa | 1.08 | B | 0 | 0 | 1 | Jacobaea | Curcuma | 0 | 0 | 0 | 0
|  |  | 1 | 0 | PL 0130 | 3; label; sensorical analysis | Osextr-Pl | 2020-07-17 | Phaseolus | vulgaris | Vaccinium | vitis-idaea | 0.79 | Chamaecyparis | lawsoniana | 0.7 | B | 0 | 0 | 1 | Vaccinium | Chamaecyparis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 370 | 3; botanical collection | Plant extr.-R | 2022-01-20 | Pinellia | ternata | Leucojum | vernum | 0.99 | Lycium | europaeum | 0.97 | B | 0 | 0 | 1 | Leucojum | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 215 | 2; determination by expert | Plant extr. | 2022-09-09 | Pinus | montezumae | Arum | italicum | 1.03 | Taxus | baccata | 1.02 | B | 0 | 0 | 1 | Arum | Taxus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 418 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper | auritum | Allium | schoenoprasum | 1.17 | Datura | stramonium | 1.1 | B | 0 | 0 | 1 | Allium | Datura | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 414 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper | betle | Coffea | arabica | 1.21 | Coffea | arabica | 1.2 | B | 0 | 0 | 1 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 411 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper | cubeba | Symphoricarpos | albus | 1.24 | Camellia | sinensis | 1.22 | B | 0 | 0 | 1 | Symphoricarpos | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 416 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper | nigrum | Convallaria | majalis | 1.27 | Convallaria | majalis | 1.17 | B | 0 | 0 | 1 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 463 | 1; label | Plant extr.-R | 2022-10-28 | Piper | sarmentosum | Apocynum | cannabinum | 1.38 | Colchicum | autumnale | 1.1 | B | 0 | 0 | 1 | Apocynum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 014.1 | 3; determination by expert, PCR on genus level | Plant extr. | 2022-02-23 | Polygonatum | multiflorum | Polygonatum | multiflorum | 1.84 | Tulbaghia | violacea | 1.39 | B | 0 | 0 | 1 | Polygonatum | Tulbaghia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 014.2 | 2; determination by expert | Plant extr. | 2022-01-27 | Polygonatum | multiflorum | Polygonatum | multiflorum | 2.6 | Coffea | arabica | 1.01 | A | 1 | 1 | 0 | Polygonatum | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 108 | 2; determination by expert | Plant extr. | 2022-01-18 | Polygonatum | multiflorum | Convallaria | majalis | 1.4 | Convallaria | majalis | 1.39 | B | 0 | 0 | 1 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 121 | 3; botanical collection | Plant extr.-R | 2022-01-18 | Polygonatum | odoratum | Viburnum | opulus | 0.95 | Moringa | stenopetala | 0.95 | B | 0 | 0 | 1 | Viburnum | Moringa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | PL 0113 | 2; label, sensory analysis | OS Extr. plant | 2018-11-22 | Prunus | armeniaca | Petroselinum | crispum | 0.78 | Lycium | europaeum | 0.75 | B | 0 | 0 | 1 | Petroselinum | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 354 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Prunus | laurocerasus | Lycium | barbarum | 1.41 | Prunus | lusitanica | 1.3 | B | 0 | 0 | 1 | Lycium | Prunus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 396.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Prunus | laurocerasus | Prunus | lusitanica | 1.25 | Colchicum | autumnale | 1.17 | B | 0 | 0 | 1 | Prunus | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 396.3 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Prunus | laurocerasus | Prunus | lusitanica | 1.42 | Allium | sativum | 1.28 | B | 0 | 0 | 1 | Prunus | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 384.1 | 2; expert opinion | Plant extr.-R | 2022-10-06 | Prunus | lusitanica | Prunus | lusitanica | 2.15 | Viburnum | opulus | 1.41 | A | 1 | 1 | 0 | Prunus | Viburnum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 384.2 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Prunus | lusitanica | Vaccinum | vitis-idaea | 0.95 | Aethusa | cynapium | 0.83 | B | 0 | 0 | 1 | Vaccinum | Aethusa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 399.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Prunus | spinosa | Viburnum | opulus | 1.28 | Symphoricarpos | albus | 1.22 | B | 0 | 0 | 1 | Viburnum | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 232 | 2; determination by expert | Plant extr. | 2022-09-09 | Pteridium | aquilinum | Jacobaea | vulgaris | 0.99 | Colchicum | autumnale | 0.95 | B | 0 | 0 | 1 | Jacobaea | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 267 | 3; botanical collection | Plant extr. | 2022-09-16 | Punica | granatum | Backhousia | citriodora | 1.34 | Symphoricarpos | albus | 1.2 | B | 0 | 0 | 1 | Backhousia | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 425.1 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Rauvolfia | tetraphylla | Spinacia | oleracea | 1.16 | Jacobaea | vulgaris | 1.08 | B | 0 | 0 | 1 | Spinacia | Jacobaea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 425.2 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Rauvolfia | tetraphylla | Spinacia | oleracea | 0.82 | Vaccinium | myrtillus | 0.8 | B | 0 | 0 | 1 | Spinacia | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 405.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rhamnus | cathartica | Viburnum | lantana | 1.05 | Coffea | arabica | 1.04 | B | 0 | 0 | 1 | Viburnum | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 405.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rhamnus | cathartica | Lycium | europaeum | 1.08 | Vaccinium | sect. | 0.89 | B | 0 | 0 | 1 | Lycium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 599.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Ribes | nigrum | Conium | maculatum | 1.22 | Petroselinum | crispum | 1.08 | B | 0 | 0 | 1 | Conium | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 599.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes | nigrum | Ribes | nigrum | 2.51 | Ribes | nigrum | 1.75 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 603.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-21 | Ribes | nigrum | Spinacia | oleracea | 0.76 | Galium | mollugo | 0.72 | B | 0 | 0 | 1 | Spinacia | Galium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 603.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-21 | Ribes | nigrum | Ribes | nigrum | 2.5 | Ribes | nigrum | 2.21 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 581.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-19 | Ribes | rubrum | Ribes | rubrum | 1.94 | Vaccinium | sect. | 0.68 | B | 0 | 0 | 1 | Ribes | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 581.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-19 | Ribes | rubrum | Ribes | rubrum | 2.45 | Ribes | rubrum | 2.34 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 582.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-16 | Ribes | rubrum | Ribes | rubrum | 2.29 | Apocynum | cannabinum | 0.91 | A | 1 | 1 | 0 | Ribes | Apocynum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 582.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-19 | Ribes | rubrum | Ribes | rubrum | 2.24 | Ribes | rubrum | 2.22 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 584.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.36 | Ribes | rubrum | 2.26 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 585.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.42 | Ribes | rubrum | 2.41 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 587.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.42 | Ribes | rubrum | 2.4 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 588.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.48 | Ribes | rubrum | 2.4 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 589.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-08 | Ribes | rubrum | Ribes | rubrum | 2.33 | Ribes | rubrum | 2.23 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 590.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.36 | Ribes | rubrum | 2.31 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 591.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.28 | Ribes | rubrum | 2.14 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 593.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.46 | Ribes | rubrum | 2.39 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 594.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.33 | Ribes | rubrum | 2.23 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 600.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Ribes | rubrum | Ribes | rubrum | 1.53 | Vaccinium | sect. | 1.07 | B | 0 | 0 | 1 | Ribes | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 600.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes | rubrum | Ribes | rubrum | 2.12 | Ribes | rubrum | 2.04 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 001 | 3; determination by expert | Plant extr. | 2021-07-22 | Ricinus | communis | Ricinus | communis | 2.65 | Ricinus | communis | 2.49 | A | 1 | 1 | 0 | Ricinus | Ricinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 002 | 3; determination by expert | Plant extr. | 2021-07-22 | Ricinus | communis | Ricinus | communis | 2.43 | Ricinus | communis | 2.02 | A | 1 | 1 | 0 | Ricinus | Ricinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 004 | 2; determination by expert | Plant extr. | 2021-07-22 | Ricinus | communis | Ricinus | communis | 2.53 | Ricinus | communis | 2.52 | A | 1 | 1 | 0 | Ricinus | Ricinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 005 | 3; determination by expert | Plant extr. | 2021-07-22 | Ricinus | communis | Ricinus | communis | 2.32 | Ricinus | communis | 2.26 | A | 1 | 1 | 0 | Ricinus | Ricinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 099.1 | 2; determination by expert | Plant extr. | 2022-01-17 | Ricinus | communis | Aethusa | cynapium | 1.21 | Lycium | europaeum | 1.03 | B | 0 | 0 | 1 | Aethusa | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 099.2 | 2; determination by expert | Plant extr. | 2022-01-21 | Ricinus | communis | Ricinus | communis | 2.18 | Ricinus | communis | 2.08 | A | 1 | 1 | 0 | Ricinus | Ricinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 112 | 2; determination by expert | Plant extr. | 2022-01-17 | Ricinus | communis | Jacobaea | vulgaris | 1.05 | Anethum | graveolens | 0.98 | B | 0 | 0 | 1 | Jacobaea | Anethum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 341.1 | 3; botanical collection | Plant extr. | 2022-09-28 | Ricinus | communis | Anethum | graveolens | 1.08 | Allium | schoenoprasum | 1 | B | 0 | 0 | 1 | Anethum | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 392 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Robinia | pseudoacacia | Leptospermum | citratum | 1.16 | Curcuma | longa | 1.14 | B | 0 | 0 | 1 | Leptospermum | Curcuma | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 387.1 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Rosa | multiflora | Aconitum | napellus | 1 | Convallaria | majalis | 1 | B | 0 | 0 | 1 | Aconitum | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 387.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Rosa | multiflora | Vaccinum | vitis-idaea | 0.85 | Aconitum | napellus | 0.74 | B | 0 | 0 | 1 | Vaccinum | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 398.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rosa | multiflora | Symphoricarpos | albus | 1.34 | Ligustrum | vulgare | 1.24 | B | 0 | 0 | 1 | Symphoricarpos | Ligustrum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 398.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rosa | multiflora | Atropa | belladonna | 1.07 | Artemisia | vulgaris | 1.01 | B | 0 | 0 | 1 | Atropa | Artemisia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 163 | 2; label | Plant extr. | 2022-02-23 | Rubus | idaeus | Vaccinium | sect. | 1.17 | Clinopodium | nepeta | 1.04 | B | 0 | 0 | 1 | Vaccinium | Clinopodium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 053.1 | 2; determination by expert | Plant extr.-R | 2022-01-24 | Salvia | pratensis | Solanum | dulcamara | 1.16 | Leucojum | vernum | 1.11 | B | 0 | 0 | 1 | Solanum | Leucojum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 053.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Salvia | pratensis | Digitalis | purpurea | 1.01 | Conium | maculatum | 0.94 | B | 0 | 0 | 1 | Digitalis | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 133 | 2; determination by expert | Plant extr.-R | 2022-02-02 | Salvia | rosmarinus | Symphoricarpos | albus | 1.27 | Digitalis | purpurea | 1.13 | B | 0 | 0 | 1 | Symphoricarpos | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 144 | 2; expert opinion | Plant extr. | 2022-10-04 | Salvia | sp | Curcuma | longa | 1.28 | Backhousia | citriodora | 1.06 | B | 0 | 0 | 1 | Curcuma | Backhousia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 266 | 3; botanical collection | Plant extr. | 2022-09-16 | Sambucus | nigra | Conium | maculatum | 0.97 | Zingiber | officinale | 0.87 | B | 0 | 0 | 1 | Conium | Zingiber | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 395.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Sambucus | nigra | Viburnum | lantana | 1.4 | Galanthus | nivalis | 1.33 | B | 0 | 0 | 1 | Viburnum | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 395.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Sambucus | nigra | Vaccinium | myrtillus | 0.75 | Allium | schoenoprasum | 0.75 | B | 0 | 0 | 1 | Vaccinium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 129.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Senecio | inaequidens | Polygonatum | multiflorum | 1.03 | Galanthus | nivalis | 0.99 | B | 0 | 0 | 1 | Polygonatum | Galanthus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 049.1 | 2; determination by expert | Plant extr. | 2022-01-14 | Senecio | jacobaea | Jacobaea | vulgaris | 2.43 | Mansoa | alliacea | 1.56 | A | 0 | 0 | 1 | Jacobaea | Mansoa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 049.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Senecio | jacobaea | Digitalis | purpurea | 1.23 | Coffea | arabica | 1.01 | B | 0 | 0 | 1 | Digitalis | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 050.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Senecio | jacobaea | Jacobaea | vulgaris | 1.32 | Clinopodium | nepeta | 1.17 | B | 0 | 0 | 1 | Jacobaea | Clinopodium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 050.3 | 2; determination by expert | Plant extr. | 2022-01-19 | Senecio | jacobaea | Jacobaea | vulgaris | 1.17 | Petroselinum | crispum | 1.15 | B | 0 | 0 | 1 | Jacobaea | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 050.4 | 2; determination by expert | Plant extr.-R | 2022-01-26 | Senecio | jacobaea | Ribes | rubrum | 1.16 | Vaccinium | corymbosum | 0.94 | B | 0 | 0 | 1 | Ribes | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 211 | 2; determination by expert | Plant extr. | 2022-09-09 | Senecio | viravira | Convallaria | majalis | 0.91 | Vaccinium | myrtillus | 0.79 | B | 0 | 0 | 1 | Convallaria | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 048.2 | 2; determination by expert | Plant extr.-R | 2022-01-26 | Senecio | vulgaris | Leucojum | vernum | 1.04 | Colchicum | autumnale | 0.99 | B | 0 | 0 | 1 | Leucojum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 218 | 2; determination by expert | Plant extr. | 2022-09-09 | Sequoiadendron | sp | Convallaria | majalis | 1.12 | Convallaria | majalis | 1.11 | B | 0 | 0 | 1 | Convallaria | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 243.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Silphium | perfoliatum | Lycium | europaeum | 1.12 | Artemisia | absinthium | 1.1 | B | 0 | 0 | 1 | Lycium | Artemisia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 243.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Silphium | perfoliatum | Convallaria | majalis | 0.97 | Ocimum | basilicum | 0.92 | B | 0 | 0 | 1 | Convallaria | Ocimum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 352.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Solanum | betaceum | Solanum | betaceum | 2.52 | Datura | stramonium | 1.04 | A | 1 | 1 | 0 | Solanum | Datura | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 238.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Solanum | dulcamara | Solanum | dulcamara | 1.78 | Prunus | lusitanica | 1.33 | B | 0 | 0 | 1 | Solanum | Prunus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 238.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Solanum | dulcamara | Ocimum | basilicum | 0.97 | Camellia | sinensis | 0.86 | B | 0 | 0 | 1 | Ocimum | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 383.1 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Solanum | dulcamara | Solanum | dulcamara | 2.22 | Lycium | europaeum | 1.25 | A | 1 | 1 | 0 | Solanum | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 383.2 | 3; botanical collection | Plant extr. | 2022-09-30 | Solanum | dulcamara | Solanum | virginianum | 1.1 | Viburnum | lantana | 1.03 | B | 0 | 0 | 1 | Solanum | Viburnum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 100.1 | 2; determination by expert | Plant extr. | 2022-02-21 | Solanum | nigrum | Solanum | dulcamara | 1.58 | Ocimum | basilicum | 1.26 | B | 0 | 0 | 1 | Solanum | Ocimum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 100.2 | 2; determination by expert | Plant extr. | 2022-01-18 | Solanum | nigrum | Allium | ursinum | 0.71 | Conium | maculatum | 0.69 | B | 0 | 0 | 1 | Allium | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 109.1 | 3; determination by expert, PCR on genus level | Plant extr. | 2022-02-21 | Solanum | nigrum | Solanum | dulcamara | 1.43 | Arum | maculatum | 1.18 | B | 0 | 0 | 1 | Solanum | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 109.2 | 3; determination by expert, PCR on genus level | Plant extr. | 2022-01-18 | Solanum | nigrum | Coffea | arabica | 0.93 | Taxus | baccata | 0.89 | B | 0 | 0 | 1 | Coffea | Taxus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 114.1 | 2; determination by expert | Plant extr. | 2022-02-21 | Solanum | nigrum | Solanum | betaceum | 1.44 | Lycium | barbarum | 1.14 | B | 0 | 0 | 1 | Solanum | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 114.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Solanum | nigrum | Convallaria | majalis | 1.01 | Coffea | canephora | 1 | B | 0 | 0 | 1 | Convallaria | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 114.3 | 2; determination by expert | Plant extr. | 2022-01-18 | Solanum | nigrum | Galium | mollugo | 1.09 | Datura | inoxia | 1.03 | B | 0 | 0 | 1 | Galium | Datura | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 264.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Solanum | nigrum | Ocimum | basilicum | 1.16 | Vaccinium | sect. | 1.12 | B | 0 | 0 | 1 | Ocimum | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 235 | 2; determination by expert | Plant extr. | 2022-09-07 | Solanum | tuberosum | Symphoricarpos | albus | 1.01 | Vaccinium | macrocarpon | 0.99 | B | 0 | 0 | 1 | Symphoricarpos | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 365 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Solanum | virginianum | Solanum | virginianum | 2.07 | Leucojum | vernum | 1.42 | A | 1 | 1 | 0 | Solanum | Leucojum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 221.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Spartium | junceum | Prunus | lusitanica | 1.07 | Conium | maculatum | 1.02 | B | 0 | 0 | 1 | Prunus | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 221.3 | 2; determination by expert | Plant extr. | 2022-09-09 | Spartium | junceum | Coffea | arabica | 1.06 | Cinnamomum | camphora | 0.9 | B | 0 | 0 | 1 | Coffea | Cinnamomum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 219 | 2; determination by expert | Plant extr. | 2022-09-05 | Spinacia | oleracea | Spinacia | oleracea | 2.13 | Allium | ursinum | 1.03 | A | 1 | 1 | 0 | Spinacia | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 386 | 2; expert opinion | Plant extr.-R | 2022-10-07 | Spinacia | oleracea | Spinacia | oleracea | 1.72 | Allium | ursinum | 1.04 | B | 0 | 0 | 1 | Spinacia | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 388.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Symphoricarpos | albus | Symphoricarpos | albus | 1.94 | Prunus | lusitanica | 1.37 | B | 0 | 0 | 1 | Symphoricarpos | Prunus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 401.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Symphoricarpos | albus | Symphoricarpos | albus | 2.27 | Viburnum | opulus | 1.31 | A | 1 | 1 | 0 | Symphoricarpos | Viburnum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 401.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Symphoricarpos | albus | Petroselinum | crispum | 1.2 | Petroselinum | crispum | 0.96 | B | 0 | 0 | 1 | Petroselinum | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 476.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Symphoricarpos | orbiculatus | Mansoa | alliacea | 1.17 | Symphoricarpos | albus | 1.03 | B | 0 | 0 | 1 | Mansoa | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 482.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Symphoricarpos | orbiculatus | Symphoricarpos | albus | 1.3 | Leucojum | vernum | 1.14 | B | 0 | 0 | 1 | Symphoricarpos | Leucojum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 008 | 2; determination by expert | Plant extr. | 2022-01-20 | Taxus | baccata | Taxus | baccata | 2.36 | Polygonatum | multiflorum | 1.06 | A | 1 | 1 | 0 | Taxus | Polygonatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 009 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus | baccata | Taxus | baccata | 2.23 | Cinnamomum | verum | 0.86 | A | 1 | 1 | 0 | Taxus | Cinnamomum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 010 | 2; determination by expert | Plant extr. | 2022-01-19 | Taxus | baccata | Clinopodium | nepeta | 0.84 | Vaccinium | myrtillus | 0.78 | B | 0 | 0 | 1 | Clinopodium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 091 | 2; determination by expert | Plant extr. | 2022-01-19 | Taxus | baccata | Vaccinium | myrtillus | 1.09 | Convallaria | majalis | 1 | B | 0 | 0 | 1 | Vaccinium | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 092 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus | baccata | Digitalis | purpurea | 1 | Allium | ursinum | 0.99 | B | 0 | 0 | 1 | Digitalis | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 095.1 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus | baccata | Coffea | arabica | 1.02 | Coffea | arabica | 0.96 | B | 0 | 0 | 1 | Coffea | Coffea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 096.1 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus | baccata | Arum | maculatum | 1.28 | Ligustrum | vulgare | 1.19 | B | 0 | 0 | 1 | Arum | Ligustrum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 096.3 | 2; determination by expert | Plant extr. | 2022-01-18 | Taxus | baccata | Taxus | baccata | 1.62 | Petroselinum | crispum | 1.17 | B | 0 | 0 | 1 | Taxus | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 096.4 | 2; determination by expert | Plant extr. | 2022-02-21 | Taxus | baccata | Viola | odorata | 0.95 | Juglans | regia | 0.92 | B | 0 | 0 | 1 | Viola | Juglans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 389.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Taxus | baccata | Allium | ursinum | 1.23 | Allium | ursinum | 1.13 | B | 0 | 0 | 1 | Allium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 389.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Taxus | baccata | Taxus | baccata | 1.54 | Nerium | oleander | 1.01 | B | 0 | 0 | 1 | Taxus | Nerium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 406 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Taxus | baccata | Taxus | baccata | 1.78 | Petroselinum | crispum | 1.09 | B | 0 | 0 | 1 | Taxus | Petroselinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 303 | 3; determination by expert | Plant extr. | 2023-11-23 | Tulbaghia | violacea | Tulbaghia | violacea | 2.34 | Convallaria | majalis | 1.37 | A | 1 | 1 | 0 | Tulbaghia | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 224.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Ulex | europaeus | Artemisia | absinthium | 1.19 | Convallaria | majalis | 1.19 | B | 0 | 0 | 1 | Artemisia | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 224.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Ulex | europaeus | Vaccinium | sect. | 1.15 | Solanum | virginianum | 1.12 | B | 0 | 0 | 1 | Vaccinium | Solanum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 664 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium | angustifolium | Vaccinium | angustifolium | 2.28 | Vaccinium | angustifolium | 1.89 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 665 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium | angustifolium | Vaccinium | angustifolium | 2.1 | Vaccinium | angustifolium | 1.63 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 666 | 3; botanical collection | Plant extr.-R | 2023-12-11 | Vaccinium | angustifolium | Vaccinium | angustifolium | 2.05 | Vaccinium | angustifolium | 1.8 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 667 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium | angustifolium | Vaccinium | angustifolium | 2.1 | Vaccinium | sect. | 2.06 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 668 | 3; botanical collection | Plant extr.-R | 2023-12-13 | Vaccinium | angustifolium | Vaccinium | angustifolium | 2.59 | Vaccinium | sect. | 1.88 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 362.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vaccinium | corymbosum | Vaccinum | vitis-idaea | 1.3 | Vaccinium | corymbosum | 1.23 | B | 0 | 0 | 1 | Vaccinum | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 362.2 | 3; botanical collection | Plant extr.-R | 2022-01-21 | Vaccinium | corymbosum | Vaccinium | corymbosum | 2.07 | Vaccinium | sect. | 1.89 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 669 | 3; botanical collection | Plant extr.-R | 2023-10-27 | Vaccinium | corymbosum | Vaccinium | corymbosum | 1.94 | Vaccinium | sect. | 1.78 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 670 | 3; botanical collection | Plant extr.-R | 2023-10-27 | Vaccinium | corymbosum | Vaccinium | sect. | 2.07 | Vaccinium | corymbosum | 2 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 671 | 3; botanical collection | Plant extr.-R | 2023-10-20 | Vaccinium | corymbosum | Vaccinium | corymbosum | 2.02 | Vaccinium | sect. | 1.75 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 680 | 3; botanical collection | Plant extr.-R | 2023-10-18 | Vaccinium | darrowii | Vaccinium | darrowii | 2.05 | Vaccinium | corymbosum | 1.57 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 677 | 3; botanical collection | Plant extr.-R | 2023-12-11 | Vaccinium | elliottii | Vaccinium | elliottii | 2.34 | Vaccinium | angustifolium | 1.7 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 672 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium | fuscatum | Vaccinium | fuscatum | 2.07 | Vaccinium | elliottii | 1.8 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 673 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium | fuscatum | Vaccinium | fuscatum | 2.18 | Vaccinium | fuscatum | 1.92 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 674 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium | fuscatum | Vaccinium | fuscatum | 2.25 | Vaccinium | fuscatum | 2.08 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 675 | 3; botanical collection | Plant extr.-R | 2023-12-12 | Vaccinium | fuscatum | Vaccinium | fuscatum | 2.16 | Vaccinium | fuscatum | 1.96 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 575.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium | macrocarpon | Vaccinium | myrtillus | 1.12 | Vaccinium | myrtillus | 1.05 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 575.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium | macrocarpon | Vaccinium | macrocarpon | 2.39 | Vaccinium | macrocarpon | 2.35 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 576.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium | macrocarpon | Vaccinium | vitis-idaea | 0.91 | Nerium | oleander | 0.86 | B | 0 | 0 | 1 | Vaccinium | Nerium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 576.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium | macrocarpon | Vaccinium | macrocarpon | 2.19 | Vaccinium | macrocarpon | 2.01 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 596.2 | 2; label, sensory analysis | Plant extr. | 2023-12-18 | Vaccinium | macrocarpon | Vaccinium | macrocarpon | 2.55 | Vaccinium | macrocarpon | 2.53 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 708.2 | 2; label, sensory analysis | Plant extr. | 2023-12-18 | Vaccinium | macrocarpon | Vaccinium | macrocarpon | 2.28 | Vaccinium | macrocarpon | 2.24 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 681 | 3; botanical collection | Plant extr.-R | 2023-12-11 | Vaccinium | myrtilloides | Vaccinium | sect. | 2.05 | Vaccinium | myrtilloides | 2.05 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 192.1 | 2; expert opinion | Plant extr.-R | 2022-11-04 | Vaccinium | myrtillus | Vaccinium | myrtillus | 1.34 | Vaccinum | vitis-idaea | 1.28 | B | 0 | 0 | 1 | Vaccinium | Vaccinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 192.2 | 2; expert opinion | Plant extr.-AceR | 2022-12-19 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.12 | Vaccinium | myrtillus | 1.92 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 192.3 | 2; expert opinion | Plant extr.-R | 2022-12-06 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.39 | Vaccinium | myrtillus | 2.15 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 469.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.34 | Vaccinium | myrtillus | 1.99 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 469.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-22 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.08 | Vaccinium | myrtillus | 2.07 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 507.1 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.2 | Vaccinium | myrtillus | 2.18 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 507.2 | 2; label; sensory analysis | Plant extr.-RL40 | 2022-12-02 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.17 | Vaccinium | myrtillus | 2.17 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 550.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-02 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.1 | Vaccinium | myrtillus | 2.06 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 550.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.32 | Vaccinium | myrtillus | 2.32 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 571.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-12 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.28 | Vaccinium | myrtillus | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 571.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-12 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.1 | Vaccinium | myrtillus | 2.09 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 572.1 | 2; label; sensory analysis; Biomarker | Plant extr.-R | 2022-12-19 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.28 | Vaccinium | myrtillus | 2.22 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 572.2 | 2; label; sensory analysis; Biomarker | Plant extr.-R | 2022-12-23 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.03 | Vaccinium | myrtillus | 1.94 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 607.1 | 2; label; sensory analysis | Plant extr.-AceR | 2023-01-19 | Vaccinium | myrtillus | Vaccinium | myrtillus | 1.98 | Vaccinium | myrtillus | 1.96 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 607.2 | 2; label; sensory analysis | Plant extr.-R | 2023-01-19 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.07 | Vaccinium | myrtillus | 2.04 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 622 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.35 | Vaccinium | myrtillus | 2.17 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 623 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.11 | Vaccinium | myrtillus | 2.1 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 624 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.66 | Vaccinium | myrtillus | 2.13 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 625 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.23 | Vaccinium | myrtillus | 2.14 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 626 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.39 | Vaccinium | myrtillus | 2.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 627 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.33 | Vaccinium | myrtillus | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 628 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.32 | Vaccinium | myrtillus | 2.28 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 629 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.44 | Vaccinium | myrtillus | 2.31 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 630 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.23 | Vaccinium | myrtillus | 2.21 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 631 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.38 | Vaccinium | myrtillus | 2.35 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 632 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.14 | Vaccinium | myrtillus | 2.13 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 633 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.45 | Vaccinium | myrtillus | 2.23 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 635 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.43 | Vaccinium | myrtillus | 2.41 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 636 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.31 | Vaccinium | myrtillus | 2.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 637 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.09 | Vaccinium | myrtillus | 2.09 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 655 | 3; botanical collection | Plant extr.-R | 2023-03-08 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.1 | Vaccinium | myrtillus | 2.04 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 656 | 3; botanical collection | Plant extr. R | 2023-03-08 | Vaccinium | myrtillus | Vaccinium | myrtillus | 1.9 | Vaccinium | myrtillus | 1.84 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 657 | 3; botanical collection | Plant extr. R | 2023-03-08 | Vaccinium | myrtillus | Arum | maculatum | 0.86 | Convallaria | majalis | 0.86 | B | 0 | 0 | 1 | Arum | Convallaria | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 658 | 3; botanical collection | Plant extr. R | 2023-03-13 | Vaccinium | myrtillus | Vaccinium | myrtillus | 1.99 | Vaccinium | myrtillus | 1.9 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 659 | 3; botanical collection | Plant extr.-R | 2023-03-13 | Vaccinium | myrtillus | Vaccinium | myrtillus | 2.14 | Vaccinium | myrtillus | 2.11 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 679 | 3; botanical collection | Plant extr.-R | 2023-10-18 | Vaccinium | parvifolium | Vaccinium | parvifolium | 2.14 | Vaccinium | myrtillus | 1.66 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 468.1 | 2; label; sensory analysis | Plant extr.-R | 2022-11-03 | Vaccinium | sect. | Vaccinium | sect. | 2.25 | Vaccinium | sect. | 2.18 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 468.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-22 | Vaccinium | sect. | Vaccinium | sect. | 2.34 | Vaccinium | sect. | 2.2 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 470.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-12 | Vaccinium | sect. | Vaccinium | sect. | 2.49 | Vaccinium | sect. | 2.19 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 470.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-12 | Vaccinium | sect. | Vaccinium | sect. | 2.38 | Vaccinium | sect. | 2.32 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 492.1 | 2; label; sensory analysis | Plant extr.-R | 2022-12-07 | Vaccinium | sect. | Vaccinium | sect. | 2.46 | Vaccinium | sect. | 2.43 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 492.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-21 | Vaccinium | sect. | Vaccinium | sect. | 2.41 | Vaccinium | sect. | 2.3 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 501.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.43 | Vaccinium | sect. | 2.01 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 501.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.22 | Vaccinium | sect. | 2.01 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 502.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.32 | Vaccinium | sect. | 2.32 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 502.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.42 | Vaccinium | sect. | 2.38 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 503.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.42 | Vaccinium | sect. | 2.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 503.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.5 | Vaccinium | sect. | 2.43 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 504.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.34 | Vaccinium | sect. | 1.99 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 504.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.36 | Vaccinium | sect. | 2.18 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 505.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.48 | Vaccinium | sect. | 2.22 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 505.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.27 | Vaccinium | sect. | 2.25 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 506.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.35 | Vaccinium | sect. | 2.27 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 506.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.46 | Vaccinium | sect. | 2.17 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 508.1 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.3 | Vaccinium | sect. | 2.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 508.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.43 | Vaccinium | sect. | 2.34 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 509.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.37 | Vaccinium | sect. | 2.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 509.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.4 | Vaccinium | sect. | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 510.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.14 | Vaccinium | sect. | 2.12 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 510.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.32 | Vaccinium | sect. | 2.31 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 511.1 | 2; label; sensory analysis | Plant extr.-AceRL42S90 | 2022-12-07 | Vaccinium | sect. | Vaccinium | sect. | 2.46 | Vaccinium | sect. | 2.44 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 511.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.31 | Vaccinium | sect. | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 512.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-06 | Vaccinium | sect. | Vaccinium | sect. | 2.39 | Vaccinium | sect. | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 512.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.45 | Vaccinium | sect. | 2.4 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 513.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-06 | Vaccinium | sect. | Vaccinium | sect. | 2.25 | Vaccinium | sect. | 2.23 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 513.2 | 2; label; sensory analysis | Plant extr.-RL40 | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.3 | Vaccinium | sect. | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 514.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-06 | Vaccinium | sect. | Vaccinium | sect. | 2.42 | Vaccinium | sect. | 2.27 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 514.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.28 | Vaccinium | sect. | 2.24 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 573.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-23 | Vaccinium | sect. | Vaccinium | sect. | 2.1 | Vaccinium | sect. | 1.77 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 573.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium | sect. | Vaccinium | sect. | 2.21 | Vaccinium | sect. | 1.92 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 604.1 | 1; label | Plant extr.-AceR | 2023-01-16 | Vaccinium | sect. | Vaccinium | sect. | 1.91 | Vaccinium | sect. | 1.89 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 604.2 | 1; label | Plant extr.-R | 2023-01-19 | Vaccinium | sect. | Vaccinium | sect. | 2.26 | Vaccinium | corymbosum | 1.95 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 495 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | sp | Vaccinium | myrtillus | 2.15 | Vaccinium | myrtillus | 2.09 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 551 | 1; label | Plant extr.-R | 2022-12-02 | Vaccinium | sp | Vaccinium | myrtillus | 1.93 | Vaccinium | myrtillus | 1.86 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 554.1 | 1; label | Plant extr.-AceR | 2022-12-06 | Vaccinium | sp | Allium | porrum | 1.08 | Arum | maculatum | 0.96 | B | 0 | 0 | 1 | Allium | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 554.2 | 1; label | Plant extr.-R | 2023-01-10 | Vaccinium | sp | Vaccinium | myrtillus | 1.76 | Vaccinium | myrtillus | 1.73 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 555.1 | 1; label | Plant extr.-AceR | 2022-12-06 | Vaccinium | sp | Vaccinium | uliginosum | 1.06 | Vaccinium | myrtillus | 0.98 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 555.2 | 1; label | Plant extr.-R | 2023-01-10 | Vaccinium | sp | Vaccinium | corymbosum | 1.82 | Vaccinium | sect. | 1.81 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 556 | 1; label | Plant extr.-R | 2023-01-10 | Vaccinium | sp | Vaccinium | myrtillus | 2.05 | Vaccinium | myrtillus | 1.92 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 558 | 1; label | Plant extr.-AceR | 2022-12-23 | Vaccinium | sp | Vaccinium | myrtillus | 1.74 | Vaccinium | myrtillus | 1.63 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 568.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Viburnum | opulus | 1.06 | Allium | ursinum | 1.02 | B | 0 | 0 | 1 | Viburnum | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 568.2 | 1; label | Plant extr.-R | 2022-12-12 | Vaccinium | sp | Vaccinium | myrtillus | 2.05 | Vaccinium | myrtillus | 1.99 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 569.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Vaccinium | myrtillus | 1.98 | Vaccinium | myrtillus | 1.97 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 569.2 | 1; label | Plant extr.-R | 2022-12-12 | Vaccinium | sp | Vaccinium | myrtillus | 2.11 | Vaccinium | myrtillus | 1.99 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 570.11 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Convallaria | majalis | 0.92 | Vaccinium | corymbosum | 0.91 | B | 0 | 0 | 1 | Convallaria | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 570.22 | 1; label | Plant extr.-R | 2022-12-23 | Vaccinium | sp | Vaccinium | sect. | 1.93 | Vaccinium | sect. | 1.83 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 574.1 | 1; label | Plant extr.-R | 2022-12-16 | Vaccinium | sp | Vaccinium | myrtillus | 1.65 | Vaccinium | myrtillus | 1.51 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 574.2 | 1; label | Plant extr.-RL42S90 | 2022-12-20 | Vaccinium | sp | Leptospermum | citratum | 0.95 | Vaccinium | myrtillus | 0.92 | B | 0 | 0 | 1 | Leptospermum | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 577.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium | sp | Vaccinium | myrtillus | 1.4 | Vaccinium | myrtillus | 1.34 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 577.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium | sp | Vaccinium | myrtillus | 1.74 | Vaccinium | sect. | 1.56 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 577.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-19 | Vaccinium | sp | Vaccinium | myrtillus | 1.5 | Vaccinium | myrtillus | 1.25 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 578.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium | sp | Vaccinium | sect. | 1.85 | Vaccinium | angustifolium | 1.83 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 578.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-23 | Vaccinium | sp | Vaccinium | myrtillus | 1.2 | Vaccinium | myrtillus | 1.03 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 579.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium | sp | Vaccinium | corymbosum | 1.85 | Vaccinium | sect. | 1.79 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 579.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium | sp | Vaccinium | angustifolium | 1.91 | Vaccinium | sect. | 1.79 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 579.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-01-02 | Vaccinium | sp | Vaccinium | sect. | 1.24 | Vaccinium | corymbosum | 1.09 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 580.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium | sp | Vaccinium | corymbosum | 2 | Vaccinium | sect. | 1.83 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 580.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium | sp | Vaccinium | angustifolium | 1.84 | Vaccinium | sect. | 1.75 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 580.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-19 | Vaccinium | sp | Vaccinium | sect. | 1.41 | Vaccinium | sect. | 1.36 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 595.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium | sp | Vaccinium | myrtillus | 1.7 | Vaccinium | myrtillus | 1.67 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 595.2 | 1; label | Plant extr.-AceR | 2022-12-19 | Vaccinium | sp | Vaccinium | myrtillus | 2.06 | Vaccinium | myrtillus | 2.03 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 595 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium | sp | Vaccinium | myrtillus | 2.02 | Vaccinium | myrtillus | 2.01 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 596.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Vaccinium | sp | Digitalis | purpurea | 1.28 | Conium | maculatum | 1.12 | B | 0 | 0 | 1 | Digitalis | Conium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 597.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Vaccinium | sp | Vaccinium | myrtillus | 1.99 | Vaccinium | myrtillus | 1.97 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 598.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Vaccinium | sp | Vaccinium | myrtillus | 2 | Vaccinium | myrtillus | 1.99 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 598.2 | 1; label | Plant extr.-R | 2022-12-21 | Vaccinium | sp | Vaccinium | myrtillus | 1.8 | Vaccinium | myrtillus | 1.79 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 606 | 1; label | Plant extr.-R | 2023-01-19 | Vaccinium | sp | Vaccinium | myrtillus | 2.2 | Vaccinium | myrtillus | 2.05 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 608 | 1; label | Plant extr.-AceR | 2022-01-03 | Vaccinium | sp | Vaccinium | myrtillus | 1.39 | Vaccinium | myrtillus | 1.14 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 609.1 | 1; label | Plant extr.-AceR 1:5 | 2022-12-22 | Vaccinium | sp | Vaccinium | sect. | 1.79 | Vaccinium | angustifolium | 1.62 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 609.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | sect. | 1.84 | Vaccinium | angustifolium | 1.8 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 609.3 | 1; label | Plant extr.-AceR | 2022-01-02 | Vaccinium | sp | Vaccinium | myrtillus | 1.09 | Vaccinium | sect. | 1.06 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 610.1 | 1; label | Plant extr.-AceR | 2022-12-22 | Vaccinium | sp | Vaccinium | sect. | 1.89 | Vaccinium | angustifolium | 1.78 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 610.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | sect. | 1.9 | Vaccinium | angustifolium | 1.9 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 610.3 | 1; label | Plant extr.-AceR 1:1 | 2022-12-23 | Vaccinium | sp | Vaccinium | sect. | 1.18 | Allium | porrum | 1.07 | B | 0 | 0 | 1 | Vaccinium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 611.11 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Ribes | rubrum | 0.9 | Vaccinium | vitis-idaea | 0.85 | B | 0 | 0 | 1 | Ribes | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 611.12 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Vaccinium | myrtillus | 1.44 | Vaccinium | myrtillus | 1.4 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 611.21 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | myrtillus | 1.99 | Vaccinium | myrtillus | 1.96 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 611.22 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | sect. | 1.78 | Vaccinium | sect. | 1.78 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 612.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Vaccinium | myrtillus | 0.98 | Vaccinium | sect. | 0.96 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 612.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | myrtillus | 1.69 | Vaccinium | myrtillus | 1.61 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 613.12 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Lycium | barbarum | 1.21 | Digitalis | purpurea | 0.98 | B | 0 | 0 | 1 | Lycium | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 613.21 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | myrtillus | 2.01 | Vaccinium | myrtillus | 1.99 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 613.22 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | corymbosum | 1.81 | Vaccinium | sect. | 1.8 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 614.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Asarum | europaeum | 0.9 | Ribes | rubrum | 0.89 | B | 0 | 0 | 1 | Asarum | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 615.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-22 | Vaccinium | sp | Vaccinium | angustifolium | 1.79 | Vaccinium | corymbosum | 1.63 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 616.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium | sp | Vaccinium | myrtillus | 1.4 | Vaccinium | myrtillus | 1.28 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 617 | 1; label | Plant extr.-AceR | 2022-01-03 | Vaccinium | sp | Digitalis | purpurea | 1.42 | Vaccinium | sect. | 1.11 | B | 0 | 0 | 1 | Digitalis | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 618 | 1; label | Plant extr.-AceR | 2022-01-03 | Vaccinium | sp | Vaccinium | myrtillus | 1.32 | Vaccinium | myrtillus | 1.25 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 619.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium | sp | Vaccinium | vitis-idaea | 1.69 | Vaccinium | vitis-idaea | 1.69 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 642 | 1; label | Plant extr.-R | 2023-02-07 | Vaccinium | sp | Vaccinium | myrtillus | 1.85 | Vaccinium | myrtillus | 1.82 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 649.2 | 1; label | Plant extr. | 2023-02-17 | Vaccinium | sp | Vaccinium | sect. | 2.38 | Vaccinium | sect. | 2.36 | A | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 650.2 | 1; label | Plant extr. | 2023-02-17 | Vaccinium | sp | Vaccinium | sect. | 1.72 | Vaccinium | sect. | 1.72 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 366 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vaccinium | uliginosum | Vaccinum | vitis-idaea | 1.42 | Camellia | sinensis | 1.38 | B | 0 | 0 | 1 | Vaccinum | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 638 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | uliginosum | Vaccinium | uliginosum | 2.02 | Vaccinium | uliginosum | 1.7 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 660 | 3; botanical collection | Plant extr. R | 2023-03-13 | Vaccinium | uliginosum | Vaccinium | uliginosum | 2.71 | Vaccinium | uliginosum | 1.49 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 369.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vaccinium | vitis-idaea | Vaccinum | vitis-idaea | 2.41 | Vaccinium | vitis-idaea | 1.5 | A | 0 | 0 | 1 | Vaccinum | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 369.2 | 3; botanical collection | Plant extr.-R | 2022-10-19 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.25 | Vaccinum | vitis-idaea | 1.67 | A | 1 | 1 | 0 | Vaccinium | Vaccinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 639 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.23 | Vaccinium | vitis-idaea | 2.16 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 640 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.4 | Vaccinium | vitis-idaea | 2.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 641 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.29 | Vaccinium | vitis-idaea | 2.22 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 643 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.35 | Vaccinium | vitis-idaea | 2.15 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 644 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.29 | Vaccinium | vitis-idaea | 2.03 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 645 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.39 | Vaccinium | vitis-idaea | 2.2 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 646 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.47 | Vaccinium | vitis-idaea | 2.1 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 647 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.29 | Vaccinium | vitis-idaea | 2.27 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 661 | 3; botanical collection | Plant extr.-R | 2023-03-13 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.32 | Vaccinium | vitis-idaea | 2.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 662 | 3; botanical collection | Plant extr. R | 2023-03-13 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.29 | Vaccinium | vitis-idaea | 2.16 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 308 | 3; botanical collection | Plant extr. | 2022-09-29 | Vanilla | planifolia | Vaccinium | corymbosum | 1.18 | Lycium | europaeum | 1.17 | B | 0 | 0 | 1 | Vaccinium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 336 | 3; botanical collection | Plant extr. | 2022-09-28 | Vanilla | planifolia | Galanthus | nivalis | 1.14 | Arum | maculatum | 1.13 | B | 0 | 0 | 1 | Galanthus | Arum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 250.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Vanilla | pompona | Lycium | europaeum | 1.12 | Anethum | graveolens | 1.06 | B | 0 | 0 | 1 | Lycium | Anethum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 339 | 3; botanical collection | Plant extr. | 2022-09-28 | Vanilla | pompona | Vaccinium | sect. | 1.18 | Tulbaghia | violacea | 1.06 | B | 0 | 0 | 1 | Vaccinium | Tulbaghia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 040 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Veratrum | album | Coffea | arabica | 0.89 | Mansoa | alliacea | 0.84 | B | 0 | 0 | 1 | Coffea | Mansoa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 381.1 | 3; botanical collection | Plant extr. | 2022-10-04 | Veratrum | album | Coffea | arabica | 0.72 | Taxus | baccata | 0.68 | B | 0 | 0 | 1 | Coffea | Taxus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 400.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Viburnum | lantana | Viburnum | lantana | 2.3 | Symphoricarpos | albus | 1.36 | A | 1 | 1 | 0 | Viburnum | Symphoricarpos | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 403.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Viburnum | opulus | Viburnum | opulus | 2.37 | Viburnum | lantana | 1.63 | A | 1 | 1 | 0 | Viburnum | Viburnum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 403.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Viburnum | opulus | Petroselinum | crispum | 1.12 | Artemisia | vulgaris | 1.1 | B | 0 | 0 | 1 | Petroselinum | Artemisia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 361 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vinca | minor | Mansoa | alliacea | 1.23 | Anethum | graveolens | 1.16 | B | 0 | 0 | 1 | Mansoa | Anethum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 069.1 | 2; label | Plant extr. | 2022-02-21 | Viola | odorata | Viola | odorata | 2.46 | Ligustrum | vulgare | 0.93 | A | 1 | 1 | 0 | Viola | Ligustrum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 069.2 | 2; label | Plant extr. | 2022-01-21 | Viola | odorata | Vaccinium | corymbosum | 0.89 | Vaccinium | sect. | 0.73 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 016 | 2; determination by expert | Plant extr. | 2022-01-26 | Viscum | album | Viscum | album | 2.27 | Allium | vineale | 0.77 | A | 1 | 1 | 0 | Viscum | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 146 | 2; determination by expert | Plant extr. | 2022-02-03 | Viscum | album | Viscum | album | 1.39 | Camellia | sinensis | 0.97 | B | 0 | 0 | 1 | Viscum | Camellia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 442 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis | tiliifolia | Camellia | sinensis | 1.35 | Backhousia | citriodora | 1.28 | B | 0 | 0 | 1 | Camellia | Backhousia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 445 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis | vinifera | Arum | maculatum | 1.34 | Aconitum | carmichaelii | 1.19 | B | 0 | 0 | 1 | Arum | Aconitum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 446.1 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis | yunnanensis | Ligustrum | vulgare | 1.21 | Colchicum | autumnale | 1.17 | B | 0 | 0 | 1 | Ligustrum | Colchicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 446.2 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis | yunnanensis | Vaccinium | sect. | 0.95 | Allium | ursinum | 0.87 | B | 0 | 0 | 1 | Vaccinium | Allium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 353.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Zanthoxylum | americanum | Prunus | lusitanica | 1.36 | Digitalis | lanata | 1.35 | B | 0 | 0 | 1 | Prunus | Digitalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TX 353.2 | 3; botanical collection | Plant extr.-R | 2022-10-21 | Zanthoxylum | americanum | Solanum | betaceum | 0.98 | Artemisia | vulgaris | 0.81 | B | 0 | 0 | 1 | Solanum | Artemisia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 428 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Zingiber | officinale | Zingiber | officinale | 2.07 | Curcuma | longa | 1.29 | A | 1 | 1 | 0 | Zingiber | Curcuma | 0 | 0 | 0 | 0

## #Parameter (Spezies)
Parameter: | Allium | ursinum | Codierung | Validität | Methode | Datum | verwendeter Organismus | Score | richtig-negativ | falsch-positiv
Test auf spektrale Nähe zum Parameter |  |  | TX 003 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus precatorius | 0.28 | 1 | 0
Anzahl Isolate: | 578 |  | TX 006 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus precatorius | 0.76 | 1 | 0
|  |  | TX 379.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Aconitum carmichaelii | 0.93 | 1 | 0
richtig-negativ: | 578 |  | TX 038.1 | 3; determination by expert | Plant extr. | 2022-01-14 | Aconitum napellus | 0.69 | 1 | 0
|  |  | TX 038.2 | 3; determination by expert | Plant extr. | 2022-01-26 | Aconitum napellus | 0.91 | 1 | 0
falsch-positiv: | 0 |  | TX 055.1 | 2; determination by expert | Plant extr. | 2022-01-20 | Aconitum napellus | 0.76 | 1 | 0
|  |  | TX 076.1 | 3; determination by expert | Plant extr. | 2022-01-14 | Aconitum napellus | 0.14 | 1 | 0
| 0% |  | TX 089 | 2; determination by expert | Plant extr. | 2022-01-14 | Aconitum napellus | 0.68 | 1 | 0
|  |  | TX 122.1 | 2; determination by expert | Plant extr. | 2022-01-14 | Aconitum napellus | 0.54 | 1 | 0
|  |  | TX 037.1 | 3; determination by expert | Plant extr. | 2022-10-07 | Aethusa cynapium | 1.12 | 1 | 0
Scorewertgrenzen: | 2 | Species | TX 037.2 | 3; determination by expert | Plant extr. | 2022-01-25 | Aethusa cynapium | 0.84 | 1 | 0
| 1.7 | Genus | TX 854.1 | 3; label, sensory analysis | Plant extr. | 2023-11-23 | Allium fistulosum | 0.99 | 1 | 0
Es wird gegen eine nur aus dem Parameter bestehende DB geprüft. |  |  | TX 851 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium porrum | 0.52 | 1 | 0
|  |  | TX 082.1 | 3; label; sensorical analysis | Plant extr. | 2021-07-30 | Allium sativum | 0.9 | 1 | 0
|  |  | TX 082.2 | 3; label; sensorical analysis | Plant extr. | 2022-01-26 | Allium sativum | 0.81 | 1 | 0
|  |  | TX 084 | 3; determination by expert | Plant extr. | 2021-07-27 | Allium schoenoprasum | 1.32 | 1 | 0
|  |  | TX 171.1 | 3; determination by expert | Plant extr. | 2023-11-15 | Allium schoenoprasum | 1.34 | 1 | 0
|  |  | TX 848 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium schoenoprasum | 1.12 | 1 | 0
|  |  | TX 849 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium schoenoprasum | 0.82 | 1 | 0
|  |  | TX 856 | 3; label; sensory analysis | Plant extr. | 2023-11-24 | Allium schoenoprasum | 1.43 | 1 | 0
|  |  | TX 847 | 3; determination by expert | Plant extr. | 2023-11-17 | Allium tuberosum | 1.13 | 1 | 0
|  |  | TX 178 | 3; determination by expert | Plant extr. | 2023-11-15 | Allium vineale | 1.36 | 1 | 0
|  |  | TX 601.1 | 2; expert opinion | Plant extr.-R | 2022-12-21 | Aloe vera | 0.52 | 1 | 0
|  |  | TX 413 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Alpinia officinarum | 0.69 | 1 | 0
|  |  | TX 402.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Amaranthus sp | 0.94 | 1 | 0
|  |  | TX 457.1 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Ampelopsis sinica | 0.81 | 1 | 0
|  |  | TX 457.2 | 3; botanical collection | Plant extr.-R | 2022-11-02 | Ampelopsis sinica | 0.86 | 1 | 0
|  |  | TX 132.1 | 2; label | Plant extr. | 2022-02-02 | Anethum graveolens | 0.74 | 1 | 0
|  |  | TX 132.2 | 2; label | Plant extr. | 2022-01-14 | Anethum graveolens | 1.11 | 1 | 0
|  |  | TX 132.3 | 2; label | Plant extr. | 2022-01-14 | Anethum graveolens | 0.68 | 1 | 0
|  |  | TX 136 | 2; label | Plant extr. | 2022-01-14 | Anethum graveolens | 0.81 | 1 | 0
|  |  | TX 363 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Apocynum cannabinum | 0.67 | 1 | 0
|  |  | TX 488.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Arctostaphylos uva-ursi | 0.69 | 1 | 0
|  |  | TX 371 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Aristolochia clematitis | 0.69 | 1 | 0
|  |  | TX 372 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Aristolochia rotunda | 0.3 | 1 | 0
|  |  | TX 028 | 3; determination by expert | Plant extr. | 2021-07-20 | Artemisia absinthium | 1.15 | 1 | 0
|  |  | TX 073.2 | 3; determination by expert | Plant extr. | 2022-01-14 | Artemisia absinthum | 0.6 | 1 | 0
|  |  | TX 152.1 | 2; label | Plant extr. | 2022-02-04 | Artemisia dracunculus | 0.75 | 1 | 0
|  |  | TX 152.2 | 2; label | Plant extr. | 2022-02-25 | Artemisia dracunculus | 1.08 | 1 | 0
|  |  | TX 152.3 | 2; label | Plant extr. | 2022-02-21 | Artemisia dracunculus | 0.73 | 1 | 0
|  |  | TX 071 | 2; Label; determination by expert | Plant extr. | 2021-07-27 | Artemisia vulgaris | 0.79 | 1 | 0
|  |  | TX 072.2 | 2; Label; determination by expert | Plant extr.-R. | 2022-01-26 | Artemisia vulgaris | 0.68 | 1 | 0
|  |  | TX 367 | 3; botanical collection | Plant extr.-R | 2022-10-07 | Arum italicum | 0.52 | 1 | 0
|  |  | TX 015 | 2; determination by expert | Plant extr. | 2021-07-20 | Arum maculatum | 0.87 | 1 | 0
|  |  | TX 368 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Arum maculatum | 0.89 | 1 | 0
|  |  | TX 695 | 3; determination by expert | Plant extr. | 2023-11-15 | Arum maculatum | 0.6 | 1 | 0
|  |  | TX 699 | 3; determination by expert | Plant extr. | 2023-11-15 | Arum maculatum | 0.68 | 1 | 0
|  |  | TX 723 | 3; determination by expert | Plant extr. | 2023-11-08 | Arum maculatum | 0.67 | 1 | 0
|  |  | TX 764 | 3; determination by expert | Plant extr. | 2023-11-09 | Arum maculatum | 0.62 | 1 | 0
|  |  | TX 765 | 3; determination by expert | Plant extr. | 2023-11-09 | Arum maculatum | 0.68 | 1 | 0
|  |  | TX 770 | 3; determination by expert | Plant extr. | 2023-10-31 | Arum maculatum | 0.49 | 1 | 0
|  |  | TX 771 | 3; determination by expert | Plant extr. | 2023-10-31 | Arum maculatum | 0.71 | 1 | 0
|  |  | TX 779 | 3; determination by expert | Plant extr. | 2023-11-06 | Arum maculatum | 1.01 | 1 | 0
|  |  | TX 785 | 3; determination by expert | Plant extr. | 2023-11-09 | Arum maculatum | 0.41 | 1 | 0
|  |  | TX 241.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Arum sp | 0.58 | 1 | 0
|  |  | TX 056.1 | 2; label | Plant extr. | 2022-02-21 | Asarum europaeum | 1.24 | 1 | 0
|  |  | TX 056.2 | 2; label | Plant extr. | 2022-01-26 | Asarum europaeum | 0.55 | 1 | 0
|  |  | TX 357.1 | 3; botanical collection | Plant extr.-R | 2022-10-07 | Asimina triloba | 0.53 | 1 | 0
|  |  | TX 011 | 3; botanical collection | Plant extr. | 2022-01-19 | Atropa belladonna | 0.64 | 1 | 0
|  |  | TX 125.2 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Atropa belladonna | 0.39 | 1 | 0
|  |  | TX 375.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Atropa belladonna | 0.65 | 1 | 0
|  |  | TX 375.2 | 3; botanical collection | Plant extr.-R | 2022-10-19 | Atropa belladonna | 0.21 | 1 | 0
|  |  | TX 448 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Backhousia citriodora | 1.11 | 1 | 0
|  |  | TX 500 | 2; label; MALDI | Plant extr.-R | 2022-11-23 | Backhousia citriodora | 0.97 | 1 | 0
|  |  | TX 475.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Berberis thunbergii | 0.8 | 1 | 0
|  |  | TX 026 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Brugmansia pittieri | 1.22 | 1 | 0
|  |  | TX 027 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Brugmansia sanguinea | 0.9 | 1 | 0
|  |  | TX 120 | 3; botanical collection | Plant extr.-R | 2022-01-27 | Bryonia cretica | 0.51 | 1 | 0
|  |  | TX 356.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Bryonia cretica | 0.77 | 1 | 0
|  |  | TX 376 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Caltha palustris | 0.38 | 1 | 0
|  |  | TX 290 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Camellia sinensis | 0.94 | 1 | 0
|  |  | TX 456.1 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Camellia sinensis | 0.81 | 1 | 0
|  |  | TX 456.2 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Camellia sinensis | 1.11 | 1 | 0
|  |  | TX 456.3 | 3; botanical collection | Plant extr.-R | 2022-10-26 | Camellia sinensis | 0.59 | 1 | 0
|  |  | TX 471 | 1; label | Plant extr.-R | 2022-10-27 | Camellia sinensis | 0.79 | 1 | 0
|  |  | TX 472 | 1; label | Plant extr.-R | 2022-11-15 | Camellia sinensis | 1.13 | 1 | 0
|  |  | TX 473 | 1; label | Plant extr.-R | 2022-11-15 | Camellia sinensis | 0.75 | 1 | 0
|  |  | TX 499 | 1; label | Plant extr.-R | 2022-11-23 | Camellia sinensis | 1.17 | 1 | 0
|  |  | TX 552 | 1; label | Plant extr.-R | 2022-12-02 | Camellia sinensis | 0.98 | 1 | 0
|  |  | TX 567 | 1; label | Plant extr.-AceRL42S90 | 2022-12-12 | Camellia sinensis | 0.73 | 1 | 0
|  |  | TX 364 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Capsicum chinense | 0.99 | 1 | 0
|  |  | TX 355 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Chamaecyparis lawsoniana | 0.84 | 1 | 0
|  |  | TX 115.2 | 3; botanical collection | Plant extr.-R | 2022-01-27 | Chelidonium majus | 0.51 | 1 | 0
|  |  | TX 046 | 3; botanical collection | Plant extr.-R | 2022-01-28 | Cicuta virosa | 0.37 | 1 | 0
|  |  | TX 412 | 3; botanical collection | Plant extr.-R | 2022-11-04 | Cinnamomum camphora | 0.71 | 1 | 0
|  |  | TX 410 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Cinnamomum pittosporoides | 0.16 | 1 | 0
|  |  | TX 408 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Cinnamomum verum | 0.46 | 1 | 0
|  |  | TX 138 | 2; expert opinion | Plant extr.-R | 2022-02-07 | Cistus sp | 0.71 | 1 | 0
|  |  | TX 263 | 3; botanical collection | Plant extr. | 2022-09-16 | Citrus cavaleriei | 0.46 | 1 | 0
|  |  | TX 078 | 2; label; determination by expert | Plant extr. | 2021-07-30 | Clinopodium nepeta | 0.88 | 1 | 0
|  |  | TX 251 | 3; botanical collection | Plant extr. | 2022-09-15 | Coffea arabica | 0.8 | 1 | 0
|  |  | TX 254 | 3; botanical collection | Plant extr. | 2022-09-15 | Coffea arabica | 0.84 | 1 | 0
|  |  | TX 256.1 | 3; botanical collection | Plant extr. | 2022-09-14 | Coffea arabica | 0.83 | 1 | 0
|  |  | TX 256.2 | 3; botanical collection | Plant extr. | 2022-10-04 | Coffea arabica | 1.02 | 1 | 0
|  |  | TX 260.1 | 3; botanical collection | Plant extr. | 2022-09-14 | Coffea arabica | 0.44 | 1 | 0
|  |  | TX 260.2 | 3; botanical collection | Plant extr. | 2022-09-29 | Coffea arabica | 0.42 | 1 | 0
|  |  | TX 328.1 | 3; botanical collection | Plant extr. | 2022-09-27 | Coffea arabica | 0.53 | 1 | 0
|  |  | TX 440 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Coffea arabica | 0.83 | 1 | 0
|  |  | TX 453 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Coffea arabica | 0.71 | 1 | 0
|  |  | TX 351 | 3; botanical collection | Plant extr. | 2022-09-28 | Coffea canephora | 0.77 | 1 | 0
|  |  | TX 018 | 3; determination by expert | Plant extr. | 2021-07-20 | Colchicum autumnale | 0.56 | 1 | 0
|  |  | TX 110.1 | 3; PCR | Plant extr. | 2023-11-17 | Colchicum autumnale | 1.07 | 1 | 0
|  |  | TX 110.3 | 3; PCR | Plant extr. | 2023-11-30 | Colchicum autumnale | 0.98 | 1 | 0
|  |  | TX 226.4 | 3; PCR | Plant extr. | 2023-11-09 | Colchicum autumnale | 1.19 | 1 | 0
|  |  | TX 226.5 | 3; PCR | Plant extr. | 2023-11-29 | Colchicum autumnale | 0.9 | 1 | 0
|  |  | TX 227.1 | 3; PCR | Plant extr. | 2023-11-29 | Colchicum autumnale | 0.84 | 1 | 0
|  |  | TX 227 | 3; PCR | Plant extr. | 2023-11-09 | Colchicum autumnale | 0.97 | 1 | 0
|  |  | TX 244.1 | 3; determination by expert | Plant extr. | 2023-11-30 | Colchicum autumnale | 0.69 | 1 | 0
|  |  | TX 244 | 3; determination by expert | Plant extr. | 2022-09-13 | Colchicum autumnale | 0.74 | 1 | 0
|  |  | TX 701.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum autumnale | 0.74 | 1 | 0
|  |  | TX 701 | 3; determination by expert | Plant extr. | 2023-11-17 | Colchicum autumnale | 0.73 | 1 | 0
|  |  | TX 724.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum autumnale | 1.04 | 1 | 0
|  |  | TX 724 | 3; determination by expert | Plant extr. | 2023-11-08 | Colchicum autumnale | 0.7 | 1 | 0
|  |  | TX 768.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum autumnale | 0.99 | 1 | 0
|  |  | TX 768 | 3; determination by expert | Plant extr. | 2023-11-09 | Colchicum autumnale | 0.72 | 1 | 0
|  |  | TX 769.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum autumnale | 0.86 | 1 | 0
|  |  | TX 769 | 3; determination by expert | Plant extr. | 2023-10-31 | Colchicum autumnale | 0.84 | 1 | 0
|  |  | TX 777.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum autumnale | 0.88 | 1 | 0
|  |  | TX 777 | 3; determination by expert | Plant extr. | 2023-11-06 | Colchicum autumnale | 1.1 | 1 | 0
|  |  | TX 778.1 | 3; determination by expert | Plant extr. | 2023-11-29 | Colchicum autumnale | 0.73 | 1 | 0
|  |  | TX 778 | 3; determination by expert | Plant extr. | 2023-11-06 | Colchicum autumnale | 0.77 | 1 | 0
|  |  | TX 041.1 | 3; determination by expert | Plant extr. | 2022-02-21 | Conium maculatum | 0.92 | 1 | 0
|  |  | TX 041.2 | 3; determination by expert | Plant extr. | 2022-01-27 | Conium maculatum | 1.01 | 1 | 0
|  |  | TX 021.1 | 3; determination by expert | Plant extr. | 2021-07-20 | Convallaria majalis | 0.64 | 1 | 0
|  |  | TX 059 | 3; label; sensory analysis | Plant extr. | 2022-01-21 | Convallaria majalis | 0.76 | 1 | 0
|  |  | TX 063 | 3; label; sensory analysis | Plant extr. | 2023-11-16 | Convallaria majalis | 0.79 | 1 | 0
|  |  | TX 128 | 3; botanical collection | Plant extr. | 2023-11-27 | Convallaria majalis | 0.81 | 1 | 0
|  |  | TX 197 | 3; determination by expert | Plant extr. | 2022-09-09 | Convallaria majalis | 0.88 | 1 | 0
|  |  | TX 220 | 3; determination by expert | Plant extr. | 2022-09-09 | Convallaria majalis | 0.66 | 1 | 0
|  |  | TX 242.1 | 3; determination by expert | Plant extr. | 2022-09-09 | Convallaria majalis | 0.88 | 1 | 0
|  |  | TX 559.1 | 3; determination by expert | Plant extr. | 2023-11-17 | Convallaria majalis | 1.42 | 1 | 0
|  |  | TX 725.1 | 3; label, sensory analysis | Plant extr. | 2023-11-08 | Convallaria majalis | 1.09 | 1 | 0
|  |  | TX 752.1 | 3; determination by expert | Plant extr. | 2023-11-17 | Convallaria majalis | 1.1 | 1 | 0
|  |  | TX 756.1 | 3; determination by expert | Plant extr. | 2023-11-17 | Convallaria majalis | 0.96 | 1 | 0
|  |  | TX 759.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria majalis | 0.59 | 1 | 0
|  |  | TX 760.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria majalis | 0.57 | 1 | 0
|  |  | TX 761.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria majalis | 0.4 | 1 | 0
|  |  | TX 762.2 | 3; determination by expert | Plant extr. | 2023-10-30 | Convallaria majalis | 0.53 | 1 | 0
|  |  | TX 774 | 3; determination by expert | Plant extr. | 2023-10-31 | Convallaria majalis | 0.74 | 1 | 0
|  |  | TX 775 | 3; determination by expert | Plant extr. | 2023-10-31 | Convallaria majalis | 0.57 | 1 | 0
|  |  | TX 776 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria majalis | 1.31 | 1 | 0
|  |  | TX 781 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria majalis | 1.16 | 1 | 0
|  |  | TX 782 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria majalis | 1.28 | 1 | 0
|  |  | TX 783 | 3; determination by expert | Plant extr. | 2023-11-06 | Convallaria majalis | 1.44 | 1 | 0
|  |  | TX 784 | 3; determination by expert | Plant extr. | 2023-11-09 | Convallaria majalis | 0.93 | 1 | 0
|  |  | TX 153.1 | 3; botanical collection | Plant extr. | 2022-02-04 | Coriandrum sativum | 0.76 | 1 | 0
|  |  | TX 153.2 | 3; botanical collection | Plant extr. | 2022-02-04 | Coriandrum sativum | 0.76 | 1 | 0
|  |  | TX 153.3 | 3; botanical collection | Plant extr. | 2022-02-25 | Coriandrum sativum | 0.9 | 1 | 0
|  |  | TX 153.4 | 3; botanical collection | Plant extr. | 2022-02-21 | Coriandrum sativum | 0.88 | 1 | 0
|  |  | TX 154 | 2; label | Plant extr. | 2022-01-21 | Coriandrum sativum | 0.78 | 1 | 0
|  |  | TX 422 | 3; botanical collection | Plant extr.-R | 2022-11-02 | Costus malortieanus | 0.96 | 1 | 0
|  |  | TX 382.1 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Cotoneaster ballatus | 0.9 | 1 | 0
|  |  | TX 397.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Crataegus monogyna | 0.9 | 1 | 0
|  |  | TX 397.2 | 2; expert opinion | Plant extr.-R-A | 2022-10-18 | Crataegus monogyna | 0.83 | 1 | 0
|  |  | TX 204.1 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Crataegus sp | 0.43 | 1 | 0
|  |  | TX 212 | 2; determination by expert | Plant extr. | 2022-09-09 | Cunninghamia konishii | 0.83 | 1 | 0
|  |  | TX 432 | 3; botanical collection | Plant extr.-R | 2022-11-04 | Curcuma longa | 0.77 | 1 | 0
|  |  | TX 394.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Datura innoxia | 0.75 | 1 | 0
|  |  | TX 034 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Datura stramonium | 1.02 | 1 | 0
|  |  | TX 378 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Datura stramonium | 1.12 | 1 | 0
|  |  | TX 393.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Datura stramonium | 0.66 | 1 | 0
|  |  | TX 377 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Digitalis grandiflora | 1.08 | 1 | 0
|  |  | TX 380 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Digitalis lanata | 1 | 1 | 0
|  |  | TX 039.1 | 3; determination by expert | Plant extr. | 2022-01-20 | Digitalis purpurea | 0.97 | 1 | 0
|  |  | TX 039.2 | 3; determination by expert | Plant extr. | 2022-01-20 | Digitalis purpurea | 0.85 | 1 | 0
|  |  | TX 054.1 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Digitalis purpurea | 0.58 | 1 | 0
|  |  | TX 054.2 | 2; determination by expert | Plant extr. | 2022-01-14 | Digitalis purpurea | 0.27 | 1 | 0
|  |  | TX 060 | 2; label | Plant extr. | 2022-01-14 | Digitalis purpurea | 0.75 | 1 | 0
|  |  | TX 061.1 | 2; label | Plant extr. | 2022-01-14 | Digitalis purpurea | 0.97 | 1 | 0
|  |  | TX 061.2 | 2; label | Plant extr. | 2022-01-27 | Digitalis purpurea | 0.66 | 1 | 0
|  |  | TX 062 | 2; label | Plant extr. | 2022-01-14 | Digitalis purpurea | 0.8 | 1 | 0
|  |  | TX 217 | 2; determination by expert | Plant extr. | 2022-09-09 | Digitalis purpurea | 0.68 | 1 | 0
|  |  | TX 103 | 2; expert opinion | Plant extr.-R | 2022-10-10 | Dryopteris filix-mas | 0.33 | 1 | 0
|  |  | TX 209 | 2; determination by expert | Plant extr. | 2022-09-09 | Echium candicans | 1.05 | 1 | 0
|  |  | TX 208 | 2; determination by expert | Plant extr. | 2022-09-09 | Echium sp | 0.71 | 1 | 0
|  |  | TX 210 | 2; determination by expert | Plant extr. | 2022-09-09 | Echium sp | 1.06 | 1 | 0
|  |  | TX 022 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Erythroxylum coca | 1.19 | 1 | 0
|  |  | TX 111 | 2; expert opinion | Plant extr. | 2022-01-28 | Euonymus europaeus | 0.68 | 1 | 0
|  |  | TX 113.2 | 2; expert opinion | Plant extr. | 2022-01-21 | Euonymus europaeus | 0.43 | 1 | 0
|  |  | TX 459.1 | 2; expert opinion | Plant extr.-R | 2022-10-28 | Euonymus europaeus | 0.45 | 1 | 0
|  |  | TX 481.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Euonymus europaeus | 0.81 | 1 | 0
|  |  | TX 477.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Euonymus fortunei | 0.85 | 1 | 0
|  |  | TX 166.1 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus nivalis | 0.43 | 1 | 0
|  |  | TX 690 | 3; determination by expert | Plant extr. | 2023-11-27 | Galanthus nivalis | 0.89 | 1 | 0
|  |  | TX 184.1 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus sp | 0.84 | 1 | 0
|  |  | TX 236.1 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus sp | 0.54 | 1 | 0
|  |  | TX 236.2 | 2; determination by expert | Plant extr. | 2022-09-07 | Galanthus sp | 0.54 | 1 | 0
|  |  | TX 086 | 1; determination by expert | Plant extr. | 2021-07-27 | Galium mollugo | 0.74 | 1 | 0
|  |  | TX 085 | 2; determination by expert | Plant extr. | 2021-07-27 | Galium odoratum | 0.71 | 1 | 0
|  |  | TX 620 | 1; label | Plant extr.-AceR | 2023-01-10 | Ganoderma lucidum | 0 | 1 | 0
|  |  | TX 087.2 | 2; expert opinion | Plant extr.-R. | 2022-01-26 | Gentiana lutea | 0.75 | 1 | 0
|  |  | TX 214 | 2; determination by expert | Plant extr. | 2022-09-09 | Gunnera sp | 1.06 | 1 | 0
|  |  | TX 391.1 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Heliotropium arborescens | 0.54 | 1 | 0
|  |  | TX 066.2 | 2; label | Plant extr.-R. | 2022-01-26 | Helleborus niger | 0.79 | 1 | 0
|  |  | TX 067.1 | 2; label | Plant extr. | 2022-01-14 | Helleborus niger | 0.55 | 1 | 0
|  |  | TX 068.1 | 2; label | OS extr.-Pl | 2018-01-23 | Helleborus niger | 0.71 | 1 | 0
|  |  | TX 118 | 3; botanical collection | OS extr.-Pl | 2018-01-23 | Helleborus niger | 0.64 | 1 | 0
|  |  | TX 119 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Humulus lupulus | 0.95 | 1 | 0
|  |  | TX 043 | 3; botanical collection | Plant extr. | 2022-01-19 | Hyoscyamus niger | 0.37 | 1 | 0
|  |  | TX 206 | 1; determination by expert | Plant extr. | 2022-09-09 | Jacobaea maritima | 0.82 | 1 | 0
|  |  | TX 245.1 | 2; determination by expert | Plant extr. | 2022-09-13 | Jacobaea maritima | 0.63 | 1 | 0
|  |  | PL 0047 | 2; label; sensorical analysis | OS extr.-Pl | 2018-01-23 | Juglans regia | 0.32 | 1 | 0
|  |  | PL 0048 | 2; label; sensorical analysis | OS extr.-Pl | 2018-01-23 | Juglans regia | 0 | 1 | 0
|  |  | TX 237 | 2; determination by expert | Plant extr. | 2022-09-09 | Juglans regia | 0.78 | 1 | 0
|  |  | TX 042.1 | 3; botanical collection | Plant extr.-R | 2022-02-21 | Juniperus sabina | 0.83 | 1 | 0
|  |  | TX 042.2 | 3; botanical collection | Plant extr. | 2022-01-18 | Juniperus sabina | 0.35 | 1 | 0
|  |  | TX 261 | 3; botanical collection | Plant extr. | 2022-09-16 | Justicia adhotoda | 0.78 | 1 | 0
|  |  | TX 047 | 3; botanical collection | Plant extr. | 2022-01-19 | Laburnum anagyroides | 0.14 | 1 | 0
|  |  | TX 231 | 2; determination by expert | Plant extr. | 2022-09-06 | Laburnum sp | 0.86 | 1 | 0
|  |  | TX 145 | 2; determination by expert | Plant extr.-R | 2022-02-03 | Laurus nobilis | 1.03 | 1 | 0
|  |  | TX 269 | 3; botanical collection | Plant extr. | 2022-09-16 | Laurus nobilis | 0.4 | 1 | 0
|  |  | TX 312 | 3; botanical collection | Plant extr. | 2022-09-15 | Lemna minor | 0.7 | 1 | 0
|  |  | TX 438 | 3; botanical collection | Plant extr.-R | 2022-10-25 | Leptospermum citratum | 0.74 | 1 | 0
|  |  | TX 698 | 3; determination by expert | Plant extr. | 2023-11-27 | Leucojum vernum | 0.42 | 1 | 0
|  |  | TX 404.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Ligustrum vulgare | 1.02 | 1 | 0
|  |  | TX 404.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Ligustrum vulgare | 0.72 | 1 | 0
|  |  | TX 460.1 | 2; expert opinion | Plant extr.-R | 2022-10-28 | Ligustrum vulgare | 0.84 | 1 | 0
|  |  | TX 491.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Ligustrum vulgare | 0.81 | 1 | 0
|  |  | TX 284.2 | 2; botanical collection | Plant extr. | 2022-10-04 | Lonicera sp | 0.84 | 1 | 0
|  |  | TX 023 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Lophophora williamsi | 1.09 | 1 | 0
|  |  | PL 0001 | 2; label; sensorical analysis | Osextr-Pl | 2017-11-03 | Lupinus sp | 0.34 | 1 | 0
|  |  | TX 443.1 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium barbarum | 0.99 | 1 | 0
|  |  | TX 443.2 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium barbarum | 0.39 | 1 | 0
|  |  | TX 443.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium barbarum | 0.12 | 1 | 0
|  |  | TX 447.1 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium chinense | 0.88 | 1 | 0
|  |  | TX 447.2 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium chinense | 0.47 | 1 | 0
|  |  | TX 447.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium chinense | 0.67 | 1 | 0
|  |  | TX 449.1 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium europaeum | 0.49 | 1 | 0
|  |  | TX 449.2 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium europaeum | 0.6 | 1 | 0
|  |  | TX 449.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium europaeum | 0.21 | 1 | 0
|  |  | TX 474.2 | 1; label | Plant extr.-R | 2022-11-21 | Lycium sp | 0.33 | 1 | 0
|  |  | TX 496.1 | 1; label | Plant extr.-R | 2022-11-23 | Lycium sp | 0.5 | 1 | 0
|  |  | TX 496.2 | 1; label | Plant extr.-R | 2022-11-23 | Lycium sp | 0.18 | 1 | 0
|  |  | TX 390.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Malus sylvestris | 0.63 | 1 | 0
|  |  | TX 124 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Mandragora officinarum | 0.51 | 1 | 0
|  |  | TX 277 | 3; botanical collection | Plant extr. | 2023-11-27 | Mansoa alliacea | 1.08 | 1 | 0
|  |  | TX 058.1 | 2; determination by expert | Plant extr. | 2022-02-21 | Mentha pulegium | 0.92 | 1 | 0
|  |  | TX 058.2 | 2; determination by expert | Plant extr. | 2022-01-18 | Mentha pulegium | 0.17 | 1 | 0
|  |  | TX 213 | 2; determination by expert | Plant extr. | 2022-09-09 | Metasequoia glyptostroboides | 0.58 | 1 | 0
|  |  | TX 257.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Microcitrus australasia | 0.58 | 1 | 0
|  |  | TX 455 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Moringa stenopetala | 1.06 | 1 | 0
|  |  | TX 012 | 2; determination by expert | Plant extr. | 2021-07-15 | Nerium oleander | 0.44 | 1 | 0
|  |  | TX 024 | 3; botanical collection | Plant extr.-R | 2022-10-10 | Nerium oleander | 0.87 | 1 | 0
|  |  | TX 360 | 3; botanical collection | Plant extr.-R | 2022-10-07 | Nicotiana sylvestris | 0.5 | 1 | 0
|  |  | TX 216 | 2; determination by expert | Plant extr. | 2022-09-09 | Nothofagus allessandrii | 1.01 | 1 | 0
|  |  | TX 131.1 | 2; label | Plant extr. | 2022-02-01 | Ocimum basilicum | 0.54 | 1 | 0
|  |  | TX 131.2 | 2; label | Plant extr. | 2022-02-03 | Ocimum basilicum | 0.84 | 1 | 0
|  |  | TX 131.3 | 2; label | Plant extr. | 2022-02-21 | Ocimum basilicum | 0.79 | 1 | 0
|  |  | TX 134 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Ocimum basilicum | 1.13 | 1 | 0
|  |  | TX 139 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Olea europaea | 0.3 | 1 | 0
|  |  | TX 135 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Origanum vulgare | 0.51 | 1 | 0
|  |  | TX 147.1 | 2; determination by expert | Plant extr. | 2022-02-03 | Origanum vulgare | 0.51 | 1 | 0
|  |  | TX 147.2 | 2; determination by expert | Plant extr. | 2022-02-03 | Origanum vulgare | 0.99 | 1 | 0
|  |  | TX 147.3 | 2; determination by expert | Plant extr. | 2022-02-21 | Origanum vulgare | 0.88 | 1 | 0
|  |  | TX 148 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Origanum vulgare | 0 | 1 | 0
|  |  | TX 036.3 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Papaver somniferum | 0.8 | 1 | 0
|  |  | TX 270.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Passiflora sp | 0.54 | 1 | 0
|  |  | TX 255 | 3; botanical collection | Plant extr. | 2022-09-16 | Pelargonium sp | 0.58 | 1 | 0
|  |  | TX 020.1 | 2; determination by expert | Plant extr. | 2022-01-24 | Petasites hybridus | 0.7 | 1 | 0
|  |  | TX 020.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Petasites hybridus | 0.8 | 1 | 0
|  |  | TX 081 | 3; label; determination by expert | Plant extr. | 2021-07-27 | Petroselinum crispum | 0.59 | 1 | 0
|  |  | TX 130.1 | 2; label | Plant extr. | 2022-02-01 | Petroselinum crispum | 0.61 | 1 | 0
|  |  | TX 130.2 | 2; label | Plant extr. | 2022-02-01 | Petroselinum crispum | 1.11 | 1 | 0
|  |  | TX 130.3 | 2; label | Plant extr. | 2022-02-21 | Petroselinum crispum | 0.95 | 1 | 0
|  |  | TX 130.4 | 2; label | Plant extr. | 2022-02-21 | Petroselinum crispum | 0.97 | 1 | 0
|  |  | TX 130.5 | 2; label | Plant extr. | 2022-02-21 | Petroselinum crispum | 0.96 | 1 | 0
|  |  | TX 137 | 2; determination by expert | Plant extr.-R | 2022-02-07 | Petroselinum crispum | 0.94 | 1 | 0
|  |  | TX 150.1 | 2; label | Plant extr. | 2022-02-04 | Petroselinum crispum | 0.79 | 1 | 0
|  |  | TX 150.2 | 2; label | Plant extr. | 2022-02-04 | Petroselinum crispum | 0.72 | 1 | 0
|  |  | TX 151.1 | 2; label | Plant extr. | 2022-02-04 | Petroselinum crispum | 0.57 | 1 | 0
|  |  | TX 151.2 | 2; label | Plant extr. | 2022-02-04 | Petroselinum crispum | 0.84 | 1 | 0
|  |  | TX 162 | 2; label | Plant extr.-R | 2022-02-23 | Petroselinum crispum | 0.54 | 1 | 0
|  |  | PL 0130 | 3; label; sensorical analysis | Osextr-Pl | 2020-07-17 | Phaseolus vulgaris | 0.39 | 1 | 0
|  |  | TX 370 | 3; botanical collection | Plant extr.-R | 2022-01-20 | Pinellia ternata | 0.8 | 1 | 0
|  |  | TX 215 | 2; determination by expert | Plant extr. | 2022-09-09 | Pinus montezumae | 0.88 | 1 | 0
|  |  | TX 418 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper auritum | 0.96 | 1 | 0
|  |  | TX 414 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper betle | 0.67 | 1 | 0
|  |  | TX 411 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper cubeba | 0.58 | 1 | 0
|  |  | TX 416 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Piper nigrum | 0.36 | 1 | 0
|  |  | TX 463 | 1; label | Plant extr.-R | 2022-10-28 | Piper sarmentosum | 0.89 | 1 | 0
|  |  | TX 014.1 | 3; determination by expert, PCR on genus level | Plant extr. | 2022-02-23 | Polygonatum multiflorum | 0.43 | 1 | 0
|  |  | TX 014.2 | 2; determination by expert | Plant extr. | 2022-01-27 | Polygonatum multiflorum | 0.46 | 1 | 0
|  |  | TX 108 | 2; determination by expert | Plant extr. | 2022-01-18 | Polygonatum multiflorum | 0.61 | 1 | 0
|  |  | TX 121 | 3; botanical collection | Plant extr.-R | 2022-01-18 | Polygonatum odoratum | 0.65 | 1 | 0
|  |  | PL 0113 | 2; label, sensory analysis | OS Extr. plant | 2018-11-22 | Prunus armeniaca | 0.56 | 1 | 0
|  |  | TX 354 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Prunus laurocerasus | 0.48 | 1 | 0
|  |  | TX 396.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Prunus laurocerasus | 0.74 | 1 | 0
|  |  | TX 396.3 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Prunus laurocerasus | 0.84 | 1 | 0
|  |  | TX 384.1 | 2; expert opinion | Plant extr.-R | 2022-10-06 | Prunus lusitanica | 0.78 | 1 | 0
|  |  | TX 384.2 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Prunus lusitanica | 0.64 | 1 | 0
|  |  | TX 399.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Prunus spinosa | 0.64 | 1 | 0
|  |  | TX 232 | 2; determination by expert | Plant extr. | 2022-09-09 | Pteridium aquilinum | 0.77 | 1 | 0
|  |  | TX 267 | 3; botanical collection | Plant extr. | 2022-09-16 | Punica granatum | 0.77 | 1 | 0
|  |  | TX 425.1 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Rauvolfia tetraphylla | 0.97 | 1 | 0
|  |  | TX 425.2 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Rauvolfia tetraphylla | 0.63 | 1 | 0
|  |  | TX 405.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rhamnus cathartica | 0.61 | 1 | 0
|  |  | TX 405.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rhamnus cathartica | 0.04 | 1 | 0
|  |  | TX 599.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Ribes nigrum | 0.43 | 1 | 0
|  |  | TX 599.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes nigrum | 0.22 | 1 | 0
|  |  | TX 603.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-21 | Ribes nigrum | 0.65 | 1 | 0
|  |  | TX 603.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-21 | Ribes nigrum | 0 | 1 | 0
|  |  | TX 581.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-19 | Ribes rubrum | 0 | 1 | 0
|  |  | TX 581.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-19 | Ribes rubrum | 0.41 | 1 | 0
|  |  | TX 582.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-16 | Ribes rubrum | 0.24 | 1 | 0
|  |  | TX 582.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-19 | Ribes rubrum | 0.2 | 1 | 0
|  |  | TX 584.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.4 | 1 | 0
|  |  | TX 585.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.5 | 1 | 0
|  |  | TX 587.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.29 | 1 | 0
|  |  | TX 588.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.33 | 1 | 0
|  |  | TX 589.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-08 | Ribes rubrum | 0.6 | 1 | 0
|  |  | TX 590.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.4 | 1 | 0
|  |  | TX 591.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.11 | 1 | 0
|  |  | TX 593.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.72 | 1 | 0
|  |  | TX 594.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.7 | 1 | 0
|  |  | TX 600.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Ribes rubrum | 0.52 | 1 | 0
|  |  | TX 600.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes rubrum | 0.75 | 1 | 0
|  |  | TX 001 | 3; determination by expert | Plant extr. | 2021-07-22 | Ricinus communis | 0.08 | 1 | 0
|  |  | TX 002 | 3; determination by expert | Plant extr. | 2021-07-22 | Ricinus communis | 0.7 | 1 | 0
|  |  | TX 004 | 2; determination by expert | Plant extr. | 2021-07-22 | Ricinus communis | 0.55 | 1 | 0
|  |  | TX 005 | 3; determination by expert | Plant extr. | 2021-07-22 | Ricinus communis | 0.49 | 1 | 0
|  |  | TX 099.1 | 2; determination by expert | Plant extr. | 2022-01-17 | Ricinus communis | 0.59 | 1 | 0
|  |  | TX 099.2 | 2; determination by expert | Plant extr. | 2022-01-21 | Ricinus communis | 0.57 | 1 | 0
|  |  | TX 112 | 2; determination by expert | Plant extr. | 2022-01-17 | Ricinus communis | 0.77 | 1 | 0
|  |  | TX 341.1 | 3; botanical collection | Plant extr. | 2022-09-28 | Ricinus communis | 0.87 | 1 | 0
|  |  | TX 392 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Robinia pseudoacacia | 0.72 | 1 | 0
|  |  | TX 387.1 | 2; expert opinion | Plant extr.-R | 2022-10-26 | Rosa multiflora | 0.35 | 1 | 0
|  |  | TX 387.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Rosa multiflora | 0.36 | 1 | 0
|  |  | TX 398.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rosa multiflora | 1.01 | 1 | 0
|  |  | TX 398.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Rosa multiflora | 0.63 | 1 | 0
|  |  | TX 163 | 2; label | Plant extr. | 2022-02-23 | Rubus idaeus | 0.97 | 1 | 0
|  |  | TX 053.1 | 2; determination by expert | Plant extr.-R | 2022-01-24 | Salvia pratensis | 0.8 | 1 | 0
|  |  | TX 053.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Salvia pratensis | 0.36 | 1 | 0
|  |  | TX 133 | 2; determination by expert | Plant extr.-R | 2022-02-02 | Salvia rosmarinus | 0.53 | 1 | 0
|  |  | TX 144 | 2; expert opinion | Plant extr. | 2022-10-04 | Salvia sp | 0.95 | 1 | 0
|  |  | TX 266 | 3; botanical collection | Plant extr. | 2022-09-16 | Sambucus nigra | 0.87 | 1 | 0
|  |  | TX 395.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Sambucus nigra | 0.56 | 1 | 0
|  |  | TX 395.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Sambucus nigra | 0.67 | 1 | 0
|  |  | TX 129.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Senecio inaequidens | 0.19 | 1 | 0
|  |  | TX 049.1 | 2; determination by expert | Plant extr. | 2022-01-14 | Senecio jacobaea | 0.95 | 1 | 0
|  |  | TX 049.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Senecio jacobaea | 0.91 | 1 | 0
|  |  | TX 050.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Senecio jacobaea | 0.84 | 1 | 0
|  |  | TX 050.3 | 2; determination by expert | Plant extr. | 2022-01-19 | Senecio jacobaea | 0.65 | 1 | 0
|  |  | TX 050.4 | 2; determination by expert | Plant extr.-R | 2022-01-26 | Senecio jacobaea | 0.4 | 1 | 0
|  |  | TX 211 | 2; determination by expert | Plant extr. | 2022-09-09 | Senecio viravira | 0.37 | 1 | 0
|  |  | TX 048.2 | 2; determination by expert | Plant extr.-R | 2022-01-26 | Senecio vulgaris | 0.63 | 1 | 0
|  |  | TX 218 | 2; determination by expert | Plant extr. | 2022-09-09 | Sequoiadendron sp | 0.87 | 1 | 0
|  |  | TX 243.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Silphium perfoliatum | 0.93 | 1 | 0
|  |  | TX 243.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Silphium perfoliatum | 0.32 | 1 | 0
|  |  | TX 352.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Solanum betaceum | 0.95 | 1 | 0
|  |  | TX 238.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Solanum dulcamara | 0.72 | 1 | 0
|  |  | TX 238.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Solanum dulcamara | 0.63 | 1 | 0
|  |  | TX 383.1 | 3; botanical collection | Plant extr.-R | 2022-10-06 | Solanum dulcamara | 0.96 | 1 | 0
|  |  | TX 383.2 | 3; botanical collection | Plant extr. | 2022-09-30 | Solanum dulcamara | 0.38 | 1 | 0
|  |  | TX 100.1 | 2; determination by expert | Plant extr. | 2022-02-21 | Solanum nigrum | 0.61 | 1 | 0
|  |  | TX 100.2 | 2; determination by expert | Plant extr. | 2022-01-18 | Solanum nigrum | 0.71 | 1 | 0
|  |  | TX 109.1 | 3; determination by expert, PCR on genus level | Plant extr. | 2022-02-21 | Solanum nigrum | 0.93 | 1 | 0
|  |  | TX 109.2 | 3; determination by expert, PCR on genus level | Plant extr. | 2022-01-18 | Solanum nigrum | 0.45 | 1 | 0
|  |  | TX 114.1 | 2; determination by expert | Plant extr. | 2022-02-21 | Solanum nigrum | 0.82 | 1 | 0
|  |  | TX 114.2 | 2; determination by expert | Plant extr. | 2022-01-17 | Solanum nigrum | 0.8 | 1 | 0
|  |  | TX 114.3 | 2; determination by expert | Plant extr. | 2022-01-18 | Solanum nigrum | 0.72 | 1 | 0
|  |  | TX 264.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Solanum nigrum | 0.7 | 1 | 0
|  |  | TX 235 | 2; determination by expert | Plant extr. | 2022-09-07 | Solanum tuberosum | 0.93 | 1 | 0
|  |  | TX 365 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Solanum virginianum | 0.66 | 1 | 0
|  |  | TX 221.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Spartium junceum | 0.91 | 1 | 0
|  |  | TX 221.3 | 2; determination by expert | Plant extr. | 2022-09-09 | Spartium junceum | 0.54 | 1 | 0
|  |  | TX 219 | 2; determination by expert | Plant extr. | 2022-09-05 | Spinacia oleracea | 1.03 | 1 | 0
|  |  | TX 386 | 2; expert opinion | Plant extr.-R | 2022-10-07 | Spinacia oleracea | 1.04 | 1 | 0
|  |  | TX 388.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Symphoricarpos albus | 0.96 | 1 | 0
|  |  | TX 401.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Symphoricarpos albus | 0.95 | 1 | 0
|  |  | TX 401.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Symphoricarpos albus | 0.2 | 1 | 0
|  |  | TX 476.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Symphoricarpos orbiculatus | 0.8 | 1 | 0
|  |  | TX 482.1 | 2; expert opinion | Plant extr.-R | 2022-11-15 | Symphoricarpos orbiculatus | 0.75 | 1 | 0
|  |  | TX 008 | 2; determination by expert | Plant extr. | 2022-01-20 | Taxus baccata | 0.47 | 1 | 0
|  |  | TX 009 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus baccata | 0.42 | 1 | 0
|  |  | TX 010 | 2; determination by expert | Plant extr. | 2022-01-19 | Taxus baccata | 0.71 | 1 | 0
|  |  | TX 091 | 2; determination by expert | Plant extr. | 2022-01-19 | Taxus baccata | 0.87 | 1 | 0
|  |  | TX 092 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus baccata | 0.99 | 1 | 0
|  |  | TX 095.1 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus baccata | 0.84 | 1 | 0
|  |  | TX 096.1 | 2; determination by expert | Plant extr. | 2022-01-28 | Taxus baccata | 0.96 | 1 | 0
|  |  | TX 096.3 | 2; determination by expert | Plant extr. | 2022-01-18 | Taxus baccata | 0.73 | 1 | 0
|  |  | TX 096.4 | 2; determination by expert | Plant extr. | 2022-02-21 | Taxus baccata | 0.26 | 1 | 0
|  |  | TX 389.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Taxus baccata | 1.23 | 1 | 0
|  |  | TX 389.2 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Taxus baccata | 0.72 | 1 | 0
|  |  | TX 406 | 2; expert opinion | Plant extr.-R | 2022-10-19 | Taxus baccata | 0.42 | 1 | 0
|  |  | TX 303 | 3; determination by expert | Plant extr. | 2023-11-23 | Tulbaghia violacea | 0.9 | 1 | 0
|  |  | TX 224.1 | 2; determination by expert | Plant extr. | 2022-09-09 | Ulex europaeus | 1.06 | 1 | 0
|  |  | TX 224.2 | 2; determination by expert | Plant extr. | 2022-09-09 | Ulex europaeus | 0.84 | 1 | 0
|  |  | TX 664 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium angustifolium | 0.6 | 1 | 0
|  |  | TX 665 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium angustifolium | 0.45 | 1 | 0
|  |  | TX 666 | 3; botanical collection | Plant extr.-R | 2023-12-11 | Vaccinium angustifolium | 0.42 | 1 | 0
|  |  | TX 667 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium angustifolium | 0.82 | 1 | 0
|  |  | TX 668 | 3; botanical collection | Plant extr.-R | 2023-12-13 | Vaccinium angustifolium | 0.32 | 1 | 0
|  |  | TX 362.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vaccinium corymbosum | 0.93 | 1 | 0
|  |  | TX 362.2 | 3; botanical collection | Plant extr.-R | 2022-01-21 | Vaccinium corymbosum | 0.91 | 1 | 0
|  |  | TX 669 | 3; botanical collection | Plant extr.-R | 2023-10-27 | Vaccinium corymbosum | 0.67 | 1 | 0
|  |  | TX 670 | 3; botanical collection | Plant extr.-R | 2023-10-27 | Vaccinium corymbosum | 0.51 | 1 | 0
|  |  | TX 671 | 3; botanical collection | Plant extr.-R | 2023-10-20 | Vaccinium corymbosum | 0.67 | 1 | 0
|  |  | TX 680 | 3; botanical collection | Plant extr.-R | 2023-10-18 | Vaccinium darrowii | 0.12 | 1 | 0
|  |  | TX 677 | 3; botanical collection | Plant extr.-R | 2023-12-11 | Vaccinium elliottii | 0.5 | 1 | 0
|  |  | TX 672 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium fuscatum | 0.44 | 1 | 0
|  |  | TX 673 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium fuscatum | 0.56 | 1 | 0
|  |  | TX 674 | 3; botanical collection | Plant extr.-R | 2023-10-19 | Vaccinium fuscatum | 0.52 | 1 | 0
|  |  | TX 675 | 3; botanical collection | Plant extr.-R | 2023-12-12 | Vaccinium fuscatum | 0.68 | 1 | 0
|  |  | TX 575.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium macrocarpon | 0.93 | 1 | 0
|  |  | TX 575.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium macrocarpon | 0.74 | 1 | 0
|  |  | TX 576.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium macrocarpon | 0.57 | 1 | 0
|  |  | TX 576.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium macrocarpon | 0.65 | 1 | 0
|  |  | TX 596.2 | 2; label, sensory analysis | Plant extr. | 2023-12-18 | Vaccinium macrocarpon | 0.69 | 1 | 0
|  |  | TX 708.2 | 2; label, sensory analysis | Plant extr. | 2023-12-18 | Vaccinium macrocarpon | 0.43 | 1 | 0
|  |  | TX 681 | 3; botanical collection | Plant extr.-R | 2023-12-11 | Vaccinium myrtilloides | 0.76 | 1 | 0
|  |  | TX 192.1 | 2; expert opinion | Plant extr.-R | 2022-11-04 | Vaccinium myrtillus | 0.64 | 1 | 0
|  |  | TX 192.2 | 2; expert opinion | Plant extr.-AceR | 2022-12-19 | Vaccinium myrtillus | 0.2 | 1 | 0
|  |  | TX 192.3 | 2; expert opinion | Plant extr.-R | 2022-12-06 | Vaccinium myrtillus | 0.63 | 1 | 0
|  |  | TX 469.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium myrtillus | 0.13 | 1 | 0
|  |  | TX 469.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-22 | Vaccinium myrtillus | 0.77 | 1 | 0
|  |  | TX 507.1 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium myrtillus | 0.47 | 1 | 0
|  |  | TX 507.2 | 2; label; sensory analysis | Plant extr.-RL40 | 2022-12-02 | Vaccinium myrtillus | 0.69 | 1 | 0
|  |  | TX 550.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-02 | Vaccinium myrtillus | 0.72 | 1 | 0
|  |  | TX 550.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium myrtillus | 0.91 | 1 | 0
|  |  | TX 571.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-12 | Vaccinium myrtillus | 0.32 | 1 | 0
|  |  | TX 571.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-12 | Vaccinium myrtillus | 0.61 | 1 | 0
|  |  | TX 572.1 | 2; label; sensory analysis; Biomarker | Plant extr.-R | 2022-12-19 | Vaccinium myrtillus | 0.33 | 1 | 0
|  |  | TX 572.2 | 2; label; sensory analysis; Biomarker | Plant extr.-R | 2022-12-23 | Vaccinium myrtillus | 0.92 | 1 | 0
|  |  | TX 607.1 | 2; label; sensory analysis | Plant extr.-AceR | 2023-01-19 | Vaccinium myrtillus | 0.52 | 1 | 0
|  |  | TX 607.2 | 2; label; sensory analysis | Plant extr.-R | 2023-01-19 | Vaccinium myrtillus | 0.74 | 1 | 0
|  |  | TX 622 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.84 | 1 | 0
|  |  | TX 623 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.63 | 1 | 0
|  |  | TX 624 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.38 | 1 | 0
|  |  | TX 625 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.87 | 1 | 0
|  |  | TX 626 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.5 | 1 | 0
|  |  | TX 627 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.37 | 1 | 0
|  |  | TX 628 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.58 | 1 | 0
|  |  | TX 629 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.65 | 1 | 0
|  |  | TX 630 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.67 | 1 | 0
|  |  | TX 631 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.55 | 1 | 0
|  |  | TX 632 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.55 | 1 | 0
|  |  | TX 633 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium myrtillus | 0.51 | 1 | 0
|  |  | TX 635 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium myrtillus | 0.34 | 1 | 0
|  |  | TX 636 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium myrtillus | 0.83 | 1 | 0
|  |  | TX 637 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium myrtillus | 0.5 | 1 | 0
|  |  | TX 655 | 3; botanical collection | Plant extr.-R | 2023-03-08 | Vaccinium myrtillus | 0.28 | 1 | 0
|  |  | TX 656 | 3; botanical collection | Plant extr. R | 2023-03-08 | Vaccinium myrtillus | 0.68 | 1 | 0
|  |  | TX 657 | 3; botanical collection | Plant extr. R | 2023-03-08 | Vaccinium myrtillus | 0.25 | 1 | 0
|  |  | TX 658 | 3; botanical collection | Plant extr. R | 2023-03-13 | Vaccinium myrtillus | 0.45 | 1 | 0
|  |  | TX 659 | 3; botanical collection | Plant extr.-R | 2023-03-13 | Vaccinium myrtillus | 0.68 | 1 | 0
|  |  | TX 679 | 3; botanical collection | Plant extr.-R | 2023-10-18 | Vaccinium parvifolium | 0.28 | 1 | 0
|  |  | TX 468.1 | 2; label; sensory analysis | Plant extr.-R | 2022-11-03 | Vaccinium sect. | 0.32 | 1 | 0
|  |  | TX 468.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-22 | Vaccinium sect. | 0.45 | 1 | 0
|  |  | TX 470.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-12 | Vaccinium sect. | 0.49 | 1 | 0
|  |  | TX 470.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-12 | Vaccinium sect. | 0.7 | 1 | 0
|  |  | TX 492.1 | 2; label; sensory analysis | Plant extr.-R | 2022-12-07 | Vaccinium sect. | 0.77 | 1 | 0
|  |  | TX 492.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-21 | Vaccinium sect. | 0.64 | 1 | 0
|  |  | TX 501.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-12-02 | Vaccinium sect. | 0.67 | 1 | 0
|  |  | TX 501.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.4 | 1 | 0
|  |  | TX 502.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.71 | 1 | 0
|  |  | TX 502.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.73 | 1 | 0
|  |  | TX 503.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.79 | 1 | 0
|  |  | TX 503.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.36 | 1 | 0
|  |  | TX 504.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.56 | 1 | 0
|  |  | TX 504.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.48 | 1 | 0
|  |  | TX 505.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.53 | 1 | 0
|  |  | TX 505.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 0.55 | 1 | 0
|  |  | TX 506.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-11-30 | Vaccinium sect. | 0.58 | 1 | 0
|  |  | TX 506.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 0.22 | 1 | 0
|  |  | TX 508.1 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 0.85 | 1 | 0
|  |  | TX 508.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 0.61 | 1 | 0
|  |  | TX 509.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-11-30 | Vaccinium sect. | 0.47 | 1 | 0
|  |  | TX 509.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 0.59 | 1 | 0
|  |  | TX 510.1 | 2; label; sensory analysis | Plant extr.-AceRL40 | 2022-11-30 | Vaccinium sect. | 0.5 | 1 | 0
|  |  | TX 510.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 0.29 | 1 | 0
|  |  | TX 511.1 | 2; label; sensory analysis | Plant extr.-AceRL42S90 | 2022-12-07 | Vaccinium sect. | 0.66 | 1 | 0
|  |  | TX 511.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium sect. | 0.32 | 1 | 0
|  |  | TX 512.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-06 | Vaccinium sect. | 0.51 | 1 | 0
|  |  | TX 512.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium sect. | 0.44 | 1 | 0
|  |  | TX 513.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-06 | Vaccinium sect. | 0.63 | 1 | 0
|  |  | TX 513.2 | 2; label; sensory analysis | Plant extr.-RL40 | 2022-12-02 | Vaccinium sect. | 0.74 | 1 | 0
|  |  | TX 514.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-06 | Vaccinium sect. | 0.67 | 1 | 0
|  |  | TX 514.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium sect. | 0.31 | 1 | 0
|  |  | TX 573.1 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-23 | Vaccinium sect. | 0.77 | 1 | 0
|  |  | TX 573.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium sect. | 0.63 | 1 | 0
|  |  | TX 604.1 | 1; label | Plant extr.-AceR | 2023-01-16 | Vaccinium sect. | 0.58 | 1 | 0
|  |  | TX 604.2 | 1; label | Plant extr.-R | 2023-01-19 | Vaccinium sect. | 0.39 | 1 | 0
|  |  | TX 495 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium sp | 0.77 | 1 | 0
|  |  | TX 551 | 1; label | Plant extr.-R | 2022-12-02 | Vaccinium sp | 0.26 | 1 | 0
|  |  | TX 554.1 | 1; label | Plant extr.-AceR | 2022-12-06 | Vaccinium sp | 0.28 | 1 | 0
|  |  | TX 554.2 | 1; label | Plant extr.-R | 2023-01-10 | Vaccinium sp | 0.62 | 1 | 0
|  |  | TX 555.1 | 1; label | Plant extr.-AceR | 2022-12-06 | Vaccinium sp | 0.35 | 1 | 0
|  |  | TX 555.2 | 1; label | Plant extr.-R | 2023-01-10 | Vaccinium sp | 0.47 | 1 | 0
|  |  | TX 556 | 1; label | Plant extr.-R | 2023-01-10 | Vaccinium sp | 0.42 | 1 | 0
|  |  | TX 558 | 1; label | Plant extr.-AceR | 2022-12-23 | Vaccinium sp | 0.5 | 1 | 0
|  |  | TX 568.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 1.02 | 1 | 0
|  |  | TX 568.2 | 1; label | Plant extr.-R | 2022-12-12 | Vaccinium sp | 0.95 | 1 | 0
|  |  | TX 569.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.5 | 1 | 0
|  |  | TX 569.2 | 1; label | Plant extr.-R | 2022-12-12 | Vaccinium sp | 0.81 | 1 | 0
|  |  | TX 570.11 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.66 | 1 | 0
|  |  | TX 570.22 | 1; label | Plant extr.-R | 2022-12-23 | Vaccinium sp | 0.58 | 1 | 0
|  |  | TX 574.1 | 1; label | Plant extr.-R | 2022-12-16 | Vaccinium sp | 0.7 | 1 | 0
|  |  | TX 574.2 | 1; label | Plant extr.-RL42S90 | 2022-12-20 | Vaccinium sp | 0.28 | 1 | 0
|  |  | TX 577.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium sp | 0.26 | 1 | 0
|  |  | TX 577.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium sp | 0.57 | 1 | 0
|  |  | TX 577.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-19 | Vaccinium sp | 0.95 | 1 | 0
|  |  | TX 578.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium sp | 0.13 | 1 | 0
|  |  | TX 578.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-23 | Vaccinium sp | 0.46 | 1 | 0
|  |  | TX 579.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium sp | 0.29 | 1 | 0
|  |  | TX 579.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium sp | 0.3 | 1 | 0
|  |  | TX 579.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-01-02 | Vaccinium sp | 0.66 | 1 | 0
|  |  | TX 580.1 | 2; label; sensory analysis | Plant extr.-AceR 1zu2 | 2022-12-19 | Vaccinium sp | 0.17 | 1 | 0
|  |  | TX 580.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium sp | 0.55 | 1 | 0
|  |  | TX 580.3 | 2; label; sensory analysis | Plant extr.-AceR | 2022-12-19 | Vaccinium sp | 0.61 | 1 | 0
|  |  | TX 595.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium sp | 0.43 | 1 | 0
|  |  | TX 595.2 | 1; label | Plant extr.-AceR | 2022-12-19 | Vaccinium sp | 0.65 | 1 | 0
|  |  | TX 595 | 1; label | Plant extr.-AceRL42S90 | 2022-12-20 | Vaccinium sp | 0.36 | 1 | 0
|  |  | TX 596.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Vaccinium sp | 0.86 | 1 | 0
|  |  | TX 597.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Vaccinium sp | 0.85 | 1 | 0
|  |  | TX 598.1 | 1; label | Plant extr.-AceR | 2022-12-21 | Vaccinium sp | 0.51 | 1 | 0
|  |  | TX 598.2 | 1; label | Plant extr.-R | 2022-12-21 | Vaccinium sp | 0.64 | 1 | 0
|  |  | TX 606 | 1; label | Plant extr.-R | 2023-01-19 | Vaccinium sp | 0.63 | 1 | 0
|  |  | TX 608 | 1; label | Plant extr.-AceR | 2022-01-03 | Vaccinium sp | 0.5 | 1 | 0
|  |  | TX 609.1 | 1; label | Plant extr.-AceR 1:5 | 2022-12-22 | Vaccinium sp | 0.7 | 1 | 0
|  |  | TX 609.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.41 | 1 | 0
|  |  | TX 609.3 | 1; label | Plant extr.-AceR | 2022-01-02 | Vaccinium sp | 0.44 | 1 | 0
|  |  | TX 610.1 | 1; label | Plant extr.-AceR | 2022-12-22 | Vaccinium sp | 0.41 | 1 | 0
|  |  | TX 610.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.56 | 1 | 0
|  |  | TX 610.3 | 1; label | Plant extr.-AceR 1:1 | 2022-12-23 | Vaccinium sp | 0.44 | 1 | 0
|  |  | TX 611.11 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.54 | 1 | 0
|  |  | TX 611.12 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.76 | 1 | 0
|  |  | TX 611.21 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.63 | 1 | 0
|  |  | TX 611.22 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.56 | 1 | 0
|  |  | TX 612.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.48 | 1 | 0
|  |  | TX 612.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.64 | 1 | 0
|  |  | TX 613.12 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.63 | 1 | 0
|  |  | TX 613.21 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.8 | 1 | 0
|  |  | TX 613.22 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.41 | 1 | 0
|  |  | TX 614.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.77 | 1 | 0
|  |  | TX 615.1 | 1; label | Plant extr.-AceRL42S90 | 2022-12-22 | Vaccinium sp | 0.61 | 1 | 0
|  |  | TX 616.1 | 1; label | Plant extr.-AceRL42S90 | 2022-01-02 | Vaccinium sp | 0.93 | 1 | 0
|  |  | TX 617 | 1; label | Plant extr.-AceR | 2022-01-03 | Vaccinium sp | 0.48 | 1 | 0
|  |  | TX 618 | 1; label | Plant extr.-AceR | 2022-01-03 | Vaccinium sp | 0.52 | 1 | 0
|  |  | TX 619.2 | 1; label | Plant extr.-R | 2022-12-22 | Vaccinium sp | 0.62 | 1 | 0
|  |  | TX 642 | 1; label | Plant extr.-R | 2023-02-07 | Vaccinium sp | 0.17 | 1 | 0
|  |  | TX 649.2 | 1; label | Plant extr. | 2023-02-17 | Vaccinium sp | 0.7 | 1 | 0
|  |  | TX 650.2 | 1; label | Plant extr. | 2023-02-17 | Vaccinium sp | 0.43 | 1 | 0
|  |  | TX 366 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vaccinium uliginosum | 0.9 | 1 | 0
|  |  | TX 638 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium uliginosum | 0.98 | 1 | 0
|  |  | TX 660 | 3; botanical collection | Plant extr. R | 2023-03-13 | Vaccinium uliginosum | 0.6 | 1 | 0
|  |  | TX 369.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vaccinium vitis-idaea | 1.26 | 1 | 0
|  |  | TX 369.2 | 3; botanical collection | Plant extr.-R | 2022-10-19 | Vaccinium vitis-idaea | 0.78 | 1 | 0
|  |  | TX 639 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium vitis-idaea | 0.59 | 1 | 0
|  |  | TX 640 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium vitis-idaea | 0.29 | 1 | 0
|  |  | TX 641 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium vitis-idaea | 0.54 | 1 | 0
|  |  | TX 643 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium vitis-idaea | 0.61 | 1 | 0
|  |  | TX 644 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium vitis-idaea | 0.82 | 1 | 0
|  |  | TX 645 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium vitis-idaea | 0.81 | 1 | 0
|  |  | TX 646 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium vitis-idaea | 0.52 | 1 | 0
|  |  | TX 647 | 3; botanical collection | Plant extr. | 2023-02-17 | Vaccinium vitis-idaea | 1.21 | 1 | 0
|  |  | TX 661 | 3; botanical collection | Plant extr.-R | 2023-03-13 | Vaccinium vitis-idaea | 0.39 | 1 | 0
|  |  | TX 662 | 3; botanical collection | Plant extr. R | 2023-03-13 | Vaccinium vitis-idaea | 0.7 | 1 | 0
|  |  | TX 308 | 3; botanical collection | Plant extr. | 2022-09-29 | Vanilla planifolia | 0.73 | 1 | 0
|  |  | TX 336 | 3; botanical collection | Plant extr. | 2022-09-28 | Vanilla planifolia | 0.43 | 1 | 0
|  |  | TX 250.1 | 3; botanical collection | Plant extr. | 2022-09-16 | Vanilla pompona | 0.62 | 1 | 0
|  |  | TX 339 | 3; botanical collection | Plant extr. | 2022-09-28 | Vanilla pompona | 0.72 | 1 | 0
|  |  | TX 040 | 3; botanical collection | Plant extr.-R | 2022-01-26 | Veratrum album | 0.44 | 1 | 0
|  |  | TX 381.1 | 3; botanical collection | Plant extr. | 2022-10-04 | Veratrum album | 0.26 | 1 | 0
|  |  | TX 400.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Viburnum lantana | 0.76 | 1 | 0
|  |  | TX 403.1 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Viburnum opulus | 0.73 | 1 | 0
|  |  | TX 403.2 | 2; expert opinion | Plant extr.-R | 2022-10-17 | Viburnum opulus | 0.71 | 1 | 0
|  |  | TX 361 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Vinca minor | 0.35 | 1 | 0
|  |  | TX 069.1 | 2; label | Plant extr. | 2022-02-21 | Viola odorata | 0.55 | 1 | 0
|  |  | TX 069.2 | 2; label | Plant extr. | 2022-01-21 | Viola odorata | 0 | 1 | 0
|  |  | TX 016 | 2; determination by expert | Plant extr. | 2022-01-26 | Viscum album | 0.64 | 1 | 0
|  |  | TX 146 | 2; determination by expert | Plant extr. | 2022-02-03 | Viscum album | 0.51 | 1 | 0
|  |  | TX 442 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis tiliifolia | 0.81 | 1 | 0
|  |  | TX 445 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis vinifera | 0.69 | 1 | 0
|  |  | TX 446.1 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis yunnanensis | 0.74 | 1 | 0
|  |  | TX 446.2 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Vitis yunnanensis | 0.87 | 1 | 0
|  |  | TX 353.1 | 3; botanical collection | Plant extr.-R | 2022-10-11 | Zanthoxylum americanum | 0.58 | 1 | 0
|  |  | TX 353.2 | 3; botanical collection | Plant extr.-R | 2022-10-21 | Zanthoxylum americanum | 0.58 | 1 | 0
|  |  | TX 428 | 3; botanical collection | Plant extr.-R | 2022-10-28 | Zingiber officinale | 0.81 | 1 | 0

## Report (Spezies)
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Parameter: Allium ursinum
Version: | 1
erstellt von: | Ruhs | geprüft von: | Rau
erstellt am: | 2024-01-16 | geprüft am: | 2024-01-18
Anwendung | Artidentifizierung von Pflanzen
Umfeld | Blätter, andere Pflanzenteile, unverarbeitet
Verwendete Datenbank / Programmversion |  | Pflanzen DB (239 MSPs) | BT 4.0
Validierungsset:
Validierungsisolate/-materialien (Parameter): 23 |  | Vergleichsisolate/-materialien (#Parameter): 578
Kultivierungsbedingungen bei Mikroorganismen: -
Zustand der Probe bei sonstigen Materialien: frisch, gefroren, getrocknet
Aufreinigungsmethode: Plant extr. (mehrere Varianten)
Es werden nur identifizierte Proben für die Ermittlung der In- und Exklusivität herangezogen.
Richtig-positiv: 1. Hit =Ziel-Parameter mit Score ≥ 2 und Score 2. Hit < 1.7 oder Score 2. Hit ≥ 1.7 & Gattung gleich 1. Hit oder Score 2. Hit ≥ 2 & Art gleich 1. Hit |  | Richtig-negativ: 1. Hit ohne Zuordnung zum Ziel-Parameter mit score ≥ 2
Falsch-negativ: Kriterien für richtig-positiv werden nicht erfüllt |  | Falsch-positiv: Kriterien für richtig-negativ werden nicht erfüllt
Validierungsergebnis | Ziel-Parameter: | Allium ursinum |  | SOLL | erfüllt
Zahl aller Spektren | 23
davon mit ID-Ergebnis | 23 | Identifiziert Rate | 100% | 20 | ja | 1
davon Eigenspektren | 23 |  |  | 4.6 | ja | 1
In Bezug auf Identifizierte: |  | Inklusivität
Positive Übereinstimmungen (PA) | 23 | Richtig-Positiv Rate | 100% | 95% | ja | 1
Negative Abweichungen (ND) | 0 | Falsch-Negativ Rate | 0% | 1% | ja | 1
Nicht-Ziel-Parameter
Zahl aller Spektren | 578
davon mit ID-Ergebnis | 261 | Identifiziert Rate | 45.16% | 30 | ja | 1
davon Eigenspektren | 261 |  |  | 52.2 | ja | 1
In Bezug auf Identifizierte: |  | Exklusivität
Negative Übereinstimmungen (NA) | 261 | Richtig-Negativ Rate | 100% | 99% | ja | 1
Positive Abweichungen (PD) | 0 | Falsch-Positiv Rate | 0% | 1% | ja | 1
Test auf spektrale Nähe | 578 | Anteil nah verwandter Sp | 0%
Von 23 identifizierten Proben des Parameters wurden unter Verwendung der vollständigen Datenbank 23 (=100%) richtig erkannt (Inklusivität). 0 (=0%) der identifizierten Proben des Parameters wurden falsch einer anderen Spezies zugeordnet.
In den 261 identifizierten Proben der Nicht-Ziel-Parameter wurden unter Verwendung der vollständigen Datenbank 100% richtig-negativ (Exklusivität) gezählt. Von diesen 261 Proben wurden 0 (=0%) fehlerhaft als Parameter identifiziert.
In den 578 Proben der Nicht-Ziel-Parameter wurden unter Verwendung einer nur den Parameter enthaltenden Datenbank bei 0 (=0%) Einträgen ein score > 2 erreicht.
Die Kriterien für die | Zielgerichtete Identifizierung | sind somit | erfüllt
Kommentar:
Freigabe an Standort | für (Abteilung / Bereich) | am (Datum) | von (Hdz.)
CVUA S | T, P, G, RK | 2024-01-18 | Rau
Das Verfahren ist für den vorgesehenen Einsatzzweck geeignet und zur Verwendung freigegeben.
als Prüfmethode für Parameter in LIMS eingepflegt
Eine unterschriebene Version dieses Validierungsreports befindet sich im erstellenden Labor.
Andere Keime, die in dieser Matrix vorkommen können.
Andere Keime, die auf diesem Agar wachsen können.
Vergleichsisolate/-materialien, die ähnliche Spektren liefern können (siehe Blatt "Ähnliche zu Parameter").
ja
nein