## Settings
Settings für die Validierung
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Institution / Labor | CVUA S
Parameter | Gattung (Genus) |  | Art (Spezies)
|  | Corynebacterium | pseudotuberculosis
Unterscheidungsgrenzen
| Score zur Unterscheidung der Gattung | 1.7 | Score zur Unterscheidung der Spezies | 2

## Parameter (Spezies)
Parameter: | Corynebacterium | pseudotuberculosis | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | nicht zugeordnet | richtig-positiv | falsch-negativ | falsch zugeordnet
|  |  | 1 | 1 | CVUAS 9749,2 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.61 | Corynebacterium | pseudotuberculosis | 2.61 | A | 1 | 0 | 1 | 0 | 0
Zahl aller Spektren | 66 |  | 1 | 1 | 992 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.5 | Corynebacterium | pseudotuberculosis | 2.25 | A | 1 | 0 | 1 | 0 | 0
davon Identifiziert | 66 |  | 1 | 1 | 993 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.55 | Corynebacterium | pseudotuberculosis | 2.44 | A | 1 | 0 | 1 | 0 | 0
nicht identifiziert: | 0 |  | 1 | 1 | CVUAS 3337 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.53 | Corynebacterium | pseudotuberculosis | 2.5 | A | 1 | 0 | 1 | 0 | 0
von allen identifizierten Isolaten |  |  | 1 | 1 | CVUAS 3338 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.56 | Corynebacterium | pseudotuberculosis | 2.51 | A | 1 | 0 | 1 | 0 | 0
richtig-positiv: | 66 |  | 1 | 1 | CVUAS 3339 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.34 | Corynebacterium | pseudotuberculosis | 2.31 | A | 1 | 0 | 1 | 0 | 0
falsch zugeordnet | 0 |  | 1 | 1 | CVUAS 3354 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | pseudotuberculosis | 2.46 | Corynebacterium | pseudotuberculosis | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3357 | 3; sequenced | EFEx | 2018-07-10 | Corynebacterium | pseudotuberculosis | 2.74 | Corynebacterium | pseudotuberculosis | 2.72 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3361 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | pseudotuberculosis | 2.61 | Corynebacterium | pseudotuberculosis | 2.54 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3362 | 3; sequenced | DT | 2013-01-09 | Corynebacterium | pseudotuberculosis | 2.48 | Corynebacterium | pseudotuberculosis | 2.47 | A | 1 | 0 | 1 | 0 | 0
Scorewertgrenzen: | 2 | Species | 1 | 1 | CVUAS 3364 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.51 | Corynebacterium | pseudotuberculosis | 2.5 | A | 1 | 0 | 1 | 0 | 0
| 1.7 | Genus | 1 | 1 | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 1 | 1 | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium | pseudotuberculosis | 2.48 | Corynebacterium | pseudotuberculosis | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL269 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.11 | Corynebacterium | pseudotuberculosis | 2.08 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | 23UOF7 | 3; sequenced | eDT | 2023-02-03 | Corynebacterium | pseudotuberculosis | 2.27 | Corynebacterium | pseudotuberculosis | 2.1 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2997 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.48 | Corynebacterium | pseudotuberculosis | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 31248,2 | 3; sequenced | eDT | 2018-11-09 | Corynebacterium | pseudotuberculosis | 2.49 | Corynebacterium | pseudotuberculosis | 2.48 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32166 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.67 | Corynebacterium | pseudotuberculosis | 2.65 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3221 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.3 | Corynebacterium | pseudotuberculosis | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32632 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.59 | Corynebacterium | pseudotuberculosis | 2.56 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32654 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.65 | Corynebacterium | pseudotuberculosis | 2.59 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32656 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.61 | Corynebacterium | pseudotuberculosis | 2.6 | A | 1 | 0 | 1 | 0 | 0
davon mit ID-Ergebnis | 66 |  | 1 | 1 | CVUAS 32689 | 3; sequenced | eDT | 2020-10-15 | Corynebacterium | pseudotuberculosis | 2.62 | Corynebacterium | pseudotuberculosis | 2.54 | A | 1 | 0 | 1 | 0 | 0
Positive Übereinstimmungen (PA) | 66 |  | 1 | 1 | CVUAS 32745 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.63 | Corynebacterium | pseudotuberculosis | 2.59 | A | 1 | 0 | 1 | 0 | 0
Negative Abweichungen (ND) | 0 |  | 1 | 1 | CVUAS 32746 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.55 | Corynebacterium | pseudotuberculosis | 2.52 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32773 | 3; sequenced (CVUAS) | eDT | 2020-11-17 | Corynebacterium | pseudotuberculosis | 2.31 | Corynebacterium | pseudotuberculosis | 2.13 | A | 1 | 0 | 1 | 0 | 0
nicht identifiziert in % | 0% |  | 1 | 1 | CVUAS 32774 | 3; sequenced (CVUAS) | eDT | 2020-11-18 | Corynebacterium | pseudotuberculosis | 2.38 | Corynebacterium | pseudotuberculosis | 2.36 | A | 1 | 0 | 1 | 0 | 0
falsch zugeordnert in % | 0% |  | 1 | 1 | CVUAS 32830 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.64 | Corynebacterium | pseudotuberculosis | 2.59 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32834,2 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.62 | Corynebacterium | pseudotuberculosis | 2.56 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32842 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.62 | Corynebacterium | pseudotuberculosis | 2.56 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.6 | Corynebacterium | pseudotuberculosis | 2.56 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3353 | 3; sequenced | EFEx | 2018-07-10 | Corynebacterium | pseudotuberculosis | 2.65 | Corynebacterium | pseudotuberculosis | 2.59 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3355 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.4 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3365 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.38 | Corynebacterium | pseudotuberculosis | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3366 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.09 | Corynebacterium | pseudotuberculosis | 2.09 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34899 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.51 | Corynebacterium | pseudotuberculosis | 2.47 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34900 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.42 | Corynebacterium | pseudotuberculosis | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34901 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.47 | Corynebacterium | pseudotuberculosis | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34902 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.5 | Corynebacterium | pseudotuberculosis | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34903 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.49 | Corynebacterium | pseudotuberculosis | 2.47 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34904 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.44 | Corynebacterium | pseudotuberculosis | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34905 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.47 | Corynebacterium | pseudotuberculosis | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34906 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.54 | Corynebacterium | pseudotuberculosis | 2.5 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34907 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34908 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.37 | Corynebacterium | pseudotuberculosis | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4258,2 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.63 | Corynebacterium | pseudotuberculosis | 2.59 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4418 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.4 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4621,2 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4987 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.23 | Corynebacterium | pseudotuberculosis | 2.21 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4988 | 3; sequenced | DT | 2012-10-08 | Corynebacterium | pseudotuberculosis | 2.46 | Corynebacterium | pseudotuberculosis | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5188 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5583,2 | 3; biochemistry; FT-IR | eDT | 2015-08-11 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | DSM 20689 CVUAS | 3; type strain | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL160 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.18 | Corynebacterium | pseudotuberculosis | 2.17 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL177 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.08 | Corynebacterium | pseudotuberculosis | 2.07 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL184 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.24 | Corynebacterium | pseudotuberculosis | 2.24 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL293 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.49 | Corynebacterium | pseudotuberculosis | 2.45 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL302 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.51 | Corynebacterium | pseudotuberculosis | 2.49 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL359 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.63 | Corynebacterium | pseudotuberculosis | 2.57 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 10301,2 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.44 | Corynebacterium | pseudotuberculosis | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32354 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 33314 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.56 | Corynebacterium | pseudotuberculosis | 2.5 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 33501 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.5 | Corynebacterium | pseudotuberculosis | 2.49 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 8472,2 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.4 | Corynebacterium | pseudotuberculosis | 2.24 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | DSM 20689 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; K6 | Corynebacterium | pseudotuberculosis | 2.55 | Corynebacterium | pseudotuberculosis | 2.5 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | LALLF 21TRD1263 | 3; sequenced | EFEx | 2022-05-09 | Corynebacterium | pseudotuberculosis | 2.34 | Corynebacterium | pseudotuberculosis | 2.29 | A | 1 | 0 | 1 | 0 | 0

## #Parameter (Spezies) DB
Parameter: | Corynebacterium pseudotuberculosis |  | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | verwendete Gattung | verwendete Art | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | richtige ID | keine richtige ID | 1. Hit | 2. Hit | 1. Hit Treffer Parameter | 2. Hit Treffer Parameter | falsch als Parameter identifiziert | PM unter den ersten 2 Hits
|  |  | 1 | 1 | DSM 44123 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; K4 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.53 | Corynebacterium | diphtheriae | 2.45 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
Anzahl Isolate: | 1278 |  | 1 | 1 | KL163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.66 | Corynebacterium | diphtheriae | 2.59 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
davon Identifiziert | 1028 |  | 1 | 1 | KL167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.58 | Corynebacterium | diphtheriae | 2.58 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
von allen identifizierten Isolaten |  |  | 1 | 1 | KL187 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.68 | Corynebacterium | diphtheriae | 2.51 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
richtige ID | 1028 |  | 1 | 1 | KL232 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.65 | Corynebacterium | diphtheriae | 2.58 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
falsch als Parameter identifiziert | 0 |  | 1 | 1 | KL237 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.63 | Corynebacterium | diphtheriae | 2.6 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL240 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-13 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.62 | Corynebacterium | diphtheriae | 2.53 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL932 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.51 | Corynebacterium | diphtheriae | 2.47 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL933 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.4 | Corynebacterium | diphtheriae | 2.38 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL935 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.4 | Corynebacterium | diphtheriae | 2.36 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
Parameter wird getroffen mit: |  |  | 0 | 0 | KL936 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.34 | Corynebacterium | diphtheriae | 2.3 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL950 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.44 | Corynebacterium | diphtheriae | 2.39 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
1. Hit Treffer Parameter | 0 |  | 0 | 0 | KL952 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.37 | Corynebacterium | diphtheriae | 2.34 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
2. Hit Treffer Parameter | 0 |  | 1 | 1 | KL173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.71 | Corynebacterium | diphtheriae | 2.65 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
PM unter den ersten 2 Hits | 0 |  | 1 | 1 | KL179 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.6 | Corynebacterium | diphtheriae | 2.49 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
in % | 0% |  | 1 | 1 | KL235 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.71 | Corynebacterium | diphtheriae | 2.61 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL940 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.53 | Corynebacterium | diphtheriae | 2.51 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL954 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.26 | Corynebacterium | diphtheriae | 2.25 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
Scorewertgrenzen: | 2 | Species | 0 | 0 | KL956 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.04 | Corynebacterium | diphtheriae | 2 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
| 1.7 | Genus | 0 | 0 | DSM 44123 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; K4 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.39 | Corynebacterium | diphtheriae | 2.39 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 0 | 0 | DSM 43988 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; I7 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.25 | Corynebacterium | diphtheriae | 2.21 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141000461 (2) | 1; 0 | 0 | 2014-01-20 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.53 | Corynebacterium | silvaticum | 2.33 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
davon mit ID-Ergebnis | 1028 |  | 0 | 0 | 141016278 | 3; Confirmation by KL Diphtheriae (LGL); WBC: FT-IR | 0 | 2014-12-22 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.46 | Corynebacterium | silvaticum | 2.43 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
Negative Übereinstimmungen (NA) | 1028 |  | 0 | 0 | 161003120-1 | 1; 0 | 0 | 2016-03-15 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.42 | Corynebacterium | silvaticum | 2.31 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
Positive Abweichungen (PD) | 0 |  | 0 | 0 | 161003120-2 | 1; 0 | 0 | 2016-03-15 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.41 | Corynebacterium | silvaticum | 2.39 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-3 | 3; Confirmation by KL Diphtheriae (LGL) | 0 | 2016-03-16 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.48 | Corynebacterium | silvaticum | 2.47 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
nicht identifiziert in % | 19.56% |  | 0 | 0 | 161003120-4 | 1; 0 | 0 | 2016-03-14 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.53 | Corynebacterium | silvaticum | 2.51 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
falsch zugeordnert in % | 0% |  | 0 | 0 | 161003120-5 | 1; 0 | 0 | 2016-03-14 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.51 | Corynebacterium | silvaticum | 2.51 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-6 | 1; 0 | 0 | 2016-03-14 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.48 | Corynebacterium | silvaticum | 2.44 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-7-D-00025 | 3; sequenced | DT | 2013-01-08 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.58 | Corynebacterium | silvaticum | 2.52 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11USF28 | 3; sequenced | DT | 2012-12-11 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.56 | Corynebacterium | silvaticum | 2.43 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11USF53 | 3; sequenced | DT | 2013-01-16 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.54 | Corynebacterium | silvaticum | 2.49 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11USF78 | 3; sequenced | DT | 2012-12-11 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.53 | Corynebacterium | silvaticum | 2.49 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11USF79 | 3; sequenced | DT | 2012-12-11 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.55 | Corynebacterium | silvaticum | 2.42 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 121017479 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.61 | Corynebacterium | silvaticum | 2.54 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131000349 | 3; sequenced | DT | 2013-01-18 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.69 | Corynebacterium | silvaticum | 2.56 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.71 | Corynebacterium | silvaticum | 2.66 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.64 | Corynebacterium | silvaticum | 2.55 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 057 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.62 | Corynebacterium | silvaticum | 2.51 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 076 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.5 | Corynebacterium | silvaticum | 2.44 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 085 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.56 | Corynebacterium | silvaticum | 2.45 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 091 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.53 | Corynebacterium | silvaticum | 2.48 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 098 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.55 | Corynebacterium | silvaticum | 2.47 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 111 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.51 | Corynebacterium | silvaticum | 2.46 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 123 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.46 | Corynebacterium | silvaticum | 2.39 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 128 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.56 | Corynebacterium | silvaticum | 2.48 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 130 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.58 | Corynebacterium | silvaticum | 2.52 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 139 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.67 | Corynebacterium | silvaticum | 2.5 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 141 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.52 | Corynebacterium | silvaticum | 2.49 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 145 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.5 | Corynebacterium | silvaticum | 2.44 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 147 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.45 | Corynebacterium | silvaticum | 2.42 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 149 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.55 | Corynebacterium | silvaticum | 2.35 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 156 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.41 | Corynebacterium | silvaticum | 2.37 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 166 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.58 | Corynebacterium | silvaticum | 2.49 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 168 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.54 | Corynebacterium | silvaticum | 2.46 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 182 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.45 | Corynebacterium | silvaticum | 2.45 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 183 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.66 | Corynebacterium | silvaticum | 2.48 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 188 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.62 | Corynebacterium | silvaticum | 2.49 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 233 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.51 | Corynebacterium | silvaticum | 2.43 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 257 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.63 | Corynebacterium | silvaticum | 2.52 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AGES WS 332 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.62 | Corynebacterium | silvaticum | 2.54 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 344 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.51 | Corynebacterium | silvaticum | 2.41 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 351 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.54 | Corynebacterium | silvaticum | 2.44 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 353 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.53 | Corynebacterium | silvaticum | 2.43 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BU125-97 | 3; sequenced | DT | 2012-10-19 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.59 | Corynebacterium | silvaticum | 2.47 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10179,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.64 | Corynebacterium | silvaticum | 2.53 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10701 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.49 | Corynebacterium | silvaticum | 2.43 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11363 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.31 | Corynebacterium | silvaticum | 2.28 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11488 | 3; FT-IR (NTTB-D) | eDT | 2017-01-11 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.25 | Corynebacterium | silvaticum | 2.15 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1400,3 | 3; FT-IR (NTTB-D) | eDT | 2018-03-02 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.42 | Corynebacterium | silvaticum | 2.33 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1717,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.18 | Corynebacterium | silvaticum | 2.09 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2040,2 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.56 | Corynebacterium | silvaticum | 2.45 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2738,2 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.63 | Corynebacterium | silvaticum | 2.47 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4292 | 3; type strain [0127] | DT | 2012-12-08 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.67 | Corynebacterium | silvaticum | 2.54 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5160 | 3; sequenced | DT | 2012-10-08 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.57 | Corynebacterium | silvaticum | 2.46 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5243,2 | 3; FT-IR (NTTB-D) | DT | 2013-11-18 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.24 | Corynebacterium | silvaticum | 2.23 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6455 | 3; sequenced, strain of the species desciption [0127] | DT | 2013-01-08 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.62 | Corynebacterium | silvaticum | 2.47 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6678,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.27 | Corynebacterium | silvaticum | 2.26 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 704,2 | 2; FT-IR (NTTB-D) | eDT | 09.12.2013 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.48 | Corynebacterium | silvaticum | 2.48 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8972,3 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.67 | Corynebacterium | silvaticum | 2.52 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 97,2 | 2; FT-IR (NTTB-D) | eDT | 27.11.2013 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.29 | Corynebacterium | silvaticum | 2.25 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-1014-16 | 3; FT-IR (NTTB-D) | eDT | 2017-01-11 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.51 | Corynebacterium | silvaticum | 2.44 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-1111-3-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.47 | Corynebacterium | silvaticum | 2.39 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-1408-1-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.47 | Corynebacterium | silvaticum | 2.44 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-1449-2-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.62 | Corynebacterium | silvaticum | 2.58 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-1514-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.47 | Corynebacterium | silvaticum | 2.36 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-1591-2-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.49 | Corynebacterium | silvaticum | 2.49 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-1627-5-12 | 3; sequenced | DT | 2012-10-08 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.57 | Corynebacterium | silvaticum | 2.44 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-241-6-18 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.61 | Corynebacterium | silvaticum | 2.55 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-241-7-18 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.56 | Corynebacterium | silvaticum | 2.54 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-28-3-13 | 3; sequenced | DT | 2013-01-18 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.67 | Corynebacterium | silvaticum | 2.6 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-308-2-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.7 | Corynebacterium | silvaticum | 2.53 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-309-1-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.71 | Corynebacterium | silvaticum | 2.59 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-310-2-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.74 | Corynebacterium | silvaticum | 2.59 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-349-18 | 3; FT-IR (NTTB-D) | eDT | 2018-04-05 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.2 | Corynebacterium | silvaticum | 2.08 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-349-3-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.57 | Corynebacterium | silvaticum | 2.45 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-427-3-15 | 3; FT-IR (NTTB-D) | eDT | 2015-03-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.36 | Corynebacterium | silvaticum | 2.25 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-434-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.49 | Corynebacterium | silvaticum | 2.36 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-443-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.52 | Corynebacterium | silvaticum | 2.46 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-450-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.49 | Corynebacterium | silvaticum | 2.39 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-468-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.24 | Corynebacterium | silvaticum | 2.23 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-619-1-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.28 | Corynebacterium | silvaticum | 2.18 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-793-3-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.58 | Corynebacterium | silvaticum | 2.52 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-803-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.6 | Corynebacterium | silvaticum | 2.46 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-884-18 | 3; FT-IR (NTTB-D) | eDT | 2018-08-21 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.68 | Corynebacterium | silvaticum | 2.52 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | W30-B05227 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.41 | Corynebacterium | silvaticum | 2.4 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | W67-B16966 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.52 | Corynebacterium | silvaticum | 2.41 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | C57 UGhent | 3; sequenced (rpoB) UGent | DT | 2020-06-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.36 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | C133 UGhent | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.45 | Corynebacterium | ulcerans | 2.32 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141005571-1 | 1; 0 | 0 | 2014-05-09 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.36 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141013039 2 | 1; 0 | 0 | 2014-05-09 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.4 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151004530 | 1; 0 | 0 | 2015-04-15 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.37 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151012433-2 LHL | 1; 0 | 0 | 2015-09-28 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161007221 | 1; 0 | 0 | 2016-06-13 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.36 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 112226 eDT | 3; type strain | eDT | 2024-08-13 | Corynebacterium | ramonii | Corynebacterium | ramonii | 2.61 | Corynebacterium | ulcerans | 2.43 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 112226 EFex | 3; type strain | EFEx | 2024-08-13 | Corynebacterium | ramonii | Corynebacterium | ramonii | 2.74 | Corynebacterium | ulcerans | 2.53 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131010012-1 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.55 | Corynebacterium | ulcerans | 2.54 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131010012-2 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.53 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131011719 | 3; sequenced | eDT | 2018-03-29 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.51 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131013415-1 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.49 | Corynebacterium | ulcerans | 2.47 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131013415-2 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.58 | Corynebacterium | ulcerans | 2.52 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131015432-10 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.5 | Corynebacterium | ulcerans | 2.47 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131015432-13 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131015432-14 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.48 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 141001018 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.56 | Corynebacterium | ulcerans | 2.55 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 141001548 | 3; sequenced | eDT | 2018-03-29 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 171008197 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.5 | Corynebacterium | ulcerans | 2.42 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C15 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.41 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C44 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.45 | Corynebacterium | ramonii | 2.43 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C55 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.44 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C64 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.61 | Corynebacterium | ulcerans | 2.59 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C78 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.41 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C86 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.45 | Corynebacterium | ulcerans | 2.44 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 2823 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.52 | Corynebacterium | ulcerans | 2.35 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.6 | Corynebacterium | ulcerans | 2.59 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34267 | 3; KL Diphteria | eDT | 2022-09-02 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.27 | Corynebacterium | ulcerans | 2.26 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 46325 CVUAS | 3; type strain | DT | 2013-01-09 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.38 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1017 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.4 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | KL199 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.49 | Corynebacterium | ulcerans | 2.49 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL853 | 3; MLST | eDT | 2016-10-27 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL941 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.4 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB921348 | 2; 0 | eDT | 2015-01-30 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.22 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C156 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.45 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C164 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.61 | Corynebacterium | ulcerans | 2.59 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C166 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.45 | Corynebacterium | ulcerans | 2.42 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C169 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C174 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.55 | Corynebacterium | ramonii | 2.49 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C25 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.38 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C92 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.44 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C95 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.52 | Corynebacterium | ulcerans | 2.49 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.53 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.34 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1015 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1025 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | KL188 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ramonii | 2.4 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL195 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.33 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL818 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.38 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL867 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.45 | Corynebacterium | ulcerans | 2.44 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL870 | 3; 0 | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.36 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL885 | 3; FT-IR no NTTB-D, MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.42 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 151012433-2 CVUAS | 3; KL Diphtheria | eDT | 2021-06-22 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.58 | Corynebacterium | ulcerans | 2.52 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 181011209 001 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.48 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 181011209 002 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.56 | Corynebacterium | ulcerans | 2.55 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C1 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.57 | Corynebacterium | ulcerans | 2.53 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C104 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.45 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C11 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.56 | Corynebacterium | ulcerans | 2.56 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C112 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.6 | Corynebacterium | ulcerans | 2.59 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C12 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.54 | Corynebacterium | ulcerans | 2.53 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C123 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.54 | Corynebacterium | ulcerans | 2.48 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C133 CVUAS | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.32 | Corynebacterium | ulcerans | 2.32 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C139 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.5 | Corynebacterium | ulcerans | 2.46 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C152 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.49 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C167 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.51 | Corynebacterium | ulcerans | 2.5 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C17 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C60 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.63 | Corynebacterium | ulcerans | 2.63 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C61 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ramonii | 2.48 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C63 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.59 | Corynebacterium | ulcerans | 2.52 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C65 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.57 | Corynebacterium | ulcerans | 2.54 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C70 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.5 | Corynebacterium | ulcerans | 2.5 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C71 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.47 | Corynebacterium | ulcerans | 2.46 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C72 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.59 | Corynebacterium | ulcerans | 2.56 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C73 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.72 | Corynebacterium | ulcerans | 2.68 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C76 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.62 | Corynebacterium | ulcerans | 2.57 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C80 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.49 | Corynebacterium | ulcerans | 2.48 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C81 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.45 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C82 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.38 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C83 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.5 | Corynebacterium | ulcerans | 2.49 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C90 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.54 | Corynebacterium | ulcerans | 2.51 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C93 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.52 | Corynebacterium | ramonii | 2.5 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1750 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.45 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10306 | 3; sequenced | EFex | 2018-04-24 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.57 | Corynebacterium | ulcerans | 2.44 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33950 | 3; KL Diphteria | eDT | 2022-06-13 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.32 | Corynebacterium | ulcerans | 2.26 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34094 | 3; KL Diphteria | eDT | 2022-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6522,4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB360523 | 3; 0 | eDT | 2013-01-30 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.35 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | KL190 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.46 | Corynebacterium | ulcerans | 2.41 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL246 | 3; WGS data LGL [0085]; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.42 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL796 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.32 | Corynebacterium | ulcerans | 2.31 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL825 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL832 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL840 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | WFL 1217-18 | 3; 0 | eDT | 2018-04-17 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.33 | Corynebacterium | ulcerans | 2.32 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C117 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.58 | Corynebacterium | ulcerans | 2.54 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C146 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.62 | Corynebacterium | ulcerans | 2.6 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C162 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.52 | Corynebacterium | ulcerans | 2.51 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C18 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.48 | Corynebacterium | ulcerans | 2.48 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C20 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.35 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C28 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ramonii | 2.46 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C31 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C32 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.4 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C39 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.63 | Corynebacterium | ulcerans | 2.6 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C41 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.4 | Corynebacterium | ulcerans | 2.38 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C51 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.39 | Corynebacterium | ulcerans | 2.38 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C52 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C53 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.45 | Corynebacterium | ulcerans | 2.42 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C57 CVUAS | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.4 | Corynebacterium | ulcerans | 2.38 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C67 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.4 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C98 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.52 | Corynebacterium | ulcerans | 2.46 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 46325 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; I8 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.6 | Corynebacterium | ulcerans | 2.42 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-477-6-21 | 3; KL Diphtheria | eDT | 2021-04-21 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.19 | Corynebacterium | amycolatum | 2.15 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2014 E8 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.3 | Corynebacterium | amycolatum | 2.29 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium | aurimucosum | Corynebacterium | aurimucosum | 2.07 | Corynebacterium | aurimucosum | 2.05 | A | 1 | 1 | 0 | Corynebacterium aurimucosum | Corynebacterium aurimucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | belfantii | Corynebacterium | belfantii | 2.75 | Corynebacterium | belfantii | 2.59 | A | 1 | 1 | 0 | Corynebacterium belfantii | Corynebacterium belfantii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0171 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | belfantii | Corynebacterium | belfantii | 2.54 | Corynebacterium | belfantii | 2.41 | A | 1 | 1 | 0 | Corynebacterium belfantii | Corynebacterium belfantii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | belfantii | Corynebacterium | belfantii | 2.66 | Corynebacterium | belfantii | 2.62 | A | 1 | 1 | 0 | Corynebacterium belfantii | Corynebacterium belfantii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium | bovis | Corynebacterium | bovis | 2.35 | Corynebacterium | bovis | 2.33 | A | 1 | 1 | 0 | Corynebacterium bovis | Corynebacterium bovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium | camporealensis | Corynebacterium | camporealensis | 2.58 | Corynebacterium | camporealensis | 1.96 | A | 1 | 1 | 0 | Corynebacterium camporealensis | Corynebacterium camporealensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium | casei | Corynebacterium | casei | 1.98 | Corynebacterium | stationis | 1.64 | B | 0 | 0 | 1 | Corynebacterium casei | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium | confusum | Corynebacterium | confusum | 2.49 | Corynebacterium | confusum | 1.93 | A | 1 | 1 | 0 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | confusum | Corynebacterium | confusum | 2.14 | Corynebacterium | confusum | 2.01 | A | 1 | 1 | 0 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium | cystitidis | Corynebacterium | cystitidis | 2.5 | Corynebacterium | cystitidis | 2.36 | A | 1 | 1 | 0 | Corynebacterium cystitidis | Corynebacterium cystitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.32 | Corynebacterium | frankenforstense | 2.2 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Corynebacterium frankenforstense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.74 | Corynebacterium | frankenforstense | 2.48 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Corynebacterium frankenforstense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium | freiburgense | Corynebacterium | freiburgense | 2.46 | Pseudomonas | gessardii | 1.27 | A | 1 | 1 | 0 | Corynebacterium freiburgense | Pseudomonas gessardii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | hansenii | Corynebacterium | hansenii | 2.35 | Corynebacterium | hansenii | 2.3 | A | 1 | 1 | 0 | Corynebacterium hansenii | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | hansenii | Corynebacterium | hansenii | 2.41 | Corynebacterium | hansenii | 2.2 | A | 1 | 1 | 0 | Corynebacterium hansenii | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | humireducens | Corynebacterium | humireducens | 2.51 | Corynebacterium | sp-CVUAS-31533 | 1.98 | A | 1 | 1 | 0 | Corynebacterium humireducens | Corynebacterium sp-CVUAS-31533 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | kutscheri | Corynebacterium | kutscheri | 2.56 | Corynebacterium | kutscheri | 2.37 | A | 1 | 1 | 0 | Corynebacterium kutscheri | Corynebacterium kutscheri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | lactis | Corynebacterium | lactis | 2.38 | Corynebacterium | lactis | 2.29 | A | 1 | 1 | 0 | Corynebacterium lactis | Corynebacterium lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | lactis | Corynebacterium | lactis | 2.7 | Corynebacterium | lactis | 2.54 | A | 1 | 1 | 0 | Corynebacterium lactis | Corynebacterium lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium | phoceense | Corynebacterium | phoceense | 2.61 | Corynebacterium | phoceense | 2.19 | A | 1 | 1 | 0 | Corynebacterium phoceense | Corynebacterium phoceense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | pilosum | Corynebacterium | pilosum | 2.3 | Corynebacterium | pilosum | 1.78 | A | 1 | 1 | 0 | Corynebacterium pilosum | Corynebacterium pilosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C138 | 3; sequenced (rpoB) UGent | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.65 | Corynebacterium | rouxii | 2.6 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.73 | Corynebacterium | rouxii | 2.71 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3559,2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.7 | Corynebacterium | rouxii | 2.66 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.74 | Corynebacterium | rouxii | 2.65 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL1306 | 3; sequenced / variant upgraded to species level Coynebacterium rouxii [0137] here: NGS data (LGA) C. rouxi | EFEx | 2020-10-20 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.75 | Corynebacterium | rouxii | 2.71 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL1663 | 3; sequenced / variant upgraded to species level Coynebacterium rouxii [0137] here: NGS data (LGA) C. rouxi | EFEx | 2020-12-15 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.69 | Corynebacterium | rouxii | 2.68 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium | sp-208-U2-6 | Corynebacterium | sp-208-U2-6 | 2.63 | Sphingomonas | yabuuchiae | 1.39 | A | 1 | 1 | 0 | Corynebacterium sp-208-U2-6 | Sphingomonas yabuuchiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32165,2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium | sp-2300500 | Corynebacterium | sp-2300500 | 2.59 | Corynebacterium | sp-2300500 | 2.56 | A | 1 | 1 | 0 | Corynebacterium sp-2300500 | Corynebacterium sp-2300500 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5775,2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium | sp-2300500 | Corynebacterium | sp-2300500 | 2.24 | Corynebacterium | sp-2300500 | 2.22 | A | 1 | 1 | 0 | Corynebacterium sp-2300500 | Corynebacterium sp-2300500 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2554,2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium | sp-2396 | Corynebacterium | sp-2396 | 2.44 | Corynebacterium | sp | 1.92 | A | 1 | 1 | 0 | Corynebacterium sp-2396 | Corynebacterium sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium | sp-243-z-3 | Corynebacterium | sp-243-z-3 | 2.57 | Corynebacterium | resistens | 1.78 | A | 1 | 1 | 0 | Corynebacterium sp-243-z-3 | Corynebacterium resistens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | sp-257-y-4 | Corynebacterium | sp-257-y-4 | 2.57 | Proteus | mirabilis | 1.36 | A | 1 | 1 | 0 | Corynebacterium sp-257-y-4 | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | sp-284z-2 | Corynebacterium | sp-284z-2 | 2.5 | Proteus | mirabilis | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-284z-2 | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium | sp-CVUAS-11664 | Corynebacterium | sp-CVUAS-11664 | 2.74 | Pseudomonas | aeruginosa | 1.32 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-11664 | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | sp-CVUAS-1888 | Corynebacterium | sp-CVUAS-1888 | 2.62 | Corynebacterium | striatum | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-1888 | Corynebacterium striatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1721,3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium | sp-CVUAS-2484.4 | Corynebacterium | sp-CVUAS-2484.4 | 2.61 | Corynebacterium | sp-CVUAS-2484.4 | 2.11 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-2484.4 | Corynebacterium sp-CVUAS-2484.4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2484,4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388; 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium | sp-CVUAS-2484.4 | Corynebacterium | sp-CVUAS-2484.4 | 2.37 | Corynebacterium | sp-CVUAS-2484.4 | 1.99 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-2484.4 | Corynebacterium sp-CVUAS-2484.4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp-CVUAS-30404 | 2.49 | Corynebacterium | sp-CVUAS-30404 | 2.27 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-30404 | Corynebacterium sp-CVUAS-30404 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 20TRD0812 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp-CVUAS-30404 | 2.31 | Corynebacterium | sp-CVUAS-30404 | 2.3 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-30404 | Corynebacterium sp-CVUAS-30404 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp-CVUAS-30404 | 2.12 | Corynebacterium | sp-CVUAS-30404 | 2.04 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-30404 | Corynebacterium sp-CVUAS-30404 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium | sp-CVUAS-31533 | Corynebacterium | sp-CVUAS-31533 | 2.4 | Corynebacterium | humireducens | 2.22 | B | 0 | 0 | 1 | Corynebacterium sp-CVUAS-31533 | Corynebacterium humireducens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32278,2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium | sp-CVUAS-32278.2 | Corynebacterium | sp-CVUAS-32278.2 | 2.72 | Shewanella | sp-RKVR22 | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-32278.2 | Shewanella sp-RKVR22 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium | sp-CVUAS-874,4 | Corynebacterium | sp-CVUAS-874,4 | 2.14 | Corynebacterium | sp-CVUAS-874,4 | 1.98 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-874,4 | Corynebacterium sp-CVUAS-874,4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 874,4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium | sp-CVUAS-874,4 | Corynebacterium | sp-CVUAS-874,4 | 2.48 | Corynebacterium | sp-CVUAS-874,4 | 2.32 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-874,4 | Corynebacterium sp-CVUAS-874,4 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 21TRB0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-LALLF-21TRB0553 | Corynebacterium | sp-LALLF-20TRB0553 | 2.39 | Cryptococcus | neoformans | 1.38 | A | 0 | 0 | 1 | Corynebacterium sp-LALLF-20TRB0553 | Cryptococcus neoformans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium | sp-LHL-0497 | Corynebacterium | sp-LHL-0497 | 2.22 | Haemophilus | influenzae | 1.26 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-0497 | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | sp-LHL-0558 | Corynebacterium | sp-LHL-0558 | 2.37 | Nocardia | cyriacigeorgica | 1.32 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-0558 | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium | sp-LHL-1322 | Corynebacterium | sp-LHL-1322 | 2.58 | Agromyces | humatus | 1.41 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-1322 | Agromyces humatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | stationis | Corynebacterium | stationis_JRA_G+R-52 | 2.37 | Corynebacterium | stationis | 2.27 | B | 0 | 0 | 1 | Corynebacterium stationis_JRA_G+R-52 | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium | testudinoris | Corynebacterium | testudinoris | 2.56 | Bacteroides | fragilis | 1.29 | A | 1 | 1 | 0 | Corynebacterium testudinoris | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5437,3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium | variabile | Corynebacterium | variabile | 2.17 | Corynebacterium | variabile | 1.62 | A | 1 | 1 | 0 | Corynebacterium variabile | Corynebacterium variabile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium | xerosis | Corynebacterium | xerosis | 2.39 | Corynebacterium | xerosis | 1.57 | A | 1 | 1 | 0 | Corynebacterium xerosis | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium | xerosis-group | Corynebacterium | xerosis-group | 2.33 | Corynebacterium | xerosis | 1.58 | A | 1 | 1 | 0 | Corynebacterium xerosis-group | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | bovis | Actinomyces | bovis | 2.05 | Actinomyces | bovis | 1.92 | A | 1 | 1 | 0 | Actinomyces bovis | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces | denticolens | Actinomyces | denticolens | 2.4 | Actinomyces | denticolens | 2.07 | A | 1 | 1 | 0 | Actinomyces denticolens | Actinomyces denticolens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces | sp-CVUAS-31303 | Actinomyces | sp-CVUAS-31303 | 2.6 | Campylobacter | jejuni | 1.26 | A | 1 | 1 | 0 | Actinomyces sp-CVUAS-31303 | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces | viscosus | Actinomyces | viscosus | 2.04 | Actinomyces | viscosus | 1.87 | A | 1 | 1 | 0 | Actinomyces viscosus | Actinomyces viscosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | weissii | Actinomyces | weissii | 2.48 | Actinomyces | weissii | 2.35 | A | 1 | 1 | 0 | Actinomyces weissii | Actinomyces weissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium | bovis | Arcanobacterium | bovis | 2.73 | Arcanobacterium | pluranimalium | 1.77 | A | 1 | 1 | 0 | Arcanobacterium bovis | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium | canis | Arcanobacterium | canis | 2.34 | Arcanobacterium | canis | 2.19 | A | 1 | 1 | 0 | Arcanobacterium canis | Arcanobacterium canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium | haemolyticum | Arcanobacterium | haemolyticum | 2.46 | Arcanobacterium | haemolyticum | 2.3 | A | 1 | 1 | 0 | Arcanobacterium haemolyticum | Arcanobacterium haemolyticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-6203 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.65 | Arcanobacterium | hippocoleae | 2.6 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.32 | Arcanobacterium | hippocoleae | 2.3 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10002 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.66 | Arcanobacterium | phocae | 2.5 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.48 | Arcanobacterium | phocae | 2.4 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2698T | 3; type strain | DT | 2013-08-07 | Arcanobacterium | phocisimile | Arcanobacterium | phocisimile | 2.15 | Arcanobacterium | phocisimile | 2.13 | A | 1 | 1 | 0 | Arcanobacterium phocisimile | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 014418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.63 | Arcanobacterium | pinnipediorum | 2.47 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium pinnipediorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.75 | Arcanobacterium | pinnipediorum | 2.56 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium pinnipediorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13483 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium | pluranimalium | Arcanobacterium | pluranimalium | 2.47 | Arcanobacterium | pluranimalium | 2.1 | A | 1 | 1 | 0 | Arcanobacterium pluranimalium | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | sp-17-2065.02 | 2.72 | Arcanobacterium | sp-17-2065.02 | 2.62 | A | 1 | 1 | 0 | Arcanobacterium sp-17-2065.02 | Arcanobacterium sp-17-2065.02 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | sp-17-2065.02 | 2.63 | Arcanobacterium | sp-17-2065.02 | 2.61 | A | 1 | 1 | 0 | Arcanobacterium sp-17-2065.02 | Arcanobacterium sp-17-2065.02 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium | sp-2701 | Arcanobacterium | sp-2701 | 2.7 | Arcanobacterium | sp-2701 | 2.57 | A | 1 | 1 | 0 | Arcanobacterium sp-2701 | Arcanobacterium sp-2701 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium | sp-2701 | Arcanobacterium | sp-2701 | 2.78 | Arcanobacterium | sp-2701 | 2.64 | A | 1 | 1 | 0 | Arcanobacterium sp-2701 | Arcanobacterium sp-2701 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1060,2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium | wilhelmae | Arcanobacterium | wilhelmae | 2.32 | Arcanobacterium | wilhelmae | 2.18 | A | 1 | 1 | 0 | Arcanobacterium wilhelmae | Arcanobacterium wilhelmae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 647 | 3; type strain | DT | 2014-09-04 | Arcanobacterium | wilhelmae | Arcanobacterium | wilhelmae | 2.43 | Arcanobacterium | wilhelmae | 2.25 | A | 1 | 1 | 0 | Arcanobacterium wilhelmae | Arcanobacterium wilhelmae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia | sp-CVUAS-30561 | Boudabousia | sp-CVUAS-30561 | 2.65 | Streptomyces | polychromogenes | 1.39 | A | 1 | 1 | 0 | Boudabousia sp-CVUAS-30561 | Streptomyces polychromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia | canis | Schaalia | canis | 2.41 | Schaalia | canis | 2.26 | A | 1 | 1 | 0 | Schaalia canis | Schaalia canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia | cardiffensis | Schaalia | cardiffensis | 2.49 | Schaalia | cardiffensis | 1.4 | A | 1 | 1 | 0 | Schaalia cardiffensis | Schaalia cardiffensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.45 | Schaalia | hyovaginalis | 2.43 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.48 | Schaalia | hyovaginalis | 2.4 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia | sp-CVUAS-32183 | Schaalia | sp-CVUAS-32183 | 2.55 | Schaalia | hyovaginalis | 1.47 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-32183 | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia | sp-CVUAS-8688 | Schaalia | sp-CVUAS-8688 | 2.57 | Schaalia | sp-CVUAS-8688 | 2.35 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-8688 | Schaalia sp-CVUAS-8688 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31845,2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia | sp-CVUAS-8688 | Schaalia | sp-CVUAS-8688 | 2.73 | Schaalia | sp-CVUAS-8688 | 2.3 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-8688 | Schaalia sp-CVUAS-8688 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella | abortisuis | Trueperella | abortisuis | 2.27 | Trueperella | abortisuis | 2.2 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | abortisuis | Trueperella | abortisuis | 2.36 | Trueperella | abortisuis | 2.16 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella | bernardiae | Trueperella | bernardiae | 2.61 | Trueperella | bernardiae | 2.11 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella | bernardiae | Trueperella | bernardiae | 2.23 | Trueperella | bernardiae | 2.15 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella | bialowiezensis | Trueperella | bialowiezensis | 2.02 | Oceanobacillus | picturae | 1.3 | A | 1 | 1 | 0 | Trueperella bialowiezensis | Oceanobacillus picturae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella | bonasi | Trueperella | bonasi | 2.47 | Trueperella | bonasi | 2.32 | A | 1 | 1 | 0 | Trueperella bonasi | Trueperella bonasi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pecoris | 2.31 | Trueperella | pecoris | 1.98 | A | 1 | 1 | 0 | Trueperella pecoris | Trueperella pecoris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pecoris | 2.37 | Trueperella | pecoris | 1.65 | A | 1 | 1 | 0 | Trueperella pecoris | Trueperella pecoris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.42 | Trueperella | pyogenes | 2.39 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.33 | Trueperella | pyogenes | 2.32 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2379,2 | 3; sequenced | EFex | 2021-04-08 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.29 | Gordonia | paraffinivorans | 2.27 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.27 | Gordonia | paraffinivorans | 2.14 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6272 | 2; API Coryne profile 0150006 | EFex | 2021-03-25 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.54 | Gordonia | paraffinivorans | 2.46 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7903,3 EFex | 3; sequenced | EFex | 2017-07-27 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.52 | Gordonia | paraffinivorans | 2.32 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7903,3 eDT | 3; sequenced | eDT | 2017-07-27 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.16 | Gordonia | paraffinivorans | 2.06 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia | rubripertincta | Gordonia | rubripertincta | 2.61 | Gordonia | rubripertincta | 2.25 | A | 1 | 1 | 0 | Gordonia rubripertincta | Gordonia rubripertincta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10185.3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium | smegmatis | Mycolicibacterium | smegmatis | 2.32 | Tsukamurella | sp | 1.49 | A | 1 | 1 | 0 | Mycolicibacterium smegmatis | Tsukamurella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31296,3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia | cyriacigeorgica | Nocardia | cyriacigeorgica | 2.45 | Nocardia | cyriacigeorgica | 2.28 | A | 1 | 1 | 0 | Nocardia cyriacigeorgica | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia | farcinica | Nocardia | farcinica | 2.36 | Nocardia | kroppenstedtii | 2.26 | B | 0 | 0 | 1 | Nocardia farcinica | Nocardia kroppenstedtii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.46 | Rhodococcus | equi | 2.42 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2755,3 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.39 | Rhodococcus | equi | 2.34 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2998 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.4 | Rhodococcus | equi | 2.36 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33761 sir | 2; 0 | EFEx | 2022-04-27 | Rhodococcus | equi | Rhodococcus | equi_JRA_R.equi-13 | 2.4 | Rhodococcus | equi_JRA_R.equi-53 | 2.37 | B | 0 | 0 | 1 | Rhodococcus equi_JRA_R.equi-13 | Rhodococcus equi_JRA_R.equi-53 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33761 EFex | 2; 0 | EFEx | 2022-04-27 | Rhodococcus | equi | Rhodococcus | equi_JRA_R.equi-53 | 2.47 | Rhodococcus | hoagii | 2.38 | B | 0 | 0 | 1 | Rhodococcus equi_JRA_R.equi-53 | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4057 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.33 | Rhodococcus | equi_JRA_R.equi-53 | 2.31 | B | 0 | 0 | 1 | Rhodococcus equi | Rhodococcus equi_JRA_R.equi-53 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5384,2 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.42 | Rhodococcus | equi_JRA_R.equi-53 | 2.3 | B | 0 | 0 | 1 | Rhodococcus equi | Rhodococcus equi_JRA_R.equi-53 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.41 | Rhodococcus | equi | 2.35 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1447,3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella | sp | Tsukamurella | sp | 2.44 | Tsukamurella | tyrosinosolvens | 1.54 | A | 1 | 1 | 0 | Tsukamurella sp | Tsukamurella tyrosinosolvens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2895,3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium | luteolum | Brevibacterium | luteolum | 2.41 | Brevibacterium | luteolum | 1.92 | A | 1 | 1 | 0 | Brevibacterium luteolum | Brevibacterium luteolum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas | flavigena | Cellulomonas | flavigena | 2.76 | Bacteroides | fragilis | 1.43 | A | 1 | 1 | 0 | Cellulomonas flavigena | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium | sp-234z-2 | Brachybacterium | sp-234z-2 | 2.32 | Glutamicibacter | protophormiae | 1.31 | A | 1 | 1 | 0 | Brachybacterium sp-234z-2 | Glutamicibacter protophormiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium | sp-CVUAS-30560 | Brachybacterium | sp-CVUAS-30560 | 2.66 | Raoultella | ornithinolytica | 1.39 | A | 1 | 1 | 0 | Brachybacterium sp-CVUAS-30560 | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus | sedentarius | Kytococcus | sedentarius | 2.1 | Kytococcus | sedentarius | 1.84 | A | 1 | 1 | 0 | Kytococcus sedentarius | Kytococcus sedentarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus | congolensis | Dermatophilus | congolensis | 2.41 | Dermatophilus | congolensis | 2.19 | A | 1 | 1 | 0 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus | congolensis | Dermatophilus | congolensis | 2.65 | Dermatophilus | congolensis | 2.48 | A | 1 | 1 | 0 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium | flaccumfaciens | Curtobacterium | flaccumfaciens | 2.42 | Curtobacterium | flaccumfaciens | 2.32 | A | 1 | 1 | 0 | Curtobacterium flaccumfaciens | Curtobacterium flaccumfaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia | sp-MUVA-A-508948 | Leifsonia | sp-MUVA-A-508948 | 2.62 | Staphylococcus | aureus | 1.46 | A | 1 | 1 | 0 | Leifsonia sp-MUVA-A-508948 | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium | sp | Microbacterium | phyllosphaerae | 1.84 | Microbacterium | foliorum | 1.84 | B | 0 | 0 | 1 | Microbacterium phyllosphaerae | Microbacterium foliorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10353 DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.09.2011 L8 | Glutamicibacter | nicotianae | Glutamicibacter | nicotianae | 2.75 | Lactobacillus | pentosus | 1.33 | A | 1 | 1 | 0 | Glutamicibacter nicotianae | Lactobacillus pentosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10432 DSM 20127 | 3; type strain | unknown | 08.09.2011, L13, [0034] | Sinomonas | atrocyanea | Sinomonas | atrocyanea | 2.52 | Sinomonas | atrocyanea | 2.24 | A | 1 | 1 | 0 | Sinomonas atrocyanea | Sinomonas atrocyanea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter | protophormiae | Glutamicibacter | protophormiae | 2.55 | Lactobacillus | sakei | 1.37 | A | 1 | 1 | 0 | Glutamicibacter protophormiae | Lactobacillus sakei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria | atrinae | Kocuria | atrinae | 2.27 | Kocuria | atrinae | 2 | A | 1 | 1 | 0 | Kocuria atrinae | Kocuria atrinae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 56a-5 | 3; sequenced | EFEx | 2020-05-20 | Kocuria | gwangalensis | Kocuria | gwangalliensis | 2.73 | Kocuria | carniphila | 1.64 | A | 0 | 0 | 1 | Kocuria gwangalliensis | Kocuria carniphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2014 E22 | Kocuria | palustris | Kocuria | palustris | 2.38 | Kocuria | palustris | 2.33 | A | 1 | 1 | 0 | Kocuria palustris | Kocuria palustris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015 K12 | Micrococcus | luteus | Micrococcus | luteus | 2.37 | Micrococcus | luteus | 2.24 | A | 1 | 1 | 0 | Micrococcus luteus | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20032 | 3; type strain | eDT | 2015-07-09 | Rothia | kristinae | Rothia | kristinae | 2.26 | Rothia | kristinae | 2.25 | A | 1 | 1 | 0 | Rothia kristinae | Rothia kristinae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium | cellulans | Cellulosimicrobium | cellulans | 2.05 | Cellulosimicrobium | cellulans | 1.97 | A | 1 | 1 | 0 | Cellulosimicrobium cellulans | Cellulosimicrobium cellulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 749,2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter | inulinus | Sanguibacter | inulinus | 2.51 | Sanguibacter | inulinus | 2.46 | A | 1 | 1 | 0 | Sanguibacter inulinus | Sanguibacter inulinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108415 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Cutibacterium acnes DSM 108415: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Cutibacterium | acnes | Cutibacterium | acnes | 2.24 | Cutibacterium | acnes | 2.17 | A | 1 | 1 | 0 | Cutibacterium acnes | Cutibacterium acnes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1079 | 3; public strain collection | EFEx | 2017-01-20 | Penicillium | roqueforti | Penicillium | roqueforti | 2.7 | Penicillium | camemberti | 1.55 | A | 1 | 1 | 0 | Penicillium roqueforti | Penicillium camemberti | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10354,2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | verrucosum | 2.35 | Trichophyton | verrucosum | 2.22 | A | 1 | 1 | 0 | Trichophyton verrucosum | Trichophyton verrucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10357.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | verrucosum | 2.65 | Trichophyton | verrucosum | 2.06 | A | 1 | 1 | 0 | Trichophyton verrucosum | Trichophyton verrucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1386 | 3; public strain collection | Fextr | 2016-03-10 | Candida | albicans | Candida | albicans | 2.2 | Candida | albicans | 2.14 | A | 1 | 1 | 0 | Candida albicans | Candida albicans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 011 IGV | 3; sequenced | EFEx | 2016-12-08 | Sarocladium | strictum | Sarocladium | strictum | 2.67 | Sarocladium | strictum | 1.54 | A | 1 | 1 | 0 | Sarocladium strictum | Sarocladium strictum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2520.4 | 3; sequenced (ITS) | F-Extr | 2018-07-26 | Fusarium | sambucinum | Fusarium | sambucinum | 2.74 | Kodamaea | ohmeri | 1.23 | A | 1 | 1 | 0 | Fusarium sambucinum | Kodamaea ohmeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus | amylolyticus | Paenibacillus | amylolyticus | 2.26 | Paenibacillus | amylolyticus | 2.25 | A | 1 | 1 | 0 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus | amylolyticus | Paenibacillus | xylanexedens | 2.22 | Paenibacillus | amylolyticus | 2.17 | B | 0 | 0 | 1 | Paenibacillus xylanexedens | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.53 | Paenibacillus | barengoltzii | 2.5 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.43 | Paenibacillus | barengoltzii | 2.4 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.18 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.25 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus | larvae | Paenibacillus | larvae | 2.46 | Paenibacillus | larvae | 2.38 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.42 | Paenibacillus | larvae | 2.39 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.38 | Paenibacillus | larvae | 2.34 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.18 | Paenibacillus | larvae | 2.18 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus | macerans | Paenibacillus | macerans | 2.28 | Paenibacillus | macerans | 2.07 | A | 1 | 1 | 0 | Paenibacillus macerans | Paenibacillus macerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.12 | Paenibacillus | pabuli | 2.02 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.22 | Paenibacillus | pabuli | 2.15 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus | pasadenensis | Paenibacillus | pasadenensis | 2.52 | Paenibacillus | pasadenensis | 2.05 | A | 1 | 1 | 0 | Paenibacillus pasadenensis | Paenibacillus pasadenensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus | validus | Paenibacillus | validus | 2.26 | Paenibacillus | validus | 2.16 | A | 1 | 1 | 0 | Paenibacillus validus | Paenibacillus validus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium | barati | Clostridium | budayi | 2.31 | Clostridium | baratii | 2.28 | B | 0 | 0 | 1 | Clostridium budayi | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium | butyricum | Clostridium | butyricum | 2.61 | Clostridium | butyricum | 2.48 | A | 1 | 1 | 0 | Clostridium butyricum | Clostridium butyricum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium | chauvoei | Clostridium | chauvoei | 2.13 | Clostridium | chauvoei | 2.01 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002729_1 coch | 2; 0 | DT | 2017-02-28 | Clostridium | cochlearium | Clostridium | cochlearium | 2.51 | Clostridium | cochlearium | 2.5 | A | 1 | 1 | 0 | Clostridium cochlearium | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium | colicanis | Clostridium | colicanis | 2.29 | Bifidobacterium | boum | 1.42 | A | 1 | 1 | 0 | Clostridium colicanis | Bifidobacterium boum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium | novyi | Clostridium | novyi | 2.03 | Clostridium | novyi | 1.96 | A | 1 | 1 | 0 | Clostridium novyi | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium | septicum | Clostridium | septicum | 2.44 | Clostridium | septicum | 2.26 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium | sporogenes | Clostridium | sporogenes | 2.34 | Clostridium | sporogenes | 2.21 | A | 1 | 1 | 0 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141015223 aerob | 2; 0 | DT | 2014-11-27 | Clostridium | tertium | Clostridium | tertium | 2.44 | Clostridium | tertium | 2.22 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium | tetani | Clostridium | tetani | 2.23 | Clostridium | tetani | 2.16 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya | limosa | Hathewaya | limosa | 2.29 | Hathewaya | limosa | 2.25 | A | 1 | 1 | 0 | Hathewaya limosa | Hathewaya limosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium | sordellii | Paeniclostridium | sordellii | 2.32 | Paeniclostridium | sordellii | 2.27 | A | 1 | 1 | 0 | Paeniclostridium sordellii | Paeniclostridium sordellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium | bifermentans | Paraclostridium | benzoelyticum | 2.05 | Paraclostridium | benzoelyticum | 2.04 | A | 0 | 0 | 1 | Paraclostridium benzoelyticum | Paraclostridium benzoelyticum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 44 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N11 | Bacillus | badius | Bacillus | badius | 2.06 | Bacillus | badius | 1.99 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.03.2017 I18 | Bacillus | ectoiniformans | Bacillus | ectoiniformans | 2.78 | Streptococcus | dysgalactiae | 1.49 | A | 1 | 1 | 0 | Bacillus ectoiniformans | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.45 | Cytobacillus | firmus | 2.24 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 37 B 88 | 3; reference strain collection | 0 | 09.02.2015, L16 | Cytobacillus | firmus | Cytobacillus | firmus | 2.06 | Cytobacillus | firmus | 1.98 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N14 | Niallia | circulans | Niallia | circulans | 1.98 | Niallia | circulans | 1.88 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 N5 | Virgibacillus | pantothenticus | Virgibacillus | panthotenticus | 2.82 | Virgibacillus | pantothenticus | 2.01 | B | 0 | 0 | 1 | Virgibacillus panthotenticus | Virgibacillus pantothenticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31071,2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 2.23 | Alkalihalobacillus | clausii | 1.69 | A | 1 | 1 | 0 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015; H10 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 2.85 | Alkalihalobacillus | clausii | 1.71 | A | 1 | 1 | 0 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 13.02.2015, K10, [0034] | Bacillus | alveayuensis | Bacillus | alveayuensis | 2.52 | Peribacillus | psychrosaccharolyticus | 1.35 | A | 1 | 1 | 0 | Bacillus alveayuensis | Peribacillus psychrosaccharolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7 RKI 2015 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCA/TA2 (A) | 16.02.2015, K16, [0034] | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.24 | Bacillus | velezensis | 2.17 | B | 0 | 0 | 1 | Bacillus amyloliquefaciens | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7 CVUAS | 3; type strain | DT | 2014-01-29 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.15 | Bacillus | amyloliquefaciens | 2.15 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E06 | Bacillus | anthracis | Bacillus | luti | 2.1 | Bacillus | cereus | 2.05 | B | 0 | 0 | 1 | Bacillus luti | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; E04 | Bacillus | anthracis | Bacillus | cereus | 2.36 | Bacillus | cereus | 2.25 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 004 | 2; 0 | DT | 2014-01-22 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.14 | Bacillus | atrophaeus | 2.14 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 L8 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.43 | Bacillus | atrophaeus | 2.36 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H20 | Bacillus | badius | Bacillus | badius | 2.45 | Bacillus | badius | 2.26 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | cereus | 2.33 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1107 | 2; 0 | eDT | 27.01.2016 | Bacillus | cereus | Bacillus | pseudomycoides | 2.45 | Bacillus | cereus | 2.41 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.05.2015; E09 | Bacillus | tropicus | Bacillus | pacificus | 2.42 | Bacillus | tropicus | 2.4 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2833 CVUAS | 4; 16S, draft genome [0009] | eDT | 2016-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.53 | Bacillus | cytotoxicus | 2.48 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NVH 391-98 | 3; type strain | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.46 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.11.2017 F14 | Bacillus | halotolerans | Bacillus | mojavensis | 2.8 | Bacillus | axarquiensis | 2.55 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus axarquiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.12.2017 I8 | Bacillus | halotolerans | Bacillus | halotolerans | 2.73 | Bacillus | mojavensis | 2.23 | B | 0 | 0 | 1 | Bacillus halotolerans | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.12.2016 [H3] | Bacillus | inaquosorum | Bacillus | inaquosorum | 2.78 | Bacillus | subtilis | 2.27 | B | 0 | 0 | 1 | Bacillus inaquosorum | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.02.2016 C20 | Bacillus | infantis | Bacillus | infantis | 2.46 | Bacillus | infantis | 2.39 | A | 1 | 1 | 0 | Bacillus infantis | Bacillus infantis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus | licheniformis | Bacillus | licheniformis | 2.17 | Bacillus | licheniformis | 2.15 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus | licheniformis | Bacillus | licheniformis | 2.27 | Bacillus | licheniformis | 1.81 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K15 | Bacillus | mojavensis | Bacillus | mojavensis | 2.78 | Bacillus | halotolerans | 2.36 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus halotolerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F4 | Bacillus | mojavensis | Bacillus | halotolerans | 2.42 | Bacillus | mojavensis | 2.33 | B | 0 | 0 | 1 | Bacillus halotolerans | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2048 CVUAS | 3; type strain | eDT | 09.12.2015 | Bacillus | mycoides | Bacillus | mycoides | 2.45 | Bacillus | mycoides | 2.22 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | mycoides | Bacillus | mycoides | 2.04 | Bacillus | mycoides | 2.01 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11778 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E8, [0034] | Bacillus | pacificus | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.06 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.53 | Bacillus | cereus | 2.48 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 12443 CVUAS | 3; public strain collection | eDT | 09.12.2015 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.64 | Bacillus | pseudomycoides | 2.4 | A | 1 | 1 | 0 | Bacillus pseudomycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B017 CVUAS | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus | pumilus | Bacillus | pumilus | 2.15 | Bacillus | pumilus | 1.96 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 340 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | pumilus | Bacillus | pumilus | 2.04 | Bacillus | pumilus | 1.95 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.05.2018, N12 | Bacillus | siamensis | Bacillus | siamensis | 2.68 | Bacillus | velezensis | 2.2 | B | 0 | 0 | 1 | Bacillus siamensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2016, K18 | Bacillus | sonorensis | Bacillus | sonorensis | 2.6 | Bacillus | sp. | 2.05 | B | 0 | 0 | 1 | Bacillus sonorensis | Bacillus sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12369 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Bacillus | sp-DSM-12369 | Bacillus | sp. | 2.83 | Bacillus | sonorensis | 2.03 | B | 0 | 0 | 1 | Bacillus sp. | Bacillus sonorensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 6633 CVUAS | 3; public strain collection | eDT | 2015-07-22 | Bacillus | spizizenii | Bacillus | subtilis | 2.08 | Bacillus | subtilis | 1.96 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 618 | 4; public strain collection | eDT | 27.01.2016 | Bacillus | spizizenii | Bacillus | subtilis | 2.09 | Bacillus | subtilis | 2.04 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 110 B068 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.33 | Bacillus | subtilis | 2.15 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1092 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.23 | Bacillus | subtilis | 2.2 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.64 | Bacillus | thuringiensis | 2.52 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.67 | Bacillus | thuringiensis | 2.64 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.55 | Bacillus | thuringiensis | 2.51 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.62 | Bacillus | thuringiensis | 2.54 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.57 | Bacillus | thuringiensis | 2.57 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6890 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2015, L4 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.42 | Bacillus | proteolyticus | 2.25 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 07.11.2006 K4 | Bacillus | thuringiensis | Bacillus | cereus | 2.33 | Bacillus | paranthracis | 2.25 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2786,2 | 3; sequenced | DT | 2013-07-31 | Bacillus | timorensis | Bacillus | timorensis | 2.59 | Bacillus | timonensis | 1.49 | A | 1 | 1 | 0 | Bacillus timorensis | Bacillus timonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.07.2017, H17 | Bacillus | vallismortis | Bacillus | vallismortis | 2.37 | Bacillus | inaquosorum | 1.97 | A | 1 | 1 | 0 | Bacillus vallismortis | Bacillus inaquosorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCEB 627 DSM 12 | 3; type strain | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.19 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 37 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.33 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.02.2016 K19 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.65 | Cytobacillus | horneckiae | 2.11 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.04.2014 H5 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.46 | Cytobacillus | horneckiae | 2.32 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus | kochii | Cytobacillus | kochii | 2.33 | Cytobacillus | kochii | 1.86 | A | 1 | 1 | 0 | Cytobacillus kochii | Cytobacillus kochii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 19.02.2015, E18 | Domibacillus | aminovorans | Domibacillus | aminovorans | 2.81 | Pseudomonas | aeruginosa | 1.37 | A | 1 | 1 | 0 | Domibacillus aminovorans | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 15.02.2015, E8 | Domibacillus | aminovorans | Domibacillus | aminovorans | 2.63 | Staphylococcus | lugdunensis | 1.46 | A | 1 | 1 | 0 | Domibacillus aminovorans | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.02.2015 H5 | Evansella | clarkii | Evansella | clarkii | 2.81 | Bacillus | clarkii | 2.35 | C | 0 | 0 | 1 | Evansella clarkii | Bacillus clarkii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F7 | Evansella | clarkii | Evansella | clarkii | 2.77 | Evansella | clarkii | 2.29 | A | 1 | 1 | 0 | Evansella clarkii | Evansella clarkii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus | stearothermophilus | Geobacillus | stearothermophilus | 2.39 | Geobacillus | stearothermophilus | 2.29 | A | 1 | 1 | 0 | Geobacillus stearothermophilus | Geobacillus stearothermophilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 I10; (single sp is included in the corresponding msp) | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.39 | Alkalihalobacillus | halodurans | 2.19 | A | 0 | 0 | 1 | Alkalihalobacillus halodurans | Alkalihalobacillus halodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 E19; (single sp is included in the corresponding msp) | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.74 | Alkalihalobacillus | okuhidensis | 2.37 | B | 0 | 0 | 1 | Alkalihalobacillus halodurans | Alkalihalobacillus okuhidensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia | sporothermodurans | Bacillus | sporothermodurans | 2.17 | Bacillus | sporothermodurans | 2.14 | A | 0 | 0 | 1 | Bacillus sporothermodurans | Bacillus sporothermodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 10599 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015, H9 | Heyndrickxia | sporothermodurans | Heyndrickxia | sporothermodurans | 2.79 | Heyndrickxia | sporothermodurans | 2.35 | A | 1 | 1 | 0 | Heyndrickxia sporothermodurans | Heyndrickxia sporothermodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017; N12 | Lederbergia | lenta | Lederbergia | lenta | 2.43 | Lederbergia | lenta | 2.33 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9 CVUAS | 3; type strain | DT | 2014-01-22 | Lederbergia | lenta | Lederbergia | lenta | 2.22 | Lederbergia | lenta | 2.08 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus | bataviensis | Neobacillus | bataviensis | 2.76 | Neobacillus | bataviensis | 2.29 | A | 1 | 1 | 0 | Neobacillus bataviensis | Neobacillus bataviensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017 N13 | Neobacillus | cucumis | Neobacillus | cucumis | 2.74 | Neobacillus | bataviensis | 2.02 | B | 0 | 0 | 1 | Neobacillus cucumis | Neobacillus bataviensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E14, [0034] | Niallia | circulans | Niallia | circulans | 1.9 | Niallia | circulans | 1.65 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2582,2 | 3; sequenced | eDT | 2015-09-23 | Niallia | nealsonii | Niallia | nealsonii | 2.34 | Niallia | taxi | 2.04 | B | 0 | 0 | 1 | Niallia nealsonii | Niallia taxi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K20 | Peribacillus | muralis | Peribacillus | muralis | 2.38 | Peribacillus | simplex | 2.09 | B | 0 | 0 | 1 | Peribacillus muralis | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016; K20 | Peribacillus | psychrosaccharolyticus | Peribacillus | psychrosaccharolyticus | 2.21 | Staphylococcus | simiae | 1.44 | A | 1 | 1 | 0 | Peribacillus psychrosaccharolyticus | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus | simplex | Peribacillus | simplex | 2.08 | Peribacillus | simplex | 1.74 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2016 H13 | Peribacillus | simplex | Peribacillus | simplex | 2.57 | Peribacillus | simplex | 1.89 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 15.12.2016, K15, [0096] | Peribacillus | simplex | Peribacillus | simplex | 2.68 | Peribacillus | simplex | 2.19 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016 K15 | Priestia | flexa | Priestia | flexa | 2.54 | Priestia | flexa | 2.17 | A | 1 | 1 | 0 | Priestia flexa | Priestia flexa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H18 | Priestia | flexa | Priestia | megaterium | 2.01 | Priestia | megaterium | 1.95 | A | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015 H21 | Priestia | megaterium | Priestia | megaterium | 1.94 | Priestia | megaterium | 1.92 | B | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia | megaterium | Priestia | megaterium | 2.06 | Priestia | megaterium | 2.06 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E3 | Salibacillus | salarius | Salibacterium | salarium | 2.05 | Aeromonas | salmonicida | 1.35 | A | 0 | 0 | 1 | Salibacterium salarium | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F13 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.21 | Sutcliffiella | cohnii | 2.13 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 K8 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.8 | Sutcliffiella | cohnii | 2.03 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus | saccharophilus | Terribacillus | saccharophilus | 2.4 | Lodderomyces | elongisporus | 1.45 | A | 1 | 1 | 0 | Terribacillus saccharophilus | Lodderomyces elongisporus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus | proomii | Virgibacillus | proomii | 2.02 | Virgibacillus | proomii | 1.99 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3616,2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus | proomii | Virgibacillus | proomii | 2.2 | Virgibacillus | proomii | 1.83 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H15 | Weizmannia | coagulans | Weizmannia | coagulans | 2.45 | Weizmannia | coagulans | 2.12 | A | 1 | 1 | 0 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 71,2 | 3; sequenced | DT | 2013-12-13 | Weizmannia | ginsengihumi | Weizmannia | ginsengihumi | 2.4 | Weizmannia | ginsengihumi | 2.19 | A | 1 | 1 | 0 | Weizmannia ginsengihumi | Weizmannia ginsengihumi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.48 | Kurthia | gibsonii | 2.47 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.34 | Kurthia | gibsonii | 2.29 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia | zopfii | Kurthia | zopfii | 2.31 | Kurthia | zopfii | 1.66 | A | 1 | 1 | 0 | Kurthia zopfii | Kurthia zopfii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.34 | Lysinibacillus | fusiformis | 2.23 | A | 1 | 1 | 0 | Lysinibacillus fusiformis | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.44 | Lysinibacillus | sphaericus | 2.36 | B | 0 | 0 | 1 | Lysinibacillus fusiformis | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7192,2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus | macroides | Lysinibacillus | macroides | 2.64 | Lysinibacillus | xylanilyticus | 2.23 | B | 0 | 0 | 1 | Lysinibacillus macroides | Lysinibacillus xylanilyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.34 | Lysinibacillus | sphaericus | 2.31 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.11 | Lysinibacillus | sphaericus | 1.96 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus | saudimassiliensis | Metalysinibacillus | saudimassiliensis | 2.53 | Lysinibacillus | saudimassiliensis | 2.01 | C | 0 | 0 | 1 | Metalysinibacillus saudimassiliensis | Lysinibacillus saudimassiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus | odysseyi | Solibacillus | odysseyi | 2.2 | Ureibacillus | sinduriensis | 1.47 | A | 1 | 1 | 0 | Solibacillus odysseyi | Ureibacillus sinduriensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium | acetylicum | Exiguobacterium | acetylicum | 2.63 | Exiguobacterium | enclense | 2.1 | B | 0 | 0 | 1 | Exiguobacterium acetylicum | Exiguobacterium enclense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium | enclense | Exiguobacterium | enclense | 2.19 | Exiguobacterium | acetylicum | 2.03 | B | 0 | 0 | 1 | Exiguobacterium enclense | Exiguobacterium acetylicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium | profundum | Exiguobacterium | profundum | 2.44 | Exiguobacterium | sp[3] | 2.41 | B | 0 | 0 | 1 | Exiguobacterium profundum | Exiguobacterium sp[3] | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.36 | Brochothrix | thermosphacta | 2.32 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.39 | Brochothrix | thermosphacta | 2.37 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria | aquatica | Listeria | aquatica | 2.33 | Listeria | aquatica | 2.28 | A | 1 | 1 | 0 | Listeria aquatica | Listeria aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria | aquatica | Listeria | aquatica | 2.5 | Listeria | aquatica | 2.43 | A | 1 | 1 | 0 | Listeria aquatica | Listeria aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.25 | Listeria | fleischmannii | 2.23 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LU2006-3 EFex | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.2 | Listeria | fleischmannii | 2.19 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105447 EFex | 3; type strain | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.32 | Listeria | grayi | 2.32 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20596 EFex | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.25 | Listeria | grayi | 2.12 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10132,2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | innocua | Listeria | innocua | 2.45 | Listeria | innocua | 2.44 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8569 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | innocua | Listeria | innocua | 2.48 | Listeria | innocua | 2.46 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria | ivanovii | Listeria | ivanovii | 2.33 | Listeria | ivanovii | 2.32 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32100 | 2; sequenced | EFex | 2019-12-13 | Listeria | ivanovii | Listeria | ivanovii | 2.41 | Listeria | ivanovii | 2.35 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 23813 EFex | 3; type strain | EFEx | 2017-02-15 | Listeria | marthii | Listeria | marthii | 2.38 | Listeria | monocytogenes | 2.28 | B | 0 | 0 | 1 | Listeria marthii | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.42 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 873,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.43 | Listeria | monocytogenes | 2.41 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 240,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.29 | Listeria | monocytogenes | 2.24 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0206 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0066 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.19 | Listeria | monocytogenes | 2.19 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.22 | Listeria | monocytogenes | 2.21 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.37 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7366,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.54 | Listeria | monocytogenes | 2.51 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria | newyorkensis | Listeria | newyorkensis | 2.47 | Listeria | sp | 2.17 | B | 0 | 0 | 1 | Listeria newyorkensis | Listeria sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 22097 EFEx | 3; type strain | EFEx | 2017-02-15 | Listeria | rocourtiae | Listeria | rocourtiae | 2.45 | Listeria | rocourtiae | 2.1 | A | 1 | 1 | 0 | Listeria rocourtiae | Listeria rocourtiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2763 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | seeligeri | Listeria | seeligeri | 2.45 | Listeria | seeligeri | 2.45 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | seeligeri | Listeria | seeligeri | 2.36 | Listeria | seeligeri | 2.3 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria | sp | Listeria | sp | 2.31 | Listeria | sp | 2.29 | A | 1 | 1 | 0 | Listeria sp | Listeria sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2412,3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | welshimeri | Listeria | welshimeri | 2.38 | Listeria | welshimeri | 2.28 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria | welshimeri | Listeria | welshimeri | 2.44 | Listeria | welshimeri | 2.41 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.31 | Brevibacillus | laterosporus | 1.95 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.42 | Brevibacillus | laterosporus | 2.25 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B210 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 K19 | Paenibacillus | alvei | Paenibacillus | alvei | 1.82 | Paenibacillus | curdlanolyticus | 1.45 | B | 0 | 0 | 1 | Paenibacillus alvei | Paenibacillus curdlanolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus | chibensis | Paenibacillus | chibensis | 2.28 | Paenibacillus | chibensis | 2.15 | A | 1 | 1 | 0 | Paenibacillus chibensis | Paenibacillus chibensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus | nuruki | Paenibacillus | nuruki | 2.54 | Ochrobactrum | anthropi | 1.29 | A | 1 | 1 | 0 | Paenibacillus nuruki | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.09.2013 C8 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 1.9 | Paenibacillus | polymyxa | 1.78 | B | 0 | 0 | 1 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.11.2017 C18 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 2.81 | Paenibacillus | polymyxa | 2.1 | A | 1 | 1 | 0 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus | sp-CVUAS-6369 | Paenibacillus | sp-CVUAS-6369 | 2.08 | Lentilactobacillus | parabuchneri | 1.35 | A | 1 | 1 | 0 | Paenibacillus sp-CVUAS-6369 | Lentilactobacillus parabuchneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus | sp-E8a | Paenibacillus | sp-E8a | 2.28 | Paenibacillus | amylolyticus | 2.24 | B | 0 | 0 | 1 | Paenibacillus sp-E8a | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus | taichungensis | Paenibacillus | taichungensis | 2.29 | Paenibacillus | taichungensis | 1.9 | A | 1 | 1 | 0 | Paenibacillus taichungensis | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | FLI003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus | taichungensis | Paenibacillus | taichungensis | 2.81 | Paenibacillus | taichungensis | 1.87 | A | 1 | 1 | 0 | Paenibacillus taichungensis | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus | tundrae | Paenibacillus | tundrae | 2.09 | Paenibacillus | taichungensis | 1.87 | A | 1 | 1 | 0 | Paenibacillus tundrae | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | halophilus | Jeotgalicoccus | halophilus | 2.45 | Listeria | innocua | 1.68 | A | 1 | 1 | 0 | Jeotgalicoccus halophilus | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | huakuii | Jeotgalicoccus | huakuii | 2.6 | Macrococcus | canis | 1.37 | A | 1 | 1 | 0 | Jeotgalicoccus huakuii | Macrococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | psychrophilus | Jeotgalicoccus | psychrophilus | 2.37 | Jeotgalicoccus | psychrophilus | 1.84 | A | 1 | 1 | 0 | Jeotgalicoccus psychrophilus | Jeotgalicoccus psychrophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus | carouselicus | Macrococcus | carouselicus | 2.59 | Staphylococcus | simulans | 1.4 | A | 1 | 1 | 0 | Macrococcus carouselicus | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 14.03.2013 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.08 | Macrococcus | caseolyticus | 1.75 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.19 | Macrococcus | caseolyticus | 1.87 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus | lentus | Mammaliicoccus | lentus | 2.4 | Staphylococcus | lentus | 1.8 | C | 0 | 0 | 1 | Mammaliicoccus lentus | Staphylococcus lentus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 2.76 | Staphylococcus | sciuri | 2.37 | C | 0 | 0 | 1 | Mammaliicoccus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus | sciuri | Staphylococcus | sciuri | 1.99 | Staphylococcus | sciuri | 1.68 | B | 0 | 0 | 1 | Staphylococcus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9937,2 EFex | 3; 16S rDNA-sequencing | EFex | 2016-04-06 | Mammaliicoccus | vitulinus | Mammaliicoccus | vitulinus | 2.31 | Staphylococcus | vitulinus | 1.87 | C | 0 | 0 | 1 | Mammaliicoccus vitulinus | Staphylococcus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 343,2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus | vitulinus | Mammaliicoccus | vitulinus | 2.23 | Mammaliicoccus | vitulinus | 2.23 | A | 1 | 1 | 0 | Mammaliicoccus vitulinus | Mammaliicoccus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5687 | 3; sequenced | DT | 2013-04-16 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | agnetis | 2.14 | B | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus agnetis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TUM 7385 | 3; sequenced | DT | 02.10.2013 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | hyicus | 1.96 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.39 | Staphylococcus | argenteus | 2.38 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.37 | Staphylococcus | argenteus | 2.28 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.01 | Staphylococcus | arlettae | 1.72 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.13 | Staphylococcus | arlettae | 1.63 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 29213 CVUAS | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9914,2 sir | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus | auricularis | Staphylococcus | auricularis | 2.14 | Staphylococcus | auricularis | 1.85 | A | 1 | 1 | 0 | Staphylococcus auricularis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20326 (T) | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.25 | Staphylococcus | capitis | 2.13 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.2 | Staphylococcus | capitis | 2.19 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20608 (T) | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.16 | Staphylococcus | caprae | 2.06 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.06 | Staphylococcus | caprae | 2.03 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.18 | Staphylococcus | carnosus | 1.55 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.24 | Staphylococcus | carnosus | 1.46 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.26 | Staphylococcus | hyicus | 1.85 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.28 | Staphylococcus | chromogenes | 1.91 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.36 | Staphylococcus | xylosus | 1.36 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.72 | Staphylococcus | cohnii | 2.25 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | x Sta | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012 K10 | Staphylococcus | delphini | Staphylococcus | delphini | 2.32 | Staphylococcus | delphini | 2.01 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus | delphini | Staphylococcus | delphini | 2.53 | Staphylococcus | delphini | 1.9 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus | devriesei | Staphylococcus | devriesei | 2.33 | Staphylococcus | capitis | 1.46 | A | 1 | 1 | 0 | Staphylococcus devriesei | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.09 | Staphylococcus | epidermidis | 2.08 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.31 | Staphylococcus | epidermidis | 1.99 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | equorum | Staphylococcus | equorum | 2.16 | Staphylococcus | equorum | 1.8 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20674 T | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | equorum | Staphylococcus | equorum | 2.11 | Staphylococcus | equorum | 1.88 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | felis | Staphylococcus | felis | 2.13 | Staphylococcus | felis | 1.96 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | felis | Staphylococcus | felis | 2.14 | Staphylococcus | felis | 2 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.08.2011 H9 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.46 | Staphylococcus | haemolyticus | 2.43 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.12 | Staphylococcus | haemolyticus | 2.04 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N6 [0034] | Staphylococcus | hominis | Staphylococcus | hominis | 2.27 | Staphylococcus | hominis | 2.23 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 789,2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus | hominis | Staphylococcus | hominis | 2.13 | Staphylococcus | hominis | 1.96 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.2 | Staphylococcus | hyicus | 2.14 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3767 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.38 | Staphylococcus | hyicus | 2.12 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, N8, [0034] | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.15 | Staphylococcus | intermedius | 2.02 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N7, [0034] | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.07 | Staphylococcus | intermedius | 1.97 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 05.04.2012, N5, [0034] | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.4 | Staphylococcus | lugdunensis | 2.4 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, H16 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.5 | Staphylococcus | lugdunensis | 2.43 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | microti | Staphylococcus | microti | 2.48 | Staphylococcus | microti | 1.85 | A | 1 | 1 | 0 | Staphylococcus microti | Staphylococcus microti | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | pasteuri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 2.17 | A | 0 | 0 | 1 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus | pasteuri | Staphylococcus | pasteuri | 2 | Staphylococcus | warneri | 1.81 | A | 1 | 1 | 0 | Staphylococcus pasteuri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.03.2016, H05, [0034] | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.76 | Staphylococcus | pseudintermedius | 2.15 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | pseudintermedius | Staphylococcus | delphini | 2.1 | Staphylococcus | pseudintermedius | 2.05 | B | 0 | 0 | 1 | Staphylococcus delphini | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.03 | Staphylococcus | kloosii | 1.43 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus kloosii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 503,2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.18 | Staphylococcus | saprophyticus | 2.06 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus | schleiferi | Staphylococcus | schleiferi | 2.15 | Staphylococcus | schleiferi | 2.02 | A | 1 | 1 | 0 | Staphylococcus schleiferi | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus | schweitzeri | Staphylococcus | schweitzeri | 2.48 | Staphylococcus | schweitzeri | 2.17 | A | 1 | 1 | 0 | Staphylococcus schweitzeri | Staphylococcus schweitzeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | simulans | Staphylococcus | simulans | 2.13 | Staphylococcus | simulans | 2.11 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | simulans | Staphylococcus | simulans | 2.17 | Staphylococcus | simulans | 2.04 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2439,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.26 | Staphylococcus | succinus | 2.23 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 761,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.28 | Staphylococcus | succinus | 2.21 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus | warneri | Staphylococcus | warneri | 2.07 | Staphylococcus | warneri | 2.02 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20316 CVUAS | 3; type strain | DT | 2013-10-02 | Staphylococcus | warneri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 1.96 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.04 | Staphylococcus | xylosus | 2.02 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.88 | Staphylococcus | xylosus | 1.76 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus | sp-185-U1-2 | Aerococcus | sp-185-U1-2 | 2.41 | Xanthomonas | hortorum | 1.38 | A | 1 | 1 | 0 | Aerococcus sp-185-U1-2 | Xanthomonas hortorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus | suis | Aerococcus | suis | 2.36 | Aerococcus | suis | 2.3 | A | 1 | 1 | 0 | Aerococcus suis | Aerococcus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus | urinaeequi | Aerococcus | urinaeequi | 2.49 | Aerococcus | viridans | 1.72 | A | 1 | 1 | 0 | Aerococcus urinaeequi | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 27.09.2016 | Aerococcus | viridans | Aerococcus | viridans | 2.33 | Aerococcus | viridans | 2.04 | A | 1 | 1 | 0 | Aerococcus viridans | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter | phocae | Atopobacter | phocae | 2.33 | Atopobacter | phocae | 2.31 | A | 1 | 1 | 0 | Atopobacter phocae | Atopobacter phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2442,2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium | divergens | Carnobacterium | divergens | 2.44 | Carnobacterium | divergens | 2.1 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium | divergens | Carnobacterium | divergens | 2.62 | Carnobacterium | divergens | 2.37 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6273 | 3; sequenced | eDT | 13.01.2016 | Carnobacterium | maltaromaticum | Carnobacterium | maltaromaticum | 2.55 | Carnobacterium | maltaromaticum | 2.54 | A | 1 | 1 | 0 | Carnobacterium maltaromaticum | Carnobacterium maltaromaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1519,3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella | adiacens | Granulicatella | adiacens | 2.3 | Granulicatella | adiacens | 2.24 | A | 1 | 1 | 0 | Granulicatella adiacens | Granulicatella adiacens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2616 | 3; sequenced | DT | 25.09.2013 | Marinilactibacillus | sp-CVUAS-2616 | Marinilactibacillus | sp-CVUAS-2616 | 2.02 | Marinilactibacillus | sp-CVUAS-2616 | 1.83 | A | 1 | 1 | 0 | Marinilactibacillus sp-CVUAS-2616 | Marinilactibacillus sp-CVUAS-2616 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8980,3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus | sp-CVUAS-2616 | Marinilactibacillus | sp-CVUAS-2616 | 2.54 | Marinilactibacillus | sp-CVUAS-2616 | 2.44 | A | 1 | 1 | 0 | Marinilactibacillus sp-CVUAS-2616 | Marinilactibacillus sp-CVUAS-2616 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16 Ec RKI | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011 H20 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.3 | Enterococcus | casseliflavus | 2.2 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H12 | Enterococcus | durans | Enterococcus | hirae | 2.09 | Enterococcus | hirae | 1.8 | A | 0 | 0 | 1 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H13 | Enterococcus | durans | Enterococcus | durans | 2 | Enterococcus | durans | 1.94 | A | 1 | 1 | 0 | Enterococcus durans | Enterococcus durans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.07.2011 N8 | Enterococcus | faecium | Enterococcus | faecium | 2.09 | Enterococcus | faecium | 2.05 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N13 | Enterococcus | gallinaru | Enterococcus | gallinarum | 2.04 | Enterococcus | gallinarum | 1.98 | A | 0 | 0 | 1 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N3 | Enterococcus | hirae | Enterococcus | hirae | 2.33 | Enterococcus | hirae | 2.16 | A | 1 | 1 | 0 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.33 | Enterococcus | casseliflavus | 2.32 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5030 | 3; sequenced | eDT | 23.07.2014 | Enterococcus | cecorum | Enterococcus | cecorum | 2.34 | Enterococcus | cecorum | 2.29 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8897 | 3; sequenced | eDT | 23.07.2014 | Enterococcus | cecorum | Enterococcus | cecorum | 2.82 | Enterococcus | cecorum | 1.41 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus | faecalis | Enterococcus | faecalis | 2.51 | Enterococcus | faecalis | 2.41 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus | faecalis | Enterococcus | faecalis | 2.39 | Enterococcus | faecalis | 2.38 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MR11-207 LVU04 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus | faecium | Enterococcus | faecium | 2.42 | Enterococcus | faecium | 2.1 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus | gallinarum | Enterococcus | gallinarum | 2.42 | Enterococcus | gallinarum | 2.38 | A | 1 | 1 | 0 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus | solitarius | Tetragenococcus | solitarius | 2.14 | Halomonas | elongata | 1.38 | A | 1 | 1 | 0 | Tetragenococcus solitarius | Halomonas elongata | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1549,2 | 3; sequenced | DT | 2016-09-27 | Vagococcus | salmoninarum | Vagococcus | salmoninarum | 2.3 | Vagococcus | salmoninarum | 2.09 | A | 1 | 1 | 0 | Vagococcus salmoninarum | Vagococcus salmoninarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus | acidophilus | Lactobacillus | acidophilus | 2.32 | Lactobacillus | acidophilus | 2.24 | A | 1 | 1 | 0 | Lactobacillus acidophilus | Lactobacillus acidophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus | johnsonii | Lactobacillus | johnsonii | 2.42 | Lactobacillus | johnsonii | 2.35 | A | 1 | 1 | 0 | Lactobacillus johnsonii | Lactobacillus johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32277,2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus | sp-CVUAS-32277,2 | Lactobacillus | sp-CVUAS-32277,2 | 2.64 | Limosilactobacillus | gorillae | 1.25 | A | 1 | 1 | 0 | Lactobacillus sp-CVUAS-32277,2 | Limosilactobacillus gorillae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus | claussenii | Pediococcus | clausenii | 2.41 | Pediococcus | claussenii | 2.25 | B | 0 | 0 | 1 | Pediococcus clausenii | Pediococcus claussenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1213 | 2; 0 | DT | 2013-07-09 | Leuconostoc | mesenteroides | Leuconostoc | mesenteroides | 2.28 | Leuconostoc | mesenteroides | 2.09 | A | 1 | 1 | 0 | Leuconostoc mesenteroides | Leuconostoc mesenteroides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1146 | 2; 0 | DT | 2013-07-09 | Lactococcus | lactis | Lactococcus | lactis | 2.42 | Lactococcus | lactis | 2.4 | A | 1 | 1 | 0 | Lactococcus lactis | Lactococcus lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.45 | Streptococcus | agalactiae | 2.44 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6784 CVUAS | 3; public strain collection | eDT | 2019-07-09 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.6 | Streptococcus | agalactiae | 2.56 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus | canis | Streptococcus | canis | 2.53 | Streptococcus | canis | 2.41 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL Mi-845_17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus | canis | Streptococcus | canis | 2.5 | Streptococcus | canis | 2.49 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.72 | Streptococcus | castoreus | 2.64 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17536 CVUAS | 3; type strain | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.68 | Streptococcus | castoreus | 2.61 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus | catagoni | Streptococcus | catagoni | 2.83 | Streptococcus | catagoni | 2.81 | A | 1 | 1 | 0 | Streptococcus catagoni | Streptococcus catagoni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 99-1-2017 | 3; type strain | EFEx | 2018-05-03 | Streptococcus | catagoni | Streptococcus | catagoni | 2.81 | Streptococcus | catagoni | 2.8 | A | 1 | 1 | 0 | Streptococcus catagoni | Streptococcus catagoni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.15 | Streptococcus | dysgalactiae | 2.14 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.26 | Streptococcus | dysgalactiae | 2.12 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.43 | Streptococcus | dysgalactiae | 2.35 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.46 | Streptococcus | dysgalactiae | 2.33 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32462,2 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi | 2.26 | Streptococcus | equi | 2.18 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32770 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi | 2.42 | Streptococcus | equi | 2.41 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus | equi | Streptococcus | equi | 2.42 | Streptococcus | equi | 2.35 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus | equi | Streptococcus | equi | 2.25 | Streptococcus | equi | 2.23 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus | iniae | Streptococcus | iniae | 2.02 | Streptococcus | iniae | 1.89 | A | 1 | 1 | 0 | Streptococcus iniae | Streptococcus iniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus | orisasini | Streptococcus | orisasini | 2.86 | Streptococcus | orisasini | 2.44 | A | 1 | 1 | 0 | Streptococcus orisasini | Streptococcus orisasini | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | parauberis | Streptococcus | parauberis_JRA_G+R-39 | 2.18 | Streptococcus | parauberis | 2.1 | B | 0 | 0 | 1 | Streptococcus parauberis_JRA_G+R-39 | Streptococcus parauberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | phocae | Streptococcus | phocae | 2.33 | Streptococcus | phocae | 2.3 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 12-13862-001-01 LVU11 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus | pluranimalium | Streptococcus | pluranimalium | 2.2 | Streptococcus | pluranimalium | 2.18 | A | 1 | 1 | 0 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.48 | Streptococcus | pneumoniae | 2.44 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.16 | Streptococcus | pneumoniae | 2.16 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.29 | Streptococcus | pneumoniae | 2.23 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.49 | Streptococcus | pneumoniae | 2.46 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus | porcinus | Streptococcus | porcinus | 2.27 | Streptococcus | porcinus | 2.1 | A | 1 | 1 | 0 | Streptococcus porcinus | Streptococcus porcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.04.2008, K11, [0034] | Streptococcus | pyogenes | Streptococcus | pyogenes | 2.43 | Streptococcus | pyogenes | 2.4 | A | 1 | 1 | 0 | Streptococcus pyogenes | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10145,3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus | salivarius | Streptococcus | salivarius | 2.13 | Streptococcus | salivarius | 2.13 | A | 1 | 1 | 0 | Streptococcus salivarius | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus | sp-CVUAS-32349 | Streptococcus | sp-CVUAS-32349 | 2.58 | Streptococcus | dysgalactiae | 1.52 | A | 1 | 1 | 0 | Streptococcus sp-CVUAS-32349 | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2260,4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus | troglodytidis | Streptococcus | troglodytidis | 2.52 | Streptococcus | troglodytidis | 2.21 | A | 1 | 1 | 0 | Streptococcus troglodytidis | Streptococcus troglodytidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5218,4 | 3; sequenced; 16S; NRZ Strep Uni AA | DT | 2015-06-11 | Streptococcus | troglodytidis | Streptococcus | troglodytidis | 2.35 | Streptococcus | troglodytidis | 2.13 | A | 1 | 1 | 0 | Streptococcus troglodytidis | Streptococcus troglodytidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.03 | Streptococcus | uberis | 1.81 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.23 | Streptococcus | uberis | 2.18 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 1.94 | Clostridioides | difficile | 1.83 | B | 0 | 0 | 1 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.2 | Clostridioides | difficile | 2.17 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.1 | Clostridium | sporogenes | 2.05 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.92 | Clostridium | sporogenes | 1.88 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 2.01 | Clostridium | botulinum | 1.99 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium botulinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B L677 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.03 | Clostridium | sporogenes | 1.95 | A | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.16 | Clostridium | sporogenes | 1.59 | A | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.99 | Clostridium | sporogenes | 1.67 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.84 | Arthrobacter | ramosus | 1.51 | B | 0 | 0 | 1 | Clostridium botulinum | Arthrobacter ramosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium | botulinum | Clostridium | botulinum | 2.36 | Bacillus | subtilis | 1.49 | A | 1 | 1 | 0 | Clostridium botulinum | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32971 | 3; sequenced | EFEx | 2021-02-25 | Clostridium | botulinum | Clostridium | botulinum | 2.32 | Clostridium | sporogenes | 2.08 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33885,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium | botulinum | Clostridium | botulinum | 1.71 | Clostridium | sporogenes | 1.71 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33886,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium | botulinum | Clostridium | sporogenes | 1.88 | Clostridium | sporogenes | 1.84 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33891,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium | botulinum | Clostridium | sporogenes | 1.86 | Clostridium | botulinum | 1.86 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium botulinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium | chauvoei | Clostridium | chauvoei | 2.38 | Clostridium | chauvoei | 2.17 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 849,2 | 3; sequenced | DT | 2014-12-17 | Clostridium | isatidis | Clostridium | isatidis | 2.13 | Clostridium | isatidis | 1.6 | A | 1 | 1 | 0 | Clostridium isatidis | Clostridium isatidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium | nigeriense | Clostridium | nigeriense | 2.58 | Clostridium | tertium | 1.75 | A | 1 | 1 | 0 | Clostridium nigeriense | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium | perfringens | Clostridium | perfringens | 2.09 | Clostridium | perfringens | 1.9 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B | 3; case rerort | DT | 2013-08-06 | Clostridium | perfringens | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.4 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium | septicum | Clostridium | septicum | 2.27 | Clostridium | septicum | 2.23 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium | sp-DSM-1985 | Clostridium | botulinum | 2.26 | Campylobacter | jejuni | 1.33 | A | 0 | 0 | 1 | Clostridium botulinum | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium | sporogenes | Clostridium | sporogenes | 1.95 | Clostridium | sporogenes | 1.79 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium | tertium | Clostridium | tertium | 2.33 | Clostridium | tertium | 2.28 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium | tetani | Clostridium | tetani | 2.55 | Clostridium | tetani | 2.47 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix | anatis | Erysipelothrix | anatis | 2.41 | Erysipelothrix | anatis | 2.35 | A | 1 | 1 | 0 | Erysipelothrix anatis | Erysipelothrix anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix | anatis | Erysipelothrix | anatis | 2.65 | Erysipelothrix | anatis | 2.43 | A | 1 | 1 | 0 | Erysipelothrix anatis | Erysipelothrix anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix | aquatica | Erysipelothrix | aquatica | 2.48 | Erysipelothrix | aquatica | 2.43 | A | 1 | 1 | 0 | Erysipelothrix aquatica | Erysipelothrix aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | VA-92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | aquatica | Erysipelothrix | larvae | 2.68 | Erysipelothrix | larvae | 2.15 | A | 0 | 0 | 1 | Erysipelothrix larvae | Erysipelothrix larvae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | larvae | Erysipelothrix | larvae | 2.55 | Erysipelothrix | larvae | 2.18 | A | 1 | 1 | 0 | Erysipelothrix larvae | Erysipelothrix larvae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 110099 CVUAS | 3; type strain | EFex [BRU03] | 2021-10-18 | Erysipelothrix | piscisicarius | Erysipelothrix | piscisicarius | 2.6 | Erysipelothrix | piscisicarius | 2.15 | A | 1 | 1 | 0 | Erysipelothrix piscisicarius | Erysipelothrix piscisicarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.7 | Erysipelothrix | rhusiopathiae | 2.54 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.25 | Erysipelothrix | rhusiopathiae | 2.22 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix | tonsillarum | Erysipelothrix | tonsillarum | 2.62 | Erysipelothrix | tonsillarum | 2.55 | A | 1 | 1 | 0 | Erysipelothrix tonsillarum | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-23 | Erysipelothrix | tonsillarum | Erysipelothrix | tonsillarum | 2.79 | Erysipelothrix | tonsillarum | 2.55 | A | 1 | 1 | 0 | Erysipelothrix tonsillarum | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus | sp-CVUAS-10169 | Anaerococcus | sp-CVUAS-10169 | 2.26 | Pseudomonas | putida | 1.34 | A | 1 | 1 | 0 | Anaerococcus sp-CVUAS-10169 | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32948 | 3; confirmed proficency test | EFEx | 2021-02-18 | Finegoldia | magna | Finegoldia | magna | 2.26 | Finegoldia | magna | 2.08 | A | 1 | 1 | 0 | Finegoldia magna | Finegoldia magna | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.25 | Helcococcus | kunzii | 2.25 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.45 | Helcococcus | kunzii | 2.4 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 433,2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus | ovis | Helcococcus | ovis | 2.41 | Helcococcus | ovis | 2.39 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus | ovis | Helcococcus | ovis | 2.28 | Helcococcus | ovis | 2.17 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus | sp-CVUAS-32321 | Helcococcus | sp-CVUAS-32321 | 2.67 | Helcococcus | sueciensis | 1.77 | A | 1 | 1 | 0 | Helcococcus sp-CVUAS-32321 | Helcococcus sueciensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 692,2 | 3; sequenced | DT | 2013-02-19 | Parvimonas | sp-CVUAS-692,2 | Parvimonas | sp-CVUAS-692,2 | 2.15 | Actinomyces | bovis | 1.44 | A | 1 | 1 | 0 | Parvimonas sp-CVUAS-692,2 | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides | finegoldii | Bacteroides | finegoldii | 2.6 | Bacteroides | finegoldii | 2.53 | A | 1 | 1 | 0 | Bacteroides finegoldii | Bacteroides finegoldii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides | fragilis | Bacteroides | fragilis | 2.53 | Bacteroides | fragilis | 2.42 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181001478 Bat | 2; 0 | DT | 2018-05-02 | Bacteroides | pyogenes | Bacteroides | pyogenes | 2.48 | Bacteroides | pyogenes | 2.48 | A | 1 | 1 | 0 | Bacteroides pyogenes | Bacteroides pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella | heparinolytica | Prevotella | heparinolytica | 2.52 | Prevotella | heparinolytica | 2.49 | A | 1 | 1 | 0 | Prevotella heparinolytica | Prevotella heparinolytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium | gleum | Chryseobacterium | gleum | 1.94 | Chryseobacterium | gleum | 1.9 | B | 0 | 0 | 1 | Chryseobacterium gleum | Chryseobacterium gleum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 2.1 | Chryseobacterium | tructae | 2.07 | A | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium tructae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 1.85 | Chryseobacterium | indologenes | 1.8 | B | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium indologenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32947 | 3; confirmed proficency test | EFEx | 2021-02-18 | Bacteroides | fragilis | Bacteroides | fragilis | 2.35 | Bacteroides | fragilis | 2.22 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium | hibernum | Flavobacterium | hibernum | 2 | Flavobacterium | saccharophilum | 1.97 | A | 1 | 1 | 0 | Flavobacterium hibernum | Flavobacterium saccharophilum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium | sp | Flavobacterium | saccharophilum | 2.04 | Flavobacterium | hibernum | 1.97 | A | 0 | 0 | 1 | Flavobacterium saccharophilum | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium | chaponense | Chryseobacterium | chaponense | 2.46 | Chryseobacterium | chaponense | 1.81 | A | 1 | 1 | 0 | Chryseobacterium chaponense | Chryseobacterium chaponense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium | hominis | Chryseobacterium | hominis | 2.55 | Chryseobacterium | hominis | 2.08 | A | 1 | 1 | 0 | Chryseobacterium hominis | Chryseobacterium hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium | manosquense | Chryseobacterium | manosquense | 2.22 | Deinococcus | wulumuqiensis | 1.45 | A | 1 | 1 | 0 | Chryseobacterium manosquense | Deinococcus wulumuqiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-6 EFex | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium | manosquense | Chryseobacterium | manosquense | 2.13 | Staphylococcus | simiae | 1.28 | A | 1 | 1 | 0 | Chryseobacterium manosquense | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11450,4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium | sp-I-7 | Chryseobacterium | sp-I-7 | 2.7 | Chryseobacterium | geocarposphaerae | 1.42 | A | 1 | 1 | 0 | Chryseobacterium sp-I-7 | Chryseobacterium geocarposphaerae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | aquidurense | Flavobacterium | aquidurense | 2.54 | Flavobacterium | frigidimaris | 1.97 | A | 1 | 1 | 0 | Flavobacterium aquidurense | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5322,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | branchiarum | Flavobacterium | branchiarum | 2.67 | Flavobacterium | pectinovorum | 1.94 | A | 1 | 1 | 0 | Flavobacterium branchiarum | Flavobacterium pectinovorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | chungangense | Flavobacterium | chungangense | 2.41 | Flavobacterium | plurextorum | 1.82 | A | 1 | 1 | 0 | Flavobacterium chungangense | Flavobacterium plurextorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium | columnare | Flavobacterium | columnare | 2.39 | Flavobacterium | columnare | 2.3 | A | 1 | 1 | 0 | Flavobacterium columnare | Flavobacterium columnare | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.41 | Flavobacterium | frigidimaris | 2.04 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4217,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.6 | Flavobacterium | frigidimaris | 1.94 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium | piscis | Flavobacterium | piscis | 2.78 | Flavobacterium | piscis | 2.29 | A | 1 | 1 | 0 | Flavobacterium piscis | Flavobacterium piscis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium | plurextorum | Flavobacterium | plurextorum | 2.62 | Flavobacterium | plurextorum | 2.35 | A | 1 | 1 | 0 | Flavobacterium plurextorum | Flavobacterium plurextorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5323,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | psychoterrae | Flavobacterium | psychoterrae | 2.41 | Flavobacterium | hydatis | 1.81 | A | 1 | 1 | 0 | Flavobacterium psychoterrae | Flavobacterium hydatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11550,4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.52 | Flavobacterium | psychrophilum | 2.4 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.3 | Flavobacterium | psychrophilum | 2.23 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | reichenbachii | Flavobacterium | reichenbachii | 2 | Flavobacterium | reichenbachii | 1.9 | A | 1 | 1 | 0 | Flavobacterium reichenbachii | Flavobacterium reichenbachii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | sp-C2 | Flavobacterium | sp-C2 | 2.4 | Paraburkholderia | phenazinium | 1.57 | A | 1 | 1 | 0 | Flavobacterium sp-C2 | Paraburkholderia phenazinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-10323 | Flavobacterium | sp-CVUAS-10323 | 2.31 | Flavobacterium | chungbukense | 1.65 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-10323 | Flavobacterium chungbukense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium | sp-CVUAS-4146 | Flavobacterium | sp-CVUAS-4146 | 2.25 | Flavobacterium | branchiarum | 2.07 | B | 0 | 0 | 1 | Flavobacterium sp-CVUAS-4146 | Flavobacterium branchiarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4216,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-4216,2 | Flavobacterium | sp-CVUAS-4216,2 | 2.07 | Flavobacterium | sp-JQMS | 1.47 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-4216,2 | Flavobacterium sp-JQMS | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-5506 | Flavobacterium | sp-CVUAS-5506 | 2.48 | Flavobacterium | johnsoniae | 1.82 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-5506 | Flavobacterium johnsoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | sp-HMD2009 | Flavobacterium | sp-HMD2009 | 2.45 | Flavobacterium | succinicans | 1.69 | A | 1 | 1 | 0 | Flavobacterium sp-HMD2009 | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium | sp-JQMS | Flavobacterium | sp-JQMS | 2.59 | Flavobacterium | succinicans | 1.67 | A | 1 | 1 | 0 | Flavobacterium sp-JQMS | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium | succinicans | Flavobacterium | succinicans | 2.49 | Flavobacterium | tiangeerense | 1.78 | A | 1 | 1 | 0 | Flavobacterium succinicans | Flavobacterium tiangeerense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | tiangeerense | Flavobacterium | tiangeerense | 2.61 | Flavobacterium | succinicans | 2.14 | B | 0 | 0 | 1 | Flavobacterium tiangeerense | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella | columbina | Riemerella | columbina | 2.56 | Lactobacillus | plantarum | 1.25 | A | 1 | 1 | 0 | Riemerella columbina | Lactobacillus plantarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2589,2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas | tenax | Epilithonimonas | tenax | 2.34 | Chryseobacterium | aquaticum | 1.4 | A | 1 | 1 | 0 | Epilithonimonas tenax | Chryseobacterium aquaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.27 | Ornithobacterium | rhinotracheale | 2.24 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.21 | Ornithobacterium | rhinotracheale | 2.04 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9288,2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca | zopfii | Prototheca | zopfii | 2.24 | Prototheca | zopfii | 2.19 | A | 1 | 1 | 0 | Prototheca zopfii | Prototheca zopfii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9876,2 | 2; 0 | Pextr | 2016-09-01 | Prototheca | zopfii | Prototheca | zopfii | 2.54 | Prototheca | zopfii | 2.53 | A | 1 | 1 | 0 | Prototheca zopfii | Prototheca zopfii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus | sp-CVUAS-11266 | Deinococcus | sp-CVUAS-11266 | 2.76 | Acidovorax | facilis | 1.23 | A | 1 | 1 | 0 | Deinococcus sp-CVUAS-11266 | Acidovorax facilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium | necrophorum | Fusobacterium | necrophorum | 2.48 | Fusobacterium | necrophorum | 2.41 | A | 1 | 1 | 0 | Fusobacterium necrophorum | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium | varium | Fusobacterium | varium | 2.09 | Fusobacterium | varium | 2.04 | A | 1 | 1 | 0 | Fusobacterium varium | Fusobacterium varium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.47 | Streptobacillus | moniliformis | 2.43 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter | abscessus | Caviibacter | abscessus | 2.32 | Caviibacter | abscessus | 2.18 | A | 1 | 1 | 0 | Caviibacter abscessus | Caviibacter abscessus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga | salmonicida | Oceanivirga | salmonicida | 2.26 | Oceanivirga | salmonicida | 2.03 | A | 1 | 1 | 0 | Oceanivirga salmonicida | Oceanivirga salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus | hongkongensis | Pseudostreptobacillus | hongkongensis | 2.21 | Pseudostreptobacillus | hongkongensis | 1.81 | A | 1 | 1 | 0 | Pseudostreptobacillus hongkongensis | Pseudostreptobacillus hongkongensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella | termitidis | Sebaldella | termitidis | 2.16 | Sebaldella | termitidis | 1.65 | A | 1 | 1 | 0 | Sebaldella termitidis | Sebaldella termitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella | termitidis | Sebaldella | termitidis | 2.47 | Sebaldella | termitidis | 1.41 | A | 1 | 1 | 0 | Sebaldella termitidis | Sebaldella termitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 38322 CVUAS | 3; public strain collection | EFEx | 16.12.2015 | Sneathia | sanguinegens | Sneathia | sanguinegens | 2.76 | Sneathia | sanguinegens | 2.27 | A | 1 | 1 | 0 | Sneathia sanguinegens | Sneathia sanguinegens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 41628 CVUAS | 3; type strain | EFEx | 2020-10-08 | Sneathia | sanguinegens | Sneathia | sanguinegens | 2.49 | Sneathia | sanguinegens | 2.45 | A | 1 | 1 | 0 | Sneathia sanguinegens | Sneathia sanguinegens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Sneathia | vaginalis | 2.61 | Sneathia | vaginalis | 2.6 | A | 1 | 1 | 0 | Sneathia vaginalis | Sneathia vaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Sneathia | vaginalis | 2.57 | Sneathia | vaginalis | 2.56 | A | 1 | 1 | 0 | Sneathia vaginalis | Sneathia vaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus | felis | Streptobacillus | felis | 2.52 | Streptobacillus | felis | 2.5 | A | 1 | 1 | 0 | Streptobacillus felis | Streptobacillus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.58 | Streptobacillus | moniliformis | 2.57 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus | notomytis | Streptobacillus | notomytis | 2.46 | Streptobacillus | notomytis | 2.11 | A | 1 | 1 | 0 | Streptobacillus notomytis | Streptobacillus notomytis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus | ratti | Streptobacillus | ratti | 2.4 | Pseudostreptobacillus | hongkongensis | 1.43 | A | 1 | 1 | 0 | Streptobacillus ratti | Pseudostreptobacillus hongkongensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012806 | 2; 0 | DT | 2017-10-04 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.18 | Ochrobactrum | intermedium | 2.11 | A | 1 | 1 | 0 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.55 | Brucella | intermedia | 2.54 | C | 0 | 0 | 1 | Ochrobactrum intermedium | Brucella intermedia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum | sp | Ochrobactrum | anthropi | 2.46 | Ochrobactrum | anthropi | 2.41 | A | 0 | 0 | 1 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum | sp | Ochrobactrum | intermedium | 1.78 | Ochrobactrum | anthropi | 1.68 | B | 0 | 0 | 1 | Ochrobactrum intermedium | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012040 Och | 2; 0 | DT | 2017-10-04 | Ochrobactrum | sp | Ochrobactrum | tritici | 1.78 | Ochrobactrum | intermedium | 1.65 | B | 0 | 0 | 1 | Ochrobactrum tritici | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia | coli | Escherichia | coli | 2.32 | Escherichia | coli | 2.29 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.4 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella | variicola | Klebsiella | variicola | 2.12 | Klebsiella | variicola | 2.08 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1816,2 | 3; sequenced | DT | 2013-07-24 | Basfia | succiniciproducens | Basfia | succiniciproducens | 2.08 | Basfia | succiniciproducens | 2.03 | A | 1 | 1 | 0 | Basfia succiniciproducens | Basfia succiniciproducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas | alcaligenes | Pseudomonas | alcaligenes | 2.44 | Pseudomonas | alcaligenes | 2.35 | A | 1 | 1 | 0 | Pseudomonas alcaligenes | Pseudomonas alcaligenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas | anguilliseptica | Pseudomonas | anguilliseptica | 2.02 | Pseudomonas | borbori | 1.95 | A | 1 | 1 | 0 | Pseudomonas anguilliseptica | Pseudomonas borbori | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas | brenneri | Pseudomonas | brenneri | 2.14 | Pseudomonas | veronii | 1.84 | A | 1 | 1 | 0 | Pseudomonas brenneri | Pseudomonas veronii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas | fulva | Pseudomonas | fulva | 2.29 | Pseudomonas | fulva | 1.9 | A | 1 | 1 | 0 | Pseudomonas fulva | Pseudomonas fulva | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas | koreensis | Pseudomonas | koreensis | 2.26 | Pseudomonas | corrugata | 1.89 | A | 1 | 1 | 0 | Pseudomonas koreensis | Pseudomonas corrugata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas | monteilii | Pseudomonas | monteilii | 2.01 | Pseudomonas | putida | 1.92 | A | 1 | 1 | 0 | Pseudomonas monteilii | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1_1 nit | 2; 0 | DT | 2017-08-31 | Pseudomonas | nitroreducens | Pseudomonas | nitroreducens | 2.07 | Pseudomonas | nitroreducens | 1.97 | A | 1 | 1 | 0 | Pseudomonas nitroreducens | Pseudomonas nitroreducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas | oryzihabitans | Pseudomonas | oryzihabitans | 2.12 | Pseudomonas | oryzihabitans | 2.09 | A | 1 | 1 | 0 | Pseudomonas oryzihabitans | Pseudomonas oryzihabitans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1_2 put | 2; 0 | DT | 2017-08-31 | Pseudomonas | putida | Pseudomonas | putida | 2.49 | Pseudomonas | putida | 2.31 | A | 1 | 1 | 0 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas | taetrolens | Pseudomonas | taetrolens | 2 | Pseudomonas | lundensis | 1.95 | A | 1 | 1 | 0 | Pseudomonas taetrolens | Pseudomonas lundensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2280,3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas | glaciistagni | Aureimonas | glaciistagni | 2.53 | Bacillus | thuringiensis | 1.31 | A | 1 | 1 | 0 | Aureimonas glaciistagni | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_08 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 02.04.2012 K8 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.42 | Ochrobactrum | anthropi | 2.39 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM373 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; D02 | Brucella | abortus | Brucella | abortus | 2.65 | Brucella | melitensis | 2.52 | B | 0 | 0 | 1 | Brucella abortus | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM369 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; E01 | Brucella | abortus | Brucella | abortus | 2.62 | Brucella | melitensis | 2.49 | B | 0 | 0 | 1 | Brucella abortus | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; I16 | Brucella | canis | Brucella | canis | 2.58 | Brucella | canis | 2.53 | A | 1 | 1 | 0 | Brucella canis | Brucella canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.03.2014; E17 | Brucella | canis | Brucella | canis | 2.56 | Brucella | melitensis | 2.47 | B | 0 | 0 | 1 | Brucella canis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM389 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; F04 | Brucella | ceti | Brucella | ceti | 2.75 | Brucella | ceti | 2.55 | A | 1 | 1 | 0 | Brucella ceti | Brucella ceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM403 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; K04 | Brucella | ceti | Brucella | ceti | 2.71 | Brucella | abortus | 2.51 | B | 0 | 0 | 1 | Brucella ceti | Brucella abortus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM374 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I02 | Brucella | melitensis | Brucella | melitensis | 2.77 | Brucella | melitensis | 2.7 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM407 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; J05 | Brucella | melitensis | Brucella | melitensis | 2.75 | Brucella | melitensis | 2.72 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM408 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; H06 | Brucella | melitensis | Brucella | melitensis | 2.73 | Brucella | melitensis | 2.62 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM410 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I06 | Brucella | melitensis | Brucella | melitensis | 2.69 | Brucella | melitensis | 2.66 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 4915 CVUAS | 3; type strain | DT | 2013-11-26 | Brucella | microti | Brucella | microti | 2.67 | Brucella | sp. | 2.39 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 4915 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.03.2015 C7 | Brucella | microti | Brucella | abortus | 2.57 | Brucella | microti | 2.52 | B | 0 | 0 | 1 | Brucella abortus | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM421 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; F07 | Brucella | neotomae | Brucella | neotomae | 2.64 | Brucella | neotomae | 2.37 | A | 1 | 1 | 0 | Brucella neotomae | Brucella neotomae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.03.2015; C15 | Brucella | neotomae | Brucella | neotomae | 2.76 | Brucella | neotomae | 2.52 | A | 1 | 1 | 0 | Brucella neotomae | Brucella neotomae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM423 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; L07 | Brucella | ovis | Brucella | ovis | 2.67 | Brucella | ovis | 2.61 | A | 1 | 1 | 0 | Brucella ovis | Brucella ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; F16 | Brucella | ovis | Brucella | ovis | 2.83 | Brucella | ovis | 2.53 | A | 1 | 1 | 0 | Brucella ovis | Brucella ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM405 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; E05 | Brucella | pinnipedialis | Brucella | pinnipedialis | 2.71 | Brucella | ceti | 2.55 | B | 0 | 0 | 1 | Brucella pinnipedialis | Brucella ceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; C03 | Brucella | pinnipedialis | Brucella | pinnipedialis | 2.76 | Brucella | pinnipedialis | 2.58 | A | 1 | 1 | 0 | Brucella pinnipedialis | Brucella pinnipedialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM431 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; B10 | Brucella | suis | Brucella | suis | 2.7 | Brucella | suis | 2.54 | A | 1 | 1 | 0 | Brucella suis | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM427 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; G08 | Brucella | suis | Brucella | suis | 2.74 | Brucella | abortus | 2.68 | B | 0 | 0 | 1 | Brucella suis | Brucella abortus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella | suis | Brucella | suis | 2.62 | Brucella | suis | 2.04 | A | 1 | 1 | 0 | Brucella suis | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM428 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I08 | Brucella | suis | Brucella | suis | 2.7 | Brucella | canis | 2.57 | B | 0 | 0 | 1 | Brucella suis | Brucella canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM430 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; I09 | Brucella | suis | Brucella | suis | 2.76 | Brucella | suis | 2.46 | A | 1 | 1 | 0 | Brucella suis | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 15.15.2016 I7 | Brucella | vulpis | Brucella | vulpis | 2.78 | Brucella | canis | 2.16 | B | 0 | 0 | 1 | Brucella vulpis | Brucella canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum | ovis | Falsochrobactrum | ovis | 2.48 | Falsochrobactrum | ovis | 2.41 | A | 1 | 1 | 0 | Falsochrobactrum ovis | Falsochrobactrum ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 03.06.2016 H12 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.54 | Ochrobactrum | anthropi | 2.42 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter | aganoensis | Aminobacter | aganoensis | 2.39 | Pseudomonas | mucidolens | 1.45 | A | 1 | 1 | 0 | Aminobacter aganoensis | Pseudomonas mucidolens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter | aminovorans | Aminobacter | aminovorans | 2.56 | Lactobacillus | curvatus | 1.5 | A | 1 | 1 | 0 | Aminobacter aminovorans | Lactobacillus curvatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2015 H22 | Paracoccus | yeei | Paracoccus | yeei | 2.45 | Paracoccus | yeei | 2.24 | A | 1 | 1 | 0 | Paracoccus yeei | Paracoccus yeei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas | mucosa | Roseomonas | mucosa | 2.53 | Roseomonas | mucosa | 2.49 | A | 1 | 1 | 0 | Roseomonas mucosa | Roseomonas mucosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 24.01.2007, H7, [0034] | Inquilinus | limosus | Inquilinus | limosus | 2.63 | Inquilinus | limosus | 2.62 | A | 1 | 1 | 0 | Inquilinus limosus | Inquilinus limosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.32 | Taylorella | asinigenitalis | 2.01 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella asinigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter | denitrificans | Achromobacter | denitrificans | 2.46 | Achromobacter | denitrificans | 2.22 | A | 1 | 1 | 0 | Achromobacter denitrificans | Achromobacter denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter | sp-DSM-30031 | Achromobacter | sp-DSM-30031 | 2.22 | Achromobacter | piechaudii | 1.95 | A | 1 | 1 | 0 | Achromobacter sp-DSM-30031 | Achromobacter piechaudii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2008 E5 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.33 | Achromobacter | xylosoxidans | 2.16 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.02.2016 N7 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.26 | Achromobacter | xylosoxidans | 2.23 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea | sp-CVUAS-31213 | Basilea | sp-CVUAS-31213 | 2.18 | Streptomyces | badius | 1.35 | A | 1 | 1 | 0 | Basilea sp-CVUAS-31213 | Streptomyces badius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella | avium | Bordetella | avium | 2.39 | Bordetella | avium | 2.29 | A | 1 | 1 | 0 | Bordetella avium | Bordetella avium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32946 | 3; confirmed proficency test | EFEx | 2021-02-18 | Kerstersia | gyiorum | Kerstersia | gyiorum | 2.48 | Kerstersia | gyiorum | 2.48 | A | 1 | 1 | 0 | Kerstersia gyiorum | Kerstersia gyiorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 2.56 | Oligella | urethralis | 2.43 | A | 1 | 1 | 0 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 2.51 | Oligella | urethralis | 2.39 | A | 1 | 1 | 0 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga | kullae | Pigmentiphaga | kullae | 2.24 | Pigmentiphaga | daeguensis | 1.57 | A | 1 | 1 | 0 | Pigmentiphaga kullae | Pigmentiphaga daeguensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.12 | Taylorella | asinigenitalis | 2.09 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella asinigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.15 | Taylorella | equigenitalis | 2.1 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.35 | Taylorella | equigenitalis | 2.24 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012; K10 | Burkholderia | pseudomallei | Burkholderia | pseudomallei | 2.34 | Burkholderia | mallei | 2.28 | B | 0 | 0 | 1 | Burkholderia pseudomallei | Burkholderia mallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; H21 | Burkholderia | pseudomallei | Burkholderia | pseudomallei | 2.49 | Burkholderia | pseudomallei | 2.47 | A | 1 | 1 | 0 | Burkholderia pseudomallei | Burkholderia pseudomallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.04.2012, K6 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.51 | Burkholderia | thailandensis | 2.38 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007, K7 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.27 | Burkholderia | ambifaria | 2.24 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16087 LMG 19182 | 3; type starin | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K19 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.36 | Burkholderia | diffusa | 2.32 | B | 0 | 0 | 1 | Burkholderia ambifaria | Burkholderia diffusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K21 | Burkholderia | anthina | Burkholderia | anthina | 2.42 | Burkholderia | anthina | 2.35 | A | 1 | 1 | 0 | Burkholderia anthina | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.08.2007, N10 | Burkholderia | anthina | Burkholderia | anthina | 2.79 | Burkholderia | pyrrocinia | 2.17 | B | 0 | 0 | 1 | Burkholderia anthina | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2010, K12 | Burkholderia | arboris | Burkholderia | arboris | 2.81 | Burkholderia | puraquae | 2.37 | B | 0 | 0 | 1 | Burkholderia arboris | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 3.4.2017; L6 | Burkholderia | arboris | Burkholderia | arboris | 2.82 | Burkholderia | arboris | 2.33 | A | 1 | 1 | 0 | Burkholderia arboris | Burkholderia arboris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K11 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.5 | Burkholderia | cenocepacia | 2.48 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.07.2009; K21 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.52 | Burkholderia | cenocepacia | 2.42 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 4.5.2006; H15 | Burkholderia | cepacia | Burkholderia | cenocepacia | 2.71 | Burkholderia | cenocepacia | 2.46 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 9.5.2007; N9 | Burkholderia | cepacia | Burkholderia | cepacia | 2.73 | Burkholderia | cepacia | 2.4 | A | 1 | 1 | 0 | Burkholderia cepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.11.2006; H10 | Burkholderia | contaminans | Burkholderia | contaminans | 2.65 | Burkholderia | contaminans | 2.62 | A | 1 | 1 | 0 | Burkholderia contaminans | Burkholderia contaminans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 01 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E10 | Burkholderia | contaminans | Burkholderia | contaminans | 2.62 | Burkholderia | cepacia | 2.42 | B | 0 | 0 | 1 | Burkholderia contaminans | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016 C13 | Burkholderia | diffusa | Burkholderia | diffusa | 2.73 | Burkholderia | ambifaria | 2.35 | B | 0 | 0 | 1 | Burkholderia diffusa | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L17 | Burkholderia | diffusa | Burkholderia | diffusa | 2.75 | Burkholderia | territorii | 2.44 | B | 0 | 0 | 1 | Burkholderia diffusa | Burkholderia territorii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K16 | Burkholderia | dolosa | Burkholderia | dolosa | 2.29 | Burkholderia | dolosa | 2.29 | A | 1 | 1 | 0 | Burkholderia dolosa | Burkholderia dolosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.11.2006; E09 | Burkholderia | dolosa | Burkholderia | dolosa | 2.08 | Burkholderia | cenocepacia | 2.07 | B | 0 | 0 | 1 | Burkholderia dolosa | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.01.2007; H22 | Burkholderia | gladioli | Burkholderia | gladioli | 2.45 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H15 | Burkholderia | gladioli | Burkholderia | gladioli | 2.42 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K03 | Burkholderia | glumae | Burkholderia | glumae | 2.62 | Burkholderia | glumae | 2.33 | A | 1 | 1 | 0 | Burkholderia glumae | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.05.2012; K22 | Burkholderia | glumae | Burkholderia | glumae | 2.79 | Burkholderia | glumae | 2.33 | A | 1 | 1 | 0 | Burkholderia glumae | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016; C10 | Burkholderia | lata | Burkholderia | lata | 2.8 | Burkholderia | puraquae | 2.49 | B | 0 | 0 | 1 | Burkholderia lata | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.08.2016; K19 | Burkholderia | lata | Burkholderia | lata | 2.83 | Burkholderia | cepacia | 2.47 | B | 0 | 0 | 1 | Burkholderia lata | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.05.2016; L06 | Burkholderia | latens | Burkholderia | latens | 2.71 | Burkholderia | latens | 2.32 | A | 1 | 1 | 0 | Burkholderia latens | Burkholderia latens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.09.2009; K20 | Burkholderia | mallei | Burkholderia | mallei | 2.23 | Burkholderia | mallei | 2.19 | A | 1 | 1 | 0 | Burkholderia mallei | Burkholderia mallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.08.2007; H13 | Burkholderia | mallei | Burkholderia | pseudomallei | 2.12 | Burkholderia | pseudomallei | 2.01 | A | 0 | 0 | 1 | Burkholderia pseudomallei | Burkholderia pseudomallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2010; H14 | Burkholderia | metallica | Burkholderia | metallica | 2.52 | Burkholderia | cepacia | 2.46 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L21 | Burkholderia | metallica | Burkholderia | metallica | 2.77 | Burkholderia | seminalis | 2.35 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia seminalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.12.2006; K19 | Burkholderia | multivorans | Burkholderia | multivorans | 2.22 | Burkholderia | multivorans | 2.21 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.11.2006; K04 | Burkholderia | multivorans | Burkholderia | multivorans | 2.13 | Burkholderia | multivorans | 2 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K17 | Burkholderia | oklahomensis | Burkholderia | oklahomensis | 2.72 | Burkholderia | oklahomensis | 2.49 | A | 1 | 1 | 0 | Burkholderia oklahomensis | Burkholderia oklahomensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K18 | Burkholderia | oklahomensis | Burkholderia | oklahomensis | 2.73 | Burkholderia | oklahomensis | 2.48 | A | 1 | 1 | 0 | Burkholderia oklahomensis | Burkholderia oklahomensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K06 | Burkholderia | plantarii | Burkholderia | plantarii | 2.49 | Burkholderia | glumae | 1.94 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K05 | Burkholderia | plantarii | Burkholderia | plantarii | 2.5 | Burkholderia | glumae | 1.95 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K07 | Burkholderia | pseudomultivorans | Burkholderia | pseudomultivorans | 2.75 | Burkholderia | ubonensis | 2.1 | B | 0 | 0 | 1 | Burkholderia pseudomultivorans | Burkholderia ubonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K08 | Burkholderia | pseudomultivorans | Burkholderia | pseudomultivorans | 2.79 | Burkholderia | pseudomultivorans | 2.28 | A | 1 | 1 | 0 | Burkholderia pseudomultivorans | Burkholderia pseudomultivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CAMPA 1043 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; K04 | Burkholderia | puraquae | Burkholderia | puraquae | 2.73 | Burkholderia | puraquae | 2.7 | A | 1 | 1 | 0 | Burkholderia puraquae | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | R-50214 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.03.2016; C04 | Burkholderia | puraquae | Burkholderia | puraquae | 2.8 | Burkholderia | puraquae | 2.79 | A | 1 | 1 | 0 | Burkholderia puraquae | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; N06 | Burkholderia | pyrrocinia | Burkholderia | pyrrocinia | 2.57 | Burkholderia | puraquae | 2.45 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; I21 | Burkholderia | pyrrocinia | Burkholderia | diffusa | 2.18 | Burkholderia | puraquae | 2.14 | B | 0 | 0 | 1 | Burkholderia diffusa | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H08 | Burkholderia | seminalis | Burkholderia | seminalis | 2.79 | Burkholderia | seminalis | 2.41 | A | 1 | 1 | 0 | Burkholderia seminalis | Burkholderia seminalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.01.2007; K12 | Burkholderia | seminalis | Burkholderia | seminalis | 2.35 | Burkholderia | cepacia | 2.29 | B | 0 | 0 | 1 | Burkholderia seminalis | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K13 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.56 | Burkholderia | stabilis | 2.51 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E08 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.51 | Burkholderia | stabilis | 2.28 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; C08 | Burkholderia | stagnalis | Burkholderia | stagnalis | 2.79 | Burkholderia | stagnalis | 2.65 | A | 1 | 1 | 0 | Burkholderia stagnalis | Burkholderia stagnalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F14 | Burkholderia | stagnalis | Burkholderia | stagnalis | 2.82 | Burkholderia | stagnalis | 2.62 | A | 1 | 1 | 0 | Burkholderia stagnalis | Burkholderia stagnalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007; H19 | Burkholderia | taxonK | Burkholderia | contaminans | 2.7 | Burkholderia | contaminans | 2.5 | A | 0 | 0 | 1 | Burkholderia contaminans | Burkholderia contaminans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.06.2009; K13 | Burkholderia | taxonK | Burkholderia | contaminans | 2.57 | Burkholderia | contaminans | 2.55 | A | 0 | 0 | 1 | Burkholderia contaminans | Burkholderia contaminans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F16 | Burkholderia | territorii | Burkholderia | territorii | 2.78 | Burkholderia | cenocepacia | 2.3 | B | 0 | 0 | 1 | Burkholderia territorii | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; E14 | Burkholderia | territorii | Burkholderia | territorii | 2.81 | Burkholderia | diffusa | 2.51 | B | 0 | 0 | 1 | Burkholderia territorii | Burkholderia diffusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13276 n RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 24.11.2010, K19 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.4 | Burkholderia | thailandensis | 2.34 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H11 | Burkholderia | ubonensis | Burkholderia | ubonensis | 2.82 | Burkholderia | ubonensis | 2.31 | A | 1 | 1 | 0 | Burkholderia ubonensis | Burkholderia ubonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; I04 | Burkholderia | ubonensis | Burkholderia | ubonensis | 2.73 | Burkholderia | ubonensis | 2.47 | A | 1 | 1 | 0 | Burkholderia ubonensis | Burkholderia ubonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9631=LMG 6999 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.05.2007 N15 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.46 | Burkholderia | vietnamiensis | 2.33 | A | 1 | 1 | 0 | Burkholderia vietnamiensis | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10929 V | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.12.2006 N9 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.78 | Burkholderia | vietnamiensis | 2.43 | A | 1 | 1 | 0 | Burkholderia vietnamiensis | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.04.2012, N07 | Caballeronia | glathei | Caballeronia | glathei | 2.55 | Caballeronia | turbans | 1.53 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H12 | Caballeronia | glathei | Caballeronia | glathei | 2.47 | Caballeronia | turbans | 1.56 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H15 | Pandoraea | apista | Pandoraea | apista | 2.36 | Pandoraea | apista | 2.28 | A | 1 | 1 | 0 | Pandoraea apista | Pandoraea apista | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H21 | Pandoraea | pnomenusa | Pandoraea | pnomenusa | 2.4 | Pandoraea | pnomenusa | 2.29 | A | 1 | 1 | 0 | Pandoraea pnomenusa | Pandoraea pnomenusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H16 | Pandoraea | pulmonicola | Pandoraea | pulmonicola | 2.49 | Pandoraea | pulmonicola | 2.25 | A | 1 | 1 | 0 | Pandoraea pulmonicola | Pandoraea pulmonicola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K7 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.52 | Paraburkholderia | caledonica | 2.45 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K8 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.46 | Paraburkholderia | caledonica | 2.37 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N9 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.71 | Paraburkholderia | caribensis | 2.25 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N10 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.8 | Paraburkholderia | caribensis | 2.23 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H03 | Paraburkholderia | hospita | Paraburkholderia | hospita | 2.42 | Paraburkholderia | azotifigens | 1.81 | A | 1 | 1 | 0 | Paraburkholderia hospita | Paraburkholderia azotifigens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H04 | Paraburkholderia | hospita | Paraburkholderia | hospita | 2.83 | Paraburkholderia | hospita | 2.07 | A | 1 | 1 | 0 | Paraburkholderia hospita | Paraburkholderia hospita | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 1.12.2010; H15 | Paraburkholderia | kururiensis | Paraburkholderia | kururiensis | 2.82 | Burkholderia | pyrrocinia | 1.64 | A | 1 | 1 | 0 | Paraburkholderia kururiensis | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.11.2010; N19 | Paraburkholderia | phymatum | Paraburkholderia | phymatum | 2.79 | Paraburkholderia | caribensis | 1.9 | A | 1 | 1 | 0 | Paraburkholderia phymatum | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.11.2010; K17 | Paraburkholderia | phytofirmans | Paraburkholderia | phytofirmans | 2.81 | Paraburkholderia | phytofirmans | 2.44 | A | 1 | 1 | 0 | Paraburkholderia phytofirmans | Paraburkholderia phytofirmans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H16 | Paraburkholderia | tuberum | Paraburkholderia | tuberum | 2.81 | Paraburkholderia | tuberum | 2.36 | A | 1 | 1 | 0 | Paraburkholderia tuberum | Paraburkholderia tuberum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.11.2010, N21 | Paraburkholderia | xenovorans | Paraburkholderia | xenovorans | 2.77 | Paraburkholderia | xenovorans | 2.49 | A | 1 | 1 | 0 | Paraburkholderia xenovorans | Paraburkholderia xenovorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 E07 | Ralstonia | insidiosa | Ralstonia | insidiosa | 2.38 | Ralstonia | insidiosa | 2.35 | A | 1 | 1 | 0 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 C06 | Ralstonia | mannitolilytica | Ralstonia | mannitolilytica | 2.46 | Ralstonia | mannitolilytica | 2.22 | A | 1 | 1 | 0 | Ralstonia mannitolilytica | Ralstonia mannitolilytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27511 DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 22.01.2007 H03 | Ralstonia | pickettii | Ralstonia | pickettii | 2.15 | Ralstonia | pickettii | 2.12 | A | 1 | 1 | 0 | Ralstonia pickettii | Ralstonia pickettii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 07.05.2009 N15 | Ralstonia | pickettii | Ralstonia | insidiosa | 2.07 | Ralstonia | insidiosa | 2.07 | A | 0 | 0 | 1 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia | acidovorans | Delftia | acidovorans | 2.54 | Delftia | tsuruhatensis | 2.44 | B | 0 | 0 | 1 | Delftia acidovorans | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia | tsuruhatensis | Delftia | tsuruhatensis | 2.45 | Delftia | tsuruhatensis | 2.31 | A | 1 | 1 | 0 | Delftia tsuruhatensis | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga | intermedia | Hydrogenophaga | intermedia | 2.45 | Hydrogenophaga | intermedia | 2.4 | A | 1 | 1 | 0 | Hydrogenophaga intermedia | Hydrogenophaga intermedia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.07.2013 L13 | Xenophilus | sp | Aromatoleum | evansii | 1.43 | Ligilactobacillus | salivarius | 1.32 | B | 0 | 0 | 1 | Aromatoleum evansii | Ligilactobacillus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24713 EFex | 3; type strain | Extr | 2016-02-02 | Deefgea | chitinilytica | Deefgea | chitinilytica | 2.72 | Mannheimia | haemolytica | 1.42 | A | 1 | 1 | 0 | Deefgea chitinilytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24713 eDT | 3; type strain | eDT | 2015-09-22 | Deefgea | chitinilytica | Deefgea | chitinilytica | 2.11 | Pichia | cactophila | 1.24 | A | 1 | 1 | 0 | Deefgea chitinilytica | Pichia cactophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea | rivuli | Deefgea | rivuli | 2.6 | Arthrobacter | woluwensis | 1.45 | A | 1 | 1 | 0 | Deefgea rivuli | Arthrobacter woluwensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 08-F17.2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Deefgea | sp-CVUAS-2703.2 | 2.03 | Deefgea | chitinilytica | 1.7 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-2703.2 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Deefgea | sp-CVUAS-2703.2 | 2.08 | Deefgea | chitinilytica | 1.65 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-2703.2 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6995,3 | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea | sp-CVUAS-6995,3 | Deefgea | sp-CVUAS-6995.3 | 2.61 | Deefgea | sp-CVUAS-9920 | 1.56 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-6995.3 | Deefgea sp-CVUAS-9920 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-6995,3 | Deefgea | sp-CVUAS-6995.3 | 2.38 | Pseudomonas | orientalis | 1.48 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-6995.3 | Pseudomonas orientalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-9920 | Deefgea | sp-CVUAS-9920 | 2.42 | Deefgea | sp-CVUAS-2703.2 | 1.53 | A | 1 | 1 | 0 | Deefgea sp-CVUAS-9920 | Deefgea sp-CVUAS-2703.2 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea | sp-CVUAS-9920 | Deefgea | sp-CVUAS-9920 | 2.21 | Deefgea | chitinilytica | 1.37 | A | 1 | 1 | 0 | Deefgea sp-CVUAS-9920 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter | limnosediminis | Iodobacter | limnosediminis | 2.6 | Iodobacter | fluviatilis | 1.45 | A | 1 | 1 | 0 | Iodobacter limnosediminis | Iodobacter fluviatilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella | kingae | Kingella | kingae | 2.49 | Kingella | kingae | 2.41 | A | 1 | 1 | 0 | Kingella kingae | Kingella kingae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10067,2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria | animaloris | Neisseria | animaloris | 2.41 | Neisseria | animaloris | 2.17 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria | animaloris | Neisseria | animaloris | 2.65 | Neisseria | animaloris | 2.42 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-4931 | Neisseria | sp-CVUAS-4931 | 2.56 | Neisseria | ornithorespiratoria' | 2.46 | B | 0 | 0 | 1 | Neisseria sp-CVUAS-4931 | Neisseria ornithorespiratoria' | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-8997 | Neisseria | sp-CVUAS-8997 | 2.27 | Neisseria_flavescens_subflava | group | 1.48 | A | 1 | 1 | 0 | Neisseria sp-CVUAS-8997 | Neisseria_flavescens_subflava group | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-LHL-1296 | Neisseria | sp-LHL-1296 | 2.59 | Neisseria | animaloris | 1.61 | A | 1 | 1 | 0 | Neisseria sp-LHL-1296 | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5699,3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria | sp-N95-16 | Neisseria | sp-N95-16 | 2.73 | Bergeriella | denitrificans | 1.99 | C | 0 | 0 | 1 | Neisseria sp-N95-16 | Bergeriella denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae | sp-CVUAS-8636 | Neisseriaceae | sp-CVUAS-8636 | 2.14 | Pseudomonas | taetrolens | 1.43 | A | 1 | 1 | 0 | Neisseriaceae sp-CVUAS-8636 | Pseudomonas taetrolens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter | fetus | Campylobacter | fetus | 2.34 | Campylobacter | fetus | 2.28 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.37 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.19 | Campylobacter | hyointestinalis | 2.12 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-07048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Burkholderia | plantarii | 1.42 | A | 1 | 1 | 0 | Campylobacter peloridis | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Campylobacter | sp-W1 | 1.48 | A | 1 | 1 | 0 | Campylobacter peloridis | Campylobacter sp-W1 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter | cryaerophilus | Arcobacter | cryaerophilus | 2.03 | Arcobacter | cryaerophilus | 2.01 | A | 0 | 0 | 1 | Arcobacter cryaerophilus | Arcobacter cryaerophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.06 | Aliarcobacter | butzleri | 1.94 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.24 | Aliarcobacter | butzleri | 2.21 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter | coli | Campylobacter | coli | 2.18 | Campylobacter | coli | 1.97 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4689 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter | coli | Campylobacter | coli | 2.24 | Campylobacter | coli | 2.21 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter | jejuni | Campylobacter | jejuni | 2.19 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter | jejuni | Campylobacter | jejuni | 2.39 | Campylobacter | jejuni | 2.32 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter | jejuni | Campylobacter | jejuni | 2.16 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | lari | Campylobacter | lari | 2.19 | Campylobacter | lari | 2.04 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-06556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.04 | Campylobacter | lari | 1.82 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.2 | Campylobacter | lari | 2.11 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 11375 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter | lari | Campylobacter | lari | 2.42 | Campylobacter | lari | 2.19 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5361 CVUAS | 3; type strain | DT | 2013-09-24 | Campylobacter | fetus | Campylobacter | fetus | 2.51 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6673,2 | 3; sequenced | DT | 2013-08-28 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.15 | Campylobacter | hyointestinalis | 2.06 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter | sputorum | Campylobacter | sputorum | 2.01 | Campylobacter | sputorum | 1.9 | A | 1 | 1 | 0 | Campylobacter sputorum | Campylobacter sputorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter | sp-W1 | Campylobacter | sp-W1 | 2.61 | Campylobacter | jejuni | 1.49 | A | 1 | 1 | 0 | Campylobacter sp-W1 | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter | upsaliensis | Campylobacter | upsaliensis | 2.05 | Campylobacter | upsaliensis | 1.96 | A | 1 | 1 | 0 | Campylobacter upsaliensis | Campylobacter upsaliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11185,3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter | canadensis | Helicobacter | canadensis | 2.72 | Clostridium | cochlearium | 1.42 | A | 1 | 1 | 0 | Helicobacter canadensis | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas | bestiarum | Aeromonas | bestiarum | 2.44 | Aeromonas | salmonicida | 2.42 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas | bivalvium | Aeromonas | veronii | 2.4 | Aeromonas | bivalvium | 2.18 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas bivalvium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas | caviae | Aeromonas | caviae | 2.29 | Aeromonas | hydrophila | 2.28 | B | 0 | 0 | 1 | Aeromonas caviae | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | hydrophila | 2.03 | Aeromonas | dhakensis | 2 | B | 0 | 0 | 1 | Aeromonas hydrophila | Aeromonas dhakensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | dhakensis | 2.42 | Aeromonas | dhakensis | 2.41 | A | 1 | 1 | 0 | Aeromonas dhakensis | Aeromonas dhakensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30187 RKI | 4; type strain, WGS sequenced (BfR) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E17 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.57 | Aeromonas | hydrophila | 2.46 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.27 | Aeromonas | hydrophila | 2.25 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas | jandaei | Aeromonas | jandaei | 2.34 | Aeromonas | jandaei | 2.05 | A | 1 | 1 | 0 | Aeromonas jandaei | Aeromonas jandaei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | media | Aeromonas | media | 2.22 | Aeromonas | media | 2.2 | A | 1 | 1 | 0 | Aeromonas media | Aeromonas media | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas | molluscorum | Aeromonas | molluscorum | 2.28 | Aeromonas | molluscorum | 2.15 | A | 1 | 1 | 0 | Aeromonas molluscorum | Aeromonas molluscorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | popoffii | Aeromonas | popoffii | 2.5 | Aeromonas | salmonicida | 2.26 | B | 0 | 0 | 1 | Aeromonas popoffii | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.22 | Aeromonas | salmonicida | 2.07 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.31 | Aeromonas | salmonicida | 2.25 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.55 | Aeromonas | salmonicida | 2.38 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.53 | Aeromonas | salmonicida | 2.22 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8400,2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.19 | Aeromonas | salmonicida | 2.07 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.38 | Aeromonas | salmonicida | 2.36 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.47 | Aeromonas | salmonicida | 2.3 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.18 | Aeromonas | salmonicida | 2.11 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas | simiae | Aeromonas | simiae | 2.31 | Aeromonas | simiae | 2.07 | A | 1 | 1 | 0 | Aeromonas simiae | Aeromonas simiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas | veronii | Aeromonas | veronii | 2.46 | Aeromonas | ichthiosmia | 2.18 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas ichthiosmia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7386 CVUAS | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas | veronii | Aeromonas | veronii | 2.46 | Aeromonas | bivalvium | 2.26 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas bivalvium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas | pectinilytica | Pseudaeromonas | pectinilytica | 2.62 | Aeromonas | eucrenophila | 1.4 | A | 1 | 1 | 0 | Pseudaeromonas pectinilytica | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella | viscosa | Moritella | viscosa | 2.32 | Moritella | viscosa | 2.22 | A | 1 | 1 | 0 | Moritella viscosa | Moritella viscosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella | viscosa | Moritella | viscosa | 2.29 | Moritella | viscosa | 2.23 | A | 1 | 1 | 0 | Moritella viscosa | Moritella viscosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | sp-CVUAS-9366 | 2.64 | Shewanella | sp-CVUAS-9366 | 2.62 | A | 1 | 1 | 0 | Shewanella sp-CVUAS-9366 | Shewanella sp-CVUAS-9366 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | sp-CVUAS-9366 | 2.55 | Shewanella | sp-CVUAS-9366 | 2.5 | A | 1 | 1 | 0 | Shewanella sp-CVUAS-9366 | Shewanella sp-CVUAS-9366 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 81,2 | 3; sequenced | EFEx | 28:11.2013 | Shewanella | sp-RKVR22 | Shewanella | sp-RKVR22 | 2.4 | Shewanella | corallii | 1.59 | A | 1 | 1 | 0 | Shewanella sp-RKVR22 | Shewanella corallii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter | condimenti | Cronobacter | condimenti | 2.49 | Cronobacter | malonaticus | 2.29 | B | 0 | 0 | 1 | Cronobacter condimenti | Cronobacter malonaticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter | dublinensis | Cronobacter | muytjensii | 2.22 | Cronobacter | sp | 2.2 | B | 0 | 0 | 1 | Cronobacter muytjensii | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18702 LALLF | 3; type strain | DT | 2019-03-05 | Cronobacter | malonaticus | Cronobacter | malonaticus | 2.56 | Cronobacter | sp | 2.46 | B | 0 | 0 | 1 | Cronobacter malonaticus | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter | turicensis | Cronobacter | turicensis | 2.52 | Cronobacter | sp | 2.36 | B | 0 | 0 | 1 | Cronobacter turicensis | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.11.2011 E6 | Enterobacter | cloacae | Enterobacter | hormaechei | 2.4 | Enterobacter | hormaechei | 2.35 | A | 0 | 0 | 1 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, H5 | Citrobacter | amalonaticus | Citrobacter | amalonaticus | 2.41 | Citrobacter | amalonaticus | 2.3 | A | 1 | 1 | 0 | Citrobacter amalonaticus | Citrobacter amalonaticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | freundii | 2.39 | Citrobacter | freundii | 2.37 | A | 1 | 1 | 0 | Citrobacter freundii | Citrobacter freundii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | braakii | 2.29 | Citrobacter | braakii | 2.23 | A | 0 | 0 | 1 | Citrobacter braakii | Citrobacter braakii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25408 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 11.04.2008, N11, [0034]: Citrobacter diversus ATCC 25408\LMTG70N11_0\0_N11\1\1SLin\fid | Citrobacter | koseri | Citrobacter | koseri | 2.52 | Citrobacter | koseri | 2.47 | A | 1 | 1 | 0 | Citrobacter koseri | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-14 | Cronobacter | malonaticus | Cronobacter | sp | 2.29 | Cronobacter | malonaticus | 2.24 | B | 0 | 0 | 1 | Cronobacter sp | Cronobacter malonaticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32945 | 3; confirmed proficency test | EFEx | 2021-02-18 | Cronobacter | sakazakii | Cronobacter | sp | 2.14 | Cronobacter | sp | 1.99 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter | sakazakii | Cronobacter | sp | 2.44 | Cronobacter | sp | 2.33 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30054 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E18, [0034] | Enterobacter | cloacae | Enterobacter | cloacae | 2.34 | Enterobacter | cloacae | 2.08 | A | 1 | 1 | 0 | Enterobacter cloacae | Enterobacter cloacae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K4, [0034] | Enterobacter | sp-ATCC-27988 | Kosakonia | radicincitans | 1.94 | Raoultella | ornithinolytica | 1.68 | B | 0 | 0 | 1 | Kosakonia radicincitans | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae | sp-CVUAS-30072 | Enterobacteriaceae | sp-CVUAS-30072 | 2.52 | Kosakonia | cowanii | 1.45 | A | 1 | 1 | 0 | Enterobacteriaceae sp-CVUAS-30072 | Kosakonia cowanii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0500,2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia | coli | Escherichia | coli | 2.41 | Escherichia | coli | 2.34 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella | grimontii | Klebsiella | oxytoca | 2.42 | Klebsiella | grimontii | 2.39 | B | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella grimontii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2486,3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella | grimontii | Klebsiella | grimontii | 2.66 | Klebsiella | michiganensis | 2.35 | B | 0 | 0 | 1 | Klebsiella grimontii | Klebsiella michiganensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11699 | 3; sequenced [BfR] | eDT | 2017-04-20 06:56 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.47 | Klebsiella | oxytoca | 2.38 | B | 0 | 0 | 1 | Klebsiella michiganensis | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2404,3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.71 | Klebsiella | grimontii | 2.46 | B | 0 | 0 | 1 | Klebsiella michiganensis | Klebsiella grimontii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10038,2 | 3; sequenced [BfR] | EFex | 2019-02-12 14:02 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.41 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11572 | 3; sequenced [BfR] | eDT | 2017-04-20 06:51 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.39 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1482,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:35 | Klebsiella | pneumoniae | Klebsiella | quasipneumoniae | 2.29 | Klebsiella | pneumoniae | 2.24 | B | 0 | 0 | 1 | Klebsiella quasipneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1616,4 | 3; sequenced [BfR] | eDT | 2017-11-22 09:55 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.42 | Klebsiella | pneumoniae | 2.39 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1678 | 3; sequenced [BfR] | eDT | 2016-10-18 06:38 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.44 | Klebsiella | pneumoniae | 2.43 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1701 | 3; sequenced [BfR] | eDT | 2016-10-18 06:41 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.5 | Klebsiella | pneumoniae | 2.44 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30499 | 3; sequenced [BfR] | eDT | 2017-11-22 09:59 | Klebsiella | pneumoniae | Klebsiella | quasipneumoniae | 2.29 | Klebsiella | pneumoniae | 2.24 | B | 0 | 0 | 1 | Klebsiella quasipneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 753,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:34 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.44 | Klebsiella | pneumoniae | 2.32 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10109 | 3; sequenced [BfR] | eDT | 2016-10-18 07:00 | Klebsiella | quasipneumoniae | Klebsiella | pneumoniae | 2.34 | Klebsiella | pneumoniae | 2.24 | A | 0 | 0 | 1 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11677 | 3; sequenced [BfR] | eDT | 2017-04-20 06:54 | Klebsiella | variicola | Klebsiella | variicola | 2.52 | Klebsiella | variicola | 2.42 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4888,2 | 3; sequenced [BfR] | EFex | 2017-07-19 11:36 | Klebsiella | variicola | Klebsiella | variicola | 2.54 | Klebsiella | variicola | 2.39 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30053 CVUAS | 3; type strain | EFEx | 2019-02-12 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.53 | Klebsiella | aerogenes | 2.46 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NM 20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E20 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.48 | Klebsiella | aerogenes | 2.29 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.67 | Klebsiella | michiganensis | 2.37 | A | 1 | 1 | 0 | Klebsiella michiganensis | Klebsiella michiganensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.11.2013 E18 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.42 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 11.04.2008 N21 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.33 | Klebsiella | pneumoniae | 2.33 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30104 | 3; type strain | EFEx | 2013-09-17 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.52 | Klebsiella | pneumoniae | 2.4 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | quasipneumoniae | 2.7 | Klebsiella | quasipneumoniae | 2.3 | A | 1 | 1 | 0 | Klebsiella quasipneumoniae | Klebsiella quasipneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | quasipneumoniae | 2.49 | Klebsiella | quasipneumoniae | 2.41 | A | 1 | 1 | 0 | Klebsiella quasipneumoniae | Klebsiella quasipneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | variicola | Klebsiella | variicola | 2.58 | Klebsiella | variicola | 2.52 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia | sacchari | Kosakonia | sacchari | 2.29 | Salmonella | sp | 1.72 | C | 0 | 0 | 1 | Kosakonia sacchari | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.36 | Plesiomonas | shigelloides | 2.34 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.42 | Plesiomonas | shigelloides | 2.41 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N17, [0034] (Enterobacter gergoviae) | Pluralibacter | gergoviae | Pluralibacter | gergoviae | 2.32 | Pluralibacter | gergoviae | 2.24 | A | 1 | 1 | 0 | Pluralibacter gergoviae | Pluralibacter gergoviae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LT21 Ib 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H9 | Salmonella | enterica | Salmonella | sp | 2.36 | Salmonella | sp | 2.36 | A | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | II 06 06820 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H6 | Shigella | boydii | Escherichia | coli | 2.43 | Escherichia | coli | 2.39 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 3A 07 02503 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H7 | Shigella | flexneri | Shigella | flexneri | 2.46 | Escherichia | coli | 2.46 | C | 0 | 0 | 1 | Shigella flexneri | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea | agglomerans | Pantoea | agglomerans | 2.52 | Pantoea | agglomerans | 2.47 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3493 CVUAS | 3; type strain | eDT | 2015-07-09 | Pantoea | agglomerans | Pantoea | agglomerans | 2.35 | Pantoea | agglomerans | 2.31 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea | sp | Pantoea | sp | 1.99 | Pantoea | agglomerans | 1.93 | B | 0 | 0 | 1 | Pantoea sp | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4263,2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea | sp-FF5 | Pantoea | sp-FF5 | 2.72 | Pantoea | septica | 1.84 | A | 1 | 1 | 0 | Pantoea sp-FF5 | Pantoea septica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella | tarda | Edwardsiella | tarda | 2.47 | Edwardsiella | tarda | 2.47 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N13, [0034] | Edwardsiella | tarda | Edwardsiella | tarda | 2.49 | Edwardsiella | tarda | 2.37 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 18.04.2008, K9, [0034] | Hafnia | alvei | Hafnia | alvei | 2.42 | Hafnia | alvei | 2.32 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia | alvei | Hafnia | alvei | 2.36 | Hafnia | alvei | 2.29 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea | myxofaciens | Cosenzaea | myxofaciens | 2.56 | Cosenzaea | myxofaciens | 2.54 | A | 1 | 1 | 0 | Cosenzaea myxofaciens | Cosenzaea myxofaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella | morganii | Morganella | morganii | 2.56 | Morganella | morganii | 2.55 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella | morganii | Morganella | morganii | 2.68 | Morganella | morganii | 2.61 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus | cibarius | Proteus | cibarius | 2.61 | Proteus | terrae | 2.5 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | hauseri | Proteus | hauseri | 2.6 | Proteus | hauseri | 2.55 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus | hauseri | Proteus | hauseri | 2.47 | Proteus | hauseri | 2.45 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus | mirabilis | Proteus | mirabilis | 2.56 | Proteus | mirabilis | 2.45 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4479 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | mirabilis | Proteus | mirabilis | 2.52 | Proteus | mirabilis | 2.46 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | penneri | Proteus | penneri | 2.48 | Proteus | penneri | 2.43 | A | 1 | 1 | 0 | Proteus penneri | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | cibarius | 2.51 | Proteus | terrae | 2.33 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | vulgaris | 2.45 | Proteus | penneri | 2.41 | B | 0 | 0 | 1 | Proteus vulgaris | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | cibarius | 2.42 | Proteus | terrae | 2.4 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | terrae | Proteus | terrae | 2.66 | Proteus | cibarius | 2.44 | B | 0 | 0 | 1 | Proteus terrae | Proteus cibarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, E5 K16, [0034]: Proteus vulgaris ATCC 33420\LMTG72K16_0\0_K16\1\1SLin\fid | Proteus | vulgaris | Proteus | vulgaris | 2.3 | Proteus | vulgaris | 2.28 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2140 | 3; strain from public collection; sequenced (rpoB) | eDT | 30.11.2017 | Proteus | vulgaris | Proteus | vulgaris | 2.38 | Proteus | vulgaris | 2.33 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia | rettgeri | Providencia | rettgeri | 2.57 | Providencia | rettgeri | 2.52 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia | rettgeri | Providencia | rettgeri | 2.54 | Providencia | rettgeri | 2.31 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25827 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K15, [0034]: Proteus stuartii ATCC 25827\LMTG72K15_0\0_K15\1\1SLin\fid in vers 2016 | Providencia | stuartii | Providencia | stuartii | 2.32 | Providencia | stuartii | 2.31 | A | 1 | 1 | 0 | Providencia stuartii | Providencia stuartii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.24 | Yersinia | enterocolitica | 2.08 | B | 0 | 0 | 1 | Yersinia frederiksenii | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia | mollaretii | Yersinia | mollaretii | 2.32 | Yersinia | mollaretii | 2.25 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 2028 | 3; public strain collection | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H13 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.48 | Yersinia | pseudotuberculosis | 2.48 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H14 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.57 | Yersinia | pseudotuberculosis | 2.43 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia | wautersii | Yersinia | wautersii | 2.29 | Yersinia | pseudotuberculosis | 2.23 | B | 0 | 0 | 1 | Yersinia wautersii | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia | fonticola | Serratia | fonticola | 2.56 | Serratia | fonticola | 2.55 | A | 1 | 1 | 0 | Serratia fonticola | Serratia fonticola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.11.2008, E11, [0034] | Serratia | grimesii | Serratia | grimesii | 2.42 | Serratia | liquefaciens | 2.03 | B | 0 | 0 | 1 | Serratia grimesii | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.19 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.13 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia | quinivorans | Serratia | proteamaculans | 2.16 | Serratia | liquefaciens | 2.12 | B | 0 | 0 | 1 | Serratia proteamaculans | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6221 | 4; sequenced | DT | 2013-01-08 | Yersinia | aldovae | Yersinia | aldovae | 2.33 | Yersinia | aldovae | 2.32 | A | 1 | 1 | 0 | Yersinia aldovae | Yersinia aldovae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 09.03.2007, H12 | Yersinia | aleksiciae | Yersinia | aleksiciae | 2.28 | Yersinia | aleksiciae | 2.28 | A | 1 | 1 | 0 | Yersinia aleksiciae | Yersinia aleksiciae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia | bercovieri | Yersinia | bercovieri | 2.33 | Yersinia | bercovieri | 2.25 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.56 | Yersinia | bercovieri | 2.45 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2012 001153 | 3; sequenced | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.49 | Yersinia | bercovieri | 2.44 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4780 CVUAS | 3; type strain | DT | 2013-05-08 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.49 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.48 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.43 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.34 | Yersinia | enterocolitica | 2.31 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.45 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.57 | Yersinia | enterocolitica | 2.53 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.ma_01\0_C13, [0034] | Yersinia | entomophaga | Yersinia | entomophaga | 2.8 | Yersinia | nurmii | 2.36 | B | 0 | 0 | 1 | Yersinia entomophaga | Yersinia nurmii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RKI 21 RKI | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 22.01.2007 H06 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.41 | Yersinia | frederiksenii | 2.36 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 21 CVUAS | 3; well described strain | DT | 2013-01-30 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.57 | Yersinia | frederiksenii | 2.31 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1824 | 2; sequenced | DT | 2012-12-05 | Yersinia | intermedia | Yersinia | intermedia | 2.46 | Yersinia | intermedia | 2.35 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia | intermedia | Yersinia | intermedia | 2.51 | Yersinia | intermedia | 2.45 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.57 | Yersinia | kristensenii | 2.53 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 302 | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.46 | Yersinia | kristensenii | 2.39 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 25.11.2015, Y.ma_01\0_C11, [0034] | Yersinia | massiliensis | Yersinia | massiliensis | 2.71 | Yersinia | massiliensis | 2.33 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB 925699 | 3; sequenced | DT | 2013-03-19 | Yersinia | massiliensis | Yersinia | massiliensis | 2.33 | Yersinia | massiliensis | 1.85 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18520 CVUAS | 3; type strain | DT | 16.01.2013 | Yersinia | mollaretii | Yersinia | mollaretii | 2.13 | Yersinia | mollaretii | 2.11 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 51 | 2; 0 | DT | 2013-01-16 | Yersinia | mollaretii | Yersinia | mollaretii | 2.42 | Yersinia | mollaretii | 2.4 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.nu_04\0_E10, [0034] | Yersinia | nurmii | Yersinia | nurmii | 2.75 | Yersinia | entomophaga | 2.33 | B | 0 | 0 | 1 | Yersinia nurmii | Yersinia entomophaga | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.pe_04\0_E8, [0034] | Yersinia | pekkanenii | Yersinia | pekkanenii | 2.56 | Yersinia | pekkanenii | 2.36 | A | 1 | 1 | 0 | Yersinia pekkanenii | Yersinia pekkanenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0304,2 DT | 2; biochemistry | DT | 2013-01-31 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.45 | Yersinia | pseudotuberculosis | 2.44 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8992 RKI | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.04.2009, E12 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.69 | Yersinia | pseudotuberculosis | 2.67 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia | rohdei | Yersinia | rohdei | 2.46 | Yersinia | rohdei | 2.41 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 02.03.2007, E07 | Yersinia | rohdei | Yersinia | rohdei | 2.83 | Yersinia | rohdei | 2.31 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2772 CVUAS | 3; sequenced | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.61 | Yersinia | ruckeri | 2.56 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.56 | Yersinia | ruckeri | 2.56 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia | similis | Yersinia | similis | 2.83 | Yersinia | similis | 2.51 | A | 1 | 1 | 0 | Yersinia similis | Yersinia similis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMB 4354 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 13.07.2007, E22 | Yersinia | similis | Yersinia | pseudotuberculosis | 2.52 | Yersinia | pseudotuberculosis | 2.49 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.wa_04\0_E5, [0034] | Yersinia | wautersii | Yersinia | wautersii | 2.69 | Yersinia | pseudotuberculosis | 2.39 | B | 0 | 0 | 1 | Yersinia wautersii | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella | londiniensis | Legionella | londiniensis | 2.27 | Legionella | londiniensis | 1.59 | A | 1 | 1 | 0 | Legionella londiniensis | Legionella londiniensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella | pneumophila | Legionella | pneumophila | 2.21 | Legionella | pneumophila | 1.73 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4122 | 3; 0 | eDT | 20.01.2016 | Legionella | pneumophila | Legionella | pneumophila | 2.33 | Legionella | pneumophila | 2.29 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella | taurinensis | Legionella | taurinensis | 2.41 | Legionella | sp | 1.86 | A | 1 | 1 | 0 | Legionella taurinensis | Legionella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10873,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | stevensii | 2.33 | Halomonas | aquamarina | 1.54 | A | 1 | 1 | 0 | Halomonas stevensii | Halomonas aquamarina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10874,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | stevensii | 2.43 | Dermacoccus | sp[3] | 1.35 | A | 1 | 1 | 0 | Halomonas stevensii | Dermacoccus sp[3] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) | caballi | Pasteurella | caballi | 2.22 | Pasteurella | caballi | 2.15 | A | 0 | 0 | 1 | Pasteurella caballi | Pasteurella caballi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) | caballi | Pasteurella | caballi | 2.54 | (Pasteurella) | caballi | 2.53 | C | 0 | 0 | 1 | Pasteurella caballi | (Pasteurella) caballi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) | testudinis | (Pasteurella) | testudinis | 2.33 | Pasteurella | testudinis | 2.01 | C | 0 | 0 | 1 | (Pasteurella) testudinis | Pasteurella testudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 957 | 3; sequenced | eDT | 02.12.2015 | (Pasteurella) | testudinis | (Pasteurella) | testudinis | 2.11 | Pasteurella | testudinis | 1.87 | C | 0 | 0 | 1 | (Pasteurella) testudinis | Pasteurella testudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | anseriformium | Actinobacillus | anseriformium | 2.48 | Actinobacillus | equuli | 2.05 | B | 0 | 0 | 1 | Actinobacillus anseriformium | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15UGF 198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus | anseriformium | Actinobacillus | anseriformium | 2.53 | Actinobacillus | anseriformium | 2.22 | A | 1 | 1 | 0 | Actinobacillus anseriformium | Actinobacillus anseriformium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus | arthritidis | Actinobacillus | arthritidis | 2.33 | Actinobacillus | arthritidis_JRA_Pasturella-4 | 2 | B | 0 | 0 | 1 | Actinobacillus arthritidis | Actinobacillus arthritidis_JRA_Pasturella-4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.42 | Actinobacillus | capsulatus | 2.2 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus capsulatus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.52 | Actinobacillus | capsulatus | 2.42 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus capsulatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.4 | Actinobacillus_equuli_subsp._equuli_ccug2041_DI05 | equuli | 2.4 | C | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus_equuli_subsp._equuli_ccug2041_DI05 equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.28 | Actinobacillus | suis | 2.22 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.36 | Actinobacillus | equuli | 2.35 | A | 1 | 1 | 0 | Actinobacillus equuli | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus | equuli | Actinobacillus | equuli | 2.52 | Actinobacillus | suis | 2.52 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus | lignieresii | Actinobacillus | pleuropneumoniae | 2.34 | Actinobacillus | lignieresii | 2.33 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9945,2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | lignieresii | Actinobacillus | pleuropneumoniae | 2.45 | Actinobacillus | lignieresii | 2.37 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus | rossii | Actinobacillus | rossii | 2.36 | Actinobacillus | rossii | 2.08 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus | rossii | Actinobacillus | rossii | 2.67 | Actinobacillus | rossii | 2.64 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32505,2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus | sp-CVUAS-32505,2 | Actinobacillus | sp-CVUAS-32505,2 | 2.54 | Actinobacillus | arthritidis_JRA_Pasturella-4 | 2.18 | B | 0 | 0 | 1 | Actinobacillus sp-CVUAS-32505,2 | Actinobacillus arthritidis_JRA_Pasturella-4 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | sp-genomospecies-1 | 2.46 | Actinobacillus_genomospecies1_ccug22231_DI05 | sp-genomospecies-1 | 2.3 | C | 0 | 0 | 1 | Actinobacillus sp-genomospecies-1 | Actinobacillus_genomospecies1_ccug22231_DI05 sp-genomospecies-1 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | sp-genomospecies-1 | 2.53 | Actinobacillus_genomospecies1_ccug22231_DI05 | sp-genomospecies-1 | 2.35 | C | 0 | 0 | 1 | Actinobacillus sp-genomospecies-1 | Actinobacillus_genomospecies1_ccug22231_DI05 sp-genomospecies-1 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus | sp-MK-2012 | Actinobacillus | sp-MK-2012 | 2.67 | Actinobacillus | sp-MK-2012 | 2.34 | A | 1 | 1 | 0 | Actinobacillus sp-MK-2012 | Actinobacillus sp-MK-2012 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-MK-2012 | Actinobacillus | sp-MK-2012 | 2.46 | Actinobacillus | sp-MK-2012 | 2.44 | A | 1 | 1 | 0 | Actinobacillus sp-MK-2012 | Actinobacillus sp-MK-2012 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus | suis | Actinobacillus | suis | 2.6 | Actinobacillus | suis | 2.55 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus | suis | Actinobacillus | suis | 2.5 | Actinobacillus | suis | 2.47 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium | gallinarum | Avibacterium | gallinarum | 2.19 | Avibacterium | gallinarum | 2.15 | A | 1 | 1 | 0 | Avibacterium gallinarum | Avibacterium gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium | gallinarum | Avibacterium | gallinarum | 2.72 | Avibacterium | gallinarum | 2.27 | A | 1 | 1 | 0 | Avibacterium gallinarum | Avibacterium gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.52 | Bibersteinia | trehalosi | 2.48 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.47 | Bibersteinia | trehalosi | 2.39 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter | oris | Chelonobacter | oris | 2.42 | Chelonobacter | oris | 2.41 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter | oris | Chelonobacter | oris | 2.11 | Chelonobacter | oris | 2.05 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3802,2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter | osteomyelitidis | Cricetibacter | osteomyelitidis | 2.54 | Haemophilus | haemolyticus | 1.43 | A | 1 | 1 | 0 | Cricetibacter osteomyelitidis | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia | canicola | Frederiksenia | canicola | 2.41 | Frederiksenia | canicola | 2.21 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia | canicola | Frederiksenia | canicola | 2.32 | Frederiksenia | canicola | 2.28 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium | anatis | Gallibacterium | anatis | 2.07 | Gallibacterium | anatis | 2.04 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium | anatis | Gallibacterium | anatis | 2.45 | Gallibacterium | anatis | 2.36 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium | sp-F151 | Gallibacterium | sp-F151 | 2.49 | Gallibacterium | anatis | 1.44 | A | 1 | 1 | 0 | Gallibacterium sp-F151 | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium | trehalosifermentans | Gallibacterium | trehalosifermentans | 2.65 | Gallibacterium | sp-F151 | 1.88 | A | 1 | 1 | 0 | Gallibacterium trehalosifermentans | Gallibacterium sp-F151 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus | haemoglobinophilus | Haemophilus | haemoglobinophilus | 2.29 | Haemophilus | sp-2269 | 1.91 | A | 1 | 1 | 0 | Haemophilus haemoglobinophilus | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | sp-2269 | 2.39 | Haemophilus | sp-2269 | 2.35 | A | 1 | 1 | 0 | Haemophilus sp-2269 | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | sp-2269 | 2.51 | Haemophilus | sp-2269 | 2.46 | A | 1 | 1 | 0 | Haemophilus sp-2269 | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P1066-2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus | somni | Histophilus | somni | 2.2 | Histophilus | somni | 1.82 | A | 1 | 1 | 0 | Histophilus somni | Histophilus somni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia | caviae | Mannheimia | caviae | 2.62 | Mannheimia | caviae | 2.22 | A | 1 | 1 | 0 | Mannheimia caviae | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1064,4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia | glucosida | Mannheimia | glucosida | 2.35 | Mannheimia | glucosida | 2.22 | A | 1 | 1 | 0 | Mannheimia glucosida | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia | glucosida | Mannheimia | glucosida | 2.59 | Mannheimia | glucosida | 2.47 | A | 1 | 1 | 0 | Mannheimia glucosida | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia | granulomatis | Mannheimia | sp-14UCF195 | 2.3 | Mannheimia | granulomatis | 2.05 | B | 0 | 0 | 1 | Mannheimia sp-14UCF195 | Mannheimia granulomatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia | granulomatis | Mannheimia | granulomatis | 2.48 | Mannheimia | sp-14UCF195 | 2.26 | B | 0 | 0 | 1 | Mannheimia granulomatis | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.53 | Mannheimia | haemolytica | 2.47 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6197,2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.57 | Mannheimia | haemolytica | 2.52 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia | pernigra | Mannheimia | pernigra | 2.51 | Mannheimia | pernigra | 2.11 | A | 1 | 1 | 0 | Mannheimia pernigra | Mannheimia pernigra | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia | pernigra | Mannheimia | pernigra | 2.32 | Mannheimia | pernigra | 2.32 | A | 1 | 1 | 0 | Mannheimia pernigra | Mannheimia pernigra | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.57 | Mannheimia | ruminalis | 2.3 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.3 | Mannheimia | ruminalis | 2.19 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10396,2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia | sp-10-12559-1 | Mannheimia | sp-10-12559-1 | 2.46 | Mannheimia | caviae | 2.03 | B | 0 | 0 | 1 | Mannheimia sp-10-12559-1 | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10183,3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia | sp-14UCF195 | Mannheimia | sp-14UCF195 | 2.64 | Mannheimia | sp-14UCF195 | 2.32 | A | 1 | 1 | 0 | Mannheimia sp-14UCF195 | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia | sp-14UCF195 | Mannheimia | sp-14UCF195 | 2.68 | Mannheimia | sp-14UCF195 | 2.29 | A | 1 | 1 | 0 | Mannheimia sp-14UCF195 | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia | sp-15USF15 | Mannheimia | sp-15USF15 | 2.68 | Mannheimia | indoligenes | 2.46 | B | 0 | 0 | 1 | Mannheimia sp-15USF15 | Mannheimia indoligenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia | sp-15USF15 | Mannheimia | sp-15USF15 | 2.49 | Mannheimia | caviae | 2.13 | B | 0 | 0 | 1 | Mannheimia sp-15USF15 | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia | sp-CVUAS-11240 | Mannheimia | sp-CVUAS-11240 | 2.65 | Mannheimia | sp-15USF15 | 1.94 | A | 1 | 1 | 0 | Mannheimia sp-CVUAS-11240 | Mannheimia sp-15USF15 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia | sp-CVUAS-1496 | Mannheimia | sp-CVUAS-1496 | 2.47 | Mannheimia | glucosida | 2.37 | B | 0 | 0 | 1 | Mannheimia sp-CVUAS-1496 | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.5 | Necropsobacter | rosorum | 2.43 | A | 1 | 1 | 0 | Necropsobacter rosorum | Necropsobacter rosorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.36 | Necropsobacter | rosorum | 2.35 | A | 1 | 1 | 0 | Necropsobacter rosorum | Necropsobacter rosorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella | semolina | Nicoletella | semolina | 2.55 | Nicoletella | semolina | 2.43 | A | 1 | 1 | 0 | Nicoletella semolina | Nicoletella semolina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella | semolina | Nicoletella | semolina | 2.46 | Nicoletella | semolina | 2.46 | A | 1 | 1 | 0 | Nicoletella semolina | Nicoletella semolina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4060,3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella | aerogenes | Pasteurella | aerogenes | 2.47 | Pasteurella | aerogenes | 2.05 | A | 1 | 1 | 0 | Pasteurella aerogenes | Pasteurella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | canis | Pasteurella | canis | 2.07 | Pasteurella | canis | 2.04 | A | 1 | 1 | 0 | Pasteurella canis | Pasteurella canis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | mairii | (Pasteurella) | mairii | 2.28 | Pasteurella | mairii | 2.11 | C | 0 | 0 | 1 | (Pasteurella) mairii | Pasteurella mairii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella | multocida | Pasteurella | multocida | 2.48 | Pasteurella | multocida | 2.33 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32055 | 3; specific PCR | eDT | 2019-11-14 | Pasteurella | multocida | Pasteurella | multocida | 2.67 | Pasteurella | multocida | 2.59 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella | multocida | Pasteurella | multocida | 2.62 | Pasteurella | multocida | 2.52 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.57 | Pasteurella | multocida | 2.56 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.51 | Pasteurella | multocida | 2.5 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella | multocida | Pasteurella | multocida | 2.57 | Pasteurella | multocida | 2.49 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32506,2 | 3; sequenced | EFEx | 2020-11-19 | Pasteurellaceae | Bisgaard-Taxon-10 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.58 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.39 | A | 1 | 1 | 0 | Pasteurellaceae Bisgaard-Taxon-10 | Pasteurellaceae Bisgaard-Taxon-10 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Pasteurellaceae | Bisgaard-Taxon-10 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.36 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.34 | A | 1 | 1 | 0 | Pasteurellaceae Bisgaard-Taxon-10 | Pasteurellaceae Bisgaard-Taxon-10 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | Bisgaard-Taxon-40 | Pasteurellaceae | Bisgaard-Taxon-40 | 2.45 | Haemophilus | parainfluenzae | 1.36 | A | 1 | 1 | 0 | Pasteurellaceae Bisgaard-Taxon-40 | Haemophilus parainfluenzae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9987,2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae | canine-oral-taxon-080 | Pasteurellaceae | canine-oral-taxon-080 | 2.56 | Frederiksenia | canicola | 1.83 | C | 0 | 0 | 1 | Pasteurellaceae canine-oral-taxon-080 | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 261,4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae | sp-CVUAS-261,4 | Pasteurellaceae | sp-CVUAS-261,4 | 2.64 | Klebsiella | oxytoca | 1.47 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-261,4 | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae | sp-CVUAS-31988 | Pasteurellaceae | sp-CVUAS-31988 | 2.61 | Haemophilus | haemolyticus | 1.65 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-31988 | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | sp-CVUAS-4868 | Pasteurellaceae | sp-CVUAS-4868 | 2.54 | Enterococcus | faecalis | 1.34 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-4868 | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 724,3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae | sp-CVUAS-724,3 | Pasteurellaceae | sp-CVUAS-724,3 | 2.6 | Rodentibacter | heylii | 1.47 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-724,3 | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae | sp-CVUAS-9305 | Pasteurellaceae | sp-CVUAS-9305 | 2.65 | Actinobacillus | equuli | 1.41 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-9305 | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9967,2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae | sp-CVUAS-9967,2 | Pasteurellaceae | sp-CVUAS-9967,2 | 2.45 | Haemophilus | haemolyticus | 1.74 | C | 0 | 0 | 1 | Pasteurellaceae sp-CVUAS-9967,2 | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae | sp-LHL-U-309 | Pasteurellaceae | sp-LHL-U-309 | 2.36 | Pasteurellaceae | sp-LHL-U-309 | 2.13 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-309 | Pasteurellaceae sp-LHL-U-309 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL U-309 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-309 | Pasteurellaceae | sp-LHL-U-309 | 2.46 | Pasteurellaceae | sp-LHL-U-309 | 2.22 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-309 | Pasteurellaceae sp-LHL-U-309 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL U-3444 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-3444 | Pasteurellaceae | sp-LHL-U-3444 | 2.54 | Chelonobacter | oris | 1.54 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-3444 | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae | sp-MCCM-2539 | Pasteurellaceae | sp-MCCM-2539 | 2.54 | Haemophilus | pittmaniae | 1.53 | A | 1 | 1 | 0 | Pasteurellaceae sp-MCCM-2539 | Haemophilus pittmaniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 308,2 | 3; sequenced | DT | 2013-11-06 | Psittacicella | hinzii | Psittacicella | hinzii | 2.57 | Agromyces | cerinus | 1.43 | A | 1 | 1 | 0 | Psittacicella hinzii | Agromyces cerinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.48 | Rodentibacter | heylii | 2.4 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.55 | Rodentibacter | heylii | 2.5 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium | mucosae | Spirabiliibacterium | mucosae | 2.39 | Spirabiliibacterium | mucosae | 2.3 | A | 1 | 1 | 0 | Spirabiliibacterium mucosae | Spirabiliibacterium mucosae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium | mucosae | Spirabiliibacterium | mucosae | 2.49 | Spirabiliibacterium | mucosae | 2.45 | A | 1 | 1 | 0 | Spirabiliibacterium mucosae | Spirabiliibacterium mucosae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8966 | 3; sequenced | DT | 23.07.2013 | Volucribacter | amazonae | Volucribacter | amazonae | 2.43 | Volucribacter | psittacicida | 2.21 | B | 0 | 0 | 1 | Volucribacter amazonae | Volucribacter psittacicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E22 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.51 | Acinetobacter | baumannii | 2.48 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2014 C05 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.46 | Acinetobacter | baumannii | 2.41 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter | defluvii | Acinetobacter | defluvii | 2.42 | Acinetobacter | defluvii | 2.16 | A | 1 | 1 | 0 | Acinetobacter defluvii | Acinetobacter defluvii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | guillouiae | Acinetobacter | guillouiae | 2.34 | Acinetobacter | guillouiae | 2.24 | A | 1 | 1 | 0 | Acinetobacter guillouiae | Acinetobacter guillouiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | johnsonii | Acinetobacter | johnsonii | 2.3 | Acinetobacter | johnsonii | 2.26 | A | 1 | 1 | 0 | Acinetobacter johnsonii | Acinetobacter johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | junii | Acinetobacter | junii | 2.33 | Acinetobacter | junii | 2.31 | A | 1 | 1 | 0 | Acinetobacter junii | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | lwoffii | Acinetobacter | lwoffii | 2.35 | Acinetobacter | lwoffii | 2.27 | A | 1 | 1 | 0 | Acinetobacter lwoffii | Acinetobacter lwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7379,3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.29 | Acinetobacter | pseudolwoffii | 2.26 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.08 | Acinetobacter | pseudolwoffii | 2.07 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | variabilis | Acinetobacter | variabilis | 2.35 | Acinetobacter | variabilis | 2.17 | A | 1 | 1 | 0 | Acinetobacter variabilis | Acinetobacter variabilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | ovis | 2.44 | Moraxella | bovoculi | 2.16 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | ovis | 2.09 | Moraxella | bovoculi | 2.06 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella | catarrhalis | Moraxella | catarrhalis | 2.53 | Moraxella | catarrhalis | 2.51 | A | 1 | 1 | 0 | Moraxella catarrhalis | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11674,3 | 3; sequenced | eDT | 2022-03-22 | Moraxella | cuniculi | Moraxella | cuniculi | 2.55 | Moraxella | catarrhalis | 1.63 | A | 1 | 1 | 0 | Moraxella cuniculi | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella | osloensis | Moraxella | osloensis | 2.53 | Moraxella | osloensis | 1.94 | A | 1 | 1 | 0 | Moraxella osloensis | Moraxella osloensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella | ovis | Moraxella | ovis | 2.39 | Moraxella | bovoculi | 2.32 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | frigidicola | Psychrobacter | frigidicola | 2.49 | Psychrobacter | pulmonis | 1.47 | A | 1 | 1 | 0 | Psychrobacter frigidicola | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter | pasteurii | Psychrobacter | pasteurii | 2.52 | Psychrobacter | pasteurii | 2.04 | A | 1 | 1 | 0 | Psychrobacter pasteurii | Psychrobacter pasteurii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.58 | Psychrobacter | pulmonis | 2.35 | A | 1 | 1 | 0 | Psychrobacter pulmonis | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.3 | Psychrobacter | faecalis | 1.83 | A | 1 | 1 | 0 | Psychrobacter pulmonis | Psychrobacter faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas | (alkylphenolica) | Pseudomonas | alkylphenolica | 2.4 | Pseudomonas | japonica | 1.59 | A | 0 | 0 | 1 | Pseudomonas alkylphenolica | Pseudomonas japonica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.33 | Pseudomonas | aeruginosa | 2.28 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007 H8 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.46 | Pseudomonas | aeruginosa | 2.38 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas | baetica | Pseudomonas | baetica | 2.66 | Pseudomonas | koreensis | 2.06 | B | 0 | 0 | 1 | Pseudomonas baetica | Pseudomonas koreensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.02.2008 N11 | Pseudomonas | chlororaphis | Pseudomonas | chlororaphis | 2.44 | Pseudomonas | chlororaphis | 2.3 | A | 1 | 1 | 0 | Pseudomonas chlororaphis | Pseudomonas chlororaphis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 02.04.2012, H21 | Francisella | tularensis | Francisella | tularensis | 2.55 | Francisella | tularensis | 2.33 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 12.12.2016, F19, [0096] | Allofrancisella | guangzhouensis | Allofrancisella | guangzhouensis | 2.69 | Paracoccus | denitrificans | 1.33 | A | 1 | 1 | 0 | Allofrancisella guangzhouensis | Paracoccus denitrificans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 18.11.2015, L16, [0096] | Francisella | hispaniensis | Francisella | hispaniensis | 2.75 | Francisella | tularensis | 1.99 | A | 1 | 1 | 0 | Francisella hispaniensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 25.11.2015, I3, [0096] | Francisella | naotunensis | Francisella | orientalis | 2.81 | Francisella | philomiragia | 1.57 | A | 0 | 0 | 1 | Francisella orientalis | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4680,2 | 3; FLI: specific PCR | DT | 2013-04-02 | Francisella | tularensis | Francisella | tularensis | 2.23 | Francisella | tularensis | 2.2 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 23.04.2013, F20, [0096] | Francisella | tularensis | Francisella | tularensis | 2.36 | Francisella | tularensis | 2.34 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 26.04.2013, I20, [0096] | Francisella | tularensis | Francisella | tularensis | 2.68 | Francisella | tularensis | 2.26 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 01.06.2017, F3, [0096] | Francisella | tularensis | Francisella | tularensis | 2.8 | Francisella | tularensis | 2.44 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio | aestuarianus | Vibrio | aestuarianus | 2.52 | Vibrio | aestuarianus | 2.5 | A | 1 | 1 | 0 | Vibrio aestuarianus | Vibrio aestuarianus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio | anguillarum | Vibrio | anguillarum | 2.34 | Vibrio | anguillarum | 2.3 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio | campbelli | Vibrio | campbellii | 2.32 | Vibrio | campbellii | 2.28 | A | 0 | 0 | 1 | Vibrio campbellii | Vibrio campbellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio | campbelli | Vibrio | campbellii | 2.38 | Vibrio | campbellii | 2.37 | A | 0 | 0 | 1 | Vibrio campbellii | Vibrio campbellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16-VB00067-0_BfR 1_eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio | cholerae | Vibrio | cholerae | 2.37 | Vibrio | cholerae | 2.37 | A | 1 | 1 | 0 | Vibrio cholerae | Vibrio cholerae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 10-VBH0202 DT-48h | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | cincinnatiensis | Vibrio | cincinnatiencis | 2.43 | Vibrio | cincinnatiensis | 2.41 | B | 0 | 0 | 1 | Vibrio cincinnatiencis | Vibrio cincinnatiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00004 LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 2.55 | Vibrio | cincinnatiencis | 2.44 | B | 0 | 0 | 1 | Vibrio cincinnatiensis | Vibrio cincinnatiencis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.08 | Vibrio | diazotrophicus | 1.98 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio | fluvialis | Vibrio | fluvialis | 2.42 | Vibrio | fluvialis | 2.16 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00036 LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio | fluvialis | Vibrio | fluvialis | 2.22 | Vibrio | fluvialis | 2.18 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 16.08.2012, K12, [0034] | Vibrio | harveyi | Vibrio | harveyi | 2.28 | Vibrio | harveyi | 2.21 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio | harveyi | Vibrio | harveyi | 2.41 | Vibrio | harveyi | 2.4 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00047 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio | jasicida | Vibrio | jasicida | 2.53 | Vibrio | harveyi | 2.49 | B | 0 | 0 | 1 | Vibrio jasicida | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00053 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio | jasicida | Vibrio | jasicida | 2.41 | Vibrio | jasicida | 2.36 | A | 1 | 1 | 0 | Vibrio jasicida | Vibrio jasicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 12-VBH0498 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metoecus | Vibrio | mimicus | 2.48 | Vibrio | metoecus | 2.41 | B | 0 | 0 | 1 | Vibrio mimicus | Vibrio metoecus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00016 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | metoecus | Vibrio | metoecus | 2.34 | Vibrio | metoecus | 2.32 | A | 1 | 1 | 0 | Vibrio metoecus | Vibrio metoecus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metschnikovii | Vibrio | metschnikovii | 2.15 | Vibrio | metschnikovii | 2.1 | A | 1 | 1 | 0 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | navarrensis | Vibrio | navarrensis | 2.16 | Vibrio | navarrensis | 1.99 | A | 1 | 1 | 0 | Vibrio navarrensis | Vibrio navarrensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio | owensii | Vibrio | owensii | 2.35 | Vibrio | jasicida | 2.27 | B | 0 | 0 | 1 | Vibrio owensii | Vibrio jasicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00254-0_BfR 18_DT | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.28 | Vibrio | parahaemolyticus | 2.26 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio | vulnificus | Vibrio | vulnificus | 2.45 | Vibrio | vulnificus | 2.35 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33180 25C | 3; sequenced | DT | 2021-08-04 | Photobacterium | damselae | Photobacterium | damselae | 2.22 | Photobacterium | damselae | 2.02 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium | damselae | Photobacterium | damselae | 2.14 | Photobacterium | damselae | 2.07 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1266,2 | 3; sequenced | EFEx | 25.09.1019 | Photobacterium | piscicola | Photobacterium | piscicola | 2.6 | Photobacterium | iliopiscarium | 1.57 | A | 1 | 1 | 0 | Photobacterium piscicola | Photobacterium iliopiscarium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 842,2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium | swingsii | Photobacterium | swingsii | 2.63 | Providencia | alcalifaciens | 1.37 | A | 1 | 1 | 0 | Photobacterium swingsii | Providencia alcalifaciens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7079 sir | 3; spezific PCR (CVFR) | eDT | 2022-06-23 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.18 | Vibrio | campbellii | 2.08 | B | 0 | 0 | 1 | Vibrio alginolyticus | Vibrio campbellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2171 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 12.07.2012 E12 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.3 | Vibrio | alginolyticus | 2.29 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | A8-12 sir | 3; gyrB sequence | eDT | 2022-06-23 | Vibrio | anguillarum | Vibrio | anguillarum | 2.22 | Vibrio | anguillarum | 2.2 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10227 | 3; sequenced | eDT | 2013-05-22 | Vibrio | cholerae | Vibrio | cholerae | 2.26 | Vibrio | cholerae | 2.14 | A | 1 | 1 | 0 | Vibrio cholerae | Vibrio cholerae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8995 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.14 | Vibrio | diazotrophicus | 2.14 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 14383 sir | 3; public strain collection, sequenced | eDT | 2022-06-23 | Vibrio | furnissi | Vibrio | furnissii | 2.21 | Vibrio | furnissii | 2.06 | A | 0 | 0 | 1 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1754 eDT | 3; sequenced | eDT | 2013-04-16 | Vibrio | furnissii | Vibrio | furnissii | 2.2 | Vibrio | furnissii | 2.05 | A | 1 | 1 | 0 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 619 sir | 2; 0 | eDT | 2022-06-23 | Vibrio | metschnikovii | Vibrio | metschnikovii | 2.35 | Vibrio | metschnikovii | 2.3 | A | 1 | 1 | 0 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7872 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio | mimicus | Vibrio | mimicus | 2.41 | Vibrio | metoecus | 2.37 | B | 0 | 0 | 1 | Vibrio mimicus | Vibrio metoecus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19130 RKI | 4; type strain; WGS sequenced | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 24.01.2011, N4, [0034] | Vibrio | mimicus | Vibrio | mimicus | 2.31 | Vibrio | mimicus | 2.28 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10027 eDT | 3; type strain | eDT | 2022-06-23 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.21 | Vibrio | parahaemolyticus | 2.16 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 871,2 | 3; sequenced | eDT | 2014-08-21 | Vibrio | sp-CVUAS-871,2 | Vibrio | sp-CVUAS-871,2 | 2.28 | Vibrio | mimicus | 1.45 | A | 1 | 1 | 0 | Vibrio sp-CVUAS-871,2 | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778,2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio | tubiashii | Vibrio | tubiashii | 2.05 | Vibrio | europaeus | 2.03 | B | 0 | 0 | 1 | Vibrio tubiashii | Vibrio europaeus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778,2 eDT | 3; sequenced | eDT | 2019-10-24 | Vibrio | tubiashii | Vibrio | tubiashii | 2.51 | Vibrio | brasiliensis | 2.18 | B | 0 | 0 | 1 | Vibrio tubiashii | Vibrio brasiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LVM05 LVU05 | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-14 | Vibrio | vulnificus | Vibrio | vulnificus | 2.46 | Vibrio | vulnificus | 2.45 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas | sp-G9077 | Wohlfahrtiimonas | sp-G9077 | 2.42 | Wohlfahrtiimonas | chitiniclastica | 2.35 | B | 0 | 0 | 1 | Wohlfahrtiimonas sp-G9077 | Wohlfahrtiimonas chitiniclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria | indica | Ignatzschineria | indica | 2.43 | Ignatzschineria | indica | 2.18 | A | 1 | 1 | 0 | Ignatzschineria indica | Ignatzschineria indica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria | ureiclastica | Ignatzschineria | ureiclastica | 2.58 | Ignatzschineria | ureiclastica | 2.42 | A | 1 | 1 | 0 | Ignatzschineria ureiclastica | Ignatzschineria ureiclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas | chelatiphaga | Stenotrophomonas | chelatiphaga | 2.14 | Stenotrophomonas | maltophilia | 1.61 | A | 1 | 1 | 0 | Stenotrophomonas chelatiphaga | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2009 K17 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.26 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.05.2009 K22 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.22 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.42 | Stenotrophomonas | rhizophila | 1.8 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas rhizophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.05.2015 C17 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.77 | Stenotrophomonas | rhizophila | 2.4 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas rhizophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4239,2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas | tumulicola | Stenotrophomonas | tumulicola | 2.47 | Stenotrophomonas | chelatiphaga | 1.69 | A | 1 | 1 | 0 | Stenotrophomonas tumulicola | Stenotrophomonas chelatiphaga | 0 | 0 | 0 | 0

## #Parameter (Spezies)
Parameter: | Corynebacterium | pseudotuberculosis | Codierung | Validität | Methode | Datum | verwendeter Organismus | Score | richtig-negativ | falsch-positiv
Test auf spektrale Nähe zum Parameter |  |  | DSM 44123 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; K4 | Corynebacterium diphtheriae | 1.47 | 1 | 0
Anzahl Isolate: | 1278 |  | KL163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.35 | 1 | 0
|  |  | KL167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.07 | 1 | 0
richtig-negativ: | 1199 |  | KL187 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.22 | 1 | 0
|  |  | KL232 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.09 | 1 | 0
falsch-positiv: | 79 |  | KL237 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.46 | 1 | 0
|  |  | KL240 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-13 | Corynebacterium diphtheriae | 1.38 | 1 | 0
| 6.182% |  | KL932 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.26 | 1 | 0
|  |  | KL933 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.49 | 1 | 0
|  |  | KL935 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.57 | 1 | 0
Scorewertgrenzen: | 2 | Species | KL936 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.41 | 1 | 0
| 1.7 | Genus | KL950 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.36 | 1 | 0
Es wird gegen eine nur aus dem Parameter bestehende DB geprüft. |  |  | KL952 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.42 | 1 | 0
|  |  | KL173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.38 | 1 | 0
|  |  | KL179 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.28 | 1 | 0
|  |  | KL235 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 1.18 | 1 | 0
|  |  | KL940 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.28 | 1 | 0
|  |  | KL954 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 0.92 | 1 | 0
|  |  | KL956 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 1.37 | 1 | 0
|  |  | DSM 44123 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; K4 | Corynebacterium diphtheriae | 1.39 | 1 | 0
|  |  | DSM 43988 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; I7 | Corynebacterium diphtheriae | 1.27 | 1 | 0
|  |  | 141000461 (2) | 1; 0 | 0 | 2014-01-20 | Corynebacterium silvaticum | 1.71 | 1 | 0
|  |  | 141016278 | 3; Confirmation by KL Diphtheriae (LGL); WBC: FT-IR | 0 | 2014-12-22 | Corynebacterium silvaticum | 1.84 | 1 | 0
|  |  | 161003120-1 | 1; 0 | 0 | 2016-03-15 | Corynebacterium silvaticum | 1.76 | 1 | 0
|  |  | 161003120-2 | 1; 0 | 0 | 2016-03-15 | Corynebacterium silvaticum | 1.78 | 1 | 0
|  |  | 161003120-3 | 3; Confirmation by KL Diphtheriae (LGL) | 0 | 2016-03-16 | Corynebacterium silvaticum | 1.91 | 1 | 0
|  |  | 161003120-4 | 1; 0 | 0 | 2016-03-14 | Corynebacterium silvaticum | 1.88 | 1 | 0
|  |  | 161003120-5 | 1; 0 | 0 | 2016-03-14 | Corynebacterium silvaticum | 1.98 | 1 | 0
|  |  | 161003120-6 | 1; 0 | 0 | 2016-03-14 | Corynebacterium silvaticum | 1.92 | 1 | 0
|  |  | 10-7-D-00025 | 3; sequenced | DT | 2013-01-08 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | 11USF28 | 3; sequenced | DT | 2012-12-11 | Corynebacterium silvaticum | 1.83 | 1 | 0
|  |  | 11USF53 | 3; sequenced | DT | 2013-01-16 | Corynebacterium silvaticum | 1.86 | 1 | 0
|  |  | 11USF78 | 3; sequenced | DT | 2012-12-11 | Corynebacterium silvaticum | 1.8 | 1 | 0
|  |  | 11USF79 | 3; sequenced | DT | 2012-12-11 | Corynebacterium silvaticum | 1.79 | 1 | 0
|  |  | 121017479 | 3; sequenced | DT | 2012-12-12 | Corynebacterium silvaticum | 1.73 | 1 | 0
|  |  | 131000349 | 3; sequenced | DT | 2013-01-18 | Corynebacterium silvaticum | 1.9 | 1 | 0
|  |  | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium silvaticum | 1.89 | 1 | 0
|  |  | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium silvaticum | 1.95 | 1 | 0
|  |  | AGES WS 057 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.8 | 1 | 0
|  |  | AGES WS 076 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.69 | 1 | 0
|  |  | AGES WS 085 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.9 | 1 | 0
|  |  | AGES WS 091 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 2.06 | 0 | 1
|  |  | AGES WS 098 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | AGES WS 111 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.85 | 1 | 0
|  |  | AGES WS 123 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.93 | 1 | 0
|  |  | AGES WS 128 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.88 | 1 | 0
|  |  | AGES WS 130 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.86 | 1 | 0
|  |  | AGES WS 139 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.92 | 1 | 0
|  |  | AGES WS 141 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.85 | 1 | 0
|  |  | AGES WS 145 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.67 | 1 | 0
|  |  | AGES WS 147 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | AGES WS 149 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.67 | 1 | 0
|  |  | AGES WS 156 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.84 | 1 | 0
|  |  | AGES WS 166 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.86 | 1 | 0
|  |  | AGES WS 168 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.81 | 1 | 0
|  |  | AGES WS 182 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.84 | 1 | 0
|  |  | AGES WS 183 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | AGES WS 188 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.84 | 1 | 0
|  |  | AGES WS 233 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | AGES WS 257 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.9 | 1 | 0
|  |  | AGES WS 332 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium silvaticum | 1.98 | 1 | 0
|  |  | AGES WS 344 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.75 | 1 | 0
|  |  | AGES WS 351 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.76 | 1 | 0
|  |  | AGES WS 353 | 2; 0 | 0 | 2019-02-27 | Corynebacterium silvaticum | 1.87 | 1 | 0
|  |  | BU125-97 | 3; sequenced | DT | 2012-10-19 | Corynebacterium silvaticum | 1.88 | 1 | 0
|  |  | CVUAS 10179,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium silvaticum | 1.94 | 1 | 0
|  |  | CVUAS 10701 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium silvaticum | 1.8 | 1 | 0
|  |  | CVUAS 11363 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium silvaticum | 1.78 | 1 | 0
|  |  | CVUAS 11488 | 3; FT-IR (NTTB-D) | eDT | 2017-01-11 | Corynebacterium silvaticum | 1.71 | 1 | 0
|  |  | CVUAS 1400,3 | 3; FT-IR (NTTB-D) | eDT | 2018-03-02 | Corynebacterium silvaticum | 1.83 | 1 | 0
|  |  | CVUAS 1717,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium silvaticum | 1.45 | 1 | 0
|  |  | CVUAS 2040,2 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | CVUAS 2738,2 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium silvaticum | 1.74 | 1 | 0
|  |  | CVUAS 4292 | 3; type strain [0127] | DT | 2012-12-08 | Corynebacterium silvaticum | 1.91 | 1 | 0
|  |  | CVUAS 5160 | 3; sequenced | DT | 2012-10-08 | Corynebacterium silvaticum | 1.78 | 1 | 0
|  |  | CVUAS 5243,2 | 3; FT-IR (NTTB-D) | DT | 2013-11-18 | Corynebacterium silvaticum | 1.83 | 1 | 0
|  |  | CVUAS 6455 | 3; sequenced, strain of the species desciption [0127] | DT | 2013-01-08 | Corynebacterium silvaticum | 1.65 | 1 | 0
|  |  | CVUAS 6678,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium silvaticum | 1.91 | 1 | 0
|  |  | CVUAS 704,2 | 2; FT-IR (NTTB-D) | eDT | 09.12.2013 | Corynebacterium silvaticum | 1.69 | 1 | 0
|  |  | CVUAS 8972,3 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | CVUAS 97,2 | 2; FT-IR (NTTB-D) | eDT | 27.11.2013 | Corynebacterium silvaticum | 1.7 | 1 | 0
|  |  | S-1014-16 | 3; FT-IR (NTTB-D) | eDT | 2017-01-11 | Corynebacterium silvaticum | 1.79 | 1 | 0
|  |  | S-1111-3-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium silvaticum | 1.88 | 1 | 0
|  |  | S-1408-1-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium silvaticum | 1.92 | 1 | 0
|  |  | S-1449-2-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium silvaticum | 1.85 | 1 | 0
|  |  | S-1514-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium silvaticum | 1.92 | 1 | 0
|  |  | S-1591-2-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium silvaticum | 1.91 | 1 | 0
|  |  | S-1627-5-12 | 3; sequenced | DT | 2012-10-08 | Corynebacterium silvaticum | 1.76 | 1 | 0
|  |  | S-241-6-18 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium silvaticum | 1.74 | 1 | 0
|  |  | S-241-7-18 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium silvaticum | 1.89 | 1 | 0
|  |  | S-28-3-13 | 3; sequenced | DT | 2013-01-18 | Corynebacterium silvaticum | 1.94 | 1 | 0
|  |  | S-308-2-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium silvaticum | 1.97 | 1 | 0
|  |  | S-309-1-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium silvaticum | 1.91 | 1 | 0
|  |  | S-310-2-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium silvaticum | 1.88 | 1 | 0
|  |  | S-349-18 | 3; FT-IR (NTTB-D) | eDT | 2018-04-05 | Corynebacterium silvaticum | 1.41 | 1 | 0
|  |  | S-349-3-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium silvaticum | 1.71 | 1 | 0
|  |  | S-427-3-15 | 3; FT-IR (NTTB-D) | eDT | 2015-03-27 | Corynebacterium silvaticum | 1.64 | 1 | 0
|  |  | S-434-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium silvaticum | 1.8 | 1 | 0
|  |  | S-443-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium silvaticum | 1.92 | 1 | 0
|  |  | S-450-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium silvaticum | 1.76 | 1 | 0
|  |  | S-468-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium silvaticum | 1.61 | 1 | 0
|  |  | S-619-1-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium silvaticum | 1.54 | 1 | 0
|  |  | S-793-3-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium silvaticum | 1.78 | 1 | 0
|  |  | S-803-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium silvaticum | 1.82 | 1 | 0
|  |  | S-884-18 | 3; FT-IR (NTTB-D) | eDT | 2018-08-21 | Corynebacterium silvaticum | 1.94 | 1 | 0
|  |  | W30-B05227 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium silvaticum | 1.8 | 1 | 0
|  |  | W67-B16966 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium silvaticum | 1.66 | 1 | 0
|  |  | C57 UGhent | 3; sequenced (rpoB) UGent | DT | 2020-06-12 | Corynebacterium ulcerans | 2.25 | 0 | 1
|  |  | C133 UGhent | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.1 | 0 | 1
|  |  | 141005571-1 | 1; 0 | 0 | 2014-05-09 | Corynebacterium ulcerans | 1.93 | 1 | 0
|  |  | 141013039 2 | 1; 0 | 0 | 2014-05-09 | Corynebacterium ulcerans | 2.03 | 0 | 1
|  |  | 151004530 | 1; 0 | 0 | 2015-04-15 | Corynebacterium ulcerans | 2.11 | 0 | 1
|  |  | 151012433-2 LHL | 1; 0 | 0 | 2015-09-28 | Corynebacterium ulcerans | 2.02 | 0 | 1
|  |  | 161007221 | 1; 0 | 0 | 2016-06-13 | Corynebacterium ulcerans | 2.09 | 0 | 1
|  |  | CIP 112226 eDT | 3; type strain | eDT | 2024-08-13 | Corynebacterium ramonii | 1.88 | 1 | 0
|  |  | CIP 112226 EFex | 3; type strain | EFEx | 2024-08-13 | Corynebacterium ramonii | 2.22 | 0 | 1
|  |  | 131010012-1 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 2.09 | 0 | 1
|  |  | 131010012-2 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 2.07 | 0 | 1
|  |  | 131011719 | 3; sequenced | eDT | 2018-03-29 | Corynebacterium ulcerans | 2 | 0 | 1
|  |  | 131013415-1 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 1.89 | 1 | 0
|  |  | 131013415-2 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 2.06 | 0 | 1
|  |  | 131015432-10 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 2.1 | 0 | 1
|  |  | 131015432-13 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 1.69 | 1 | 0
|  |  | 131015432-14 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 1.97 | 1 | 0
|  |  | 141001018 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium ulcerans | 2.03 | 0 | 1
|  |  | 141001548 | 3; sequenced | eDT | 2018-03-29 | Corynebacterium ulcerans | 1.88 | 1 | 0
|  |  | 171008197 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium ulcerans | 1.94 | 1 | 0
|  |  | C15 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.84 | 1 | 0
|  |  | C44 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.97 | 1 | 0
|  |  | C55 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.02 | 0 | 1
|  |  | C64 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.19 | 0 | 1
|  |  | C78 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.8 | 1 | 0
|  |  | C86 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.91 | 1 | 0
|  |  | CCM 2823 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 2.02 | 0 | 1
|  |  | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium ulcerans | 2.03 | 0 | 1
|  |  | CVUAS 34267 | 3; KL Diphteria | eDT | 2022-09-02 | Corynebacterium ulcerans | 2.1 | 0 | 1
|  |  | DSM 46325 CVUAS | 3; type strain | DT | 2013-01-09 | Corynebacterium ulcerans | 1.96 | 1 | 0
|  |  | KL1017 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.16 | 0 | 1
|  |  | KL199 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium ulcerans | 2.01 | 0 | 1
|  |  | KL853 | 3; MLST | eDT | 2016-10-27 | Corynebacterium ulcerans | 2.22 | 0 | 1
|  |  | KL941 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.19 | 0 | 1
|  |  | VB921348 | 2; 0 | eDT | 2015-01-30 | Corynebacterium ulcerans | 1.78 | 1 | 0
|  |  | C156 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.02 | 0 | 1
|  |  | C164 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.16 | 0 | 1
|  |  | C166 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2 | 0 | 1
|  |  | C169 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.04 | 0 | 1
|  |  | C174 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.04 | 0 | 1
|  |  | C25 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.94 | 1 | 0
|  |  | C92 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.96 | 1 | 0
|  |  | C95 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.07 | 0 | 1
|  |  | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium ulcerans | 2.05 | 0 | 1
|  |  | JB 362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium ulcerans | 1.95 | 1 | 0
|  |  | KL1015 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.1 | 0 | 1
|  |  | KL1025 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.08 | 0 | 1
|  |  | KL188 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium ulcerans | 1.9 | 1 | 0
|  |  | KL195 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium ulcerans | 1.92 | 1 | 0
|  |  | KL818 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.09 | 0 | 1
|  |  | KL867 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.1 | 0 | 1
|  |  | KL870 | 3; 0 | eDT | 0 | Corynebacterium ulcerans | 2.15 | 0 | 1
|  |  | KL885 | 3; FT-IR no NTTB-D, MLST | eDT | 0 | Corynebacterium ulcerans | 1.98 | 1 | 0
|  |  | 151012433-2 CVUAS | 3; KL Diphtheria | eDT | 2021-06-22 | Corynebacterium ulcerans | 1.98 | 1 | 0
|  |  | 181011209 001 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium ulcerans | 2.04 | 0 | 1
|  |  | 181011209 002 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium ulcerans | 1.98 | 1 | 0
|  |  | C1 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.06 | 0 | 1
|  |  | C104 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.02 | 0 | 1
|  |  | C11 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.11 | 0 | 1
|  |  | C112 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.08 | 0 | 1
|  |  | C12 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.02 | 0 | 1
|  |  | C123 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.13 | 0 | 1
|  |  | C133 CVUAS | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 1.85 | 1 | 0
|  |  | C139 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2 | 0 | 1
|  |  | C152 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.04 | 0 | 1
|  |  | C167 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.08 | 0 | 1
|  |  | C17 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 1.87 | 1 | 0
|  |  | C60 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.21 | 0 | 1
|  |  | C61 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 1.99 | 1 | 0
|  |  | C63 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.17 | 0 | 1
|  |  | C65 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.13 | 0 | 1
|  |  | C70 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.16 | 0 | 1
|  |  | C71 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.13 | 0 | 1
|  |  | C72 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.1 | 0 | 1
|  |  | C73 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.22 | 0 | 1
|  |  | C76 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.1 | 0 | 1
|  |  | C80 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.06 | 0 | 1
|  |  | C81 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.04 | 0 | 1
|  |  | C82 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.11 | 0 | 1
|  |  | C83 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.08 | 0 | 1
|  |  | C90 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.14 | 0 | 1
|  |  | C93 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium ulcerans | 2.09 | 0 | 1
|  |  | CCM 1750 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 2.04 | 0 | 1
|  |  | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 2 | 0 | 1
|  |  | CVUAS 10306 | 3; sequenced | EFex | 2018-04-24 | Corynebacterium ulcerans | 1.97 | 1 | 0
|  |  | CVUAS 33950 | 3; KL Diphteria | eDT | 2022-06-13 | Corynebacterium ulcerans | 1.96 | 1 | 0
|  |  | CVUAS 34094 | 3; KL Diphteria | eDT | 2022-08-01 | Corynebacterium ulcerans | 2 | 0 | 1
|  |  | CVUAS 6522,4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium ulcerans | 1.98 | 1 | 0
|  |  | JB360523 | 3; 0 | eDT | 2013-01-30 | Corynebacterium ulcerans | 1.87 | 1 | 0
|  |  | KL190 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium ulcerans | 1.97 | 1 | 0
|  |  | KL246 | 3; WGS data LGL [0085]; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium ulcerans | 2.07 | 0 | 1
|  |  | KL796 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.13 | 0 | 1
|  |  | KL825 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.09 | 0 | 1
|  |  | KL832 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2.17 | 0 | 1
|  |  | KL840 | 3; MLST | eDT | 0 | Corynebacterium ulcerans | 2 | 0 | 1
|  |  | WFL 1217-18 | 3; 0 | eDT | 2018-04-17 | Corynebacterium ulcerans | 1.94 | 1 | 0
|  |  | C117 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.03 | 0 | 1
|  |  | C146 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.14 | 0 | 1
|  |  | C162 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.06 | 0 | 1
|  |  | C18 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.04 | 0 | 1
|  |  | C20 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.96 | 1 | 0
|  |  | C28 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.01 | 0 | 1
|  |  | C31 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.05 | 0 | 1
|  |  | C32 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.88 | 1 | 0
|  |  | C39 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.14 | 0 | 1
|  |  | C41 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.03 | 0 | 1
|  |  | C51 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.89 | 1 | 0
|  |  | C52 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.05 | 0 | 1
|  |  | C53 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.01 | 0 | 1
|  |  | C57 CVUAS | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.99 | 1 | 0
|  |  | C67 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 1.94 | 1 | 0
|  |  | C98 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium ulcerans | 2.02 | 0 | 1
|  |  | DSM 46325 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 27.05.2022; I8 | Corynebacterium ulcerans | 2.24 | 0 | 1
|  |  | S-477-6-21 | 3; KL Diphtheria | eDT | 2021-04-21 | Corynebacterium ulcerans | 1.92 | 1 | 0
|  |  | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium amycolatum | 1.1 | 1 | 0
|  |  | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2014 E8 | Corynebacterium amycolatum | 0.98 | 1 | 0
|  |  | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium aurimucosum | 0.93 | 1 | 0
|  |  | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium belfantii | 1.53 | 1 | 0
|  |  | KL0171 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium belfantii | 1.21 | 1 | 0
|  |  | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium belfantii | 1.35 | 1 | 0
|  |  | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium bovis | 1.07 | 1 | 0
|  |  | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium camporealensis | 0.93 | 1 | 0
|  |  | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium casei | 0.84 | 1 | 0
|  |  | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium confusum | 0.71 | 1 | 0
|  |  | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium confusum | 0.45 | 1 | 0
|  |  | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium cystitidis | 0.87 | 1 | 0
|  |  | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium frankenforstense | 0.56 | 1 | 0
|  |  | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium frankenforstense | 0.19 | 1 | 0
|  |  | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium freiburgense | 1 | 1 | 0
|  |  | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium hansenii | 1.16 | 1 | 0
|  |  | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium hansenii | 1.24 | 1 | 0
|  |  | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium humireducens | 0.78 | 1 | 0
|  |  | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium kutscheri | 0.76 | 1 | 0
|  |  | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium lactis | 0.8 | 1 | 0
|  |  | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium lactis | 0.72 | 1 | 0
|  |  | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium phoceense | 1.11 | 1 | 0
|  |  | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium pilosum | 0.92 | 1 | 0
|  |  | C138 | 3; sequenced (rpoB) UGent | EFEx | 2020-07-30 | Corynebacterium rouxii | 1.09 | 1 | 0
|  |  | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium rouxii | 1.11 | 1 | 0
|  |  | CVUAS 3559,2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium rouxii | 1.25 | 1 | 0
|  |  | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium rouxii | 1.3 | 1 | 0
|  |  | KL1306 | 3; sequenced / variant upgraded to species level Coynebacterium rouxii [0137] here: NGS data (LGA) C. rouxi | EFEx | 2020-10-20 | Corynebacterium rouxii | 1.27 | 1 | 0
|  |  | KL1663 | 3; sequenced / variant upgraded to species level Coynebacterium rouxii [0137] here: NGS data (LGA) C. rouxi | EFEx | 2020-12-15 | Corynebacterium rouxii | 1.42 | 1 | 0
|  |  | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium sp-208-U2-6 | 0.66 | 1 | 0
|  |  | CVUAS 32165,2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium sp-2300500 | 0.6 | 1 | 0
|  |  | CVUAS 5775,2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium sp-2300500 | 0.96 | 1 | 0
|  |  | CVUAS 2554,2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium sp-2396 | 0.81 | 1 | 0
|  |  | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium sp-243-z-3 | 0.79 | 1 | 0
|  |  | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium sp-257-y-4 | 0.76 | 1 | 0
|  |  | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium sp-284z-2 | 0.88 | 1 | 0
|  |  | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium sp-CVUAS-11664 | 0.63 | 1 | 0
|  |  | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium sp-CVUAS-1888 | 0.75 | 1 | 0
|  |  | CVUAS 1721,3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium sp-CVUAS-2484.4 | 1.16 | 1 | 0
|  |  | CVUAS 2484,4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388; 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium sp-CVUAS-2484.4 | 0.46 | 1 | 0
|  |  | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium sp-CVUAS-30404 | 0.31 | 1 | 0
|  |  | LALLF 20TRD0812 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-CVUAS-30404 | 0.8 | 1 | 0
|  |  | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-CVUAS-30404 | 0.79 | 1 | 0
|  |  | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium sp-CVUAS-31533 | 0.93 | 1 | 0
|  |  | CVUAS 32278,2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium sp-CVUAS-32278.2 | 0.55 | 1 | 0
|  |  | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium sp-CVUAS-874,4 | 1.07 | 1 | 0
|  |  | CVUAS 874,4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium sp-CVUAS-874,4 | 0.74 | 1 | 0
|  |  | LALLF 21TRB0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-LALLF-21TRB0553 | 1.21 | 1 | 0
|  |  | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium sp-LHL-0497 | 0.29 | 1 | 0
|  |  | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium sp-LHL-0558 | 0.74 | 1 | 0
|  |  | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium sp-LHL-1322 | 0.78 | 1 | 0
|  |  | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium stationis | 0.82 | 1 | 0
|  |  | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium testudinoris | 0.34 | 1 | 0
|  |  | CVUAS 5437,3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium variabile | 0.41 | 1 | 0
|  |  | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium xerosis | 0.94 | 1 | 0
|  |  | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium xerosis-group | 1.23 | 1 | 0
|  |  | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces bovis | 0.2 | 1 | 0
|  |  | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces denticolens | 0.39 | 1 | 0
|  |  | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces sp-CVUAS-31303 | 0.58 | 1 | 0
|  |  | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces viscosus | 0.6 | 1 | 0
|  |  | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces weissii | 0.37 | 1 | 0
|  |  | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium bovis | 0.83 | 1 | 0
|  |  | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium canis | 0.92 | 1 | 0
|  |  | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium haemolyticum | 0.74 | 1 | 0
|  |  | 17-6203 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium hippocoleae | 1.05 | 1 | 0
|  |  | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium hippocoleae | 1.01 | 1 | 0
|  |  | DSM 10002 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium phocae | 0.66 | 1 | 0
|  |  | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium phocae | 0.76 | 1 | 0
|  |  | 2698T | 3; type strain | DT | 2013-08-07 | Arcanobacterium phocisimile | 0.88 | 1 | 0
|  |  | 014418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium pinnipediorum | 1.02 | 1 | 0
|  |  | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium pinnipediorum | 0.48 | 1 | 0
|  |  | DSM 13483 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium pluranimalium | 0.48 | 1 | 0
|  |  | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 0.25 | 1 | 0
|  |  | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 0.64 | 1 | 0
|  |  | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium sp-2701 | 0.7 | 1 | 0
|  |  | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium sp-2701 | 0.63 | 1 | 0
|  |  | CVUAS 1060,2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium wilhelmae | 0.58 | 1 | 0
|  |  | CVUAS 647 | 3; type strain | DT | 2014-09-04 | Arcanobacterium wilhelmae | 0.52 | 1 | 0
|  |  | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia sp-CVUAS-30561 | 0.38 | 1 | 0
|  |  | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia canis | 0.36 | 1 | 0
|  |  | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia cardiffensis | 0 | 1 | 0
|  |  | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia hyovaginalis | 0.98 | 1 | 0
|  |  | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia hyovaginalis | 0.51 | 1 | 0
|  |  | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia sp-CVUAS-32183 | 0.52 | 1 | 0
|  |  | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia sp-CVUAS-8688 | 0.99 | 1 | 0
|  |  | CVUAS 31845,2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia sp-CVUAS-8688 | 0.21 | 1 | 0
|  |  | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella abortisuis | 0.59 | 1 | 0
|  |  | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella abortisuis | 0.54 | 1 | 0
|  |  | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella bernardiae | 0.86 | 1 | 0
|  |  | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella bernardiae | 0.7 | 1 | 0
|  |  | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella bialowiezensis | 0.95 | 1 | 0
|  |  | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella bonasi | 0.43 | 1 | 0
|  |  | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.43 | 1 | 0
|  |  | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.85 | 1 | 0
|  |  | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella pyogenes | 0.68 | 1 | 0
|  |  | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella pyogenes | 0.61 | 1 | 0
|  |  | CVUAS 2379,2 | 3; sequenced | EFex | 2021-04-08 | Gordonia paraffinivorans | 0.05 | 1 | 0
|  |  | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia paraffinivorans | 0.43 | 1 | 0
|  |  | CVUAS 6272 | 2; API Coryne profile 0150006 | EFex | 2021-03-25 | Gordonia paraffinivorans | 0.88 | 1 | 0
|  |  | CVUAS 7903,3 EFex | 3; sequenced | EFex | 2017-07-27 | Gordonia paraffinivorans | 0.51 | 1 | 0
|  |  | CVUAS 7903,3 eDT | 3; sequenced | eDT | 2017-07-27 | Gordonia paraffinivorans | 0.81 | 1 | 0
|  |  | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia rubripertincta | 1.28 | 1 | 0
|  |  | CVUAS 10185.3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium smegmatis | 0.75 | 1 | 0
|  |  | CVUAS 31296,3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia cyriacigeorgica | 0.54 | 1 | 0
|  |  | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia farcinica | 0.56 | 1 | 0
|  |  | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus equi | 0.17 | 1 | 0
|  |  | CVUAS 2755,3 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 0.33 | 1 | 0
|  |  | CVUAS 2998 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 0.15 | 1 | 0
|  |  | CVUAS 33761 sir | 2; 0 | EFEx | 2022-04-27 | Rhodococcus equi | 0.28 | 1 | 0
|  |  | CVUAS 33761 EFex | 2; 0 | EFEx | 2022-04-27 | Rhodococcus equi | 0.63 | 1 | 0
|  |  | CVUAS 4057 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 0.54 | 1 | 0
|  |  | CVUAS 5384,2 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 0.21 | 1 | 0
|  |  | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 0.39 | 1 | 0
|  |  | CVUAS 1447,3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella sp | 0.52 | 1 | 0
|  |  | CVUAS 2895,3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium luteolum | 0.29 | 1 | 0
|  |  | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas flavigena | 0.2 | 1 | 0
|  |  | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium sp-234z-2 | 0.45 | 1 | 0
|  |  | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium sp-CVUAS-30560 | 0.69 | 1 | 0
|  |  | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus sedentarius | 0.7 | 1 | 0
|  |  | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus congolensis | 0.77 | 1 | 0
|  |  | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus congolensis | 0.7 | 1 | 0
|  |  | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium flaccumfaciens | 0 | 1 | 0
|  |  | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia sp-MUVA-A-508948 | 0.88 | 1 | 0
|  |  | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium sp | 0.65 | 1 | 0
|  |  | IMET 10353 DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.09.2011 L8 | Glutamicibacter nicotianae | 0.85 | 1 | 0
|  |  | IMET 10432 DSM 20127 | 3; type strain | unknown | 08.09.2011, L13, [0034] | Sinomonas atrocyanea | 0.85 | 1 | 0
|  |  | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter protophormiae | 0.37 | 1 | 0
|  |  | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria atrinae | 0.28 | 1 | 0
|  |  | 56a-5 | 3; sequenced | EFEx | 2020-05-20 | Kocuria gwangalensis | 0.36 | 1 | 0
|  |  | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2014 E22 | Kocuria palustris | 0.5 | 1 | 0
|  |  | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015 K12 | Micrococcus luteus | 0.67 | 1 | 0
|  |  | DSM 20032 | 3; type strain | eDT | 2015-07-09 | Rothia kristinae | 1.02 | 1 | 0
|  |  | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium cellulans | 0.66 | 1 | 0
|  |  | CVUAS 749,2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter inulinus | 0.84 | 1 | 0
|  |  | DSM 108415 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Cutibacterium acnes DSM 108415: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Cutibacterium acnes | 0.21 | 1 | 0
|  |  | DSM 1079 | 3; public strain collection | EFEx | 2017-01-20 | Penicillium roqueforti | 0.55 | 1 | 0
|  |  | CVUAS 10354,2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton verrucosum | 0 | 1 | 0
|  |  | CVUAS 10357.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton verrucosum | 0.66 | 1 | 0
|  |  | DSM 1386 | 3; public strain collection | Fextr | 2016-03-10 | Candida albicans | 0.49 | 1 | 0
|  |  | IGV 011 IGV | 3; sequenced | EFEx | 2016-12-08 | Sarocladium strictum | 0.65 | 1 | 0
|  |  | CVUAS 2520.4 | 3; sequenced (ITS) | F-Extr | 2018-07-26 | Fusarium sambucinum | 0.1 | 1 | 0
|  |  | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus amylolyticus | 0.56 | 1 | 0
|  |  | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus amylolyticus | 0.87 | 1 | 0
|  |  | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus barengoltzii | 0.76 | 1 | 0
|  |  | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus barengoltzii | 0.67 | 1 | 0
|  |  | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus glucanolyticus | 0.58 | 1 | 0
|  |  | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus glucanolyticus | 0.17 | 1 | 0
|  |  | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus larvae | 0.09 | 1 | 0
|  |  | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus larvae | 0.54 | 1 | 0
|  |  | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 0 | 1 | 0
|  |  | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 0.34 | 1 | 0
|  |  | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus macerans | 0.74 | 1 | 0
|  |  | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus pabuli | 0.71 | 1 | 0
|  |  | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus pabuli | 0.54 | 1 | 0
|  |  | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus pasadenensis | 0.2 | 1 | 0
|  |  | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus validus | 0.67 | 1 | 0
|  |  | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium barati | 0.58 | 1 | 0
|  |  | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium butyricum | 0.33 | 1 | 0
|  |  | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium chauvoei | 0.08 | 1 | 0
|  |  | 171002729_1 coch | 2; 0 | DT | 2017-02-28 | Clostridium cochlearium | 0.36 | 1 | 0
|  |  | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium colicanis | 0.56 | 1 | 0
|  |  | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium novyi | 0.68 | 1 | 0
|  |  | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium septicum | 0.15 | 1 | 0
|  |  | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium sporogenes | 0.76 | 1 | 0
|  |  | 141015223 aerob | 2; 0 | DT | 2014-11-27 | Clostridium tertium | 0.06 | 1 | 0
|  |  | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium tetani | 0.23 | 1 | 0
|  |  | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya limosa | 0.37 | 1 | 0
|  |  | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium sordellii | 0.59 | 1 | 0
|  |  | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium bifermentans | 0.6 | 1 | 0
|  |  | B 44 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N11 | Bacillus badius | 0.53 | 1 | 0
|  |  | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.03.2017 I18 | Bacillus ectoiniformans | 0.35 | 1 | 0
|  |  | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F22 | Cytobacillus firmus | 0.42 | 1 | 0
|  |  | CCM 37 B 88 | 3; reference strain collection | 0 | 09.02.2015, L16 | Cytobacillus firmus | 0.66 | 1 | 0
|  |  | B 46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N14 | Niallia circulans | 1.02 | 1 | 0
|  |  | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 N5 | Virgibacillus pantothenticus | 1.03 | 1 | 0
|  |  | CVUAS 31071,2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus clausii | 0.17 | 1 | 0
|  |  | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015; H10 | Alkalihalobacillus clausii | 0.4 | 1 | 0
|  |  | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 13.02.2015, K10, [0034] | Bacillus alveayuensis | 0.27 | 1 | 0
|  |  | DSM 7 RKI 2015 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCA/TA2 (A) | 16.02.2015, K16, [0034] | Bacillus amyloliquefaciens | 0.38 | 1 | 0
|  |  | DSM 7 CVUAS | 3; type strain | DT | 2014-01-29 | Bacillus amyloliquefaciens | 0.1 | 1 | 0
|  |  | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E06 | Bacillus anthracis | 0.39 | 1 | 0
|  |  | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; E04 | Bacillus anthracis | 0.62 | 1 | 0
|  |  | B 004 | 2; 0 | DT | 2014-01-22 | Bacillus atrophaeus | 0.12 | 1 | 0
|  |  | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 L8 | Bacillus atrophaeus | 0.36 | 1 | 0
|  |  | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H20 | Bacillus badius | 0.66 | 1 | 0
|  |  | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus cereus | 0.62 | 1 | 0
|  |  | CVUAS 1107 | 2; 0 | eDT | 27.01.2016 | Bacillus cereus | 0.24 | 1 | 0
|  |  | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.05.2015; E09 | Bacillus tropicus | 0.82 | 1 | 0
|  |  | CVUAS 2833 CVUAS | 4; 16S, draft genome [0009] | eDT | 2016-01-27 | Bacillus cytotoxicus | 0.29 | 1 | 0
|  |  | NVH 391-98 | 3; type strain | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.51 | 1 | 0
|  |  | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.11.2017 F14 | Bacillus halotolerans | 0.26 | 1 | 0
|  |  | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.12.2017 I8 | Bacillus halotolerans | 0.06 | 1 | 0
|  |  | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.12.2016 [H3] | Bacillus inaquosorum | 0.32 | 1 | 0
|  |  | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.02.2016 C20 | Bacillus infantis | 0.53 | 1 | 0
|  |  | B084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus licheniformis | 0.34 | 1 | 0
|  |  | DSM 13 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus licheniformis | 0.29 | 1 | 0
|  |  | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K15 | Bacillus mojavensis | 0.39 | 1 | 0
|  |  | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F4 | Bacillus mojavensis | 0.45 | 1 | 0
|  |  | DSM 2048 CVUAS | 3; type strain | eDT | 09.12.2015 | Bacillus mycoides | 0.28 | 1 | 0
|  |  | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus mycoides | 0.5 | 1 | 0
|  |  | ATCC 11778 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E8, [0034] | Bacillus pacificus | 0.89 | 1 | 0
|  |  | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus pseudomycoides | 0.54 | 1 | 0
|  |  | DSM 12443 CVUAS | 3; public strain collection | eDT | 09.12.2015 | Bacillus pseudomycoides | 0.49 | 1 | 0
|  |  | B017 CVUAS | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus pumilus | 0.45 | 1 | 0
|  |  | CCM 340 | 3; reference strain collection | DT | 2014-01-22 | Bacillus pumilus | 0.51 | 1 | 0
|  |  | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.05.2018, N12 | Bacillus siamensis | 0.32 | 1 | 0
|  |  | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2016, K18 | Bacillus sonorensis | 0.08 | 1 | 0
|  |  | DSM 12369 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Bacillus sp-DSM-12369 | 0.38 | 1 | 0
|  |  | ATCC 6633 CVUAS | 3; public strain collection | eDT | 2015-07-22 | Bacillus spizizenii | 0.55 | 1 | 0
|  |  | DSM 618 | 4; public strain collection | eDT | 27.01.2016 | Bacillus spizizenii | 0.73 | 1 | 0
|  |  | CCM 110 B068 | 3; reference strain collection | DT | 2014-01-22 | Bacillus subtilis | 0.08 | 1 | 0
|  |  | DSM 1092 | 3; reference strain collection | DT | 2014-01-22 | Bacillus subtilis | 0.43 | 1 | 0
|  |  | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.21 | 1 | 0
|  |  | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.57 | 1 | 0
|  |  | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus thuringiensis | 0.46 | 1 | 0
|  |  | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.48 | 1 | 0
|  |  | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.57 | 1 | 0
|  |  | DSM 6890 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2015, L4 | Bacillus thuringiensis | 0.67 | 1 | 0
|  |  | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 07.11.2006 K4 | Bacillus thuringiensis | 0.52 | 1 | 0
|  |  | CVUAS 2786,2 | 3; sequenced | DT | 2013-07-31 | Bacillus timorensis | 0.6 | 1 | 0
|  |  | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.07.2017, H17 | Bacillus vallismortis | 0.46 | 1 | 0
|  |  | CCEB 627 DSM 12 | 3; type strain | DT | 2014-01-22 | Cytobacillus firmus | 0.73 | 1 | 0
|  |  | CCM 37 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Cytobacillus firmus | 0.43 | 1 | 0
|  |  | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.02.2016 K19 | Cytobacillus horneckiae | 0.92 | 1 | 0
|  |  | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.04.2014 H5 | Cytobacillus horneckiae | 0.65 | 1 | 0
|  |  | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus kochii | 0.52 | 1 | 0
|  |  | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 19.02.2015, E18 | Domibacillus aminovorans | 0.45 | 1 | 0
|  |  | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 15.02.2015, E8 | Domibacillus aminovorans | 0.01 | 1 | 0
|  |  | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.02.2015 H5 | Evansella clarkii | 0.48 | 1 | 0
|  |  | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F7 | Evansella clarkii | 0.4 | 1 | 0
|  |  | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus stearothermophilus | 0.34 | 1 | 0
|  |  | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 I10; (single sp is included in the corresponding msp) | Halalkalihalobacillus halodurans | 0.85 | 1 | 0
|  |  | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 E19; (single sp is included in the corresponding msp) | Halalkalihalobacillus halodurans | 0.33 | 1 | 0
|  |  | B041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia sporothermodurans | 0.61 | 1 | 0
|  |  | DSM 10599 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015, H9 | Heyndrickxia sporothermodurans | 0.61 | 1 | 0
|  |  | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017; N12 | Lederbergia lenta | 0.31 | 1 | 0
|  |  | DSM 9 CVUAS | 3; type strain | DT | 2014-01-22 | Lederbergia lenta | 0.5 | 1 | 0
|  |  | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus bataviensis | 0.16 | 1 | 0
|  |  | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017 N13 | Neobacillus cucumis | 0.21 | 1 | 0
|  |  | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E14, [0034] | Niallia circulans | 0.57 | 1 | 0
|  |  | CVUAS 2582,2 | 3; sequenced | eDT | 2015-09-23 | Niallia nealsonii | 0.83 | 1 | 0
|  |  | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K20 | Peribacillus muralis | 0.5 | 1 | 0
|  |  | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016; K20 | Peribacillus psychrosaccharolyticus | 0.49 | 1 | 0
|  |  | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus simplex | 0.32 | 1 | 0
|  |  | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2016 H13 | Peribacillus simplex | 0.74 | 1 | 0
|  |  | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 15.12.2016, K15, [0096] | Peribacillus simplex | 0.74 | 1 | 0
|  |  | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016 K15 | Priestia flexa | 0.72 | 1 | 0
|  |  | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H18 | Priestia flexa | 0.56 | 1 | 0
|  |  | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015 H21 | Priestia megaterium | 0.85 | 1 | 0
|  |  | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia megaterium | 0.38 | 1 | 0
|  |  | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E3 | Salibacillus salarius | 0.46 | 1 | 0
|  |  | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F13 | Sutcliffiella cohnii | 0.55 | 1 | 0
|  |  | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 K8 | Sutcliffiella cohnii | 0.5 | 1 | 0
|  |  | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus saccharophilus | 0.4 | 1 | 0
|  |  | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus proomii | 0.83 | 1 | 0
|  |  | CVUAS 3616,2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus proomii | 0.28 | 1 | 0
|  |  | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H15 | Weizmannia coagulans | 0.29 | 1 | 0
|  |  | CVUAS 71,2 | 3; sequenced | DT | 2013-12-13 | Weizmannia ginsengihumi | 0.57 | 1 | 0
|  |  | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.65 | 1 | 0
|  |  | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.11 | 1 | 0
|  |  | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia zopfii | 0.26 | 1 | 0
|  |  | B092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus fusiformis | 0.52 | 1 | 0
|  |  | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus fusiformis | 1.04 | 1 | 0
|  |  | CVUAS 7192,2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus macroides | 0.55 | 1 | 0
|  |  | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus sphaericus | 0.26 | 1 | 0
|  |  | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus sphaericus | 0.77 | 1 | 0
|  |  | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus saudimassiliensis | 0.28 | 1 | 0
|  |  | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus odysseyi | 0.23 | 1 | 0
|  |  | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium acetylicum | 0 | 1 | 0
|  |  | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium enclense | 0.24 | 1 | 0
|  |  | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium profundum | 0.46 | 1 | 0
|  |  | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix thermosphacta | 0.63 | 1 | 0
|  |  | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix thermosphacta | 0.46 | 1 | 0
|  |  | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria aquatica | 0.44 | 1 | 0
|  |  | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria aquatica | 0.17 | 1 | 0
|  |  | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria fleischmannii | 0.38 | 1 | 0
|  |  | LU2006-3 EFex | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria fleischmannii | 0.18 | 1 | 0
|  |  | CIP 105447 EFex | 3; type strain | EFEx | 2017-02-15 | Listeria grayi | 0.59 | 1 | 0
|  |  | DSM 20596 EFex | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria grayi | 0.83 | 1 | 0
|  |  | CVUAS 10132,2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria innocua | 0.65 | 1 | 0
|  |  | CVUAS 8569 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria innocua | 0.69 | 1 | 0
|  |  | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria ivanovii | 0.44 | 1 | 0
|  |  | CVUAS 32100 | 2; sequenced | EFex | 2019-12-13 | Listeria ivanovii | 0.67 | 1 | 0
|  |  | DSM 23813 EFex | 3; type strain | EFEx | 2017-02-15 | Listeria marthii | 0.4 | 1 | 0
|  |  | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.56 | 1 | 0
|  |  | CVUAS 873,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.37 | 1 | 0
|  |  | CVUAS 240,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.75 | 1 | 0
|  |  | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.46 | 1 | 0
|  |  | 10BfR L0206 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.46 | 1 | 0
|  |  | 10BfR L0066 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.71 | 1 | 0
|  |  | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.33 | 1 | 0
|  |  | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.85 | 1 | 0
|  |  | CVUAS 7366,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.44 | 1 | 0
|  |  | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria newyorkensis | 0.38 | 1 | 0
|  |  | DSM 22097 EFEx | 3; type strain | EFEx | 2017-02-15 | Listeria rocourtiae | 0.74 | 1 | 0
|  |  | CVUAS 2763 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria seeligeri | 0.08 | 1 | 0
|  |  | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria seeligeri | 0.19 | 1 | 0
|  |  | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria sp | 0.51 | 1 | 0
|  |  | CVUAS 2412,3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria welshimeri | 0.83 | 1 | 0
|  |  | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria welshimeri | 0.73 | 1 | 0
|  |  | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus laterosporus | 0.59 | 1 | 0
|  |  | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus laterosporus | 0.64 | 1 | 0
|  |  | B210 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 K19 | Paenibacillus alvei | 0.73 | 1 | 0
|  |  | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus chibensis | 1.07 | 1 | 0
|  |  | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus nuruki | 0.5 | 1 | 0
|  |  | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.09.2013 C8 | Paenibacillus polymyxa | 0.69 | 1 | 0
|  |  | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.11.2017 C18 | Paenibacillus polymyxa | 0.55 | 1 | 0
|  |  | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus sp-CVUAS-6369 | 0.63 | 1 | 0
|  |  | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus sp-E8a | 0.28 | 1 | 0
|  |  | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus taichungensis | 0.48 | 1 | 0
|  |  | FLI003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus taichungensis | 0.69 | 1 | 0
|  |  | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus tundrae | 0.59 | 1 | 0
|  |  | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus halophilus | 0.65 | 1 | 0
|  |  | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus huakuii | 0.54 | 1 | 0
|  |  | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus psychrophilus | 0.64 | 1 | 0
|  |  | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus carouselicus | 0.58 | 1 | 0
|  |  | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 14.03.2013 | Macrococcus caseolyticus | 0.19 | 1 | 0
|  |  | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 0.64 | 1 | 0
|  |  | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus lentus | 0.52 | 1 | 0
|  |  | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus sciuri | 0.41 | 1 | 0
|  |  | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus sciuri | 0.57 | 1 | 0
|  |  | CVUAS 9937,2 EFex | 3; 16S rDNA-sequencing | EFex | 2016-04-06 | Mammaliicoccus vitulinus | 0.45 | 1 | 0
|  |  | CVUAS 343,2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus vitulinus | 0.58 | 1 | 0
|  |  | CVUAS 5687 | 3; sequenced | DT | 2013-04-16 | Staphylococcus agnetis | 0.34 | 1 | 0
|  |  | TUM 7385 | 3; sequenced | DT | 02.10.2013 | Staphylococcus agnetis | 0.69 | 1 | 0
|  |  | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus argenteus | 0.42 | 1 | 0
|  |  | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus argenteus | 0.2 | 1 | 0
|  |  | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus arlettae | 0.42 | 1 | 0
|  |  | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus arlettae | 0.36 | 1 | 0
|  |  | ATCC 29213 CVUAS | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus aureus | 0.54 | 1 | 0
|  |  | CVUAS 9914,2 sir | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus aureus | 0.41 | 1 | 0
|  |  | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus auricularis | 0.23 | 1 | 0
|  |  | DSM 20326 (T) | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus capitis | 0.63 | 1 | 0
|  |  | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus capitis | 0.35 | 1 | 0
|  |  | DSM 20608 (T) | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus caprae | 0.51 | 1 | 0
|  |  | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus caprae | 0.15 | 1 | 0
|  |  | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus carnosus | 0.66 | 1 | 0
|  |  | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus carnosus | 0.64 | 1 | 0
|  |  | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus chromogenes | 0.46 | 1 | 0
|  |  | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus chromogenes | 0.42 | 1 | 0
|  |  | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus cohnii | 0.12 | 1 | 0
|  |  | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus cohnii | 0.3 | 1 | 0
|  |  | x Sta | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012 K10 | Staphylococcus delphini | 0.42 | 1 | 0
|  |  | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus delphini | 0.74 | 1 | 0
|  |  | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus devriesei | 0.78 | 1 | 0
|  |  | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 0.59 | 1 | 0
|  |  | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus epidermidis | 0.84 | 1 | 0
|  |  | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus equorum | 1.17 | 1 | 0
|  |  | DSM 20674 T | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus equorum | 0.53 | 1 | 0
|  |  | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus felis | 0.3 | 1 | 0
|  |  | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus felis | 0.71 | 1 | 0
|  |  | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.08.2011 H9 | Staphylococcus haemolyticus | 0.69 | 1 | 0
|  |  | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus haemolyticus | 0.15 | 1 | 0
|  |  | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N6 [0034] | Staphylococcus hominis | 0.94 | 1 | 0
|  |  | CVUAS 789,2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus hominis | 0.47 | 1 | 0
|  |  | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus hyicus | 0.79 | 1 | 0
|  |  | CVUAS 3767 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus hyicus | 0.8 | 1 | 0
|  |  | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, N8, [0034] | Staphylococcus intermedius | 0.56 | 1 | 0
|  |  | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N7, [0034] | Staphylococcus intermedius | 0.8 | 1 | 0
|  |  | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 05.04.2012, N5, [0034] | Staphylococcus lugdunensis | 0.56 | 1 | 0
|  |  | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, H16 | Staphylococcus lugdunensis | 0.35 | 1 | 0
|  |  | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus microti | 0.62 | 1 | 0
|  |  | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus pasteuri | 0.2 | 1 | 0
|  |  | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus pasteuri | 0.46 | 1 | 0
|  |  | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.03.2016, H05, [0034] | Staphylococcus pseudintermedius | 0.7 | 1 | 0
|  |  | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus pseudintermedius | 0.41 | 1 | 0
|  |  | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus saprophyticus | 0.22 | 1 | 0
|  |  | CVUAS 503,2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus saprophyticus | 0.54 | 1 | 0
|  |  | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus schleiferi | 0.82 | 1 | 0
|  |  | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus schweitzeri | 0.39 | 1 | 0
|  |  | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus simulans | 0.43 | 1 | 0
|  |  | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus simulans | 0.74 | 1 | 0
|  |  | CVUAS 2439,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 0.99 | 1 | 0
|  |  | CVUAS 761,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 0.45 | 1 | 0
|  |  | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus warneri | 0.46 | 1 | 0
|  |  | DSM 20316 CVUAS | 3; type strain | DT | 2013-10-02 | Staphylococcus warneri | 0.67 | 1 | 0
|  |  | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.3 | 1 | 0
|  |  | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.25 | 1 | 0
|  |  | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus sp-185-U1-2 | 0.36 | 1 | 0
|  |  | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus suis | 0.8 | 1 | 0
|  |  | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus urinaeequi | 0.48 | 1 | 0
|  |  | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 27.09.2016 | Aerococcus viridans | 0.51 | 1 | 0
|  |  | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter phocae | 0.73 | 1 | 0
|  |  | CVUAS 2442,2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium divergens | 0.31 | 1 | 0
|  |  | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium divergens | 0.15 | 1 | 0
|  |  | CVUAS 6273 | 3; sequenced | eDT | 13.01.2016 | Carnobacterium maltaromaticum | 0.27 | 1 | 0
|  |  | CVUAS 1519,3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella adiacens | 0.38 | 1 | 0
|  |  | CVUAS 2616 | 3; sequenced | DT | 25.09.2013 | Marinilactibacillus sp-CVUAS-2616 | 0.55 | 1 | 0
|  |  | CVUAS 8980,3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus sp-CVUAS-2616 | 0.51 | 1 | 0
|  |  | 16 Ec RKI | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011 H20 | Enterococcus casseliflavus | 0.76 | 1 | 0
|  |  | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H12 | Enterococcus durans | 0.56 | 1 | 0
|  |  | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H13 | Enterococcus durans | 0 | 1 | 0
|  |  | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.07.2011 N8 | Enterococcus faecium | 0.79 | 1 | 0
|  |  | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N13 | Enterococcus gallinaru | 0.53 | 1 | 0
|  |  | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N3 | Enterococcus hirae | 0.2 | 1 | 0
|  |  | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus casseliflavus | 0.81 | 1 | 0
|  |  | CVUAS 5030 | 3; sequenced | eDT | 23.07.2014 | Enterococcus cecorum | 0.79 | 1 | 0
|  |  | CVUAS 8897 | 3; sequenced | eDT | 23.07.2014 | Enterococcus cecorum | 0.72 | 1 | 0
|  |  | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus faecalis | 0.29 | 1 | 0
|  |  | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus faecalis | 0.46 | 1 | 0
|  |  | MR11-207 LVU04 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus faecium | 0.64 | 1 | 0
|  |  | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus gallinarum | 0.6 | 1 | 0
|  |  | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus solitarius | 1.05 | 1 | 0
|  |  | CVUAS 1549,2 | 3; sequenced | DT | 2016-09-27 | Vagococcus salmoninarum | 0.52 | 1 | 0
|  |  | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus acidophilus | 0.63 | 1 | 0
|  |  | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus johnsonii | 0.72 | 1 | 0
|  |  | CVUAS 32277,2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus sp-CVUAS-32277,2 | 0 | 1 | 0
|  |  | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus claussenii | 0.69 | 1 | 0
|  |  | LTH 1213 | 2; 0 | DT | 2013-07-09 | Leuconostoc mesenteroides | 0.67 | 1 | 0
|  |  | LTH 1146 | 2; 0 | DT | 2013-07-09 | Lactococcus lactis | 0.2 | 1 | 0
|  |  | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus agalactiae | 0.55 | 1 | 0
|  |  | DSM 6784 CVUAS | 3; public strain collection | eDT | 2019-07-09 | Streptococcus agalactiae | 0.22 | 1 | 0
|  |  | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus canis | 0.47 | 1 | 0
|  |  | LHL Mi-845_17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus canis | 0.15 | 1 | 0
|  |  | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus castoreus | 0.13 | 1 | 0
|  |  | DSM 17536 CVUAS | 3; type strain | EFEx | 2018-05-03 | Streptococcus castoreus | 0.04 | 1 | 0
|  |  | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus catagoni | 0.89 | 1 | 0
|  |  | 99-1-2017 | 3; type strain | EFEx | 2018-05-03 | Streptococcus catagoni | 0.52 | 1 | 0
|  |  | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus dysgalactiae | 0.52 | 1 | 0
|  |  | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0.34 | 1 | 0
|  |  | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus dysgalactiae | 0.36 | 1 | 0
|  |  | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0.53 | 1 | 0
|  |  | CVUAS 32462,2 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus equi | 0.43 | 1 | 0
|  |  | CVUAS 32770 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus equi | 0.36 | 1 | 0
|  |  | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus equi | 0.35 | 1 | 0
|  |  | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus equi | 0.36 | 1 | 0
|  |  | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus iniae | 0.48 | 1 | 0
|  |  | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus orisasini | 0.41 | 1 | 0
|  |  | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus parauberis | 0.32 | 1 | 0
|  |  | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus phocae | 0.41 | 1 | 0
|  |  | 12-13862-001-01 LVU11 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus pluranimalium | 0.62 | 1 | 0
|  |  | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.66 | 1 | 0
|  |  | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.27 | 1 | 0
|  |  | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.04 | 1 | 0
|  |  | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus pneumoniae | 0.4 | 1 | 0
|  |  | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus porcinus | 0.26 | 1 | 0
|  |  | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.04.2008, K11, [0034] | Streptococcus pyogenes | 0.52 | 1 | 0
|  |  | CVUAS 10145,3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus salivarius | 0.34 | 1 | 0
|  |  | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus sp-CVUAS-32349 | 0.45 | 1 | 0
|  |  | CVUAS 2260,4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus troglodytidis | 0.46 | 1 | 0
|  |  | CVUAS 5218,4 | 3; sequenced; 16S; NRZ Strep Uni AA | DT | 2015-06-11 | Streptococcus troglodytidis | 0.52 | 1 | 0
|  |  | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0 | 1 | 0
|  |  | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0.23 | 1 | 0
|  |  | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.55 | 1 | 0
|  |  | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.38 | 1 | 0
|  |  | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.08 | 1 | 0
|  |  | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.68 | 1 | 0
|  |  | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.47 | 1 | 0
|  |  | C bot B L677 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.57 | 1 | 0
|  |  | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.31 | 1 | 0
|  |  | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0 | 1 | 0
|  |  | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.41 | 1 | 0
|  |  | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium botulinum | 0.7 | 1 | 0
|  |  | CVUAS 32971 | 3; sequenced | EFEx | 2021-02-25 | Clostridium botulinum | 0.61 | 1 | 0
|  |  | CVUAS 33885,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium botulinum | 0.67 | 1 | 0
|  |  | CVUAS 33886,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium botulinum | 0.38 | 1 | 0
|  |  | CVUAS 33891,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium botulinum | 0.48 | 1 | 0
|  |  | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium chauvoei | 0.3 | 1 | 0
|  |  | CVUAS 849,2 | 3; sequenced | DT | 2014-12-17 | Clostridium isatidis | 0.33 | 1 | 0
|  |  | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium nigeriense | 0.38 | 1 | 0
|  |  | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium perfringens | 0.68 | 1 | 0
|  |  | B | 3; case rerort | DT | 2013-08-06 | Clostridium perfringens | 0.33 | 1 | 0
|  |  | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium septicum | 0.61 | 1 | 0
|  |  | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium sp-DSM-1985 | 0.76 | 1 | 0
|  |  | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium sporogenes | 0.54 | 1 | 0
|  |  | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium tertium | 0.31 | 1 | 0
|  |  | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium tetani | 0 | 1 | 0
|  |  | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix anatis | 0.44 | 1 | 0
|  |  | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix anatis | 0.7 | 1 | 0
|  |  | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix aquatica | 0.73 | 1 | 0
|  |  | VA-92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix aquatica | 0 | 1 | 0
|  |  | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix larvae | 0 | 1 | 0
|  |  | DSM 110099 CVUAS | 3; type strain | EFex [BRU03] | 2021-10-18 | Erysipelothrix piscisicarius | 0.42 | 1 | 0
|  |  | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix rhusiopathiae | 0.22 | 1 | 0
|  |  | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix rhusiopathiae | 0.23 | 1 | 0
|  |  | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix tonsillarum | 0.8 | 1 | 0
|  |  | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-23 | Erysipelothrix tonsillarum | 0.59 | 1 | 0
|  |  | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus sp-CVUAS-10169 | 0.11 | 1 | 0
|  |  | CVUAS 32948 | 3; confirmed proficency test | EFEx | 2021-02-18 | Finegoldia magna | 0.35 | 1 | 0
|  |  | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus kunzii | 0.7 | 1 | 0
|  |  | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus kunzii | 0.42 | 1 | 0
|  |  | CVUAS 433,2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus ovis | 0.79 | 1 | 0
|  |  | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus ovis | 0.58 | 1 | 0
|  |  | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus sp-CVUAS-32321 | 0.29 | 1 | 0
|  |  | CVUAS 692,2 | 3; sequenced | DT | 2013-02-19 | Parvimonas sp-CVUAS-692,2 | 0.47 | 1 | 0
|  |  | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides finegoldii | 0.29 | 1 | 0
|  |  | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides fragilis | 0.64 | 1 | 0
|  |  | 181001478 Bat | 2; 0 | DT | 2018-05-02 | Bacteroides pyogenes | 0.34 | 1 | 0
|  |  | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella heparinolytica | 0.27 | 1 | 0
|  |  | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium gleum | 0.41 | 1 | 0
|  |  | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium indologenes | 0.21 | 1 | 0
|  |  | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium indologenes | 0.61 | 1 | 0
|  |  | CVUAS 32947 | 3; confirmed proficency test | EFEx | 2021-02-18 | Bacteroides fragilis | 0.26 | 1 | 0
|  |  | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium hibernum | 0.83 | 1 | 0
|  |  | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium sp | 0.63 | 1 | 0
|  |  | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium chaponense | 0.22 | 1 | 0
|  |  | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium hominis | 0.2 | 1 | 0
|  |  | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium manosquense | 0.93 | 1 | 0
|  |  | 185-U1-6 EFex | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium manosquense | 0.67 | 1 | 0
|  |  | CVUAS 11450,4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium sp-I-7 | 0.64 | 1 | 0
|  |  | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium aquidurense | 0.31 | 1 | 0
|  |  | CVUAS 5322,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium branchiarum | 0.12 | 1 | 0
|  |  | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium chungangense | 0.27 | 1 | 0
|  |  | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium columnare | 1.01 | 1 | 0
|  |  | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium frigidimaris | 1.07 | 1 | 0
|  |  | CVUAS 4217,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium frigidimaris | 0.96 | 1 | 0
|  |  | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium piscis | 0.52 | 1 | 0
|  |  | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium plurextorum | 0.59 | 1 | 0
|  |  | CVUAS 5323,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium psychoterrae | 0.38 | 1 | 0
|  |  | CVUAS 11550,4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium psychrophilum | 0.86 | 1 | 0
|  |  | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium psychrophilum | 0.77 | 1 | 0
|  |  | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium reichenbachii | 1.04 | 1 | 0
|  |  | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium sp-C2 | 0.6 | 1 | 0
|  |  | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-10323 | 0.49 | 1 | 0
|  |  | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium sp-CVUAS-4146 | 0.69 | 1 | 0
|  |  | CVUAS 4216,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-4216,2 | 0.92 | 1 | 0
|  |  | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-5506 | 0.42 | 1 | 0
|  |  | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium sp-HMD2009 | 0.48 | 1 | 0
|  |  | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium sp-JQMS | 0 | 1 | 0
|  |  | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium succinicans | 0.32 | 1 | 0
|  |  | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium tiangeerense | 0.39 | 1 | 0
|  |  | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella columbina | 0 | 1 | 0
|  |  | CVUAS 2589,2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas tenax | 0.47 | 1 | 0
|  |  | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium rhinotracheale | 0.55 | 1 | 0
|  |  | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium rhinotracheale | 0.57 | 1 | 0
|  |  | CVUAS 9288,2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca zopfii | 0.54 | 1 | 0
|  |  | CVUAS 9876,2 | 2; 0 | Pextr | 2016-09-01 | Prototheca zopfii | 0.46 | 1 | 0
|  |  | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus sp-CVUAS-11266 | 0.33 | 1 | 0
|  |  | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium necrophorum | 0.27 | 1 | 0
|  |  | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium varium | 0.41 | 1 | 0
|  |  | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus moniliformis | 0.04 | 1 | 0
|  |  | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter abscessus | 0.8 | 1 | 0
|  |  | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga salmonicida | 0.63 | 1 | 0
|  |  | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus hongkongensis | 0.49 | 1 | 0
|  |  | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella termitidis | 0.29 | 1 | 0
|  |  | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella termitidis | 0.4 | 1 | 0
|  |  | CCUG 38322 CVUAS | 3; public strain collection | EFEx | 16.12.2015 | Sneathia sanguinegens | 0.83 | 1 | 0
|  |  | CCUG 41628 CVUAS | 3; type strain | EFEx | 2020-10-08 | Sneathia sanguinegens | 0.7 | 1 | 0
|  |  | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.78 | 1 | 0
|  |  | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.72 | 1 | 0
|  |  | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus felis | 0.18 | 1 | 0
|  |  | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus moniliformis | 0.19 | 1 | 0
|  |  | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus notomytis | 0 | 1 | 0
|  |  | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus ratti | 0.25 | 1 | 0
|  |  | 171012806 | 2; 0 | DT | 2017-10-04 | Ochrobactrum intermedium | 0.23 | 1 | 0
|  |  | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum intermedium | 0.62 | 1 | 0
|  |  | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum sp | 0.63 | 1 | 0
|  |  | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum sp | 0.39 | 1 | 0
|  |  | 171012040 Och | 2; 0 | DT | 2017-10-04 | Ochrobactrum sp | 0.32 | 1 | 0
|  |  | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia coli | 0.72 | 1 | 0
|  |  | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella oxytoca | 0.66 | 1 | 0
|  |  | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella variicola | 0.16 | 1 | 0
|  |  | CVUAS 1816,2 | 3; sequenced | DT | 2013-07-24 | Basfia succiniciproducens | 0.7 | 1 | 0
|  |  | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas alcaligenes | 0.44 | 1 | 0
|  |  | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas anguilliseptica | 0.74 | 1 | 0
|  |  | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas brenneri | 0 | 1 | 0
|  |  | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas fulva | 0.44 | 1 | 0
|  |  | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas koreensis | 0.13 | 1 | 0
|  |  | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas monteilii | 0.5 | 1 | 0
|  |  | 171011722-1_1 nit | 2; 0 | DT | 2017-08-31 | Pseudomonas nitroreducens | 0.35 | 1 | 0
|  |  | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas oryzihabitans | 0.27 | 1 | 0
|  |  | 171011722-1_2 put | 2; 0 | DT | 2017-08-31 | Pseudomonas putida | 0.42 | 1 | 0
|  |  | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas taetrolens | 0.61 | 1 | 0
|  |  | CVUAS 2280,3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas glaciistagni | 0.47 | 1 | 0
|  |  | A-269_08 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 02.04.2012 K8 | Ochrobactrum anthropi | 0.06 | 1 | 0
|  |  | BM373 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; D02 | Brucella abortus | 0.24 | 1 | 0
|  |  | BM369 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; E01 | Brucella abortus | 0.11 | 1 | 0
|  |  | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; I16 | Brucella canis | 0.75 | 1 | 0
|  |  | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.03.2014; E17 | Brucella canis | 0.47 | 1 | 0
|  |  | BM389 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; F04 | Brucella ceti | 0.33 | 1 | 0
|  |  | BM403 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; K04 | Brucella ceti | 0.2 | 1 | 0
|  |  | BM374 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I02 | Brucella melitensis | 0.32 | 1 | 0
|  |  | BM407 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; J05 | Brucella melitensis | 0.5 | 1 | 0
|  |  | BM408 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; H06 | Brucella melitensis | 0.17 | 1 | 0
|  |  | BM410 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I06 | Brucella melitensis | 0.3 | 1 | 0
|  |  | CCM 4915 CVUAS | 3; type strain | DT | 2013-11-26 | Brucella microti | 0.26 | 1 | 0
|  |  | CCM 4915 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.03.2015 C7 | Brucella microti | 0.77 | 1 | 0
|  |  | BM421 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; F07 | Brucella neotomae | 0.22 | 1 | 0
|  |  | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.03.2015; C15 | Brucella neotomae | 0.35 | 1 | 0
|  |  | BM423 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; L07 | Brucella ovis | 0.44 | 1 | 0
|  |  | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; F16 | Brucella ovis | 0.54 | 1 | 0
|  |  | BM405 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; E05 | Brucella pinnipedialis | 0.54 | 1 | 0
|  |  | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; C03 | Brucella pinnipedialis | 0.54 | 1 | 0
|  |  | BM431 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; B10 | Brucella suis | 0.62 | 1 | 0
|  |  | BM427 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; G08 | Brucella suis | 0.09 | 1 | 0
|  |  | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella suis | 0.66 | 1 | 0
|  |  | BM428 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I08 | Brucella suis | 0.14 | 1 | 0
|  |  | BM430 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; I09 | Brucella suis | 0.12 | 1 | 0
|  |  | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 15.15.2016 I7 | Brucella vulpis | 0.21 | 1 | 0
|  |  | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum ovis | 0.26 | 1 | 0
|  |  | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 03.06.2016 H12 | Ochrobactrum anthropi | 0.77 | 1 | 0
|  |  | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter aganoensis | 0.83 | 1 | 0
|  |  | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter aminovorans | 0.68 | 1 | 0
|  |  | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2015 H22 | Paracoccus yeei | 0.85 | 1 | 0
|  |  | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas mucosa | 0.75 | 1 | 0
|  |  | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 24.01.2007, H7, [0034] | Inquilinus limosus | 0.41 | 1 | 0
|  |  | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella asinigenitalis | 0.38 | 1 | 0
|  |  | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter denitrificans | 0.52 | 1 | 0
|  |  | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter sp-DSM-30031 | 0.83 | 1 | 0
|  |  | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2008 E5 | Achromobacter xylosoxidans | 0.3 | 1 | 0
|  |  | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.02.2016 N7 | Achromobacter xylosoxidans | 0.95 | 1 | 0
|  |  | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea sp-CVUAS-31213 | 0.86 | 1 | 0
|  |  | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella avium | 0.68 | 1 | 0
|  |  | CVUAS 32946 | 3; confirmed proficency test | EFEx | 2021-02-18 | Kerstersia gyiorum | 0.74 | 1 | 0
|  |  | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.54 | 1 | 0
|  |  | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.35 | 1 | 0
|  |  | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga kullae | 0.83 | 1 | 0
|  |  | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella asinigenitalis | 0.09 | 1 | 0
|  |  | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella equigenitalis | 0 | 1 | 0
|  |  | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella equigenitalis | 0.29 | 1 | 0
|  |  | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012; K10 | Burkholderia pseudomallei | 1.03 | 1 | 0
|  |  | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; H21 | Burkholderia pseudomallei | 0.92 | 1 | 0
|  |  | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.04.2012, K6 | Burkholderia thailandensis | 0.32 | 1 | 0
|  |  | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007, K7 | Burkholderia ambifaria | 0.38 | 1 | 0
|  |  | DSM 16087 LMG 19182 | 3; type starin | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K19 | Burkholderia ambifaria | 0.31 | 1 | 0
|  |  | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K21 | Burkholderia anthina | 0.46 | 1 | 0
|  |  | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.08.2007, N10 | Burkholderia anthina | 0.03 | 1 | 0
|  |  | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2010, K12 | Burkholderia arboris | 0.61 | 1 | 0
|  |  | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 3.4.2017; L6 | Burkholderia arboris | 0.46 | 1 | 0
|  |  | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K11 | Burkholderia cenocepacia | 0.79 | 1 | 0
|  |  | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.07.2009; K21 | Burkholderia cenocepacia | 0.75 | 1 | 0
|  |  | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 4.5.2006; H15 | Burkholderia cepacia | 0.67 | 1 | 0
|  |  | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 9.5.2007; N9 | Burkholderia cepacia | 0.65 | 1 | 0
|  |  | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.11.2006; H10 | Burkholderia contaminans | 0.61 | 1 | 0
|  |  | BC 01 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E10 | Burkholderia contaminans | 0.74 | 1 | 0
|  |  | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016 C13 | Burkholderia diffusa | 0.58 | 1 | 0
|  |  | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L17 | Burkholderia diffusa | 0.27 | 1 | 0
|  |  | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K16 | Burkholderia dolosa | 0.35 | 1 | 0
|  |  | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.11.2006; E09 | Burkholderia dolosa | 0.31 | 1 | 0
|  |  | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.01.2007; H22 | Burkholderia gladioli | 0.23 | 1 | 0
|  |  | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H15 | Burkholderia gladioli | 0.51 | 1 | 0
|  |  | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K03 | Burkholderia glumae | 0.82 | 1 | 0
|  |  | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.05.2012; K22 | Burkholderia glumae | 0.68 | 1 | 0
|  |  | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016; C10 | Burkholderia lata | 0.52 | 1 | 0
|  |  | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.08.2016; K19 | Burkholderia lata | 0.47 | 1 | 0
|  |  | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.05.2016; L06 | Burkholderia latens | 0.32 | 1 | 0
|  |  | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.09.2009; K20 | Burkholderia mallei | 0.57 | 1 | 0
|  |  | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.08.2007; H13 | Burkholderia mallei | 0.84 | 1 | 0
|  |  | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2010; H14 | Burkholderia metallica | 0.69 | 1 | 0
|  |  | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L21 | Burkholderia metallica | 0.61 | 1 | 0
|  |  | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.12.2006; K19 | Burkholderia multivorans | 0.43 | 1 | 0
|  |  | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.11.2006; K04 | Burkholderia multivorans | 0.3 | 1 | 0
|  |  | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K17 | Burkholderia oklahomensis | 0.73 | 1 | 0
|  |  | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K18 | Burkholderia oklahomensis | 0.4 | 1 | 0
|  |  | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K06 | Burkholderia plantarii | 0.26 | 1 | 0
|  |  | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K05 | Burkholderia plantarii | 0.2 | 1 | 0
|  |  | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K07 | Burkholderia pseudomultivorans | 0.27 | 1 | 0
|  |  | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K08 | Burkholderia pseudomultivorans | 0.36 | 1 | 0
|  |  | CAMPA 1043 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; K04 | Burkholderia puraquae | 0.73 | 1 | 0
|  |  | R-50214 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.03.2016; C04 | Burkholderia puraquae | 0.66 | 1 | 0
|  |  | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; N06 | Burkholderia pyrrocinia | 0.77 | 1 | 0
|  |  | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; I21 | Burkholderia pyrrocinia | 0.44 | 1 | 0
|  |  | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H08 | Burkholderia seminalis | 0.63 | 1 | 0
|  |  | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.01.2007; K12 | Burkholderia seminalis | 0.72 | 1 | 0
|  |  | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K13 | Burkholderia stabilis | 0.95 | 1 | 0
|  |  | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E08 | Burkholderia stabilis | 0.77 | 1 | 0
|  |  | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; C08 | Burkholderia stagnalis | 0.48 | 1 | 0
|  |  | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F14 | Burkholderia stagnalis | 0.77 | 1 | 0
|  |  | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007; H19 | Burkholderia taxonK | 0.74 | 1 | 0
|  |  | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.06.2009; K13 | Burkholderia taxonK | 0.61 | 1 | 0
|  |  | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F16 | Burkholderia territorii | 0.53 | 1 | 0
|  |  | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; E14 | Burkholderia territorii | 0.26 | 1 | 0
|  |  | DSM 13276 n RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 24.11.2010, K19 | Burkholderia thailandensis | 0.33 | 1 | 0
|  |  | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H11 | Burkholderia ubonensis | 0.34 | 1 | 0
|  |  | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; I04 | Burkholderia ubonensis | 0.13 | 1 | 0
|  |  | CCUG 9631=LMG 6999 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.05.2007 N15 | Burkholderia vietnamiensis | 0.44 | 1 | 0
|  |  | LMG 10929 V | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.12.2006 N9 | Burkholderia vietnamiensis | 0.7 | 1 | 0
|  |  | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.04.2012, N07 | Caballeronia glathei | 0.62 | 1 | 0
|  |  | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H12 | Caballeronia glathei | 0.86 | 1 | 0
|  |  | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H15 | Pandoraea apista | 0.16 | 1 | 0
|  |  | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H21 | Pandoraea pnomenusa | 0.75 | 1 | 0
|  |  | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H16 | Pandoraea pulmonicola | 0.88 | 1 | 0
|  |  | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K7 | Paraburkholderia caledonica | 0.63 | 1 | 0
|  |  | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K8 | Paraburkholderia caledonica | 0.61 | 1 | 0
|  |  | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N9 | Paraburkholderia caribensis | 0.76 | 1 | 0
|  |  | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N10 | Paraburkholderia caribensis | 0.82 | 1 | 0
|  |  | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H03 | Paraburkholderia hospita | 0.69 | 1 | 0
|  |  | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H04 | Paraburkholderia hospita | 0.49 | 1 | 0
|  |  | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 1.12.2010; H15 | Paraburkholderia kururiensis | 0.66 | 1 | 0
|  |  | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.11.2010; N19 | Paraburkholderia phymatum | 0.77 | 1 | 0
|  |  | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.11.2010; K17 | Paraburkholderia phytofirmans | 0.62 | 1 | 0
|  |  | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H16 | Paraburkholderia tuberum | 0.65 | 1 | 0
|  |  | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.11.2010, N21 | Paraburkholderia xenovorans | 0.66 | 1 | 0
|  |  | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 E07 | Ralstonia insidiosa | 0.5 | 1 | 0
|  |  | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 C06 | Ralstonia mannitolilytica | 0.76 | 1 | 0
|  |  | ATCC 27511 DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 22.01.2007 H03 | Ralstonia pickettii | 0.1 | 1 | 0
|  |  | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 07.05.2009 N15 | Ralstonia pickettii | 0.87 | 1 | 0
|  |  | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia acidovorans | 0.49 | 1 | 0
|  |  | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia tsuruhatensis | 0.72 | 1 | 0
|  |  | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga intermedia | 0.21 | 1 | 0
|  |  | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.07.2013 L13 | Xenophilus sp | 0.19 | 1 | 0
|  |  | DSM 24713 EFex | 3; type strain | Extr | 2016-02-02 | Deefgea chitinilytica | 0 | 1 | 0
|  |  | DSM 24713 eDT | 3; type strain | eDT | 2015-09-22 | Deefgea chitinilytica | 0.36 | 1 | 0
|  |  | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea rivuli | 0.84 | 1 | 0
|  |  | 08-F17.2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.11 | 1 | 0
|  |  | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.64 | 1 | 0
|  |  | CVUAS 6995,3 | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea sp-CVUAS-6995,3 | 0.6 | 1 | 0
|  |  | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-6995,3 | 1.01 | 1 | 0
|  |  | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-9920 | 1.05 | 1 | 0
|  |  | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea sp-CVUAS-9920 | 0.79 | 1 | 0
|  |  | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter limnosediminis | 0.19 | 1 | 0
|  |  | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella kingae | 0.97 | 1 | 0
|  |  | CVUAS 10067,2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria animaloris | 0.73 | 1 | 0
|  |  | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria animaloris | 0.56 | 1 | 0
|  |  | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-4931 | 0.31 | 1 | 0
|  |  | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-8997 | 0.44 | 1 | 0
|  |  | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-LHL-1296 | 0.4 | 1 | 0
|  |  | CVUAS 5699,3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria sp-N95-16 | 0.45 | 1 | 0
|  |  | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae sp-CVUAS-8636 | 0.41 | 1 | 0
|  |  | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter fetus | 0.26 | 1 | 0
|  |  | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter fetus | 0.62 | 1 | 0
|  |  | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter fetus | 0.53 | 1 | 0
|  |  | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter hyointestinalis | 0.12 | 1 | 0
|  |  | BfR-CA-07048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter peloridis | 0.19 | 1 | 0
|  |  | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter peloridis | 0.66 | 1 | 0
|  |  | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter cryaerophilus | 0.39 | 1 | 0
|  |  | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter butzleri | 0.45 | 1 | 0
|  |  | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter butzleri | 0.79 | 1 | 0
|  |  | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter coli | 0.32 | 1 | 0
|  |  | DSM 4689 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter coli | 0.23 | 1 | 0
|  |  | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter jejuni | 0.95 | 1 | 0
|  |  | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter jejuni | 0.71 | 1 | 0
|  |  | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter jejuni | 0.63 | 1 | 0
|  |  | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter lari | 0.44 | 1 | 0
|  |  | BfR-CA-06556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.6 | 1 | 0
|  |  | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.76 | 1 | 0
|  |  | DSM 11375 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter lari | 0.48 | 1 | 0
|  |  | DSM 5361 CVUAS | 3; type strain | DT | 2013-09-24 | Campylobacter fetus | 0.43 | 1 | 0
|  |  | CVUAS 6673,2 | 3; sequenced | DT | 2013-08-28 | Campylobacter hyointestinalis | 0.5 | 1 | 0
|  |  | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter sputorum | 0.46 | 1 | 0
|  |  | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter sp-W1 | 0.5 | 1 | 0
|  |  | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter upsaliensis | 0.35 | 1 | 0
|  |  | CVUAS 11185,3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter canadensis | 0.52 | 1 | 0
|  |  | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas bestiarum | 0.69 | 1 | 0
|  |  | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas bivalvium | 0.56 | 1 | 0
|  |  | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas caviae | 0.38 | 1 | 0
|  |  | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.1 | 1 | 0
|  |  | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.18 | 1 | 0
|  |  | DSM 30187 RKI | 4; type strain, WGS sequenced (BfR) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E17 | Aeromonas hydrophila | 0.6 | 1 | 0
|  |  | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas hydrophila | 0.69 | 1 | 0
|  |  | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas jandaei | 0.58 | 1 | 0
|  |  | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas media | 0.06 | 1 | 0
|  |  | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas molluscorum | 0.95 | 1 | 0
|  |  | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas popoffii | 0.7 | 1 | 0
|  |  | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas salmonicida | 0.98 | 1 | 0
|  |  | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas salmonicida | 0.71 | 1 | 0
|  |  | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas salmonicida | 0.55 | 1 | 0
|  |  | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas salmonicida | 0.83 | 1 | 0
|  |  | CVUAS 8400,2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas salmonicida | 0.74 | 1 | 0
|  |  | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas salmonicida | 0.54 | 1 | 0
|  |  | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas salmonicida | 0.32 | 1 | 0
|  |  | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas salmonicida | 0.85 | 1 | 0
|  |  | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas simiae | 0.18 | 1 | 0
|  |  | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas veronii | 0.15 | 1 | 0
|  |  | DSM 7386 CVUAS | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas veronii | 0.68 | 1 | 0
|  |  | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas pectinilytica | 0.24 | 1 | 0
|  |  | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella viscosa | 0.48 | 1 | 0
|  |  | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella viscosa | 0.77 | 1 | 0
|  |  | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.39 | 1 | 0
|  |  | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.32 | 1 | 0
|  |  | CVUAS 81,2 | 3; sequenced | EFEx | 28:11.2013 | Shewanella sp-RKVR22 | 0.86 | 1 | 0
|  |  | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter condimenti | 0.98 | 1 | 0
|  |  | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter dublinensis | 0.49 | 1 | 0
|  |  | DSM 18702 LALLF | 3; type strain | DT | 2019-03-05 | Cronobacter malonaticus | 0 | 1 | 0
|  |  | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter turicensis | 0.39 | 1 | 0
|  |  | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.11.2011 E6 | Enterobacter cloacae | 0.27 | 1 | 0
|  |  | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, H5 | Citrobacter amalonaticus | 0.22 | 1 | 0
|  |  | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter freundii | 0.13 | 1 | 0
|  |  | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter freundii | 0.16 | 1 | 0
|  |  | ATCC 25408 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 11.04.2008, N11, [0034]: Citrobacter diversus ATCC 25408\LMTG70N11_0\0_N11\1\1SLin\fid | Citrobacter koseri | 0.41 | 1 | 0
|  |  | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-14 | Cronobacter malonaticus | 0 | 1 | 0
|  |  | CVUAS 32945 | 3; confirmed proficency test | EFEx | 2021-02-18 | Cronobacter sakazakii | 0.37 | 1 | 0
|  |  | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter sakazakii | 0.01 | 1 | 0
|  |  | DSM 30054 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E18, [0034] | Enterobacter cloacae | 0.83 | 1 | 0
|  |  | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K4, [0034] | Enterobacter sp-ATCC-27988 | 0.95 | 1 | 0
|  |  | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae sp-CVUAS-30072 | 0.61 | 1 | 0
|  |  | CVUAS 0500,2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia coli | 0.78 | 1 | 0
|  |  | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella grimontii | 0.31 | 1 | 0
|  |  | CVUAS 2486,3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella grimontii | 0.23 | 1 | 0
|  |  | CVUAS 11699 | 3; sequenced [BfR] | eDT | 2017-04-20 06:56 | Klebsiella michiganensis | 0.78 | 1 | 0
|  |  | CVUAS 2404,3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella michiganensis | 0.27 | 1 | 0
|  |  | CVUAS 10038,2 | 3; sequenced [BfR] | EFex | 2019-02-12 14:02 | Klebsiella oxytoca | 0.64 | 1 | 0
|  |  | CVUAS 11572 | 3; sequenced [BfR] | eDT | 2017-04-20 06:51 | Klebsiella oxytoca | 0.59 | 1 | 0
|  |  | CVUAS 1482,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:35 | Klebsiella pneumoniae | 0.58 | 1 | 0
|  |  | CVUAS 1616,4 | 3; sequenced [BfR] | eDT | 2017-11-22 09:55 | Klebsiella pneumoniae | 0.3 | 1 | 0
|  |  | CVUAS 1678 | 3; sequenced [BfR] | eDT | 2016-10-18 06:38 | Klebsiella pneumoniae | 0.19 | 1 | 0
|  |  | CVUAS 1701 | 3; sequenced [BfR] | eDT | 2016-10-18 06:41 | Klebsiella pneumoniae | 0.36 | 1 | 0
|  |  | CVUAS 30499 | 3; sequenced [BfR] | eDT | 2017-11-22 09:59 | Klebsiella pneumoniae | 0.23 | 1 | 0
|  |  | CVUAS 753,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:34 | Klebsiella pneumoniae | 0.13 | 1 | 0
|  |  | CVUAS 10109 | 3; sequenced [BfR] | eDT | 2016-10-18 07:00 | Klebsiella quasipneumoniae | 0.52 | 1 | 0
|  |  | CVUAS 11677 | 3; sequenced [BfR] | eDT | 2017-04-20 06:54 | Klebsiella variicola | 0.33 | 1 | 0
|  |  | CVUAS 4888,2 | 3; sequenced [BfR] | EFex | 2017-07-19 11:36 | Klebsiella variicola | 0.59 | 1 | 0
|  |  | DSM 30053 CVUAS | 3; type strain | EFEx | 2019-02-12 | Klebsiella aerogenes | 0.47 | 1 | 0
|  |  | NM 20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E20 | Klebsiella aerogenes | 0.57 | 1 | 0
|  |  | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella michiganensis | 0.55 | 1 | 0
|  |  | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.11.2013 E18 | Klebsiella oxytoca | 0.39 | 1 | 0
|  |  | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 11.04.2008 N21 | Klebsiella pneumoniae | 0.45 | 1 | 0
|  |  | DSM 30104 | 3; type strain | EFEx | 2013-09-17 | Klebsiella pneumoniae | 0.91 | 1 | 0
|  |  | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.93 | 1 | 0
|  |  | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.38 | 1 | 0
|  |  | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella variicola | 0.28 | 1 | 0
|  |  | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia sacchari | 0.74 | 1 | 0
|  |  | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.26 | 1 | 0
|  |  | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.11 | 1 | 0
|  |  | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N17, [0034] (Enterobacter gergoviae) | Pluralibacter gergoviae | 0.5 | 1 | 0
|  |  | LT21 Ib 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H9 | Salmonella enterica | 0.21 | 1 | 0
|  |  | II 06 06820 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H6 | Shigella boydii | 0.32 | 1 | 0
|  |  | 3A 07 02503 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H7 | Shigella flexneri | 0.4 | 1 | 0
|  |  | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea agglomerans | 0.5 | 1 | 0
|  |  | DSM 3493 CVUAS | 3; type strain | eDT | 2015-07-09 | Pantoea agglomerans | 0.33 | 1 | 0
|  |  | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea sp | 0.02 | 1 | 0
|  |  | CVUAS 4263,2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea sp-FF5 | 0.28 | 1 | 0
|  |  | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella tarda | 0.73 | 1 | 0
|  |  | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N13, [0034] | Edwardsiella tarda | 0.89 | 1 | 0
|  |  | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 18.04.2008, K9, [0034] | Hafnia alvei | 0.55 | 1 | 0
|  |  | LHL1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia alvei | 0.13 | 1 | 0
|  |  | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea myxofaciens | 0.43 | 1 | 0
|  |  | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella morganii | 0.32 | 1 | 0
|  |  | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella morganii | 0.55 | 1 | 0
|  |  | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus cibarius | 0.58 | 1 | 0
|  |  | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus hauseri | 0.26 | 1 | 0
|  |  | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus hauseri | 0.32 | 1 | 0
|  |  | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus mirabilis | 1.03 | 1 | 0
|  |  | DSM 4479 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus mirabilis | 1.02 | 1 | 0
|  |  | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus penneri | 0.36 | 1 | 0
|  |  | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.65 | 1 | 0
|  |  | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.82 | 1 | 0
|  |  | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.72 | 1 | 0
|  |  | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus terrae | 0.84 | 1 | 0
|  |  | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, E5 K16, [0034]: Proteus vulgaris ATCC 33420\LMTG72K16_0\0_K16\1\1SLin\fid | Proteus vulgaris | 0.87 | 1 | 0
|  |  | DSM 2140 | 3; strain from public collection; sequenced (rpoB) | eDT | 30.11.2017 | Proteus vulgaris | 0.3 | 1 | 0
|  |  | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia rettgeri | 0.81 | 1 | 0
|  |  | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia rettgeri | 0.49 | 1 | 0
|  |  | ATCC 25827 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K15, [0034]: Proteus stuartii ATCC 25827\LMTG72K15_0\0_K15\1\1SLin\fid in vers 2016 | Providencia stuartii | 0.52 | 1 | 0
|  |  | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia frederiksenii | 0.68 | 1 | 0
|  |  | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia mollaretii | 0.24 | 1 | 0
|  |  | NCTC 2028 | 3; public strain collection | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H13 | Yersinia pestis | 0.41 | 1 | 0
|  |  | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H14 | Yersinia pestis | 0.49 | 1 | 0
|  |  | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia wautersii | 0.35 | 1 | 0
|  |  | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia fonticola | 0.44 | 1 | 0
|  |  | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.11.2008, E11, [0034] | Serratia grimesii | 0 | 1 | 0
|  |  | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia marcescens | 0.66 | 1 | 0
|  |  | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia marcescens | 0.73 | 1 | 0
|  |  | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia quinivorans | 0.58 | 1 | 0
|  |  | CVUAS 6221 | 4; sequenced | DT | 2013-01-08 | Yersinia aldovae | 0.01 | 1 | 0
|  |  | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 09.03.2007, H12 | Yersinia aleksiciae | 0.51 | 1 | 0
|  |  | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia bercovieri | 0.57 | 1 | 0
|  |  | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia bercovieri | 0.77 | 1 | 0
|  |  | SL2012 001153 | 3; sequenced | DT | 2013-01-30 | Yersinia bercovieri | 0.55 | 1 | 0
|  |  | DSM 4780 CVUAS | 3; type strain | DT | 2013-05-08 | Yersinia enterocolitica | 0.52 | 1 | 0
|  |  | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.45 | 1 | 0
|  |  | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia enterocolitica | 0.39 | 1 | 0
|  |  | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.24 | 1 | 0
|  |  | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia enterocolitica | 0.5 | 1 | 0
|  |  | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia enterocolitica | 0.78 | 1 | 0
|  |  | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.ma_01\0_C13, [0034] | Yersinia entomophaga | 0.58 | 1 | 0
|  |  | RKI 21 RKI | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 22.01.2007 H06 | Yersinia frederiksenii | 0.29 | 1 | 0
|  |  | RKI 21 CVUAS | 3; well described strain | DT | 2013-01-30 | Yersinia frederiksenii | 0.77 | 1 | 0
|  |  | CVUAS 1824 | 2; sequenced | DT | 2012-12-05 | Yersinia intermedia | 0.6 | 1 | 0
|  |  | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia intermedia | 0.39 | 1 | 0
|  |  | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia kristensenii | 0.6 | 1 | 0
|  |  | CVUAS 302 | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia kristensenii | 0.66 | 1 | 0
|  |  | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 25.11.2015, Y.ma_01\0_C11, [0034] | Yersinia massiliensis | 0.51 | 1 | 0
|  |  | VB 925699 | 3; sequenced | DT | 2013-03-19 | Yersinia massiliensis | 0.61 | 1 | 0
|  |  | DSM 18520 CVUAS | 3; type strain | DT | 16.01.2013 | Yersinia mollaretii | 0.17 | 1 | 0
|  |  | RKI 51 | 2; 0 | DT | 2013-01-16 | Yersinia mollaretii | 0.14 | 1 | 0
|  |  | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.nu_04\0_E10, [0034] | Yersinia nurmii | 0.3 | 1 | 0
|  |  | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.pe_04\0_E8, [0034] | Yersinia pekkanenii | 0.51 | 1 | 0
|  |  | CVUAS 0304,2 DT | 2; biochemistry | DT | 2013-01-31 | Yersinia pseudotuberculosis | 0.8 | 1 | 0
|  |  | DSM 8992 RKI | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.04.2009, E12 | Yersinia pseudotuberculosis | 0.56 | 1 | 0
|  |  | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia rohdei | 0.63 | 1 | 0
|  |  | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 02.03.2007, E07 | Yersinia rohdei | 0.29 | 1 | 0
|  |  | CVUAS 2772 CVUAS | 3; sequenced | DT | 2013-01-30 | Yersinia ruckeri | 0 | 1 | 0
|  |  | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia ruckeri | 0.76 | 1 | 0
|  |  | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia similis | 0.26 | 1 | 0
|  |  | IMB 4354 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 13.07.2007, E22 | Yersinia similis | 0.35 | 1 | 0
|  |  | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.wa_04\0_E5, [0034] | Yersinia wautersii | 0.62 | 1 | 0
|  |  | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella londiniensis | 0.79 | 1 | 0
|  |  | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella pneumophila | 0.73 | 1 | 0
|  |  | CVUAS 4122 | 3; 0 | eDT | 20.01.2016 | Legionella pneumophila | 0.69 | 1 | 0
|  |  | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella taurinensis | 0.5 | 1 | 0
|  |  | CVUAS 10873,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0.59 | 1 | 0
|  |  | CVUAS 10874,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0.31 | 1 | 0
|  |  | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) caballi | 0.61 | 1 | 0
|  |  | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) caballi | 0.67 | 1 | 0
|  |  | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) testudinis | 0.51 | 1 | 0
|  |  | LHL 957 | 3; sequenced | eDT | 02.12.2015 | (Pasteurella) testudinis | 0.75 | 1 | 0
|  |  | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus anseriformium | 0.76 | 1 | 0
|  |  | 15UGF 198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus anseriformium | 0.33 | 1 | 0
|  |  | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus arthritidis | 0.51 | 1 | 0
|  |  | 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus capsulatus | 0.69 | 1 | 0
|  |  | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus capsulatus | 0.67 | 1 | 0
|  |  | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.55 | 1 | 0
|  |  | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.72 | 1 | 0
|  |  | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.45 | 1 | 0
|  |  | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus equuli | 0.79 | 1 | 0
|  |  | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus lignieresii | 0.84 | 1 | 0
|  |  | CVUAS 9945,2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus lignieresii | 0.87 | 1 | 0
|  |  | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus rossii | 0.68 | 1 | 0
|  |  | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus rossii | 0.66 | 1 | 0
|  |  | CVUAS 32505,2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus sp-CVUAS-32505,2 | 0.63 | 1 | 0
|  |  | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus sp-genomospecies-1 | 0.85 | 1 | 0
|  |  | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-genomospecies-1 | 0.58 | 1 | 0
|  |  | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus sp-MK-2012 | 0.08 | 1 | 0
|  |  | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-MK-2012 | 0.44 | 1 | 0
|  |  | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus suis | 0.62 | 1 | 0
|  |  | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus suis | 0.65 | 1 | 0
|  |  | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium gallinarum | 0.38 | 1 | 0
|  |  | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium gallinarum | 0.4 | 1 | 0
|  |  | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 0.5 | 1 | 0
|  |  | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 0.66 | 1 | 0
|  |  | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter oris | 0.78 | 1 | 0
|  |  | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter oris | 0.49 | 1 | 0
|  |  | CVUAS 3802,2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter osteomyelitidis | 0.75 | 1 | 0
|  |  | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia canicola | 0.56 | 1 | 0
|  |  | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia canicola | 0.65 | 1 | 0
|  |  | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium anatis | 0.74 | 1 | 0
|  |  | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium anatis | 0.18 | 1 | 0
|  |  | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium sp-F151 | 0.29 | 1 | 0
|  |  | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium trehalosifermentans | 0.5 | 1 | 0
|  |  | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus haemoglobinophilus | 1.06 | 1 | 0
|  |  | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus sp-2269 | 0.88 | 1 | 0
|  |  | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus sp-2269 | 0.75 | 1 | 0
|  |  | P1066-2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus somni | 0.3 | 1 | 0
|  |  | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia caviae | 1.02 | 1 | 0
|  |  | CVUAS 1064,4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia glucosida | 1.02 | 1 | 0
|  |  | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia glucosida | 0.44 | 1 | 0
|  |  | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia granulomatis | 0.69 | 1 | 0
|  |  | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia granulomatis | 0.7 | 1 | 0
|  |  | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia haemolytica | 0.72 | 1 | 0
|  |  | CVUAS 6197,2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia haemolytica | 0.64 | 1 | 0
|  |  | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia pernigra | 0.72 | 1 | 0
|  |  | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia pernigra | 0.55 | 1 | 0
|  |  | LLBB 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia ruminalis | 0.6 | 1 | 0
|  |  | LLBB 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia ruminalis | 0.52 | 1 | 0
|  |  | CVUAS 10396,2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia sp-10-12559-1 | 0.59 | 1 | 0
|  |  | CVUAS 10183,3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia sp-14UCF195 | 0.38 | 1 | 0
|  |  | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia sp-14UCF195 | 0.52 | 1 | 0
|  |  | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia sp-15USF15 | 0.67 | 1 | 0
|  |  | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia sp-15USF15 | 0.58 | 1 | 0
|  |  | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia sp-CVUAS-11240 | 0.49 | 1 | 0
|  |  | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia sp-CVUAS-1496 | 0.67 | 1 | 0
|  |  | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter rosorum | 0.5 | 1 | 0
|  |  | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter rosorum | 0.74 | 1 | 0
|  |  | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella semolina | 1 | 1 | 0
|  |  | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella semolina | 0.77 | 1 | 0
|  |  | CVUAS 4060,3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella aerogenes | 0.85 | 1 | 0
|  |  | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella canis | 0.74 | 1 | 0
|  |  | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | Pasteurella mairii | 0.93 | 1 | 0
|  |  | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella multocida | 0.79 | 1 | 0
|  |  | CVUAS 32055 | 3; specific PCR | eDT | 2019-11-14 | Pasteurella multocida | 0.79 | 1 | 0
|  |  | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella multocida | 0.59 | 1 | 0
|  |  | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella multocida | 0.68 | 1 | 0
|  |  | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella multocida | 0.85 | 1 | 0
|  |  | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella multocida | 0.46 | 1 | 0
|  |  | CVUAS 32506,2 | 3; sequenced | EFEx | 2020-11-19 | Pasteurellaceae Bisgaard-Taxon-10 | 0.72 | 1 | 0
|  |  | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Pasteurellaceae Bisgaard-Taxon-10 | 0.88 | 1 | 0
|  |  | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae Bisgaard-Taxon-40 | 1.06 | 1 | 0
|  |  | CVUAS 9987,2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae canine-oral-taxon-080 | 1.03 | 1 | 0
|  |  | CVUAS 261,4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae sp-CVUAS-261,4 | 0.75 | 1 | 0
|  |  | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae sp-CVUAS-31988 | 0.87 | 1 | 0
|  |  | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae sp-CVUAS-4868 | 0.36 | 1 | 0
|  |  | CVUAS 724,3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae sp-CVUAS-724,3 | 0.46 | 1 | 0
|  |  | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae sp-CVUAS-9305 | 0.61 | 1 | 0
|  |  | CVUAS 9967,2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae sp-CVUAS-9967,2 | 0.56 | 1 | 0
|  |  | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae sp-LHL-U-309 | 0.57 | 1 | 0
|  |  | LHL U-309 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-309 | 0.69 | 1 | 0
|  |  | LHL U-3444 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-3444 | 0.71 | 1 | 0
|  |  | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae sp-MCCM-2539 | 0.47 | 1 | 0
|  |  | CVUAS 308,2 | 3; sequenced | DT | 2013-11-06 | Psittacicella hinzii | 0.53 | 1 | 0
|  |  | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.49 | 1 | 0
|  |  | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.64 | 1 | 0
|  |  | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium mucosae | 1 | 1 | 0
|  |  | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium mucosae | 0.58 | 1 | 0
|  |  | CVUAS 8966 | 3; sequenced | DT | 23.07.2013 | Volucribacter amazonae | 0.79 | 1 | 0
|  |  | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E22 | Acinetobacter baumannii | 0.01 | 1 | 0
|  |  | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2014 C05 | Acinetobacter baumannii | 0.48 | 1 | 0
|  |  | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter defluvii | 0.12 | 1 | 0
|  |  | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter guillouiae | 0.38 | 1 | 0
|  |  | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter johnsonii | 0.41 | 1 | 0
|  |  | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter junii | 0 | 1 | 0
|  |  | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter lwoffii | 0 | 1 | 0
|  |  | CVUAS 7379,3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter pseudolwoffii | 0.54 | 1 | 0
|  |  | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter pseudolwoffii | 0.69 | 1 | 0
|  |  | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter variabilis | 0.11 | 1 | 0
|  |  | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 0.61 | 1 | 0
|  |  | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 0 | 1 | 0
|  |  | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella catarrhalis | 0.77 | 1 | 0
|  |  | CVUAS 11674,3 | 3; sequenced | eDT | 2022-03-22 | Moraxella cuniculi | 0.63 | 1 | 0
|  |  | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella osloensis | 0 | 1 | 0
|  |  | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella ovis | 0.26 | 1 | 0
|  |  | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter frigidicola | 0 | 1 | 0
|  |  | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter pasteurii | 0.45 | 1 | 0
|  |  | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.51 | 1 | 0
|  |  | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.41 | 1 | 0
|  |  | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas (alkylphenolica) | 0.48 | 1 | 0
|  |  | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas aeruginosa | 0.11 | 1 | 0
|  |  | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007 H8 | Pseudomonas aeruginosa | 0.49 | 1 | 0
|  |  | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas baetica | 0.63 | 1 | 0
|  |  | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.02.2008 N11 | Pseudomonas chlororaphis | 0.27 | 1 | 0
|  |  | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 02.04.2012, H21 | Francisella tularensis | 0.33 | 1 | 0
|  |  | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 12.12.2016, F19, [0096] | Allofrancisella guangzhouensis | 0 | 1 | 0
|  |  | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 18.11.2015, L16, [0096] | Francisella hispaniensis | 0.38 | 1 | 0
|  |  | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 25.11.2015, I3, [0096] | Francisella naotunensis | 0.32 | 1 | 0
|  |  | CVUAS 4680,2 | 3; FLI: specific PCR | DT | 2013-04-02 | Francisella tularensis | 0.12 | 1 | 0
|  |  | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 23.04.2013, F20, [0096] | Francisella tularensis | 0.43 | 1 | 0
|  |  | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 26.04.2013, I20, [0096] | Francisella tularensis | 0.74 | 1 | 0
|  |  | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 01.06.2017, F3, [0096] | Francisella tularensis | 0.26 | 1 | 0
|  |  | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio aestuarianus | 0.23 | 1 | 0
|  |  | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio anguillarum | 0.68 | 1 | 0
|  |  | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio campbelli | 0.43 | 1 | 0
|  |  | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio campbelli | 0.42 | 1 | 0
|  |  | 16-VB00067-0_BfR 1_eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio cholerae | 0.58 | 1 | 0
|  |  | 10-VBH0202 DT-48h | 4; WGS sequenced | DT | 2021-02-02 | Vibrio cincinnatiensis | 0.79 | 1 | 0
|  |  | BfR-VB-00004 LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio cincinnatiensis | 0.7 | 1 | 0
|  |  | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio diazotrophicus | 0.72 | 1 | 0
|  |  | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio fluvialis | 0.67 | 1 | 0
|  |  | BfR-VB-00036 LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio fluvialis | 0.4 | 1 | 0
|  |  | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 16.08.2012, K12, [0034] | Vibrio harveyi | 0.49 | 1 | 0
|  |  | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio harveyi | 0.09 | 1 | 0
|  |  | 20-VB00047 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio jasicida | 0.4 | 1 | 0
|  |  | 20-VB00053 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio jasicida | 0.74 | 1 | 0
|  |  | 12-VBH0498 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio metoecus | 0.57 | 1 | 0
|  |  | 20-VB00016 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio metoecus | 0.53 | 1 | 0
|  |  | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio metschnikovii | 0.7 | 1 | 0
|  |  | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio navarrensis | 0.12 | 1 | 0
|  |  | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio owensii | 0.56 | 1 | 0
|  |  | 17-VB00254-0_BfR 18_DT | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio parahaemolyticus | 0.55 | 1 | 0
|  |  | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio vulnificus | 0.28 | 1 | 0
|  |  | CVUAS 33180 25C | 3; sequenced | DT | 2021-08-04 | Photobacterium damselae | 0.81 | 1 | 0
|  |  | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium damselae | 0.94 | 1 | 0
|  |  | CVUAS 1266,2 | 3; sequenced | EFEx | 25.09.1019 | Photobacterium piscicola | 0.41 | 1 | 0
|  |  | CVUAS 842,2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium swingsii | 0.26 | 1 | 0
|  |  | CVUAS 7079 sir | 3; spezific PCR (CVFR) | eDT | 2022-06-23 | Vibrio alginolyticus | 1.02 | 1 | 0
|  |  | DSM 2171 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 12.07.2012 E12 | Vibrio alginolyticus | 0.4 | 1 | 0
|  |  | A8-12 sir | 3; gyrB sequence | eDT | 2022-06-23 | Vibrio anguillarum | 0.47 | 1 | 0
|  |  | CVUAS 10227 | 3; sequenced | eDT | 2013-05-22 | Vibrio cholerae | 0.44 | 1 | 0
|  |  | CVUAS 8995 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio diazotrophicus | 0.65 | 1 | 0
|  |  | DSM 14383 sir | 3; public strain collection, sequenced | eDT | 2022-06-23 | Vibrio furnissi | 0.62 | 1 | 0
|  |  | CVUAS 1754 eDT | 3; sequenced | eDT | 2013-04-16 | Vibrio furnissii | 0.33 | 1 | 0
|  |  | CVUAS 619 sir | 2; 0 | eDT | 2022-06-23 | Vibrio metschnikovii | 0.51 | 1 | 0
|  |  | CVUAS 7872 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio mimicus | 0.5 | 1 | 0
|  |  | DSM 19130 RKI | 4; type strain; WGS sequenced | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 24.01.2011, N4, [0034] | Vibrio mimicus | 0.42 | 1 | 0
|  |  | DSM 10027 eDT | 3; type strain | eDT | 2022-06-23 | Vibrio parahaemolyticus | 0.64 | 1 | 0
|  |  | CVUAS 871,2 | 3; sequenced | eDT | 2014-08-21 | Vibrio sp-CVUAS-871,2 | 0.08 | 1 | 0
|  |  | CVUAS 2778,2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio tubiashii | 0.32 | 1 | 0
|  |  | CVUAS 2778,2 eDT | 3; sequenced | eDT | 2019-10-24 | Vibrio tubiashii | 0.5 | 1 | 0
|  |  | LVM05 LVU05 | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-14 | Vibrio vulnificus | 0.71 | 1 | 0
|  |  | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas sp-G9077 | 0.71 | 1 | 0
|  |  | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria indica | 0.61 | 1 | 0
|  |  | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria ureiclastica | 0.47 | 1 | 0
|  |  | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas chelatiphaga | 0.72 | 1 | 0
|  |  | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2009 K17 | Stenotrophomonas maltophilia | 0.53 | 1 | 0
|  |  | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.05.2009 K22 | Stenotrophomonas maltophilia | 0.85 | 1 | 0
|  |  | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas rhizophila | 0.37 | 1 | 0
|  |  | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.05.2015 C17 | Stenotrophomonas rhizophila | 0.01 | 1 | 0
|  |  | CVUAS 4239,2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas tumulicola | 0.35 | 1 | 0

## Report (Spezies)
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Parameter: Corynebacterium pseudotuberculosis
Version: | 3
erstellt von: | Dyk | geprüft von: | Rau
erstellt am: | 2024-09-03 | geprüft am: | 2024-09-03
Anwendung | Bakterienisolate aus Lebensmitteln & der veterinärmedizinischen Diagnostik
Umfeld | Lebensmittel-Mikrobiologie, Mastitis, Veterinärmedizinische Diagnostik
Verwendete Datenbank / Programmversion |  | BRU 11897 + UA BW 1826 | BT 4.0
Validierungsset:
Validierungsisolate/-materialien (Parameter): 66 |  | Vergleichsisolate/-materialien (#Parameter): 1278
Kultivierungsbedingungen bei Mikroorganismen: SBA, 37° C, 24h, aerob
Zustand der Probe bei sonstigen Materialien:
Aufreinigungsmethode: DT, eDT, EFex, 80%TFA
Es werden nur identifizierte Proben für die Ermittlung der In- und Exklusivität herangezogen.
Richtig-positiv: 1. Hit =Ziel-Parameter mit Score ≥ 2 und Score 2. Hit < 1.7 oder Score 2. Hit ≥ 1.7 & Gattung gleich 1. Hit oder Score 2. Hit ≥ 2 & Art gleich 1. Hit |  | Richtig-negativ: 1. Hit ohne Zuordnung zum Ziel-Parameter mit score ≥ 2
Falsch-negativ: Kriterien für richtig-positiv werden nicht erfüllt |  | Falsch-positiv: Kriterien für richtig-negativ werden nicht erfüllt
Validierungsergebnis | Ziel-Parameter: | Corynebacterium pseudotuberculosis |  | SOLL | erfüllt
Zahl aller Spektren | 66
davon mit ID-Ergebnis | 66 | Identifiziert Rate | 100% | 20 | ja | 1
davon Eigenspektren | 65 |  |  | 13.2 | ja | 1
In Bezug auf Identifizierte: |  | Inklusivität
Positive Übereinstimmungen (PA) | 66 | Richtig-Positiv Rate | 100% | 95% | ja | 1
Negative Abweichungen (ND) | 0 | Falsch-Negativ Rate | 0% | 1% | ja | 1
Nicht-Ziel-Parameter
Zahl aller Spektren | 1278
davon mit ID-Ergebnis | 1028 | Identifiziert Rate | 80.44% | 30 | ja | 1
davon Eigenspektren | 743 |  |  | 205.6 | ja | 1
In Bezug auf Identifizierte: |  | Exklusivität
Negative Übereinstimmungen (NA) | 1028 | Richtig-Negativ Rate | 100% | 99% | ja | 1
Positive Abweichungen (PD) | 0 | Falsch-Positiv Rate | 0% | 1% | ja | 1
Test auf spektrale Nähe | 1278 | Anteil nah verwandter Sp | 6.182%
Von 66 identifizierten Proben des Parameters wurden unter Verwendung der vollständigen Datenbank 66 (=100%) richtig erkannt (Inklusivität). 0 (=0%) der identifizierten Proben des Parameters wurden falsch einer anderen Spezies zugeordnet.
In den 1028 identifizierten Proben der Nicht-Ziel-Parameter wurden unter Verwendung der vollständigen Datenbank 100% richtig-negativ (Exklusivität) gezählt. Von diesen 1028 Proben wurden 0 (=0%) fehlerhaft als Parameter identifiziert.
In den 1278 Proben der Nicht-Ziel-Parameter wurden unter Verwendung einer nur den Parameter enthaltenden Datenbank bei 79 (=6.2%) Einträgen ein score > 2 erreicht.
Die Kriterien für die | Zielgerichtete Identifizierung | sind somit | erfüllt
Kommentar:
Freigabe an Standort | für (Abteilung / Bereich) | am (Datum) | von (Hdz.)
CVUA S | T, D, MT | 2024-09-03 | Rau
Das Verfahren ist für den vorgesehenen Einsatzzweck geeignet und zur Verwendung freigegeben.
als Prüfmethode für Parameter in LIMS eingepflegt
Eine unterschriebene Version dieses Validierungsreports befindet sich im erstellenden Labor.
Andere Keime, die in dieser Matrix vorkommen können.
Andere Keime, die auf diesem Agar wachsen können.
Vergleichsisolate/-materialien, die ähnliche Spektren liefern können (siehe Blatt "Ähnliche zu Parameter").
ja
nein