## Settings
Settings for the Validation
Targeted Identification
following the Guidelines for Validating Species Identifications using MALDI-TOF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022
Institution / Labor | CVUA Stuttgart / Germany
Parameter | Genus |  | Species
|  | Bacillus | cytotoxicus
Differentiation limits
| Score for differentiation on genus level | 1.7 | Score for differentiation on species level | 2
The data shown in this validation report is linked to
Rau J, Etter D, Frentzel H, Lasch P, Contzen M. 2024.
Reliable delineation of Bacillus cytotoxicus from other members of the Bacillus cereus group by MALDI-TOF MS – an extensive validation study.
Food Control accepted manuscript
This validation report provides all functionalities in excel.

## Target parameter (Species)
Parameter: | Bacillus | cytotoxicus | own spectrum | own ID | Isolate ID | Validity | preparation method (if given) | date | Genus | Species | Score | Genus | Species | Score | Consistence | identified | not assigned | correct | false negative | wrong species identified | second spectra of isolate: | 13
|  |  | 1 | 1 | 46776 | 3; specific PCR, growth at 52°C | eDT | 2017-05-31 | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
Total number of spectra | 183 |  | 1 | 1 | 46778 | 3; specific PCR, growth at 52°C | eDT | 2017-05-31 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0 |  | 1
of which with ID result | 182 |  | 0 | 0 | 46776 MVZ | 3; specific PCR, growth at 52°C | DT | 2017-05-23 | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.39 | A | 1 | 0 | 1 | 0 | 0 |  | 1
not identified | 1 |  | 0 | 0 | 46778 MVZ | 3; specific PCR, growth at 52°C | DT | 2017-05-23 | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
with respect to identified spectra: |  |  | 1 | 1 | Ad 2163 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.43 | A | 1 | 0 | 1 | 0 | 0
positive accordance (PA) | 182 |  | 1 | 1 | Ad 2164 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
negative deviation (ND) | 0 |  | 1 | 1 | Ad 2466 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.45 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | Ad 2469 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | 2.25 | Bacillus | cytotoxicus | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | BfR-BA-01360 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | 2.1 | Bacillus | cytotoxicus | 2.09 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | BfR-BA-01430 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | 1.85 | Bacillus | cytotoxicus | 1.74 | B | 0 | 1 | 0 | 0 | 0
score limit | 2 | Species | 0 | 0 | BfR-BA-01431 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | 2.21 | Bacillus | cytotoxicus | 2.1 | A | 1 | 0 | 1 | 0 | 0
settings | 1.7 | Genus | 0 | 0 | BfR-BA-01492 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | 2.11 | Bacillus | cytotoxicus | 2.1 | A | 1 | 0 | 1 | 0 | 0
The complete database to be checked is used. |  |  | 0 | 0 | BfR-BA-01522 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | 2.01 | Bacillus | cytotoxicus | 1.98 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH1 | 3; WGS | 0 | 0 | Bacillus | cytotoxicus | 2.51 | Bacillus | cytotoxicus | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH13 | 3; WGS | 0 | 0 | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus | 2.4 | A | 1 | 0 | 1 | 0 | 0
number of individual isolates | 170 |  | 0 | 0 | CH134 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.34 | A | 1 | 0 | 1 | 0 | 0
mean score first hit | 2.3768 |  | 0 | 0 | CH15 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH16 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.59 | Bacillus | cytotoxicus | 2.47 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH2 | 3; WGS | 0 | 0 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH212 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH213 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH214 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.39 | A | 1 | 0 | 1 | 0 | 0
of which with ID result | 182 |  | 0 | 0 | CH216 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
positive accordance (PA) | 182 |  | 0 | 0 | CH217 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.57 | Bacillus | cytotoxicus | 2.43 | A | 1 | 0 | 1 | 0 | 0
negative deviation (ND) | 0 |  | 0 | 0 | CH218 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH23 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.41 | A | 1 | 0 | 1 | 0 | 0
not identified in % | 0.5464% |  | 0 | 0 | CH25 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.44 | A | 1 | 0 | 1 | 0 | 0
mismatched in % | 0% |  | 0 | 0 | CH3 | 3; WGS | 0 | 0 | Bacillus | cytotoxicus | 2.51 | Bacillus | cytotoxicus | 2.47 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH348 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.39 | A | 1 | 0 | 1 | 0 | 0
thereof |  |  | 0 | 0 | CH38 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.34 | A | 1 | 0 | 1 | 0 | 0
own spectra (CVUAS) | 93 |  | 0 | 0 | CH39 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
own spectra [%] | 50.82% |  | 0 | 0 | CH4 | 3; WGS | 0 | 0 | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.27 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH647 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH755 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH834 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.45 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH835 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.27 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH836 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.2 | Bacillus | cytotoxicus | 2.12 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH837 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH838 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH839 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH840 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH841 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH842 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.38 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH843 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.3 | Bacillus | cytotoxicus | 2.24 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH844 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.29 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH845 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH846 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.31 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH847 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH848 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH849 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH850 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH851 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.06 | Bacillus | cytotoxicus | 1.97 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH852 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH853 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.31 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH854 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH855 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH856 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.27 | Bacillus | cytotoxicus | 2.27 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH857 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.22 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH858 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH859 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH860 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH861 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH862 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH863 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.21 | Bacillus | cytotoxicus | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH864 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.24 | Bacillus | cytotoxicus | 2.21 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH865 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH866 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH867 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.47 | Bacillus | cytotoxicus | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH868 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH869 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.51 | Bacillus | cytotoxicus | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH870 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH871 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH872 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH873 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH874 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH875 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH876 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH877 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH878 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH879 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.21 | Bacillus | cytotoxicus | 2.1 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH880 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.29 | Bacillus | cytotoxicus | 2.25 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH881 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.27 | Bacillus | cytotoxicus | 2.25 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH882 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH883 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.24 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH884 | 3; WGS | 0 | 44952 | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CH886 | 3; PCR | 0 | 44952 | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 0644 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.11 | Bacillus | cytotoxicus | 2.08 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 0 | 0 | CVUAS 0644 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.79 | Bacillus | cytotoxicus | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 10162 | 3; sequenced | eDT | 42338 | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | DSM 22905 CVUAS 10202 | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 30.05.2016, F14, [0034] | Bacillus | cytotoxicus | 2.68 | Bacillus | cytotoxicus | 2.4 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 1 | 1 | CVUAS 10577 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.38 | Bacillus | cytotoxicus | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 11025 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 11326,2 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2831 | 3; sequenced; Wt 55°C, cytK1 | DT | 41292 | Bacillus | cytotoxicus | 2.28 | Bacillus | cytotoxicus | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2833 | 4; 16S, draft genome [0009] | eDT | 42396 | Bacillus | cytotoxicus | 2.53 | Bacillus | cytotoxicus | 2.48 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CVUAS 2833 | 4; 16S, draft genome [0009] | eDT | 42396 | Bacillus | cytotoxicus | 2.55 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 1 | 1 | CVUAS 2894 | 3; sequenced; Wt 55°C, cytK1 | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2905 | 3; sequenced; Wt 55°C, cytK1 | DT | 41547 | Bacillus | cytotoxicus | 2.3 | Bacillus | cytotoxicus | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2979 | 3; 16S | eDT | 42396 | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.45 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CVUAS 2979 | 3; 16S | eDT | 42396 | Bacillus | cytotoxicus | 2.77 | Bacillus | cytotoxicus | 2.41 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 1 | 1 | CVUAS 2982 | 3; sequenced; Wt 55°C, cytK1 | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3017,3 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.32 | Bacillus | cytotoxicus | 2.27 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 30572 | 1; 0 | EFex [BRU03] | 43061 | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.51 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 30605 | 1; 0 | eDT | 43070 | Bacillus | cytotoxicus | 2.55 | Bacillus | cytotoxicus | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 30626 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 31073 | 3; specific PCR | eDT | 44883 | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 31094 | 1; 0 | eDT | 43326 | Bacillus | cytotoxicus | 2.56 | Bacillus | cytotoxicus | 2.5 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 355,3 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 356,3 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 363,4 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.38 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3827 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CVUAS 3827 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.56 | Bacillus | cytotoxicus | 2.19 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 1 | 1 | CVUAS 3828 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus | 2.42 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 0 | 0 | CVUAS 3828 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.68 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3896 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 42018 | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4117 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 42018 | Bacillus | cytotoxicus | 2.54 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4545 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.29 | Bacillus | cytotoxicus | 2.1 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4546,4 | 3; cyt K1 | eDT | 42144 | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5334 | 3; Wt 55°C, cytK1 | eDT | 44883 | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5335 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.32 | Bacillus | cytotoxicus | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5403 | 2; Wt 55°C | eDT | 44883 | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.45 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5404 | 2; Wt 55°C | eDT | 42696 | Bacillus | cytotoxicus | 2.16 | Bacillus | cytotoxicus | 2.09 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5405 | 2; Wt 55°C | eDT | 42696 | Bacillus | cytotoxicus | 2.15 | Bacillus | cytotoxicus | 2.11 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6081 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.2 | Bacillus | cytotoxicus | 2.19 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6082 | 3; Wt 50°C, cytK1 | eDT | 42690 | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CVUAS 6082 | 3; Wt 50°C, cytK1 | eDT | 42690 | Bacillus | cytotoxicus | 2.7 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 1 | 1 | CVUAS 6083 | 3; Wt 50°C, cytK1 | eDT | 42690 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6092 | 3; Wt 50°C, cytK1 | eDT | 44883 | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6128 | 3; Wt 50°C, cytK1 | EFEx | 42017 | Bacillus | cytotoxicus | 2.17 | Bacillus | cytotoxicus | 2.09 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6129 | 3; Wt 50°C, cytK1 | EFEx | 42019 | Bacillus | cytotoxicus | 2.5 | Bacillus | cytotoxicus | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6130 | 3; Wt 50°C, cytK1 | EFEx | 42017 | Bacillus | cytotoxicus | 2.23 | Bacillus | cytotoxicus | 2.15 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6183 | 3; Wt 55°C, cytK1 | eDT | 42690 | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 0 | 0 | CVUAS 6183 | 3; Wt 55°C, cytK1 | eDT | 42690 | Bacillus | cytotoxicus | 2.66 | Bacillus | cytotoxicus | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6204 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.23 | Bacillus | cytotoxicus | 2.19 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6205 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42018 | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6206 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6207 | 3; Wt 50°C, cytK1 | eDT | 42690 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6208 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.25 | Bacillus | cytotoxicus | 2.21 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6209 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42080 | Bacillus | cytotoxicus | 2.2 | Bacillus | cytotoxicus | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6210 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42031 | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6211 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.24 | Bacillus | cytotoxicus | 2.1 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6238 | 3; sequenced; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.21 | Bacillus | cytotoxicus | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6239 | 3; sequenced; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.32 | Bacillus | cytotoxicus | 2.21 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6240 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6241 | 3; Wt 50°C, cytK1 | eDT | 44883 | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6242 | 3; Wt 50°C, cytK1 | eDT | 44883 | Bacillus | cytotoxicus | 2.54 | Bacillus | cytotoxicus | 2.51 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6243 | 3; Wt 50°C, cytK1 | eDT | 44883 | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6244 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CVUAS 6244 | 3; 16S | eDT | 42690 | Bacillus | cytotoxicus | 2.63 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 1 | 1 | CVUAS 6245 | 3; Wt 50°C, cytK1 | eDT | 44888 | Bacillus | cytotoxicus | 2.48 | Bacillus | cytotoxicus | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6292 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6336 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42031 | Bacillus | cytotoxicus | 2.18 | Bacillus | cytotoxicus | 2.16 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6433,2 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.47 | Bacillus | cytotoxicus | 2.47 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6435 | 3; Wt 50°C, cytK1 | eDT | 44883 | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6475 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42031 | Bacillus | cytotoxicus | 2.12 | Bacillus | cytotoxicus | 2.05 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6608 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42031 | Bacillus | cytotoxicus | 2.12 | Bacillus | cytotoxicus | 2.07 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6655 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6656 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42074 | Bacillus | cytotoxicus | 2.32 | Bacillus | cytotoxicus | 2.22 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6657 | 3; Wt 50°C, cytK1 | eDT | 42088 | Bacillus | cytotoxicus | 2.29 | Bacillus | cytotoxicus | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6658 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6659 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42486 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.45 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6660 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.23 | Bacillus | cytotoxicus | 2.13 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6661 | 3; Wt 55°C, cytK1 | eDT | 42347 | Bacillus | cytotoxicus | 2.54 | Bacillus | cytotoxicus | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6688 | 3; Wt 55°C, cytK1 | eDT | 43733 | Bacillus | cytotoxicus | 2.47 | Bacillus | cytotoxicus | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | CVUAS 6688 | 3; Wt 55°C, cytK1 | eDT | 43733 | Bacillus | cytotoxicus | 2.72 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0 |  | 1
|  |  | 1 | 1 | CVUAS 6689 | 3; cytK1 | EFex [BRU03] | 42074 | Bacillus | cytotoxicus | 2.48 | Bacillus | cytotoxicus | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6690 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.14 | Bacillus | cytotoxicus | 2.12 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6691 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.18 | Bacillus | cytotoxicus | 2.12 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6692 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.27 | Bacillus | cytotoxicus | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6693 | 3; Wt 50°C, cytK1 | eDT | 42025 | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6694 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 42031 | Bacillus | cytotoxicus | 2.2 | Bacillus | cytotoxicus | 2.17 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 701,3 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 7208 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.53 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 7209 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.28 | Bacillus | cytotoxicus | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 7500 | 3; specific PCR | EFex [BRU03] | 42017 | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 8754 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 8755 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 9297 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 9298 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 9299 | 1; 0 | eDT | 44883 | Bacillus | cytotoxicus | 2.38 | Bacillus | cytotoxicus | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | DSM 22905 | 3; type strain | eDT | 42396 | Bacillus | cytotoxicus | 2.56 | Bacillus | cytotoxicus | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | IGV 003 | 3; sequenced | eDT | 42594 | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | NVH 391-98 | 3; type strain | eDT | 42690 | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.46 | A | 1 | 0 | 1 | 0 | 0 |  | 1

## non-target parameter (Species)
Parameter: | Bacillus cytotoxicus |  | own spectrum | own ID | Isolate ID | Validity | preparation method (if given) | date | Genus | Species | Genus | Species | Score | Genus | Species | Score | Consistence | identified | correct ID | not assigned on sepcies level | 1. Hit | 2. Hit | 1. Hit is met with the parameter | 2. Hit is met with the parameter | wrong identified as parameter | parameter in the first or the second hit | Correct ID as Bacillus cereus group
252
|  |  | 1 | 1 | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | bovis | Actinomyces | bovis | 2.05 | Actinomyces | bovis | 1.92 | A | 1 | 1 | 0 | Actinomyces bovis | Actinomyces bovis | 0 | 0 | 0 | 0
total number of spectra | 1408 |  | 1 | 1 | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces | denticolens | Actinomyces | denticolens | 2.4 | Actinomyces | denticolens | 2.07 | A | 1 | 1 | 0 | Actinomyces denticolens | Actinomyces denticolens | 0 | 0 | 0 | 0
of which with ID result | 916 |  | 1 | 1 | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces | sp-CVUAS-31303 | Actinomyces | sp-CVUAS-31303 | 2.6 | Campylobacter | jejuni | 1.26 | A | 1 | 1 | 0 | Actinomyces sp-CVUAS-31303 | Campylobacter jejuni | 0 | 0 | 0 | 0
with respect to identified spectra: |  |  | 1 | 1 | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces | viscosus | Actinomyces | viscosus | 2.04 | Actinomyces | viscosus | 1.87 | A | 1 | 1 | 0 | Actinomyces viscosus | Actinomyces viscosus | 0 | 0 | 0 | 0
negative accordance (NA) (correct identified not as the parameter) | 916 |  | 1 | 1 | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | weissii | Actinomyces | weissii | 2.48 | Actinomyces | weissii | 2.35 | A | 1 | 1 | 0 | Actinomyces weissii | Actinomyces weissii | 0 | 0 | 0 | 0
wrong identified as parameter (PD) | 0 |  | 1 | 1 | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium | bovis | Arcanobacterium | bovis | 2.73 | Arcanobacterium | pluranimalium | 1.77 | A | 1 | 1 | 0 | Arcanobacterium bovis | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium | canis | Arcanobacterium | canis | 2.34 | Arcanobacterium | canis | 2.19 | A | 1 | 1 | 0 | Arcanobacterium canis | Arcanobacterium canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium | haemolyticum | Arcanobacterium | haemolyticum | 2.46 | Arcanobacterium | haemolyticum | 2.3 | A | 1 | 1 | 0 | Arcanobacterium haemolyticum | Arcanobacterium haemolyticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-6203 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.65 | Arcanobacterium | hippocoleae | 2.6 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.32 | Arcanobacterium | hippocoleae | 2.3 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
Parameter is assigned as: |  |  | 1 | 1 | DSM 10002 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.66 | Arcanobacterium | phocae | 2.5 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.48 | Arcanobacterium | phocae | 2.4 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
1. Hit is met with the parameter | 0 |  | 1 | 1 | 2698T | 3; type strain | DT | 2013-08-07 | Arcanobacterium | phocisimile | Arcanobacterium | phocisimile | 2.15 | Arcanobacterium | phocisimile | 2.13 | A | 1 | 1 | 0 | Arcanobacterium phocisimile | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
2. Hit is met with the parameter | 0 |  | 1 | 1 | 14418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.63 | Arcanobacterium | pinnipediorum | 2.47 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium pinnipediorum | 0 | 0 | 0 | 0
PM is among the first 2 hits | 0 |  | 1 | 1 | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.75 | Arcanobacterium | pinnipediorum | 2.56 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium pinnipediorum | 0 | 0 | 0 | 0
in % | 0% |  | 1 | 1 | DSM 13483 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium | pluranimalium | Arcanobacterium | pluranimalium | 2.47 | Arcanobacterium | pluranimalium | 2.1 | A | 1 | 1 | 0 | Arcanobacterium pluranimalium | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | sp-17-2065.02 | 2.72 | Arcanobacterium | sp-17-2065.02 | 2.62 | A | 1 | 1 | 0 | Arcanobacterium sp-17-2065.02 | Arcanobacterium sp-17-2065.02 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | sp-17-2065.02 | 2.63 | Arcanobacterium | sp-17-2065.02 | 2.61 | A | 1 | 1 | 0 | Arcanobacterium sp-17-2065.02 | Arcanobacterium sp-17-2065.02 | 0 | 0 | 0 | 0
score limit | 2 | Species | 1 | 1 | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium | sp-2701 | Arcanobacterium | sp-2701 | 2.7 | Arcanobacterium | sp-2701 | 2.57 | A | 1 | 1 | 0 | Arcanobacterium sp-2701 | Arcanobacterium sp-2701 | 0 | 0 | 0 | 0
settings | 1.7 | Genus | 1 | 1 | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium | sp-2701 | Arcanobacterium | sp-2701 | 2.78 | Arcanobacterium | sp-2701 | 2.64 | A | 1 | 1 | 0 | Arcanobacterium sp-2701 | Arcanobacterium sp-2701 | 0 | 0 | 0 | 0
The complete database to be checked is used. |  |  | 1 | 1 | CVUAS 1060,2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium | wilhelmae | Arcanobacterium | wilhelmae | 2.32 | Arcanobacterium | wilhelmae | 2.18 | A | 1 | 1 | 0 | Arcanobacterium wilhelmae | Arcanobacterium wilhelmae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 647 | 3; type strain | DT | 2014-09-04 | Arcanobacterium | wilhelmae | Arcanobacterium | wilhelmae | 2.43 | Arcanobacterium | wilhelmae | 2.25 | A | 1 | 1 | 0 | Arcanobacterium wilhelmae | Arcanobacterium wilhelmae | 0 | 0 | 0 | 0
of which with ID result | 916 |  | 1 | 1 | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia | sp-CVUAS-30561 | Boudabousia | sp-CVUAS-30561 | 2.65 | Streptomyces | polychromogenes | 1.39 | A | 1 | 1 | 0 | Boudabousia sp-CVUAS-30561 | Streptomyces polychromogenes | 0 | 0 | 0 | 0
positive accordance (PA) | 916 |  | 1 | 1 | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia | canis | Schaalia | canis | 2.41 | Schaalia | canis | 2.26 | A | 1 | 1 | 0 | Schaalia canis | Schaalia canis | 0 | 0 | 0 | 0
negative deviation (ND) | 0 |  | 1 | 1 | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia | cardiffensis | Schaalia | cardiffensis | 2.49 | Schaalia | cardiffensis | 1.4 | A | 1 | 1 | 0 | Schaalia cardiffensis | Schaalia cardiffensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.45 | Schaalia | hyovaginalis | 2.43 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
not identified in % | 34.94% |  | 1 | 1 | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.48 | Schaalia | hyovaginalis | 2.4 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
mismatched in % | 0% |  | 1 | 1 | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia | sp-CVUAS-32183 | Schaalia | sp-CVUAS-32183 | 2.55 | Schaalia | hyovaginalis | 1.47 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-32183 | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia | sp-CVUAS-8688 | Schaalia | sp-CVUAS-8688 | 2.57 | Schaalia | sp-CVUAS-8688 | 2.35 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-8688 | Schaalia sp-CVUAS-8688 | 0 | 0 | 0 | 0
own spectra (CVUAS) | 749 |  | 1 | 1 | CVUAS 31845,2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia | sp-CVUAS-8688 | Schaalia | sp-CVUAS-8688 | 2.73 | Schaalia | sp-CVUAS-8688 | 2.3 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-8688 | Schaalia sp-CVUAS-8688 | 0 | 0 | 0 | 0
own spectra [%] | 53.12% |  | 1 | 1 | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella | abortisuis | Trueperella | abortisuis | 2.27 | Trueperella | abortisuis | 2.2 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | abortisuis | Trueperella | abortisuis | 2.36 | Trueperella | abortisuis | 2.16 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
own spectra identified | 352 |  | 1 | 1 | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella | bernardiae | Trueperella | bernardiae | 2.61 | Trueperella | bernardiae | 2.11 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
own spectra identified [%] | 55.87% |  | 1 | 1 | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella | bernardiae | Trueperella | bernardiae | 2.23 | Trueperella | bernardiae | 2.15 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella | bialowiezensis | Trueperella | bialowiezensis | 2.02 | Oceanobacillus | picturae | 1.3 | A | 1 | 1 | 0 | Trueperella bialowiezensis | Oceanobacillus picturae | 0 | 0 | 0 | 0
external spectra (CVUAS) | 660 |  | 1 | 1 | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella | bonasi | Trueperella | bonasi | 2.47 | Trueperella | bonasi | 2.32 | A | 1 | 1 | 0 | Trueperella bonasi | Trueperella bonasi | 0 | 0 | 0 | 0
external spectra [%] | 46.88% |  | 1 | 1 | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pecoris | 2.31 | Trueperella | pecoris | 1.98 | A | 1 | 1 | 0 | Trueperella pecoris | Trueperella pecoris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pecoris | 2.37 | Trueperella | pecoris | 1.65 | A | 1 | 1 | 0 | Trueperella pecoris | Trueperella pecoris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.42 | Trueperella | pyogenes | 2.39 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.33 | Trueperella | pyogenes | 2.32 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.19 | Corynebacterium | amycolatum | 2.15 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2014 E8 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.3 | Corynebacterium | amycolatum | 2.29 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium | aurimucosum | Corynebacterium | aurimucosum | 2.07 | Corynebacterium | aurimucosum | 2.05 | A | 1 | 1 | 0 | Corynebacterium aurimucosum | Corynebacterium aurimucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | belfantii | Corynebacterium | belfantii | 2.75 | Corynebacterium | belfantii | 2.59 | A | 1 | 1 | 0 | Corynebacterium belfantii | Corynebacterium belfantii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | belfantii | Corynebacterium | belfantii | 2.66 | Corynebacterium | belfantii | 2.62 | A | 1 | 1 | 0 | Corynebacterium belfantii | Corynebacterium belfantii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium | bovis | Corynebacterium | bovis | 2.35 | Corynebacterium | bovis | 2.33 | A | 1 | 1 | 0 | Corynebacterium bovis | Corynebacterium bovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium | camporealensis | Corynebacterium | camporealensis | 2.58 | Corynebacterium | camporealensis | 1.96 | A | 1 | 1 | 0 | Corynebacterium camporealensis | Corynebacterium camporealensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium | casei | Corynebacterium | casei | 1.98 | Corynebacterium | stationis | 1.64 | B | 0 | 0 | 1 | Corynebacterium casei | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium | confusum | Corynebacterium | confusum | 2.49 | Corynebacterium | confusum | 1.93 | A | 1 | 1 | 0 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | confusum | Corynebacterium | confusum | 2.14 | Corynebacterium | confusum | 2.01 | A | 1 | 1 | 0 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium | cystitidis | Corynebacterium | cystitidis | 2.5 | Corynebacterium | cystitidis | 2.36 | A | 1 | 1 | 0 | Corynebacterium cystitidis | Corynebacterium cystitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.66 | Corynebacterium | diphtheriae | 2.59 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.58 | Corynebacterium | diphtheriae | 2.58 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.71 | Corynebacterium | diphtheriae | 2.65 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0956 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.04 | Corynebacterium | diphtheriae | 2 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.32 | Corynebacterium | frankenforstense | 2.2 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Corynebacterium frankenforstense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.74 | Corynebacterium | frankenforstense | 2.48 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Corynebacterium frankenforstense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium | freiburgense | Corynebacterium | freiburgense | 2.46 | Pseudomonas | gessardii | 1.27 | A | 1 | 1 | 0 | Corynebacterium freiburgense | Pseudomonas gessardii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | hansenii | Corynebacterium | hansenii | 2.35 | Corynebacterium | hansenii | 2.3 | A | 1 | 1 | 0 | Corynebacterium hansenii | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | hansenii | Corynebacterium | hansenii | 2.41 | Corynebacterium | hansenii | 2.2 | A | 1 | 1 | 0 | Corynebacterium hansenii | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | humireducens | Corynebacterium | humireducens | 2.51 | Corynebacterium | sp-CVUAS-31533 | 1.98 | A | 1 | 1 | 0 | Corynebacterium humireducens | Corynebacterium sp-CVUAS-31533 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | kutscheri | Corynebacterium | kutscheri | 2.56 | Corynebacterium | kutscheri | 2.37 | A | 1 | 1 | 0 | Corynebacterium kutscheri | Corynebacterium kutscheri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | lactis | Corynebacterium | lactis | 2.38 | Corynebacterium | lactis | 2.29 | A | 1 | 1 | 0 | Corynebacterium lactis | Corynebacterium lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | lactis | Corynebacterium | lactis | 2.7 | Corynebacterium | lactis | 2.54 | A | 1 | 1 | 0 | Corynebacterium lactis | Corynebacterium lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium | phoceense | Corynebacterium | phoceense | 2.61 | Corynebacterium | phoceense | 2.19 | A | 1 | 1 | 0 | Corynebacterium phoceense | Corynebacterium phoceense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | pilosum | Corynebacterium | pilosum | 2.3 | Corynebacterium | pilosum | 1.78 | A | 1 | 1 | 0 | Corynebacterium pilosum | Corynebacterium pilosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.48 | Corynebacterium | pseudotuberculosis | 2.46 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.6 | Corynebacterium | pseudotuberculosis | 2.56 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20689 | 3; type strain | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C138 | 3; sequenced (rpoB) UGent | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.79 | Corynebacterium | rouxii | 2.65 | B | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.73 | Corynebacterium | rouxii | 2.71 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3559,2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.7 | Corynebacterium | rouxii | 2.66 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.74 | Corynebacterium | rouxii | 2.65 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.71 | Corynebacterium | silvaticum | 2.66 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.64 | Corynebacterium | silvaticum | 2.55 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium | sp-208-U2-6 | Corynebacterium | sp-208-U2-6 | 2.63 | Sphingomonas | yabuuchiae | 1.39 | A | 1 | 1 | 0 | Corynebacterium sp-208-U2-6 | Sphingomonas yabuuchiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32165,2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium | sp-2300500 | Corynebacterium | sp-2300500 | 2.59 | Corynebacterium | sp-2300500 | 2.56 | A | 1 | 1 | 0 | Corynebacterium sp-2300500 | Corynebacterium sp-2300500 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5775,2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium | sp-2300500 | Corynebacterium | sp-2300500 | 2.24 | Corynebacterium | sp-2300500 | 2.22 | A | 1 | 1 | 0 | Corynebacterium sp-2300500 | Corynebacterium sp-2300500 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2554,2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium | sp-2396 | Corynebacterium | sp-2396 | 2.44 | Corynebacterium | sp | 1.92 | A | 1 | 1 | 0 | Corynebacterium sp-2396 | Corynebacterium sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium | sp-243-z-3 | Corynebacterium | sp-243-z-3 | 2.57 | Corynebacterium | resistens | 1.78 | A | 1 | 1 | 0 | Corynebacterium sp-243-z-3 | Corynebacterium resistens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | sp-257-y-4 | Corynebacterium | sp-257-y-4 | 2.57 | Proteus | mirabilis | 1.36 | A | 1 | 1 | 0 | Corynebacterium sp-257-y-4 | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | sp-284z-2 | Corynebacterium | sp-284z-2 | 2.5 | Proteus | mirabilis | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-284z-2 | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium | sp-CVUAS-11664 | Corynebacterium | sp-CVUAS-11664 | 2.74 | Pseudomonas | aeruginosa | 1.32 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-11664 | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | sp-CVUAS-1888 | Corynebacterium | sp-CVUAS-1888 | 2.62 | Corynebacterium | striatum | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-1888 | Corynebacterium striatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1721,3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium | sp-CVUAS-2484.4 | Corynebacterium | sp-CVUAS-2484.4 | 2.61 | Corynebacterium | sp-CVUAS-2484.4 | 2.11 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-2484.4 | Corynebacterium sp-CVUAS-2484.4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2484,4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388 = 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium | sp-CVUAS-2484.4 | Corynebacterium | sp-CVUAS-2484.4 | 2.37 | Corynebacterium | sp-CVUAS-2484.4 | 1.99 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-2484.4 | Corynebacterium sp-CVUAS-2484.4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp-CVUAS-30404 | 2.49 | Corynebacterium | sp-CVUAS-30404 | 2.27 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-30404 | Corynebacterium sp-CVUAS-30404 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp-CVUAS-30404 | 2.12 | Corynebacterium | sp-CVUAS-30404 | 2.04 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-30404 | Corynebacterium sp-CVUAS-30404 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium | sp-CVUAS-31533 | Corynebacterium | sp-CVUAS-31533 | 2.4 | Corynebacterium | humireducens | 2.22 | B | 0 | 0 | 1 | Corynebacterium sp-CVUAS-31533 | Corynebacterium humireducens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32278,2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium | sp-CVUAS-32278.2 | Corynebacterium | sp-CVUAS-32278.2 | 2.72 | Shewanella | sp-RKVR22 | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-32278.2 | Shewanella sp-RKVR22 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium | sp-CVUAS-874,4 | Corynebacterium | sp-CVUAS-874,4 | 2.14 | Corynebacterium | sp-CVUAS-874,4 | 1.98 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-874,4 | Corynebacterium sp-CVUAS-874,4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 874,4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium | sp-CVUAS-874,4 | Corynebacterium | sp-CVUAS-874,4 | 2.48 | Corynebacterium | sp-CVUAS-874,4 | 2.32 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-874,4 | Corynebacterium sp-CVUAS-874,4 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 21TRD0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-LALLF-21TRB0553 | Corynebacterium | sp-LALLF-20TRB0553 | 2.39 | Cryptococcus | neoformans | 1.38 | A | 0 | 0 | 1 | Corynebacterium sp-LALLF-20TRB0553 | Cryptococcus neoformans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium | sp-LHL-0497 | Corynebacterium | sp-LHL-0497 | 2.22 | Haemophilus | influenzae | 1.26 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-0497 | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | sp-LHL-0558 | Corynebacterium | sp-LHL-0558 | 2.37 | Nocardia | cyriacigeorgica | 1.32 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-0558 | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium | sp-LHL-1322 | Corynebacterium | sp-LHL-1322 | 2.58 | Agromyces | humatus | 1.41 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-1322 | Agromyces humatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | stationis | Corynebacterium | stationis_JRA_G+R-52 | 2.37 | Corynebacterium | stationis | 2.27 | B | 0 | 0 | 1 | Corynebacterium stationis_JRA_G+R-52 | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium | testudinoris | Corynebacterium | testudinoris | 2.56 | Bacteroides | fragilis | 1.29 | A | 1 | 1 | 0 | Corynebacterium testudinoris | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 2823 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.31 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.6 | Corynebacterium | ulcerans | 2.59 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.53 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.34 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6522,4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5437,3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium | variabile | Corynebacterium | variabile | 2.17 | Corynebacterium | variabile | 1.62 | A | 1 | 1 | 0 | Corynebacterium variabile | Corynebacterium variabile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium | xerosis | Corynebacterium | xerosis | 2.39 | Corynebacterium | xerosis | 1.57 | A | 1 | 1 | 0 | Corynebacterium xerosis | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium | xerosis-group | Corynebacterium | xerosis-group | 2.33 | Corynebacterium | xerosis | 1.58 | A | 1 | 1 | 0 | Corynebacterium xerosis-group | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2379,2 | 3; sequenced | EFex | 2021-04-08 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.29 | Gordonia | paraffinivorans | 2.27 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.27 | Gordonia | paraffinivorans | 2.14 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia | rubripertincta | Gordonia | rubripertincta | 2.61 | Gordonia | rubripertincta | 2.25 | A | 1 | 1 | 0 | Gordonia rubripertincta | Gordonia rubripertincta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10185,3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium | smegmatis | Mycolicibacterium | smegmatis | 2.32 | Tsukamurella | sp | 1.49 | A | 1 | 1 | 0 | Mycolicibacterium smegmatis | Tsukamurella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31296,3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia | cyriacigeorgica | Nocardia | cyriacigeorgica | 2.45 | Nocardia | cyriacigeorgica | 2.28 | A | 1 | 1 | 0 | Nocardia cyriacigeorgica | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia | farcinica | Nocardia | farcinica | 2.36 | Nocardia | kroppenstedtii | 2.26 | B | 0 | 0 | 1 | Nocardia farcinica | Nocardia kroppenstedtii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.46 | Rhodococcus | equi | 2.42 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.41 | Rhodococcus | equi | 2.35 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1447,3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella | sp | Tsukamurella | sp | 2.44 | Tsukamurella | tyrosinosolvens | 1.54 | A | 1 | 1 | 0 | Tsukamurella sp | Tsukamurella tyrosinosolvens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2895,3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium | luteolum | Brevibacterium | luteolum | 2.41 | Brevibacterium | luteolum | 1.92 | A | 1 | 1 | 0 | Brevibacterium luteolum | Brevibacterium luteolum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas | flavigena | Cellulomonas | flavigena | 2.76 | Bacteroides | fragilis | 1.43 | A | 1 | 1 | 0 | Cellulomonas flavigena | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium | sp-234z-2 | Brachybacterium | sp-234z-2 | 2.32 | Glutamicibacter | protophormiae | 1.31 | A | 1 | 1 | 0 | Brachybacterium sp-234z-2 | Glutamicibacter protophormiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium | sp-CVUAS-30560 | Brachybacterium | sp-CVUAS-30560 | 2.66 | Raoultella | ornithinolytica | 1.39 | A | 1 | 1 | 0 | Brachybacterium sp-CVUAS-30560 | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus | sedentarius | Kytococcus | sedentarius | 2.1 | Kytococcus | sedentarius | 1.84 | A | 1 | 1 | 0 | Kytococcus sedentarius | Kytococcus sedentarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus | congolensis | Dermatophilus | congolensis | 2.41 | Dermatophilus | congolensis | 2.19 | A | 1 | 1 | 0 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus | congolensis | Dermatophilus | congolensis | 2.65 | Dermatophilus | congolensis | 2.48 | A | 1 | 1 | 0 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium | flaccumfaciens | Curtobacterium | flaccumfaciens | 2.42 | Curtobacterium | flaccumfaciens | 2.32 | A | 1 | 1 | 0 | Curtobacterium flaccumfaciens | Curtobacterium flaccumfaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia | sp-MUVA-A-508948 | Leifsonia | sp-MUVA-A-508948 | 2.62 | Staphylococcus | aureus | 1.46 | A | 1 | 1 | 0 | Leifsonia sp-MUVA-A-508948 | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium | sp | Microbacterium | phyllosphaerae | 1.84 | Microbacterium | foliorum | 1.84 | B | 0 | 0 | 1 | Microbacterium phyllosphaerae | Microbacterium foliorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10353 DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.09.2011 L8 | Glutamicibacter | nicotianae | Lactobacillus | pentosus | 1.33 | Burkholderia | cepacia | 1.26 | B | 0 | 0 | 1 | Lactobacillus pentosus | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter | protophormiae | Glutamicibacter | protophormiae | 2.55 | Lactobacillus | sakei | 1.37 | A | 1 | 1 | 0 | Glutamicibacter protophormiae | Lactobacillus sakei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria | atrinae | Kocuria | atrinae | 2.27 | Kocuria | atrinae | 2 | A | 1 | 1 | 0 | Kocuria atrinae | Kocuria atrinae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 56a-5 | 3; sequenced | OSextr | 2020-05-20 | Kocuria | gwangalliensis | Kocuria | gwangalliensis | 2.73 | Kocuria | carniphila | 1.64 | A | 1 | 1 | 0 | Kocuria gwangalliensis | Kocuria carniphila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 20032 | 3; type strain | eDT | 2015-07-09 | Kocuria | kristinae | Rothia | kristinae | 2.26 | Rothia | kristinae | 2.25 | A | 0 | 0 | 1 | Rothia kristinae | Rothia kristinae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2014 E22 | Kocuria | palustris | Kocuria | palustris | 2.38 | Kocuria | palustris | 2.33 | A | 1 | 1 | 0 | Kocuria palustris | Kocuria palustris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015 K12 | Micrococcus | luteus | Micrococcus | luteus | 2.37 | Micrococcus | luteus | 2.24 | A | 1 | 1 | 0 | Micrococcus luteus | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10432 DSM 20127 | 3; type strain | unknown | 08.09.2011, L13, [0034] | Sinomonas | atrocyanea | Sinomonas | atrocyanea | 2.52 | Sinomonas | atrocyanea | 2.24 | A | 1 | 1 | 0 | Sinomonas atrocyanea | Sinomonas atrocyanea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium | cellulans | Cellulosimicrobium | cellulans | 2.05 | Cellulosimicrobium | cellulans | 1.97 | A | 1 | 1 | 0 | Cellulosimicrobium cellulans | Cellulosimicrobium cellulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 749,2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter | inulinus | Sanguibacter | inulinus | 2.51 | Sanguibacter | inulinus | 2.46 | A | 1 | 1 | 0 | Sanguibacter inulinus | Sanguibacter inulinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A281_1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2012 N7 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 1.98 | Alkalihalobacillus | clausii | 1.69 | B | 0 | 0 | 1 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31071,2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 2.23 | Alkalihalobacillus | clausii | 1.69 | A | 1 | 1 | 0 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.02.2015 N6 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 1.71 | Alkalihalobacillus | clausii | 1.7 | B | 0 | 0 | 1 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Oman 24 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2012 N9 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 1.6 | Alkalihalobacillus | clausii | 1.57 | B | 0 | 0 | 1 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 13.02.2015, K10, [0034] | Bacillus | alveayuensis | Bacillus | alveayuensis | 2.52 | Peribacillus | psychrosaccharolyticus | 1.35 | A | 1 | 1 | 0 | Bacillus alveayuensis | Peribacillus psychrosaccharolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 015 | 3; sequenced | DT | 2016-11-30 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.33 | Bacillus | siamensis | 1.73 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus siamensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 016 | 3; sequenced | DT | 2016-11-30 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 1.93 | Bacillus | amyloliquefaciens | 1.58 | B | 0 | 0 | 1 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7 B159 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 01.03.2016 E10 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.23 | Bacillus | amyloliquefaciens | 2.06 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCA/TA2 (A) | 16.02.2015, K16, [0034] | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.24 | Bacillus | siamensis | 2.01 | B | 0 | 0 | 1 | Bacillus amyloliquefaciens | Bacillus siamensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7 | 3; type strain | DT | 2014-01-29 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.15 | Bacillus | amyloliquefaciens | 2.15 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A10 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E13 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.18 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A12 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E15 | Bacillus | anthracis | Bacillus | cereus | 2.18 | Bacillus | cereus | 1.95 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A18 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.04.2006; H12 | Bacillus | anthracis | Bacillus | cereus | 2.15 | Bacillus | cereus | 2.08 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A25 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.04.2006; H17 | Bacillus | anthracis | Bacillus | cereus | 1.76 | Bacillus | cereus | 1.68 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A30 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.04.2006; H21 | Bacillus | anthracis | Bacillus | cereus | 1.9 | Bacillus | thuringiensis | 1.7 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A4 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E07 | Bacillus | anthracis | Bacillus | cereus | 1.88 | Bacillus | cereus | 1.85 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A5 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E08 | Bacillus | anthracis | Bacillus | cereus | 2.02 | Bacillus | cereus | 2.02 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A7 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E10 | Bacillus | anthracis | Bacillus | cereus | 2.17 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A9 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E12 | Bacillus | anthracis | Bacillus | cereus | 2.2 | Bacillus | cereus | 2.08 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 14185 A2 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E05 | Bacillus | anthracis | Bacillus | cereus | 2.02 | Bacillus | cereus | 1.97 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A269-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.04.2012; H15 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 1.94 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.05.2015; I20 | Bacillus | anthracis | Bacillus | cereus | 2.25 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A100 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; E16 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 2 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A101 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E08 | Bacillus | anthracis | Bacillus | cereus | 2.01 | Bacillus | cereus | 1.91 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A102 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E09 | Bacillus | anthracis | Bacillus | cereus | 2.14 | Bacillus | cereus | 2.01 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A103 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E10 | Bacillus | anthracis | Bacillus | cereus | 1.98 | Bacillus | cereus | 1.96 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A106 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2016; E17 | Bacillus | anthracis | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A107 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2016; C19 | Bacillus | anthracis | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.19 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A108 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E13 | Bacillus | anthracis | Bacillus | cereus | 2.01 | Bacillus | cereus | 1.97 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A11 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.05.2015; C08 | Bacillus | anthracis | Bacillus | cereus | 1.59 | Bacillus | thuringiensis | 1.58 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A110 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.11.2015; F11 | Bacillus | anthracis | Bacillus | cereus | 1.92 | Bacillus | cereus | 1.89 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A111 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E15 | Bacillus | anthracis | Bacillus | cereus | 2.07 | Bacillus | cereus | 2.03 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A112 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.02.2016; F09 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 1.98 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A112-1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.02.2013; C19 | Bacillus | anthracis | Bacillus | cereus | 1.83 | Bacillus | thuringiensis | 1.78 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A113 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E04 | Bacillus | anthracis | Bacillus | cereus | 2.06 | Bacillus | cereus | 2.03 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A118 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H13 | Bacillus | anthracis | Bacillus | cereus | 2.16 | Bacillus | cereus | 2.08 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A120 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H15 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.17 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A122 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H17 | Bacillus | anthracis | Bacillus | cereus | 2.15 | Bacillus | cereus | 1.98 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A123 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H18 | Bacillus | anthracis | Bacillus | cereus | 2.07 | Bacillus | cereus | 1.91 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A124 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H19 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | cereus | 1.98 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A125 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H20 | Bacillus | anthracis | Bacillus | cereus | 2.06 | Bacillus | cereus | 1.79 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A126 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; H21 | Bacillus | anthracis | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A127 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.11.2015; F22 | Bacillus | anthracis | Bacillus | cereus | 1.95 | Bacillus | cereus | 1.92 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A128 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.02.2016; H13 | Bacillus | anthracis | Bacillus | cereus | 2.16 | Bacillus | cereus | 2.03 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A130 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.11.2015; K08 | Bacillus | anthracis | Bacillus | cereus | 1.88 | Bacillus | cereus | 1.88 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A131 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H13 | Bacillus | anthracis | Bacillus | cereus | 1.82 | Bacillus | cereus | 1.8 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A132 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2006; K21 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | cereus | 2.07 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A133 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.02.2016; H20 | Bacillus | anthracis | Bacillus | cereus | 2.14 | Bacillus | cereus | 2.09 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A135 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H14 | Bacillus | anthracis | Bacillus | cereus | 2.12 | Bacillus | cereus | 1.91 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A136 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2006; N07 | Bacillus | anthracis | Bacillus | cereus | 2.02 | Bacillus | thuringiensis | 1.55 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A137 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.11.2015; L07 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | cereus | 2.01 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A138-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.03.2013; F14 | Bacillus | anthracis | Bacillus | cereus | 1.96 | Bacillus | cereus | 1.72 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A15 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.06.2015; H04 | Bacillus | anthracis | Bacillus | cereus | 2.08 | Bacillus | cereus | 2 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A16 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.06.2015; H06 | Bacillus | anthracis | Bacillus | cereus | 2.11 | Bacillus | cereus | 2.02 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A173-2 (Ringversuch A938-2) | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.05.2017; C04 | Bacillus | anthracis | Bacillus | cereus | 2.18 | Bacillus | cereus | 1.96 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; L09 | Bacillus | anthracis | Bacillus | cereus | 2 | Bacillus | cereus | 1.92 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A22 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I07 | Bacillus | anthracis | Bacillus | cereus | 1.94 | Bacillus | cereus | 1.83 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A23 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I09 | Bacillus | anthracis | Bacillus | cereus | 1.8 | Bacillus | cereus | 1.79 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A24 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I11 | Bacillus | anthracis | Bacillus | cereus | 2.01 | Bacillus | cereus | 1.93 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A27 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; K16 | Bacillus | anthracis | Bacillus | cereus | 2.23 | Bacillus | cereus | 2.21 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A28 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; K18 | Bacillus | anthracis | Bacillus | cereus | 2.24 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I20 | Bacillus | anthracis | Bacillus | cereus | 2.14 | Bacillus | cereus | 2.03 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A294 1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; H10 | Bacillus | anthracis | Bacillus | cereus | 1.82 | Bacillus | cereus | 1.8 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A3 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.05.2015; H08 | Bacillus | anthracis | Bacillus | cereus | 2.1 | Bacillus | cereus | 2.06 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A300-1a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.03.2013; N20 | Bacillus | anthracis | Bacillus | cereus | 1.81 | Bacillus | cereus | 1.62 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A300-1b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.03.2013; I04 | Bacillus | anthracis | Bacillus | cereus | 1.77 | Bacillus | cereus | 1.73 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A300-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.10.2012; E16 | Bacillus | anthracis | Bacillus | pseudomycoides | 1.69 | Janthinobacterium | lividum | 1.57 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Janthinobacterium lividum | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A300-2b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.03.2013; F10 | Bacillus | anthracis | Bacillus | cereus | 2.02 | Bacillus | cereus | 1.85 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A302-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.03.2013; C13 | Bacillus | anthracis | Bacillus | cereus | 1.94 | Bacillus | cereus | 1.86 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A306-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.03.2013; E19 | Bacillus | anthracis | Bacillus | cereus | 1.74 | Bacillus | cereus | 1.68 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A315-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.03.2013; F05 | Bacillus | anthracis | Bacillus | cereus | 1.68 | Bacillus | cereus | 1.68 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A315-2 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.03.2013; E16 | Bacillus | anthracis | Bacillus | cereus | 1.58 | Bacillus | pseudomycoides | 1.5 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A32 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; K22 | Bacillus | anthracis | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.04 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A33 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; L13 | Bacillus | anthracis | Bacillus | cereus | 1.86 | Bacillus | cereus | 1.81 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A34 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L06 | Bacillus | anthracis | Bacillus | cereus | 2.04 | Bacillus | cereus | 2.02 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A35 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.04.2006; K19 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.09 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A36 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.04.2006; K20 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 2.02 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A37 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L10 | Bacillus | anthracis | Bacillus | cereus | 1.95 | Bacillus | cereus | 1.93 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A38 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; L15 | Bacillus | anthracis | Bacillus | cereus | 2.08 | Bacillus | cereus | 2.03 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A39 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L14 | Bacillus | anthracis | Bacillus | cereus | 2.16 | Bacillus | cereus | 2.08 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A40 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.04.2006; N06 | Bacillus | anthracis | Bacillus | cereus | 1.92 | Bacillus | cereus | 1.88 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A41 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; N18 | Bacillus | anthracis | Bacillus | cereus | 2.14 | Bacillus | cereus | 2.13 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A42 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L20 | Bacillus | anthracis | Bacillus | cereus | 2.02 | Bacillus | cereus | 1.99 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A43 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L21 | Bacillus | anthracis | Bacillus | cereus | 1.92 | Bacillus | cereus | 1.89 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A44 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; E03 | Bacillus | anthracis | Bacillus | mycoides | 1.99 | Bacillus | cereus | 1.96 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A45 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; E06 | Bacillus | anthracis | Bacillus | cereus | 1.88 | Bacillus | cereus | 1.85 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A46 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; E07 | Bacillus | anthracis | Bacillus | cereus | 1.96 | Bacillus | cereus | 1.9 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A49 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; C12 | Bacillus | anthracis | Bacillus | cereus | 2.03 | Bacillus | mycoides | 1.96 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A58 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; E13 | Bacillus | anthracis | Bacillus | cereus | 2.04 | Bacillus | cereus | 2.03 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A59 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; E15 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 2 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A6 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.05.2015; F13 | Bacillus | anthracis | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A60 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; C17 | Bacillus | anthracis | Bacillus | cereus | 1.96 | Bacillus | cereus | 1.96 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A61 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; C20 | Bacillus | anthracis | Bacillus | cereus | 2.17 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A62 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; C21 | Bacillus | anthracis | Bacillus | cereus | 2.14 | Bacillus | cereus | 2.13 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A63 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.04.2006; E14 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 1.95 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A64 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; H05 | Bacillus | anthracis | Bacillus | cereus | 2.09 | Bacillus | cereus | 2.03 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A65 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; F07 | Bacillus | anthracis | Bacillus | cereus | 1.68 | Bacillus | cereus | 1.65 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A66 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; F09 | Bacillus | anthracis | Bacillus | cereus | 2.16 | Bacillus | cereus | 2.07 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A67 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; H12 | Bacillus | anthracis | Bacillus | cereus | 2.01 | Bacillus | cereus | 1.99 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A68 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2006; H13 | Bacillus | anthracis | Bacillus | cereus | 2.21 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A69 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; F16 | Bacillus | anthracis | Bacillus | cereus | 2.16 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A72 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2006; H17 | Bacillus | anthracis | Bacillus | cereus | 2.08 | Bacillus | cereus | 1.94 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A73 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2006; H18 | Bacillus | anthracis | Bacillus | cereus | 2.3 | Bacillus | cereus | 2.08 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A74 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; I07 | Bacillus | anthracis | Bacillus | cereus | 2.06 | Bacillus | cereus | 2.05 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A8 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.05.2015; H18 | Bacillus | anthracis | Bacillus | cereus | 2.2 | Bacillus | cereus | 2.17 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A80 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; K09 | Bacillus | anthracis | Bacillus | cereus | 1.82 | Bacillus | cereus | 1.78 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A81 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; K12 | Bacillus | anthracis | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A82 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; K13 | Bacillus | anthracis | Bacillus | cereus | 2.11 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A83 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; K17 | Bacillus | anthracis | Bacillus | cereus | 2.11 | Bacillus | cereus | 2.09 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A84 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; N20 | Bacillus | anthracis | Bacillus | cereus | 2.07 | Bacillus | cereus | 2.05 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A85 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; K19 | Bacillus | anthracis | Bacillus | cereus | 1.72 | Janthinobacterium | lividum | 1.5 | B | 0 | 0 | 1 | Bacillus cereus | Janthinobacterium lividum | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A86 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2016; E09 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | cereus | 2.13 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A87 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; K21 | Bacillus | anthracis | Bacillus | cereus | 1.9 | Bacillus | cereus | 1.9 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A88 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; N05 | Bacillus | anthracis | Bacillus | cereus | 2.22 | Bacillus | cereus | 2.12 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A88 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; N04 | Bacillus | anthracis | Bacillus | cereus | 2.18 | Bacillus | cereus | 2.11 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A89 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; N07 | Bacillus | anthracis | Bacillus | cereus | 2.21 | Bacillus | thuringiensis | 1.89 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A89 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; C10 | Bacillus | anthracis | Bacillus | cereus | 2.03 | Bacillus | cereus | 2.02 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | A90 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; C09 | Bacillus | anthracis | Bacillus | cereus | 1.95 | Bacillus | cereus | 1.92 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A90 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; E12 | Bacillus | anthracis | Bacillus | cereus | 1.98 | Bacillus | cereus | 1.87 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A91 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E04 | Bacillus | anthracis | Bacillus | cereus | 1.8 | Bacillus | cereus | 1.79 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A92 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; C13 | Bacillus | anthracis | Bacillus | cereus | 1.93 | Bacillus | cereus | 1.87 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A93 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; E15 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 2.04 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | AF 28 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H07 | Bacillus | anthracis | Bacillus | cereus | 2.12 | Bacillus | thuringiensis | 1.84 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | AF 29 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H08 | Bacillus | anthracis | Bacillus | cereus | 2.15 | Bacillus | cereus | 1.85 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | AF 31 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H09 | Bacillus | anthracis | Bacillus | cereus | 2.18 | Bacillus | cereus | 1.98 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | AF 71 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H10 | Bacillus | anthracis | Bacillus | cereus | 2.32 | Bacillus | cereus | 1.88 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | AF 96 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H11 | Bacillus | anthracis | Bacillus | cereus | 2.09 | Bacillus | cereus | 1.83 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E06 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 2.02 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 14578 A70 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; H18 | Bacillus | anthracis | Bacillus | cereus | 2.15 | Bacillus | mycoides | 2.04 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; E04 | Bacillus | anthracis | Bacillus | cereus | 2.36 | Bacillus | cereus | 2.25 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | SA 20 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H12 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | cereus | 1.84 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | Sterne | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015; E03 | Bacillus | anthracis | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.06 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | vollum | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2007; E22 | Bacillus | anthracis | Bacillus | cereus | 1.89 | Bacillus | thuringiensis | 1.67 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Wirt | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015; C05 | Bacillus | anthracis | Bacillus | cereus | 2.14 | Bacillus | cereus | 1.86 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 9372 WIS-St. Nr. 39 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007 E10 | Bacillus | atrophaeus | Bacillus | atrophaeus | 1.9 | Bacillus | atrophaeus | 1.69 | B | 0 | 0 | 1 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 004 | 2; 0 | DT | 2014-01-22 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.14 | Bacillus | atrophaeus | 2.14 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 4 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 E7 | Bacillus | atrophaeus | Bacillus | atrophaeus | 1.96 | Enterococcus | faecalis | 1.56 | B | 0 | 0 | 1 | Bacillus atrophaeus | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 93 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 N3 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.13 | Bacillus | atrophaeus | 2.11 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2277 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E13, [0034] | Bacillus | atrophaeus | Bacillus | atrophaeus | 1.86 | Bacillus | atrophaeus | 1.79 | B | 0 | 0 | 1 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 675 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 13.04.2007, H17, [0034] | Bacillus | atrophaeus | Bacillus | atrophaeus | 1.51 | Bacillus | atrophaeus | 1.51 | B | 0 | 0 | 1 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 L8 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.43 | Bacillus | atrophaeus | 2.36 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 44 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N11 | Bacillus | badius | Bacillus | badius | 2.06 | Bacillus | badius | 1.99 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H20 | Bacillus | badius | Bacillus | badius | 2.45 | Bacillus | badius | 2.26 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004439-1 | 2; 0 | DT | 2016-04-12 | Bacillus | cereus | Bacillus | thuringiensis | 2.51 | Bacillus | thuringiensis | 2.51 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | 161004439-2 | 2; 0 | DT | 2016-04-12 | Bacillus | cereus | Bacillus | cereus | 2.34 | Bacillus | thuringiensis | 2.21 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | 161005557-1 | 2; 0 | DT | 2016-08-11 | Bacillus | cereus | Bacillus | cereus | 2.24 | Bacillus | cereus | 2.18 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | 171012589-5 | 2; 0 | DT | 2017-09-19 | Bacillus | cereus | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.21 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | 181015909-7 | 2; 0 | DT | 2018-11-02 | Bacillus | cereus | Bacillus | cereus | 2.13 | Bacillus | cereus | 2.07 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 11778 LHL | 3; public strain collection | DT | 2015-06-23 | Bacillus | cereus | Bacillus | cereus | 2.21 | Bacillus | cereus | 2.19 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 14579 IGV | 3; type strain | DT | 2016-04-06 | Bacillus | cereus | Bacillus | cereus | 2.79 | Bacillus | thuringiensis | 2.43 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 4312 IGV | 3; public strain collection | DT | 2016-04-13 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | cereus | 2.2 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 4313 IGV | 3; public strain collection | DT | 2016-04-06 | Bacillus | cereus | Bacillus | cereus | 2.32 | Bacillus | cereus | 2.21 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 11778 IGV | 3; public strain collection | DT | 2016-02-03 | Bacillus | cereus | Bacillus | cereus | 2.25 | Bacillus | cereus | 2.21 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | AaCol | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2009 H3 | Bacillus | cereus | Bacillus | cereus | 1.83 | Bacillus | cereus | 1.82 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SK1b | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2009 D21 | Bacillus | cereus | Bacillus | cereus | 2.05 | Bacillus | cereus | 1.94 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B194 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H10 | Bacillus | cereus | Bacillus | cereus | 2.62 | Bacillus | cereus | 2.24 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B195 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H11 | Bacillus | cereus | Bacillus | cereus | 2.6 | Bacillus | cereus | 2.14 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B196 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H12 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | cereus | 2.05 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | CVUAS 34551 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus | cereus | Bacillus | cereus | 2.49 | Bacillus | cereus | 2.44 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | CVUAS 34552 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus | cereus | Bacillus | thuringiensis | 2.01 | Bacillus | cereus | 2 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | 301101RA0432 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2007; E20 | Bacillus | cereus | Bacillus | cereus | 2.41 | Bacillus | cereus | 1.89 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | 43 Paciflor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; E16 | Bacillus | cereus | Bacillus | cereus | 1.65 | Bacillus | cereus | 1.62 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 8 Toyocerin | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.06.2006 E14 | Bacillus | cereus | Bacillus | cereus | 2.04 | Bacillus | cereus | 1.99 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 10876 B270 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I09 | Bacillus | cereus | Bacillus | cereus | 2.25 | Bacillus | thuringiensis | 2.11 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 11778 RKI | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 23.10.2015, E6, [0034] | Bacillus | cereus | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.2 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 11778 B23 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.08.2006 E21 | Bacillus | cereus | Bacillus | cereus | 2.5 | Bacillus | cereus | 2.26 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 31 B69 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 06.02.2015 I21 | Bacillus | cereus | Bacillus | cereus | 2.31 | Bacillus | cereus | 2.28 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 14579 RKI | 3; type strain | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E7, [0034] | Bacillus | cereus | Bacillus | cereus | 2.13 | Bacillus | cereus | 1.94 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | cereus | 2.33 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 33018 B245 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2016; H22 | Bacillus | cereus | Bacillus | thuringiensis | 1.92 | Bacillus | thuringiensis | 1.88 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33019 B244 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.04.2015; H15 | Bacillus | cereus | Bacillus | cereus | 2.45 | Bacillus | thuringiensis | 2.21 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 6464 B331 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; K14 | Bacillus | cereus | Bacillus | thuringiensis | 2.16 | Bacillus | thuringiensis | 2.12 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 7004 B284 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; L06 | Bacillus | cereus | Bacillus | cereus | 2.5 | Bacillus | cereus | 2.29 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B10 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 1.12.2014 C18 | Bacillus | cereus | Bacillus | thuringiensis | 2.08 | Bacillus | cereus | 1.93 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B108 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 15.08.2006 N6 | Bacillus | cereus | Bacillus | cereus | 2.42 | Bacillus | cereus | 2.38 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B109 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.03.2015 N20 | Bacillus | cereus | Bacillus | cereus | 2.24 | Bacillus | cereus | 2.03 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B110 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 15.08.2006 N8 | Bacillus | cereus | Bacillus | cereus | 2.29 | Bacillus | cereus | 2.12 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B146 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 27.03.2015 I11 | Bacillus | cereus | Bacillus | thuringiensis | 2.2 | Bacillus | cereus | 2.16 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B16 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.12.2014 F6 | Bacillus | cereus | Bacillus | thuringiensis | 2.12 | Bacillus | cereus | 1.75 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B197 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.04.2015; H22 | Bacillus | cereus | Bacillus | cereus | 2.49 | Bacillus | cereus | 2.11 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B198 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2015; K09 | Bacillus | cereus | Bacillus | pseudomycoides | 2.38 | Bacillus | cereus | 2.29 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B199 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H15 | Bacillus | cereus | Bacillus | cereus | 2.32 | Bacillus | cereus | 2.09 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B200 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H16 | Bacillus | cereus | Bacillus | cereus | 2.54 | Bacillus | cereus | 2.09 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B232 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.08.2006; K10 | Bacillus | cereus | Bacillus | cereus | 2.35 | Bacillus | thuringiensis | 2.25 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B24 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H08 | Bacillus | cereus | Bacillus | cereus | 2.09 | Bacillus | cereus | 2.09 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B252 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; K03 | Bacillus | cereus | Bacillus | thuringiensis | 2.2 | Bacillus | thuringiensis | 1.96 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B253 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H16 | Bacillus | cereus | Bacillus | thuringiensis | 1.93 | Bacillus | thuringiensis | 1.88 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B254 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H17 | Bacillus | cereus | Bacillus | cereus | 1.9 | Bacillus | cereus | 1.84 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B255 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; F20 | Bacillus | cereus | Bacillus | cereus | 2.45 | Bacillus | cereus | 2.11 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B256 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; K06 | Bacillus | cereus | Bacillus | thuringiensis | 2.07 | Bacillus | thuringiensis | 2.06 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B257 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; K05 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | thuringiensis | 2.22 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B258 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; I08 | Bacillus | cereus | Bacillus | cereus | 2.39 | Bacillus | thuringiensis | 2.21 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B259 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; K09 | Bacillus | cereus | Bacillus | cereus | 2.43 | Bacillus | thuringiensis | 2.27 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B260 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; I07 | Bacillus | cereus | Bacillus | cereus | 2.1 | Bacillus | cereus | 2.05 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B261 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; I14 | Bacillus | cereus | Bacillus | cereus | 2.32 | Bacillus | cereus | 1.9 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B262 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.09.2006; H06 | Bacillus | cereus | Bacillus | cereus | 2.25 | Bacillus | cereus | 2.01 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B272 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I13 | Bacillus | cereus | Bacillus | cereus | 2.27 | Bacillus | thuringiensis | 2.04 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B273 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I15 | Bacillus | cereus | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.11 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B274 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; K17 | Bacillus | cereus | Bacillus | mycoides | 2.05 | Bacillus | pseudomycoides | 1.97 | A | 0 | 0 | 1 | Bacillus mycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B275 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; K19 | Bacillus | cereus | Bacillus | cereus | 2.19 | Bacillus | thuringiensis | 2.05 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B277 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I20 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | cereus | 2.15 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B278 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; L04 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | thuringiensis | 2.12 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B282 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; N07 | Bacillus | cereus | Bacillus | cereus | 2.38 | Bacillus | thuringiensis | 2 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B293 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.09.2006; K06 | Bacillus | cereus | Bacillus | cereus | 2.44 | Bacillus | thuringiensis | 2.27 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B294 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; L14 | Bacillus | cereus | Bacillus | cereus | 2.38 | Bacillus | cereus | 2.32 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B295 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; L16 | Bacillus | cereus | Bacillus | cereus | 2.56 | Bacillus | cereus | 2.07 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B296 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; N18 | Bacillus | cereus | Bacillus | cereus | 2.62 | Bacillus | cereus | 2.25 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B297 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C07 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | cereus | 2.18 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B298 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; E09 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | cereus | 2.13 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B299 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.09.2006; K12 | Bacillus | cereus | Bacillus | cereus | 2.49 | Bacillus | cereus | 2.09 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B3 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.08.2006 E6 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | cereus | 2.18 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B300 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C14 | Bacillus | cereus | Bacillus | cereus | 2.6 | Bacillus | cereus | 2.22 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B302 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C18 | Bacillus | cereus | Bacillus | cereus | 2.54 | Bacillus | cereus | 2.08 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B303 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C20 | Bacillus | cereus | Bacillus | cereus | 2.59 | Bacillus | cereus | 2.27 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B304 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; E21 | Bacillus | cereus | Bacillus | cereus | 2.38 | Bacillus | cereus | 2.29 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B305 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; H04 | Bacillus | cereus | Bacillus | cereus | 2.53 | Bacillus | cereus | 2.18 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B306 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; H06 | Bacillus | cereus | Bacillus | cereus | 2.56 | Bacillus | cereus | 2.16 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B308 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; H10 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | cereus | 2.25 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B309 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.09.2006; E10 | Bacillus | cereus | Bacillus | cereus | 2.3 | Bacillus | thuringiensis | 2.14 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B310 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; F13 | Bacillus | cereus | Bacillus | cereus | 2.37 | Bacillus | cereus | 2.24 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B311 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; H16 | Bacillus | cereus | Bacillus | cereus | 2.37 | Bacillus | cereus | 2.14 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B316 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; H22 | Bacillus | cereus | Bacillus | cereus | 2.38 | Bacillus | thuringiensis | 2.2 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B33 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.12.2014 I20 | Bacillus | cereus | Bacillus | cereus | 2.54 | Bacillus | cereus | 2.24 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B332 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; I16 | Bacillus | cereus | Bacillus | thuringiensis | 2.22 | Bacillus | thuringiensis | 2.21 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B74 CCM 88 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 14.08.2006 K9 | Bacillus | cereus | Bacillus | thuringiensis | 2.32 | Bacillus | cereus | 2.25 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | BW-B | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015; E07 | Bacillus | cereus | Bacillus | cereus | 2 | Bacillus | cereus | 1.9 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | CCM 99 B249 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H10 | Bacillus | cereus | Bacillus | cereus | 2.2 | Bacillus | cereus | 2.07 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | CVUAS 1107 | 2; 0 | eDT | 27.01.2016 | Bacillus | cereus | Bacillus | pseudomycoides | 2.45 | Bacillus | cereus | 2.41 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | CVUAS 3901 | 3; 0 | eDT | 26.04.2016 | Bacillus | cereus | Bacillus | cereus | 2.5 | Bacillus | cereus | 2.48 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 1274 B247 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; F06 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | thuringiensis | 2.15 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 2299 B190 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.08.2006; H08 | Bacillus | cereus | Bacillus | cereus | 2.31 | Bacillus | cereus | 2.03 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 2302 B192 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.04.2015; F13 | Bacillus | cereus | Bacillus | cereus | 2.39 | Bacillus | cereus | 2.35 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 31 B181 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 30.03.2015 E21 | Bacillus | cereus | Bacillus | cereus | 2.42 | Bacillus | cereus | 2.38 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | DSM 31 | 3; type strain | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | cereus | 2.55 | Bacillus | cereus | 2.43 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 31 RKI | 3; type strain | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 21.09.2015, E13, [0034] | Bacillus | cereus | Bacillus | cereus | 2.31 | Bacillus | cereus | 2.21 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 3101 B237 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2015; C20 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | thuringiensis | 2.23 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 318 B233 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2015; C13 | Bacillus | cereus | Bacillus | cereus | 2.28 | Bacillus | cereus | 2.06 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 336 B234 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.08.2006; K12 | Bacillus | cereus | Bacillus | cereus | 2.31 | Bacillus | cereus | 2.06 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 345 B170 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 30.03.2015 C11 | Bacillus | cereus | Bacillus | cereus | 2.36 | Bacillus | cereus | 2.19 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | DSM 345 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | cereus | 2.43 | Bacillus | cereus | 2.37 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 351 B235 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.08.2006; K13 | Bacillus | cereus | Bacillus | thuringiensis | 2.44 | Bacillus | thuringiensis | 2.29 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 4218 B238 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2015; C21 | Bacillus | cereus | Bacillus | thuringiensis | 2.04 | Bacillus | thuringiensis | 1.96 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 4312 B269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.09.2006; H11 | Bacillus | cereus | Bacillus | cereus | 2.52 | Bacillus | cereus | 2.15 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B329 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; K11 | Bacillus | cereus | Bacillus | cereus | 2.53 | Bacillus | cereus | 2.13 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | DSM 4312 | 3; public strain collection | eDT | 27.01.2016 | Bacillus | cereus | Bacillus | cereus | 2.5 | Bacillus | pseudomycoides | 2.47 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 4312 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E5, [0034] | Bacillus | cereus | Bacillus | cereus | 2.44 | Bacillus | cereus | 2.02 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | DSM 4313 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.62 | Bacillus | cereus | 2.34 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 4490 | 3; public strain collection | in 80% TFA, Zentrifugation und Sterilfiltration, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 21.09.2015, E9, [0034] | Bacillus | cereus | Bacillus | cereus | 2.34 | Bacillus | cereus | 2.29 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 487 B169 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006 E15 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | cereus | 2.13 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 487 B185 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006 H4 | Bacillus | cereus | Bacillus | cereus | 2.43 | Bacillus | cereus | 2.06 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 508 B186 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 7.04.2015 F6 | Bacillus | cereus | Bacillus | cereus | 2.42 | Bacillus | cereus | 2.3 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 609 | 3; strain from public collection | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 21.09.2015, C17, [0034] | Bacillus | cereus | Bacillus | thuringiensis | 2.24 | Bacillus | thuringiensis | 2.19 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 626 B187 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.04.2015; F07 | Bacillus | cereus | Bacillus | cereus | 2.33 | Bacillus | thuringiensis | 2.16 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 679 B242 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.09.2006; E06 | Bacillus | cereus | Bacillus | thuringiensis | 2.23 | Bacillus | cereus | 2.19 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 689 B243 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.09.2006; E07 | Bacillus | cereus | Bacillus | cereus | 2.37 | Bacillus | cereus | 2.08 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | Hohenheim | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2007; K04 | Bacillus | cereus | Bacillus | cereus | 2.06 | Bacillus | cereus | 1.86 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 9139 B289 | 3; reference strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; N07 | Bacillus | cereus | Bacillus | cereus | 2.42 | Bacillus | thuringiensis | 2.19 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | NCIMB 11925 | 3; reference strain collection | DT | 2015-07-22 | Bacillus | cereus | Bacillus | cereus | 2.33 | Bacillus | cereus | 2.22 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | NCTC 10320 B148 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006 E5 | Bacillus | cereus | Bacillus | cereus | 2.55 | Bacillus | cereus | 2.21 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | NCTC 10989 B268 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; N04 | Bacillus | cereus | Bacillus | cereus | 2.29 | Bacillus | cereus | 2.23 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | NCTC 6474 B265 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; K20 | Bacillus | cereus | Bacillus | thuringiensis | 2.44 | Bacillus | thuringiensis | 2.33 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | NCTC 9939 B267 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; I22 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | cereus | 2.1 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | NiMe_1 | 3; 16S, ces-pos | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.59 | Bacillus | cereus | 2.49 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | Paciflor | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2007; K04 | Bacillus | cereus | Bacillus | mycoides | 1.64 | Bacillus | thuringiensis | 1.59 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2009001164 | 3; 0 | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | cereus | 2.43 | Bacillus | cereus | 2.4 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | SL2009005808 | 2; 0 | eDT | 2019-09-23 | Bacillus | cereus | Bacillus | cereus | 2.59 | Bacillus | cereus | 2.45 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | SL2009007190 | 2; 0 | eDT | 2019-09-25 | Bacillus | cereus | Bacillus | cereus | 2.53 | Bacillus | cereus | 2.39 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | SL2014002835 | 2; CytK2 pos, Krist neg | eDT | 2019-09-25 | Bacillus | cereus | Bacillus | cereus | 2.46 | Bacillus | pseudomycoides | 2.32 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | Toyocerin | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; H07 | Bacillus | cereus | Bacillus | cereus | 1.8 | Bacillus | cytotoxicus | 1.72 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cytotoxicus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Wasserprobe | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; H06 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | thuringiensis | 2.23 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WIS-St. Nr. 29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007; E09 | Bacillus | cereus | Bacillus | thuringiensis | 1.68 | Bacillus | thuringiensis | 1.48 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WIS-St. Nr. 311 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007; E12 | Bacillus | cereus | Bacillus | cereus | 2.13 | Bacillus | pseudomycoides | 1.96 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10028 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H07 | Bacillus | cereus | Bacillus | cereus | 2.46 | Bacillus | thuringiensis | 2.16 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10030 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H08 | Bacillus | cereus | Bacillus | cereus | 2.32 | Bacillus | cereus | 2.19 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10286 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; E09 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | cereus | 2.22 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10288 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H10 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | thuringiensis | 2.18 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10441 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H11 | Bacillus | cereus | Bacillus | cereus | 2.14 | Bacillus | cereus | 2.14 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | WSBC 10925 | 3; well defined isolat, ces-pos | eDT | 27.01.2016 | Bacillus | cereus | Bacillus | pseudomycoides | 2.53 | Bacillus | cereus | 2.44 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.03.2017 I18 | Bacillus | ectoiniformans | Bacillus | ectoiniformans | 2.78 | Streptococcus | dysgalactiae | 1.49 | A | 1 | 1 | 0 | Bacillus ectoiniformans | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.11.2017 F14 | Bacillus | halotolerans | Bacillus | mojavensis | 1.96 | Bacillus | mojavensis | 1.83 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.12.2017 I8 | Bacillus | halotolerans | Bacillus | mojavensis | 2.23 | Bacillus | subtilis | 2.1 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.12.2016 [H3] | Bacillus | inaquosorum | Bacillus | subtilis | 2.27 | Bacillus | subtilis | 2.12 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.02.2016 C20 | Bacillus | infantis | Bacillus | infantis | 2.46 | Bacillus | infantis | 2.39 | A | 1 | 1 | 0 | Bacillus infantis | Bacillus infantis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141014442-3 | 3; VETQAS | DT | 2014-11-07 | Bacillus | licheniformis | Bacillus | licheniformis | 2.02 | Bacillus | licheniformis | 1.99 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141015909-3 | 2; 0 | DT | 2014-12-12 | Bacillus | licheniformis | Bacillus | licheniformis | 2.21 | Bacillus | licheniformis | 2.15 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004445 | 2; 0 | DT | 2016-04-12 | Bacillus | licheniformis | Bacillus | licheniformis | 2.09 | Bacillus | licheniformis | 2.05 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161007249-3 | 2; 0 | DT | 2016-06-14 | Bacillus | licheniformis | Bacillus | licheniformis | 2.15 | Bacillus | licheniformis | 2.11 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14580 | 3; type strain | DT | 2015-06-23 | Bacillus | licheniformis | Bacillus | licheniformis | 2.11 | Bacillus | licheniformis | 2.11 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 62 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 H21 | Bacillus | licheniformis | Bacillus | licheniformis | 2.18 | Bacillus | licheniformis | 1.71 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015, F12 | Bacillus | licheniformis | Bacillus | licheniformis | 2.19 | Bacillus | licheniformis | 1.53 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 118 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015, F18 | Bacillus | licheniformis | Bacillus | licheniformis | 2.3 | Bacillus | licheniformis | 1.47 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 72 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.08.2006 K8 | Bacillus | licheniformis | Bacillus | licheniformis | 2.19 | Bacillus | licheniformis | 1.76 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus | licheniformis | Bacillus | licheniformis | 2.17 | Bacillus | licheniformis | 2.15 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B116 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.03.2015 F13 | Bacillus | licheniformis | Bacillus | licheniformis | 1.87 | Bacillus | licheniformis | 1.49 | B | 0 | 0 | 1 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BRLK08_04 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.04.2015 I16 | Bacillus | licheniformis | Bacillus | licheniformis | 2.3 | Bacillus | licheniformis | 1.89 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13 B178 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006; E20 | Bacillus | licheniformis | Bacillus | licheniformis | 2.26 | Bacillus | licheniformis | 1.73 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13 | 3; type strain | eDT | 2016-01-27 | Bacillus | licheniformis | Bacillus | licheniformis | 2.27 | Bacillus | licheniformis | 1.81 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2007 E15 | Bacillus | licheniformis | Bacillus | licheniformis | 1.74 | Bacillus | licheniformis | 1.61 | B | 0 | 0 | 1 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K15 | Bacillus | mojavensis | Bacillus | subtilis | 2.13 | Bacillus | subtilis | 2.08 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F4 | Bacillus | mojavensis | Bacillus | mojavensis | 2.33 | Bacillus | subtilis | 2.27 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 299 | 3; public strain collection | DT | 2016-04-08 | Bacillus | mycoides | Bacillus | mycoides | 2.57 | Bacillus | mycoides | 2.41 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 12826 WIS-St. Nr. 283 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007 E11 | Bacillus | mycoides | Bacillus | cereus | 1.9 | Bacillus | cereus | 1.74 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B103 | 2; 0 | DT | 2015-07-22 | Bacillus | mycoides | Bacillus | cereus | 2.31 | Bacillus | cereus | 2.27 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 104119 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.05.2018 I15 | Bacillus | mycoides | Bacillus | mycoides | 2.24 | Bacillus | mycoides | 2.2 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | DSM 11821 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | mycoides | Bacillus | mycoides | 2.18 | Bacillus | mycoides | 2.17 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 11821 RKI | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 03.05.2007, E4, [0034] | Bacillus | mycoides | Bacillus | mycoides | 1.97 | Bacillus | mycoides | 1.95 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2048 | 3; typestrain | eDT | 09.12.2015 | Bacillus | mycoides | Bacillus | mycoides | 2.45 | Bacillus | mycoides | 2.22 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 2048 RKI | 3; typestrain | EFex [BRU03] | 2022-01-13 | Bacillus | mycoides | Bacillus | mycoides | 2.2 | Bacillus | mycoides | 2.15 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | mycoides | Bacillus | mycoides | 2.04 | Bacillus | mycoides | 2.01 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10204 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K5, [0034] | Bacillus | mycoides | Bacillus | mycoides | 2.43 | Bacillus | mycoides | 2.37 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10207 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K6, [0034] | Bacillus | mycoides | Bacillus | mycoides | 2.31 | Bacillus | mycoides | 2.29 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10277 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H14 | Bacillus | mycoides | Bacillus | mycoides | 2.16 | Bacillus | mycoides | 2.02 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10279 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H15 | Bacillus | mycoides | Bacillus | mycoides | 2.17 | Bacillus | mycoides | 2.11 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10291 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H16 | Bacillus | mycoides | Bacillus | pseudomycoides | 2.15 | Bacillus | pseudomycoides | 1.99 | A | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10295 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K7, [0034] | Bacillus | mycoides | Bacillus | mycoides | 2.05 | Bacillus | mycoides | 1.97 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10360 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H17 | Bacillus | mycoides | Bacillus | pseudomycoides | 2.08 | Bacillus | cereus | 1.87 | A | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 10363 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K8, [0034] | Bacillus | mycoides | Bacillus | mycoides | 1.94 | Bacillus | mycoides | 1.92 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10387 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K9, [0034] | Bacillus | mycoides | Bacillus | mycoides | 2.05 | Bacillus | mycoides | 1.99 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 10897 | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E8, [0034] | Bacillus | pacificus | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.06 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 12759 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007 E10 | Bacillus | paralicheniformis | Staphylococcus | epidermidis | 1.45 | Staphylococcus | carnosus | 1.45 | B | 0 | 0 | 1 | Staphylococcus epidermidis | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.53 | Bacillus | cereus | 2.48 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | DSM 12442 | 3; type strain | eDT | 2016-01-27 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.3 | Bacillus | pseudomycoides | 2.2 | A | 1 | 1 | 0 | Bacillus pseudomycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 12442 RKI | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 03.05.2007, N9, [0034] | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.23 | Bacillus | pseudomycoides | 2.14 | A | 1 | 1 | 0 | Bacillus pseudomycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | DSM 12443 | 3; public strain collection | eDT | 09.12.2015 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.64 | Bacillus | pseudomycoides | 2.4 | A | 1 | 1 | 0 | Bacillus pseudomycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B 17 | 3; sequenced (CVUAS) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F7 | Bacillus | pumilus | Bacillus | pumilus | 2.12 | Bacillus | pumilus | 1.67 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 031 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2006, H6 | Bacillus | pumilus | Bacillus | pumilus | 1.93 | Bacillus | safensis | 1.81 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus safensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 91 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.08.2006 K15 | Bacillus | pumilus | Bacillus | pumilus | 1.69 | Bacillus | altitudinis | 1.63 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B017 | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus | pumilus | Bacillus | pumilus | 2.15 | Bacillus | pumilus | 1.96 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 340 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | pumilus | Bacillus | pumilus | 2.04 | Bacillus | pumilus | 1.95 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13835 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2007, N7 | Bacillus | pumilus | Bacillus | pumilus | 1.83 | Bacillus | pumilus | 1.68 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27 B165 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E7 | Bacillus | pumilus | Bacillus | pumilus | 2.18 | Bacillus | pumilus | 2.02 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27 Neukauf2017 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 , K22 | Bacillus | pumilus | Bacillus | pumilus | 2.27 | Bacillus | pumilus | 2.09 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 27 | 3; type strain | DT | 2014-01-22 | Bacillus | pumilus | Bacillus | pumilus | 2.22 | Bacillus | pumilus | 2.02 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 355 B161 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; N10 | Bacillus | pumilus | Bacillus | pumilus | 1.86 | Bacillus | safensis | 1.66 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus safensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 361 B162 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E6 | Bacillus | pumilus | Bacillus | pumilus | 1.88 | Bacillus | pumilus | 1.82 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 492 B160 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.08.2006, E11 | Bacillus | pumilus | Bacillus | pumilus | 1.74 | Bacillus | pumilus | 1.62 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MM 171 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007 H6 | Bacillus | pumilus | Bacillus | pumilus | 1.6 | Bacillus | pumilus | 1.55 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NVH-819-05 | 3; well described isolate [0017] | DT | 2015-04-21 | Bacillus | pumilus | Bacillus | pumilus | 2.11 | Bacillus | altitudinis | 1.56 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.05.2018, N12 | Bacillus | siamensis | Bacillus | siamensis | 2.68 | Bacillus | amyloliquefaciens | 1.87 | A | 1 | 1 | 0 | Bacillus siamensis | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2016, K18 | Bacillus | sonorensis | Bacillus | sonorensis | 1.95 | Listeria | grayi | 1.56 | B | 0 | 0 | 1 | Bacillus sonorensis | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12369 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.04.2007, K15 | Bacillus | sp-DSM-12369 | Bacillus | sonorensis | 1.53 | Thauera | mechernichensis | 1.39 | B | 0 | 0 | 1 | Bacillus sonorensis | Thauera mechernichensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 6633 B13 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2014; E22 | Bacillus | spizizenii | Bacillus | subtilis | 2.12 | Bacillus | subtilis | 1.85 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 6633 | 3; public strain collection | DT | 2016-02-11 | Bacillus | spizizenii | Bacillus | subtilis | 2.08 | Bacillus | subtilis | 1.96 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 6633 DSM 347 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007 E9 | Bacillus | spizizenii | Bacillus | subtilis | 1.63 | Bacillus | subtilis | 1.54 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1087 B156 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; K19 | Bacillus | spizizenii | Bacillus | subtilis | 2.37 | Bacillus | subtilis | 2.08 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 1087 | 3; public strain collection | DT | 2014-01-22 | Bacillus | spizizenii | Bacillus | subtilis | 2.02 | Bacillus | subtilis | 1.96 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1087 RKI | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 18.12.2017; N5 | Bacillus | spizizenii | Bacillus | subtilis | 2.17 | Bacillus | subtilis | 2.05 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 347 | 3; public strain collection | DT | 2014-01-10 | Bacillus | spizizenii | Bacillus | subtilis | 2.04 | Bacillus | subtilis | 1.58 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 618 | 4; strain copy from public strain collection | eDT | 27.01.2016 | Bacillus | spizizenii | Bacillus | subtilis | 2.09 | Bacillus | subtilis | 2.04 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 020 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2.12.2014, H11 | Bacillus | subtilis | Bacillus | subtilis | 1.88 | Bacillus | siamensis | 1.55 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus siamensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 032 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014, K18 | Bacillus | subtilis | Bacillus | siamensis | 1.99 | Bacillus | amyloliquefaciens | 1.57 | B | 0 | 0 | 1 | Bacillus siamensis | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 013 | 3; sequenced | DT | 2016-11-09 | Bacillus | subtilis | Bacillus | subtilis | 2.27 | Bacillus | subtilis | 1.9 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 130 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015, I10 | Bacillus | subtilis | Bacillus | subtilis | 1.96 | Bacillus | subtilis | 1.65 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 402 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.11.2017 F4 | Bacillus | subtilis | Bacillus | subtilis | 2.45 | Bacillus | subtilis | 2.37 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 6051 B158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; N03 | Bacillus | subtilis | Bacillus | subtilis | 1.9 | Bacillus | subtilis | 1.64 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 061 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.01.2015; F20 | Bacillus | subtilis | Bacillus | subtilis | 1.82 | Bacillus | subtilis | 1.61 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2007; N14 | Bacillus | subtilis | Bacillus | subtilis | 1.98 | Bacillus | subtilis | 1.76 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B129 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; N19 | Bacillus | subtilis | Bacillus | subtilis | 1.87 | Bacillus | siamensis | 1.49 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus siamensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B14 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2006; E14 | Bacillus | subtilis | Bacillus | subtilis | 1.62 | Bacillus | subtilis | 1.46 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B154 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; I18 | Bacillus | subtilis | Bacillus | subtilis | 2.38 | Bacillus | subtilis | 2.29 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2014; H09 | Bacillus | subtilis | Bacillus | subtilis | 2.21 | Bacillus | subtilis | 2.08 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BRLK08_02 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015; I15 | Bacillus | subtilis | Bacillus | subtilis | 2.1 | Bacillus | subtilis | 1.97 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 110 B068 | 3; reference strain collection | DT | 2015-07-22 | Bacillus | subtilis | Bacillus | subtilis | 2.33 | Bacillus | subtilis | 2.15 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 10 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.03.2016 C20 | Bacillus | subtilis | Bacillus | subtilis | 2.28 | Bacillus | subtilis | 2.19 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1092 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.23 | Bacillus | subtilis | 2.2 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | E rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2007; N15 | Bacillus | subtilis | Bacillus | subtilis | 1.99 | Bacillus | subtilis | 1.83 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | I rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.04.2007; E09 | Bacillus | subtilis | Bacillus | subtilis | 1.89 | Clostridium | novyi | 1.51 | B | 0 | 0 | 1 | Bacillus subtilis | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LVM03 LVU 03 | 3; sequenced (LGL) | eDT | 2015-07-09 | Bacillus | subtilis | Bacillus | subtilis | 2.1 | Bacillus | subtilis | 1.94 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MM 508 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007; H11 | Bacillus | subtilis | Bacillus | siamensis | 1.88 | Staphylococcus | fleurettii | 1.55 | B | 0 | 0 | 1 | Bacillus siamensis | Staphylococcus fleurettii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RKI-w2gr | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2008; H16 | Bacillus | subtilis | Bacillus | subtilis | 2.04 | Bacillus | subtilis | 1.87 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WIS-St. Nr. 491 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007; E14 | Bacillus | subtilis | Bacillus | subtilis | 1.81 | Bacillus | subtilis | 1.79 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6022 | 3; public strain collection | DT | 2016-04-06 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.43 | Bacillus | thuringiensis | 2.28 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.64 | Bacillus | thuringiensis | 2.52 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.67 | Bacillus | thuringiensis | 2.64 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.55 | Bacillus | thuringiensis | 2.51 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.62 | Bacillus | thuringiensis | 2.54 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.57 | Bacillus | thuringiensis | 2.57 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B 8 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 1.12.2014 E15 | Bacillus | thuringiensis | Bacillus | cereus | 2.46 | Bacillus | thuringiensis | 2.31 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | CVUAS 10203 | 4; Cry-pos, Cyt-K2, Krist. Pos | eDT | 27.01.2016 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.53 | Bacillus | thuringiensis | 2.51 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | CVUAS 5760 | 4; Cry-pos, Krist. Pos | eDT | 27.01.2016 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.6 | Bacillus | thuringiensis | 2.58 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | CVUAS 5761 | 4; Cry-pos, Krist. Pos | eDT | 27.01.2016 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.54 | Bacillus | thuringiensis | 2.48 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 0 | DSM 2046 | 3; type strain | DT | 2014-01-22 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.15 | Bacillus | cereus | 2.13 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 2046 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015 F5 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.02 | Bacillus | thuringiensis | 1.99 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2015 K6 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.08 | Bacillus | mycoides | 2.03 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus mycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 5815 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015 C21 | Bacillus | thuringiensis | Bacillus | thuringiensis | 1.96 | Bacillus | cereus | 1.83 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6073 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.12.2015 E18 | Bacillus | thuringiensis | Bacillus | cereus | 2.13 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 6890 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2015, L4 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.18 | Bacillus | thuringiensis | 2.16 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 07.11.2006 K4 | Bacillus | thuringiensis | Bacillus | cereus | 2.33 | Bacillus | thuringiensis | 2.22 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0 | 1
|  |  | 1 | 1 | CVUAS 2786,2 | 3; sequenced | DT | 2013-07-31 | Bacillus | timorensis | Bacillus | timorensis | 2.59 | Bacillus | timonensis | 1.49 | A | 1 | 1 | 0 | Bacillus timorensis | Bacillus timonensis | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.05.2015; E09 | Bacillus | tropicus | Bacillus | cereus | 2.28 | Bacillus | pseudomycoides | 2.28 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pseudomycoides | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.07.2017, H17 | Bacillus | vallismortis | Bacillus | vallismortis | 2.37 | Bacillus | subtilis | 1.89 | A | 1 | 1 | 0 | Bacillus vallismortis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.45 | Cytobacillus | firmus | 2.24 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0 | 1
|  |  | 0 | 0 | B 48 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.01.2015 H18 | Cytobacillus | firmus | Cytobacillus | firmus | 2.33 | Cytobacillus | firmus | 2.2 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 70 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.02.2015 N3 | Cytobacillus | firmus | Cytobacillus | firmus | 2.02 | Cytobacillus | firmus | 1.78 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 88 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.02.2015 L16 | Cytobacillus | firmus | Cytobacillus | firmus | 2.06 | Cytobacillus | firmus | 1.98 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCEB 627 | 3; type strain | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.19 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 37 | 3; reference strain collection | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.33 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.02.2016 K19 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.11 | Cytobacillus | horneckiae | 2.01 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.04.2014 H5 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.32 | Cytobacillus | horneckiae | 2.18 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus | kochii | Cytobacillus | kochii | 2.33 | Cytobacillus | kochii | 1.86 | A | 1 | 1 | 0 | Cytobacillus kochii | Cytobacillus kochii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 19.02.2015, E18 | Domibacillus | aminovorans | Domibacillus | aminovorans | 2.81 | Pseudomonas | aeruginosa | 1.37 | A | 1 | 1 | 0 | Domibacillus aminovorans | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 15.02.2015, E8 | Domibacillus | aminovorans | Domibacillus | aminovorans | 2.63 | Staphylococcus | lugdunensis | 1.46 | A | 1 | 1 | 0 | Domibacillus aminovorans | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.02.2015 H5 | Evansella | clarkii | Bacillus | clarkii | 2.35 | Bacillus | mycoides | 1.45 | A | 0 | 0 | 1 | Bacillus clarkii | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F7 | Evansella | clarkii | Bacillus | clarkii | 2.04 | Alkalihalobacillus | krulwichiae | 1.46 | A | 0 | 0 | 1 | Bacillus clarkii | Alkalihalobacillus krulwichiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus | stearothermophilus | Geobacillus | stearothermophilus | 2.39 | Geobacillus | stearothermophilus | 2.29 | A | 1 | 1 | 0 | Geobacillus stearothermophilus | Geobacillus stearothermophilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 I10 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.19 | Alkalihalobacillus | okuhidensis | 1.98 | A | 0 | 0 | 1 | Alkalihalobacillus halodurans | Alkalihalobacillus okuhidensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6939 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 C18 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | okuhidensis | 2.12 | Alkalihalobacillus | halodurans | 1.76 | A | 0 | 0 | 1 | Alkalihalobacillus okuhidensis | Alkalihalobacillus halodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 E19 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | okuhidensis | 2.14 | Alkalihalobacillus | halodurans | 1.92 | A | 0 | 0 | 1 | Alkalihalobacillus okuhidensis | Alkalihalobacillus halodurans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia | sporothermodurans | Bacillus | sporothermodurans | 2.17 | Bacillus | sporothermodurans | 2.14 | A | 0 | 0 | 1 | Bacillus sporothermodurans | Bacillus sporothermodurans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B106 | 3; sequenced | DT | 2015-07-22 | Heyndrickxia | sporothermodurans | Bacillus | sporothermodurans | 2.22 | Bacillus | sporothermodurans | 2.2 | A | 0 | 0 | 1 | Bacillus sporothermodurans | Bacillus sporothermodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 10599 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015, H9 | Heyndrickxia | sporothermodurans | Heyndrickxia | sporothermodurans | 2.35 | Bacillus | sporothermodurans | 2.14 | C | 0 | 0 | 1 | Heyndrickxia sporothermodurans | Bacillus sporothermodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017; N12 | Lederbergia | lenta | Lederbergia | lenta | 2.43 | Lederbergia | lenta | 2.33 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9 | 3; type strain | DT | 2014-01-22 | Lederbergia | lenta | Lederbergia | lenta | 2.22 | Lederbergia | lenta | 2.08 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2007 H7 | Lederbergia | lenta | Pichia | fermentans | 1.51 | Lactobacillus | crispatus | 1.39 | B | 0 | 0 | 1 | Pichia fermentans | Lactobacillus crispatus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.34 | Lysinibacillus | fusiformis | 2.23 | A | 1 | 1 | 0 | Lysinibacillus fusiformis | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.44 | Lysinibacillus | sphaericus | 2.36 | B | 0 | 0 | 1 | Lysinibacillus fusiformis | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7192,2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus | macroides | Lysinibacillus | macroides | 2.64 | Lysinibacillus | xylanilyticus | 2.23 | B | 0 | 0 | 1 | Lysinibacillus macroides | Lysinibacillus xylanilyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.34 | Lysinibacillus | sphaericus | 2.31 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.11 | Lysinibacillus | sphaericus | 1.96 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus | bataviensis | Neobacillus | bataviensis | 2.76 | Neobacillus | bataviensis | 2.29 | A | 1 | 1 | 0 | Neobacillus bataviensis | Neobacillus bataviensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017 N13 | Neobacillus | cucumis | Neobacillus | bataviensis | 2.02 | Neobacillus | bataviensis | 1.69 | A | 0 | 0 | 1 | Neobacillus bataviensis | Neobacillus bataviensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N14 | Niallia | circulans | Niallia | circulans | 1.98 | Niallia | circulans | 1.88 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 216 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 L4 | Niallia | circulans | Niallia | circulans | 2.13 | Niallia | circulans | 1.94 | A | 1 | 1 | 0 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 11 | 3; type strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 13.04.2007, H11, [0034] | Niallia | circulans | Niallia | circulans | 2.05 | Niallia | circulans | 1.61 | A | 1 | 1 | 0 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E14, [0034] | Niallia | circulans | Niallia | circulans | 1.9 | Niallia | circulans | 1.65 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2582,2 | 3; sequenced | eDT | 2015-09-23 | Niallia | nealsonii | Niallia | nealsonii | 2.34 | Niallia | taxi | 2.04 | B | 0 | 0 | 1 | Niallia nealsonii | Niallia taxi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33253 | 2; sequenced | EFex | 2023-01-12 | Niallia | sp | Niallia | nealsonii | 2.62 | Niallia | taxi | 2.1 | B | 0 | 0 | 1 | Niallia nealsonii | Niallia taxi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K20 | Peribacillus | muralis | Peribacillus | muralis | 2.38 | Peribacillus | simplex | 1.81 | A | 1 | 1 | 0 | Peribacillus muralis | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016; K20 | Peribacillus | psychrosaccharolyticus | Peribacillus | psychrosaccharolyticus | 2.21 | Staphylococcus | simiae | 1.44 | A | 1 | 1 | 0 | Peribacillus psychrosaccharolyticus | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus | simplex | Peribacillus | simplex | 2.08 | Peribacillus | simplex | 1.67 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2016 H13 | Peribacillus | simplex | Peribacillus | simplex | 1.88 | Peribacillus | simplex | 1.78 | B | 0 | 0 | 1 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 15.12.2016, K15, [0096] | Peribacillus | simplex | Peribacillus | simplex | 2.19 | Peribacillus | simplex | 1.64 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016 K15 | Priestia | flexa | Priestia | flexa | 2.54 | Priestia | flexa | 2.17 | A | 1 | 1 | 0 | Priestia flexa | Priestia flexa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H18 | Priestia | flexa | Priestia | megaterium | 2.01 | Priestia | megaterium | 1.8 | A | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 2 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2014 E5 | Priestia | megaterium | Priestia | megaterium | 2.33 | Priestia | megaterium | 1.71 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 14581 | 3; type strain | DT | 2014-01-29 | Priestia | megaterium | Priestia | megaterium | 2.32 | Bacillus | arybhattai | 2.28 | C | 0 | 0 | 1 | Priestia megaterium | Bacillus arybhattai | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015 H21 | Priestia | megaterium | Priestia | megaterium | 1.92 | Priestia | megaterium | 1.92 | B | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 333 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.05.2015 K18 | Priestia | megaterium | Priestia | megaterium | 2.42 | Priestia | megaterium | 1.78 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia | megaterium | Priestia | megaterium | 2.06 | Priestia | megaterium | 2.06 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 32 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.04.2007 K17 | Priestia | megaterium | Bacillus | arybhattai | 2.05 | Priestia | megaterium | 1.84 | C | 0 | 0 | 1 | Bacillus arybhattai | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 333 B175 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015 C16 | Priestia | megaterium | Priestia | megaterium | 2.31 | Bacillus | arybhattai | 1.72 | C | 0 | 0 | 1 | Priestia megaterium | Bacillus arybhattai | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 509 B172 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015 E14 | Priestia | megaterium | Bacillus | arybhattai | 1.9 | Priestia | megaterium | 1.87 | C | 0 | 0 | 1 | Bacillus arybhattai | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 90 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2007 H10 | Priestia | megaterium | Ligilactobacillus | agilis | 1.55 | Priestia | megaterium | 1.5 | B | 0 | 0 | 1 | Ligilactobacillus agilis | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_01 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.05.2015 K18 | Priestia | megaterium | Priestia | megaterium | 2.44 | Bacillus | arybhattai | 2.08 | C | 0 | 0 | 1 | Priestia megaterium | Bacillus arybhattai | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E3 | Salibacterium | salarium | Salibacterium | salarium | 2.05 | Aeromonas | salmonicida | 1.35 | A | 1 | 1 | 0 | Salibacterium salarium | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6307 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 H11 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.24 | Sutcliffiella | cohnii | 1.82 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F13 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.13 | Sutcliffiella | cohnii | 1.86 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 K8 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.03 | Sutcliffiella | cohnii | 1.79 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus | saccharophilus | Terribacillus | saccharophilus | 2.4 | Lodderomyces | elongisporus | 1.45 | A | 1 | 1 | 0 | Terribacillus saccharophilus | Lodderomyces elongisporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 N5 | Virgibacillus | pantothenticus | Virgibacillus | pantothenticus | 2.01 | Virgibacillus | pantothenticus | 1.84 | A | 1 | 1 | 0 | Virgibacillus pantothenticus | Virgibacillus pantothenticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus | proomii | Virgibacillus | proomii | 2.02 | Virgibacillus | proomii | 1.99 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3616,2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus | proomii | Virgibacillus | proomii | 2.2 | Virgibacillus | proomii | 1.83 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H15 | Weizmannia | coagulans | Weizmannia | coagulans | 2.45 | Weizmannia | coagulans | 2.12 | A | 1 | 1 | 0 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 71,2 | 3; sequenced | DT | 2013-12-13 | Weizmannia | ginsengihumi | Weizmannia | ginsengihumi | 2.4 | Weizmannia | ginsengihumi | 2.19 | A | 1 | 1 | 0 | Weizmannia ginsengihumi | Weizmannia ginsengihumi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium | acetylicum | Exiguobacterium | acetylicum | 2.63 | Exiguobacterium | enclense | 2.1 | B | 0 | 0 | 1 | Exiguobacterium acetylicum | Exiguobacterium enclense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium | enclense | Exiguobacterium | enclense | 2.19 | Exiguobacterium | acetylicum | 2.03 | B | 0 | 0 | 1 | Exiguobacterium enclense | Exiguobacterium acetylicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium | profundum | Exiguobacterium | profundum | 2.44 | Exiguobacterium | sp[3] | 2.41 | B | 0 | 0 | 1 | Exiguobacterium profundum | Exiguobacterium sp[3] | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.36 | Brochothrix | thermosphacta | 2.32 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.39 | Brochothrix | thermosphacta | 2.37 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria | aquatica | Listeria | aquatica | 2.33 | Listeria | aquatica | 2.28 | A | 1 | 1 | 0 | Listeria aquatica | Listeria aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria | aquatica | Listeria | aquatica | 2.5 | Listeria | aquatica | 2.43 | A | 1 | 1 | 0 | Listeria aquatica | Listeria aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.25 | Listeria | fleischmannii | 2.23 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LU2006 3 | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.2 | Listeria | fleischmannii | 2.19 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105447 | 3; type strain | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.32 | Listeria | grayi | 2.32 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20596 | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.25 | Listeria | grayi | 2.12 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10132,2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | innocua | Listeria | innocua | 2.45 | Listeria | innocua | 2.44 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8569 | 3; 16S Sequenz | DT | 19.02.2013 | Listeria | innocua | Listeria | innocua | 2.48 | Listeria | innocua | 2.46 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria | ivanovii | Listeria | ivanovii | 2.33 | Listeria | ivanovii | 2.32 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32100 | 2; sequenced | EFex | 2019-12-13 | Listeria | ivanovii | Listeria | ivanovii | 2.41 | Listeria | ivanovii | 2.35 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 23813 | 3; type strain | eDT | 2015-02-18 | Listeria | marthii | Listeria | marthii | 2.38 | Listeria | monocytogenes | 2.28 | B | 0 | 0 | 1 | Listeria marthii | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.42 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 873,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.43 | Listeria | monocytogenes | 2.41 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 240,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.29 | Listeria | monocytogenes | 2.24 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0206 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0066 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.19 | Listeria | monocytogenes | 2.19 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.22 | Listeria | monocytogenes | 2.21 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.37 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7366,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.54 | Listeria | monocytogenes | 2.51 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria | newyorkensis | Listeria | newyorkensis | 2.47 | Listeria | sp | 2.17 | B | 0 | 0 | 1 | Listeria newyorkensis | Listeria sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 22097 | 3; type strain | eDT | 2015-11-20 | Listeria | rocourtiae | Listeria | rocourtiae | 2.45 | Listeria | rocourtiae | 2.1 | A | 1 | 1 | 0 | Listeria rocourtiae | Listeria rocourtiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2763 | 3; 16S Sequenz | DT | 19.02.2013 | Listeria | seeligeri | Listeria | seeligeri | 2.45 | Listeria | seeligeri | 2.45 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | seeligeri | Listeria | seeligeri | 2.36 | Listeria | seeligeri | 2.3 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria | sp | Listeria | sp | 2.31 | Listeria | sp | 2.29 | A | 1 | 1 | 0 | Listeria sp | Listeria sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2412,3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | welshimeri | Listeria | welshimeri | 2.38 | Listeria | welshimeri | 2.28 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria | welshimeri | Listeria | welshimeri | 2.44 | Listeria | welshimeri | 2.41 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.31 | Brevibacillus | laterosporus | 1.95 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.42 | Brevibacillus | laterosporus | 2.25 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B210 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 K19 | Paenibacillus | alvei | Paenibacillus | alvei | 1.82 | Paenibacillus | curdlanolyticus | 1.45 | B | 0 | 0 | 1 | Paenibacillus alvei | Paenibacillus curdlanolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus | amylolyticus | Paenibacillus | amylolyticus | 2.26 | Paenibacillus | amylolyticus | 2.25 | A | 1 | 1 | 0 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus | amylolyticus | Paenibacillus | xylanexedens | 2.22 | Paenibacillus | amylolyticus | 2.17 | B | 0 | 0 | 1 | Paenibacillus xylanexedens | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.53 | Paenibacillus | barengoltzii | 2.5 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.43 | Paenibacillus | barengoltzii | 2.4 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus | chibensis | Paenibacillus | chibensis | 2.28 | Paenibacillus | chibensis | 2.15 | A | 1 | 1 | 0 | Paenibacillus chibensis | Paenibacillus chibensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.18 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.25 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus | larvae | Paenibacillus | larvae | 2.46 | Paenibacillus | larvae | 2.38 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.42 | Paenibacillus | larvae | 2.39 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.38 | Paenibacillus | larvae | 2.34 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.18 | Paenibacillus | larvae | 2.18 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus | macerans | Paenibacillus | macerans | 2.28 | Paenibacillus | macerans | 2.07 | A | 1 | 1 | 0 | Paenibacillus macerans | Paenibacillus macerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus | nuruki | Paenibacillus | nuruki | 2.54 | Ochrobactrum | anthropi | 1.29 | A | 1 | 1 | 0 | Paenibacillus nuruki | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.12 | Paenibacillus | pabuli | 2.02 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.22 | Paenibacillus | pabuli | 2.15 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus | pasadenensis | Paenibacillus | pasadenensis | 2.52 | Paenibacillus | pasadenensis | 2.05 | A | 1 | 1 | 0 | Paenibacillus pasadenensis | Paenibacillus pasadenensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.09.2013 C8 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 1.9 | Paenibacillus | polymyxa | 1.78 | B | 0 | 0 | 1 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.11.2017 C18 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 2.1 | Paenibacillus | polymyxa | 2.02 | A | 1 | 1 | 0 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus | sp-CVUAS-6369 | Paenibacillus | sp-CVUAS-6369 | 2.08 | Lentilactobacillus | parabuchneri | 1.35 | A | 1 | 1 | 0 | Paenibacillus sp-CVUAS-6369 | Lentilactobacillus parabuchneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus | sp-E8a | Paenibacillus | sp-E8a | 2.28 | Paenibacillus | amylolyticus | 2.24 | B | 0 | 0 | 1 | Paenibacillus sp-E8a | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus | taichungensis | Paenibacillus | taichungensis | 2.29 | Paenibacillus | taichungensis | 1.9 | A | 1 | 1 | 0 | Paenibacillus taichungensis | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | FLI003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus | taichungensis | Paenibacillus | taichungensis | 2.81 | Paenibacillus | taichungensis | 1.87 | A | 1 | 1 | 0 | Paenibacillus taichungensis | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus | tundrae | Paenibacillus | tundrae | 2.09 | Paenibacillus | taichungensis | 1.87 | A | 1 | 1 | 0 | Paenibacillus tundrae | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus | validus | Paenibacillus | validus | 2.26 | Paenibacillus | validus | 2.16 | A | 1 | 1 | 0 | Paenibacillus validus | Paenibacillus validus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus | saudimassiliensis | Metalysinibacillus | saudimassiliensis | 2.53 | Lysinibacillus | saudimassiliensis | 2.01 | C | 0 | 0 | 1 | Metalysinibacillus saudimassiliensis | Lysinibacillus saudimassiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus | odysseyi | Solibacillus | odysseyi | 2.2 | Ureibacillus | sinduriensis | 1.47 | A | 1 | 1 | 0 | Solibacillus odysseyi | Ureibacillus sinduriensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.48 | Kurthia | gibsonii | 2.47 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.34 | Kurthia | gibsonii | 2.29 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia | zopfii | Kurthia | zopfii | 2.31 | Kurthia | zopfii | 1.66 | A | 1 | 1 | 0 | Kurthia zopfii | Kurthia zopfii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | halophilus | Jeotgalicoccus | halophilus | 2.45 | Listeria | innocua | 1.68 | A | 1 | 1 | 0 | Jeotgalicoccus halophilus | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | huakuii | Jeotgalicoccus | huakuii | 2.6 | Macrococcus | canis | 1.37 | A | 1 | 1 | 0 | Jeotgalicoccus huakuii | Macrococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | psychrophilus | Jeotgalicoccus | psychrophilus | 2.37 | Jeotgalicoccus | psychrophilus | 1.84 | A | 1 | 1 | 0 | Jeotgalicoccus psychrophilus | Jeotgalicoccus psychrophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus | carouselicus | Macrococcus | carouselicus | 2.59 | Staphylococcus | simulans | 1.4 | A | 1 | 1 | 0 | Macrococcus carouselicus | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 14.03.2013 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.08 | Macrococcus | caseolyticus | 1.75 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.19 | Macrococcus | caseolyticus | 1.87 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus | lentus | Mammaliicoccus | lentus | 2.4 | Staphylococcus | lentus | 1.8 | C | 0 | 0 | 1 | Mammaliicoccus lentus | Staphylococcus lentus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 2.76 | Staphylococcus | sciuri | 2.37 | C | 0 | 0 | 1 | Mammaliicoccus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus | sciuri | Staphylococcus | sciuri | 1.99 | Staphylococcus | sciuri | 1.68 | B | 0 | 0 | 1 | Staphylococcus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9937,2 | 3; 16S rDNA-sequencing | DT | 01.12.2015 | Mammaliicoccus | vitulinus | Mammaliicoccus | vitulinus | 2.31 | Staphylococcus | vitulinus | 1.87 | C | 0 | 0 | 1 | Mammaliicoccus vitulinus | Staphylococcus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 343,2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus | vitulinus | Mammaliicoccus | vitulinus | 2.23 | Mammaliicoccus | vitulinus | 2.23 | A | 1 | 1 | 0 | Mammaliicoccus vitulinus | Mammaliicoccus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5687 | 3; sequenced | DT | 2013-04-16 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | agnetis | 2.14 | B | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus agnetis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TUM 7385 | 3; sequenced | DT | 02.10.2013 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | hyicus | 1.96 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.39 | Staphylococcus | argenteus | 2.38 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.37 | Staphylococcus | argenteus | 2.28 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.01 | Staphylococcus | arlettae | 1.72 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.13 | Staphylococcus | arlettae | 1.63 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 29213 | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9914,2 | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus | auricularis | Staphylococcus | auricularis | 2.14 | Staphylococcus | auricularis | 1.85 | A | 1 | 1 | 0 | Staphylococcus auricularis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20326 (T) | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.25 | Staphylococcus | capitis | 2.13 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.2 | Staphylococcus | capitis | 2.19 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20608 (T) | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.16 | Staphylococcus | caprae | 2.06 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.06 | Staphylococcus | caprae | 2.03 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.18 | Staphylococcus | carnosus | 1.55 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.24 | Staphylococcus | carnosus | 1.46 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.26 | Staphylococcus | hyicus | 1.85 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.28 | Staphylococcus | chromogenes | 1.91 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.36 | Staphylococcus | xylosus | 1.36 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.72 | Staphylococcus | cohnii | 2.25 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | x | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 15.11.2011 E20 | Staphylococcus | delphini | Staphylococcus | delphini | 2.32 | Staphylococcus | delphini | 2.01 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus | delphini | Staphylococcus | delphini | 2.53 | Staphylococcus | delphini | 1.9 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus | devriesei | Staphylococcus | devriesei | 2.33 | Staphylococcus | capitis | 1.46 | A | 1 | 1 | 0 | Staphylococcus devriesei | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.09 | Staphylococcus | epidermidis | 2.08 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.31 | Staphylococcus | epidermidis | 1.99 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | equorum | Staphylococcus | equorum | 2.16 | Staphylococcus | equorum | 1.8 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20674 T | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | equorum | Staphylococcus | equorum | 2.11 | Staphylococcus | equorum | 1.88 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | felis | Staphylococcus | felis | 2.13 | Staphylococcus | felis | 1.96 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | felis | Staphylococcus | felis | 2.14 | Staphylococcus | felis | 2 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.08.2011 H9 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.46 | Staphylococcus | haemolyticus | 2.43 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.12 | Staphylococcus | haemolyticus | 2.04 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N6 [0034] | Staphylococcus | hominis | Staphylococcus | hominis | 2.27 | Staphylococcus | hominis | 2.23 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 789,2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus | hominis | Staphylococcus | hominis | 2.13 | Staphylococcus | hominis | 1.96 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.2 | Staphylococcus | hyicus | 2.14 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3767 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.38 | Staphylococcus | hyicus | 2.12 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, N8, [0034] | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.15 | Staphylococcus | intermedius | 2.02 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N7, [0034] | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.07 | Staphylococcus | intermedius | 1.97 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 05.04.2012, N5, [0034] | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.4 | Staphylococcus | lugdunensis | 2.4 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, H16 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.5 | Staphylococcus | lugdunensis | 2.43 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | microti | Staphylococcus | microti | 2.48 | Staphylococcus | microti | 1.85 | A | 1 | 1 | 0 | Staphylococcus microti | Staphylococcus microti | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | pasteuri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 2.17 | A | 0 | 0 | 1 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus | pasteuri | Staphylococcus | pasteuri | 2 | Staphylococcus | warneri | 1.81 | A | 1 | 1 | 0 | Staphylococcus pasteuri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.03.2016, H05, [0034] | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.15 | Staphylococcus | pseudintermedius | 2.1 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | pseudintermedius | Staphylococcus | delphini | 2.1 | Staphylococcus | pseudintermedius | 2.03 | B | 0 | 0 | 1 | Staphylococcus delphini | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.03 | Staphylococcus | kloosii | 1.43 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus kloosii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 503,2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.18 | Staphylococcus | saprophyticus | 2.06 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus | schleiferi | Staphylococcus | schleiferi | 2.15 | Staphylococcus | schleiferi | 2.02 | A | 1 | 1 | 0 | Staphylococcus schleiferi | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus | schweitzeri | Staphylococcus | schweitzeri | 2.48 | Staphylococcus | schweitzeri | 2.17 | A | 1 | 1 | 0 | Staphylococcus schweitzeri | Staphylococcus schweitzeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | simulans | Staphylococcus | simulans | 2.13 | Staphylococcus | simulans | 2.11 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | simulans | Staphylococcus | simulans | 2.17 | Staphylococcus | simulans | 2.04 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2439,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.26 | Staphylococcus | succinus | 2.23 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 761,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.28 | Staphylococcus | succinus | 2.21 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus | warneri | Staphylococcus | warneri | 2.07 | Staphylococcus | warneri | 2.02 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20316 | 3; type strain | DT | 2013-10-02 | Staphylococcus | warneri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 1.96 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.04 | Staphylococcus | xylosus | 2.02 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.88 | Staphylococcus | xylosus | 1.76 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus | sp-185-U1-2 | Aerococcus | sp-185-U1-2 | 2.41 | Xanthomonas | hortorum | 1.38 | A | 1 | 1 | 0 | Aerococcus sp-185-U1-2 | Xanthomonas hortorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus | suis | Aerococcus | suis | 2.36 | Aerococcus | suis | 2.3 | A | 1 | 1 | 0 | Aerococcus suis | Aerococcus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus | urinaeequi | Aerococcus | urinaeequi | 2.49 | Aerococcus | viridans | 1.72 | A | 1 | 1 | 0 | Aerococcus urinaeequi | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 27.09.2016 | Aerococcus | viridans | Aerococcus | viridans | 2.33 | Aerococcus | viridans | 2.04 | A | 1 | 1 | 0 | Aerococcus viridans | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter | phocae | Atopobacter | phocae | 2.33 | Atopobacter | phocae | 2.31 | A | 1 | 1 | 0 | Atopobacter phocae | Atopobacter phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2442,2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium | divergens | Carnobacterium | divergens | 2.44 | Carnobacterium | divergens | 2.1 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium | divergens | Carnobacterium | divergens | 2.62 | Carnobacterium | divergens | 2.37 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6273 | 3; sequenced | eDT | 13.01.2016 | Carnobacterium | maltaromaticum | Carnobacterium | maltaromaticum | 2.55 | Carnobacterium | maltaromaticum | 2.54 | A | 1 | 1 | 0 | Carnobacterium maltaromaticum | Carnobacterium maltaromaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1519,3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella | adiacens | Granulicatella | adiacens | 2.3 | Granulicatella | adiacens | 2.24 | A | 1 | 1 | 0 | Granulicatella adiacens | Granulicatella adiacens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2616 | 3; sequenced | DT | 25.09.2013 | Marinilactibacillus | sp-CVUAS-2616 | Marinilactibacillus | sp-CVUAS-2616 | 2.02 | Marinilactibacillus | sp-CVUAS-2616 | 1.83 | A | 1 | 1 | 0 | Marinilactibacillus sp-CVUAS-2616 | Marinilactibacillus sp-CVUAS-2616 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8980,3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus | sp-CVUAS-2616 | Marinilactibacillus | sp-CVUAS-2616 | 2.54 | Marinilactibacillus | sp-CVUAS-2616 | 2.44 | A | 1 | 1 | 0 | Marinilactibacillus sp-CVUAS-2616 | Marinilactibacillus sp-CVUAS-2616 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16 | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011 H20 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.3 | Enterococcus | casseliflavus | 2.2 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.33 | Enterococcus | casseliflavus | 2.32 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5030 | 3; sequenced | eDT | 23.07.2014 | Enterococcus | cecorum | Enterococcus | cecorum | 2.34 | Enterococcus | cecorum | 2.29 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8897 | 3; sequenced | eDT | 23.07.2014 | Enterococcus | cecorum | Enterococcus | cecorum | 2.82 | Enterococcus | cecorum | 1.41 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H12 | Enterococcus | durans | Enterococcus | hirae | 2.09 | Enterococcus | hirae | 1.8 | A | 0 | 0 | 1 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H13 | Enterococcus | durans | Enterococcus | durans | 2 | Enterococcus | durans | 1.94 | A | 1 | 1 | 0 | Enterococcus durans | Enterococcus durans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus | faecalis | Enterococcus | faecalis | 2.51 | Enterococcus | faecalis | 2.41 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus | faecalis | Enterococcus | faecalis | 2.39 | Enterococcus | faecalis | 2.38 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.07.2011 N8 | Enterococcus | faecium | Enterococcus | faecium | 2.09 | Enterococcus | faecium | 2.05 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MR11-207 LVU04 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus | faecium | Enterococcus | faecium | 2.42 | Enterococcus | faecium | 2.1 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N13 | Enterococcus | gallinaru | Enterococcus | gallinarum | 2.04 | Enterococcus | gallinarum | 1.98 | A | 0 | 0 | 1 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus | gallinarum | Enterococcus | gallinarum | 2.42 | Enterococcus | gallinarum | 2.38 | A | 1 | 1 | 0 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N3 | Enterococcus | hirae | Enterococcus | hirae | 2.33 | Enterococcus | hirae | 2.16 | A | 1 | 1 | 0 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus | solitarius | Tetragenococcus | solitarius | 2.14 | Halomonas | elongata | 1.38 | A | 1 | 1 | 0 | Tetragenococcus solitarius | Halomonas elongata | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1549,2 | 3; sequenced | DT | 2016-09-27 | Vagococcus | salmoninarum | Vagococcus | salmoninarum | 2.3 | Vagococcus | salmoninarum | 2.09 | A | 1 | 1 | 0 | Vagococcus salmoninarum | Vagococcus salmoninarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus | acidophilus | Lactobacillus | acidophilus | 2.32 | Lactobacillus | acidophilus | 2.24 | A | 1 | 1 | 0 | Lactobacillus acidophilus | Lactobacillus acidophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus | johnsonii | Lactobacillus | johnsonii | 2.42 | Lactobacillus | johnsonii | 2.35 | A | 1 | 1 | 0 | Lactobacillus johnsonii | Lactobacillus johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32277,2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus | sp-CVUAS-32277,2 | Lactobacillus | sp-CVUAS-32277,2 | 2.64 | Limosilactobacillus | gorillae | 1.25 | A | 1 | 1 | 0 | Lactobacillus sp-CVUAS-32277,2 | Limosilactobacillus gorillae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus | claussenii | Pediococcus | clausenii | 2.41 | Pediococcus | claussenii | 2.25 | B | 0 | 0 | 1 | Pediococcus clausenii | Pediococcus claussenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1213 | 2; 0 | DT | 2015-07-09 | Leuconostoc | mesenteroides | Leuconostoc | mesenteroides | 2.28 | Leuconostoc | mesenteroides | 2.09 | A | 1 | 1 | 0 | Leuconostoc mesenteroides | Leuconostoc mesenteroides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1146 | 2; 0 | DT | 2015-07-09 | Lactococcus | lactis | Lactococcus | lactis | 2.42 | Lactococcus | lactis | 2.4 | A | 1 | 1 | 0 | Lactococcus lactis | Lactococcus lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.45 | Streptococcus | agalactiae | 2.44 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6784 | 3; public strain collection | eDT | 2015-03-11 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.6 | Streptococcus | agalactiae | 2.56 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus | canis | Streptococcus | canis | 2.53 | Streptococcus | canis | 2.41 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL Mi-845_17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus | canis | Streptococcus | canis | 2.5 | Streptococcus | canis | 2.49 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.72 | Streptococcus | castoreus | 2.64 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17536 | 3; type strain | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.68 | Streptococcus | castoreus | 2.61 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus | catagoni | Streptococcus | catagoni | 2.83 | Streptococcus | catagoni | 2.81 | A | 1 | 1 | 0 | Streptococcus catagoni | Streptococcus catagoni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 99-1-2017 | 3; typestrain | EFEx | 2018-05-03 | Streptococcus | catagoni | Streptococcus | catagoni | 2.81 | Streptococcus | catagoni | 2.8 | A | 1 | 1 | 0 | Streptococcus catagoni | Streptococcus catagoni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.15 | Streptococcus | dysgalactiae | 2.14 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.26 | Streptococcus | dysgalactiae | 2.12 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.43 | Streptococcus | dysgalactiae | 2.35 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.46 | Streptococcus | dysgalactiae | 2.33 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32462,2 | 3; specific PCR | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi_ssp_equi | 2.12 | Streptococcus | equi_ssp_equi | 2.06 | A | 0 | 0 | 1 | Streptococcus equi_ssp_equi | Streptococcus equi_ssp_equi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32770 | 3; specific PCR | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi_ssp_equi | 2.29 | Streptococcus | equi_ssp_equi | 2.24 | A | 0 | 0 | 1 | Streptococcus equi_ssp_equi | Streptococcus equi_ssp_equi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus | equi | Streptococcus | equi_ssp_zooepidemicus | 2.34 | Streptococcus | equi_ssp_zooepidemicus | 2.34 | A | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus | equi | Streptococcus | equi | 2.22 | Streptococcus | equi_ssp_zooepidemicus | 2.21 | B | 0 | 0 | 1 | Streptococcus equi | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus | iniae | Streptococcus | iniae | 2.02 | Streptococcus | iniae | 1.89 | A | 1 | 1 | 0 | Streptococcus iniae | Streptococcus iniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus | orisasini | Streptococcus | orisasini | 2.86 | Streptococcus | orisasini | 2.44 | A | 1 | 1 | 0 | Streptococcus orisasini | Streptococcus orisasini | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | parauberis | Streptococcus | parauberis_JRA_G+R-39 | 2.18 | Streptococcus | parauberis | 2.1 | B | 0 | 0 | 1 | Streptococcus parauberis_JRA_G+R-39 | Streptococcus parauberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | phocae | Streptococcus | phocae | 2.33 | Streptococcus | phocae | 2.3 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 12-13862-001-01 LVU11 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus | pluranimalium | Streptococcus | pluranimalium | 2.2 | Streptococcus | pluranimalium | 2.18 | A | 1 | 1 | 0 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.48 | Streptococcus | pneumoniae | 2.44 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.16 | Streptococcus | pneumoniae | 2.16 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.29 | Streptococcus | pneumoniae | 2.23 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.49 | Streptococcus | pneumoniae | 2.46 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus | porcinus | Streptococcus | porcinus | 2.27 | Streptococcus | porcinus | 2.1 | A | 1 | 1 | 0 | Streptococcus porcinus | Streptococcus porcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.04.2008, K11, [0034] | Streptococcus | pyogenes | Streptococcus | pyogenes | 2.43 | Streptococcus | pyogenes | 2.4 | A | 1 | 1 | 0 | Streptococcus pyogenes | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10145,3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus | salivarius | Streptococcus | salivarius | 2.13 | Streptococcus | salivarius | 2.13 | A | 1 | 1 | 0 | Streptococcus salivarius | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus | sp-CVUAS-32349 | Streptococcus | sp-CVUAS-32349 | 2.58 | Streptococcus | dysgalactiae | 1.52 | A | 1 | 1 | 0 | Streptococcus sp-CVUAS-32349 | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2260,4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus | troglodytidis | Streptococcus | troglodytidis | 2.52 | Streptococcus | troglodytidis | 2.21 | A | 1 | 1 | 0 | Streptococcus troglodytidis | Streptococcus troglodytidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5218,4 | 3; sequenced; 16S; NRZ Strep Uni AA | DT | 2015-06-11 | Streptococcus | troglodytidis | Streptococcus | troglodytidis | 2.35 | Streptococcus | troglodytidis | 2.13 | A | 1 | 1 | 0 | Streptococcus troglodytidis | Streptococcus troglodytidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.03 | Streptococcus | uberis | 1.81 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.23 | Streptococcus | uberis | 2.18 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 1.94 | Clostridioides | difficile | 1.83 | B | 0 | 0 | 1 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.2 | Clostridioides | difficile | 2.17 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.1 | Clostridium | sporogenes | 2.05 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.92 | Clostridium | sporogenes | 1.88 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 2.01 | Clostridium | botulinum | 1.99 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium botulinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B L677 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.03 | Clostridium | sporogenes | 1.95 | A | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.16 | Clostridium | sporogenes | 1.59 | A | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.99 | Clostridium | sporogenes | 1.67 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.84 | Arthrobacter | ramosus | 1.51 | B | 0 | 0 | 1 | Clostridium botulinum | Arthrobacter ramosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium | barati | Clostridium | budayi | 2.31 | Clostridium | baratii | 2.28 | B | 0 | 0 | 1 | Clostridium budayi | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium | botulinum | Clostridium | botulinum | 2.36 | Bacillus | subtilis | 1.49 | A | 1 | 1 | 0 | Clostridium botulinum | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32971 | 3; sequenced | EFEx | 2021-02-25 | Clostridium | botulinum | Clostridium | botulinum | 2.32 | Clostridium | sporogenes | 2.08 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium | butyricum | Clostridium | butyricum | 2.61 | Clostridium | butyricum | 2.48 | A | 1 | 1 | 0 | Clostridium butyricum | Clostridium butyricum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium | chauvoei | Clostridium | chauvoei | 2.13 | Clostridium | chauvoei | 2.01 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium | chauvoei | Clostridium | chauvoei | 2.38 | Clostridium | chauvoei | 2.17 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002729 | 2; 0 | DT | 2017-02-28 | Clostridium | colchlearium | Clostridium | cochlearium | 2.51 | Clostridium | cochlearium | 2.5 | A | 0 | 0 | 1 | Clostridium cochlearium | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium | colicanis | Clostridium | colicanis | 2.29 | Bifidobacterium | boum | 1.42 | A | 1 | 1 | 0 | Clostridium colicanis | Bifidobacterium boum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 849,2 | 3; sequenced | DT | 2014-12-17 | Clostridium | isatidis | Clostridium | isatidis | 2.13 | Clostridium | isatidis | 1.6 | A | 1 | 1 | 0 | Clostridium isatidis | Clostridium isatidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium | nigeriense | Clostridium | nigeriense | 2.58 | Clostridium | tertium | 1.75 | A | 1 | 1 | 0 | Clostridium nigeriense | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium | novyi | Clostridium | novyi | 2.03 | Clostridium | novyi | 1.96 | A | 1 | 1 | 0 | Clostridium novyi | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium | perfringens | Clostridium | perfringens | 2.09 | Clostridium | perfringens | 1.9 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B | 3; case rerort | DT | 2013-08-06 | Clostridium | perfringens | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.4 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium | septicum | Clostridium | septicum | 2.44 | Clostridium | septicum | 2.26 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium | septicum | Clostridium | septicum | 2.27 | Clostridium | septicum | 2.23 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium | sp-DSM-1985 | Clostridium | botulinum | 2.26 | Campylobacter | jejuni | 1.33 | A | 0 | 0 | 1 | Clostridium botulinum | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium | sporogenes | Clostridium | sporogenes | 2.34 | Clostridium | sporogenes | 2.21 | A | 1 | 1 | 0 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium | sporogenes | Clostridium | sporogenes | 1.95 | Clostridium | sporogenes | 1.79 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141015223 aerob | 2; 0 | DT | 2014-11-27 | Clostridium | tertium | Clostridium | tertium | 2.44 | Clostridium | tertium | 2.22 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium | tertium | Clostridium | tertium | 2.33 | Clostridium | tertium | 2.28 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium | tetani | Clostridium | tetani | 2.23 | Clostridium | tetani | 2.16 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium | tetani | Clostridium | tetani | 2.55 | Clostridium | tetani | 2.47 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya | limosa | Hathewaya | limosa | 2.29 | Hathewaya | limosa | 2.25 | A | 1 | 1 | 0 | Hathewaya limosa | Hathewaya limosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium | sordellii | Paeniclostridium | sordellii | 2.32 | Paeniclostridium | sordellii | 2.27 | A | 1 | 1 | 0 | Paeniclostridium sordellii | Paeniclostridium sordellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium | bifermentans | Paraclostridium | bifermentans | 2.02 | Paraclostridium | bifermentans | 1.92 | A | 1 | 1 | 0 | Paraclostridium bifermentans | Paraclostridium bifermentans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix | anatis | Erysipelothrix | anatis | 2.41 | Erysipelothrix | anatis | 2.35 | A | 1 | 1 | 0 | Erysipelothrix anatis | Erysipelothrix anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix | anatis | Erysipelothrix | anatis | 2.65 | Erysipelothrix | anatis | 2.43 | A | 1 | 1 | 0 | Erysipelothrix anatis | Erysipelothrix anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix | aquatica | Erysipelothrix | aquatica | 2.48 | Erysipelothrix | aquatica | 2.43 | A | 1 | 1 | 0 | Erysipelothrix aquatica | Erysipelothrix aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | VA-92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | aquatica | Erysipelothrix | larvae | 2.68 | Erysipelothrix | larvae | 2.15 | A | 0 | 0 | 1 | Erysipelothrix larvae | Erysipelothrix larvae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | larvae | Erysipelothrix | larvae | 2.55 | Erysipelothrix | larvae | 2.18 | A | 1 | 1 | 0 | Erysipelothrix larvae | Erysipelothrix larvae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 110099 | 3; type strain | DT | 2021-10-28 | Erysipelothrix | piscisicarius | Erysipelothrix | piscisicarius | 2.6 | Erysipelothrix | piscisicarius | 2.15 | A | 1 | 1 | 0 | Erysipelothrix piscisicarius | Erysipelothrix piscisicarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.7 | Erysipelothrix | rhusiopathiae | 2.54 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.25 | Erysipelothrix | rhusiopathiae | 2.22 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix | tonsillarum | Erysipelothrix | tonsillarum | 2.62 | Erysipelothrix | tonsillarum | 2.55 | A | 1 | 1 | 0 | Erysipelothrix tonsillarum | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-27 | Erysipelothrix | tonsillarum | Erysipelothrix | tonsillarum | 2.79 | Erysipelothrix | tonsillarum | 2.55 | A | 1 | 1 | 0 | Erysipelothrix tonsillarum | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus | sp-CVUAS-10169 | Anaerococcus | sp-CVUAS-10169 | 2.26 | Pseudomonas | putida | 1.34 | A | 1 | 1 | 0 | Anaerococcus sp-CVUAS-10169 | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32948 | 2; 0 | EFEx | 2021-02-18 | Finegoldia | magna | Finegoldia | magna | 2.26 | Finegoldia | magna | 2.08 | A | 1 | 1 | 0 | Finegoldia magna | Finegoldia magna | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.25 | Helcococcus | kunzii | 2.25 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.45 | Helcococcus | kunzii | 2.4 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 433,2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus | ovis | Helcococcus | ovis | 2.41 | Helcococcus | ovis | 2.39 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus | ovis | Helcococcus | ovis | 2.28 | Helcococcus | ovis | 2.17 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus | sp-CVUAS-32321 | Helcococcus | sp-CVUAS-32321 | 2.67 | Helcococcus | sueciensis | 1.77 | A | 1 | 1 | 0 | Helcococcus sp-CVUAS-32321 | Helcococcus sueciensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 692,2 | 3; sequenced | DT | 2013-02-19 | Parvimonas | sp-CVUAS-692,2 | Parvimonas | sp-CVUAS-692,2 | 2.15 | Actinomyces | bovis | 1.44 | A | 1 | 1 | 0 | Parvimonas sp-CVUAS-692,2 | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides | finegoldii | Bacteroides | finegoldii | 2.6 | Bacteroides | finegoldii | 2.53 | A | 1 | 1 | 0 | Bacteroides finegoldii | Bacteroides finegoldii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides | fragilis | Bacteroides | fragilis | 2.53 | Bacteroides | fragilis | 2.42 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32947 | 2; 0 | EFEx | 2021-02-18 | Bacteroides | fragilis | Bacteroides | fragilis | 2.35 | Bacteroides | fragilis | 2.22 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181001478 | 2; 0 | DT | 2018-05-02 | Bacteroides | pyogenes | Bacteroides | pyogenes | 2.48 | Bacteroides | pyogenes | 2.48 | A | 1 | 1 | 0 | Bacteroides pyogenes | Bacteroides pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella | heparinolytica | Prevotella | heparinolytica | 2.52 | Prevotella | heparinolytica | 2.49 | A | 1 | 1 | 0 | Prevotella heparinolytica | Prevotella heparinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella | columbina | Riemerella | columbina | 2.56 | Lactobacillus | plantarum | 1.25 | A | 1 | 1 | 0 | Riemerella columbina | Lactobacillus plantarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium | chaponense | Chryseobacterium | chaponense | 2.46 | Chryseobacterium | chaponense | 1.81 | A | 1 | 1 | 0 | Chryseobacterium chaponense | Chryseobacterium chaponense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium | gleum | Chryseobacterium | gleum | 1.94 | Chryseobacterium | gleum | 1.9 | B | 0 | 0 | 1 | Chryseobacterium gleum | Chryseobacterium gleum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium | hominis | Chryseobacterium | hominis | 2.55 | Chryseobacterium | hominis | 2.08 | A | 1 | 1 | 0 | Chryseobacterium hominis | Chryseobacterium hominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 2.1 | Chryseobacterium | tructae | 2.07 | A | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium tructae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 1.85 | Chryseobacterium | indologenes | 1.8 | B | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium indologenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium | manosquense | Chryseobacterium | manosquense | 2.22 | Deinococcus | wulumuqiensis | 1.45 | A | 1 | 1 | 0 | Chryseobacterium manosquense | Deinococcus wulumuqiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-6 | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium | manosquense | Chryseobacterium | manosquense | 2.13 | Staphylococcus | simiae | 1.28 | A | 1 | 1 | 0 | Chryseobacterium manosquense | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11450,4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium | sp-I-7 | Chryseobacterium | sp-I-7 | 2.7 | Chryseobacterium | geocarposphaerae | 1.42 | A | 1 | 1 | 0 | Chryseobacterium sp-I-7 | Chryseobacterium geocarposphaerae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | aquidurense | Flavobacterium | aquidurense | 2.54 | Flavobacterium | frigidimaris | 1.97 | A | 1 | 1 | 0 | Flavobacterium aquidurense | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5322,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | branchiarum | Flavobacterium | branchiarum | 2.67 | Flavobacterium | pectinovorum | 1.94 | A | 1 | 1 | 0 | Flavobacterium branchiarum | Flavobacterium pectinovorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | chungangense | Flavobacterium | chungangense | 2.41 | Flavobacterium | plurextorum | 1.82 | A | 1 | 1 | 0 | Flavobacterium chungangense | Flavobacterium plurextorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium | columnare | Flavobacterium | columnare | 2.39 | Flavobacterium | columnare | 2.3 | A | 1 | 1 | 0 | Flavobacterium columnare | Flavobacterium columnare | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.41 | Flavobacterium | frigidimaris | 2.04 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4217,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.6 | Flavobacterium | frigidimaris | 1.94 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium | hibernum | Flavobacterium | hibernum | 2 | Flavobacterium | saccharophilum | 1.97 | A | 1 | 1 | 0 | Flavobacterium hibernum | Flavobacterium saccharophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium | piscis | Flavobacterium | piscis | 2.78 | Flavobacterium | piscis | 2.29 | A | 1 | 1 | 0 | Flavobacterium piscis | Flavobacterium piscis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium | plurextorum | Flavobacterium | plurextorum | 2.62 | Flavobacterium | plurextorum | 2.35 | A | 1 | 1 | 0 | Flavobacterium plurextorum | Flavobacterium plurextorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5323,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | psychoterrae | Flavobacterium | psychoterrae | 2.41 | Flavobacterium | hydatis | 1.81 | A | 1 | 1 | 0 | Flavobacterium psychoterrae | Flavobacterium hydatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11550,4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.52 | Flavobacterium | psychrophilum | 2.4 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.3 | Flavobacterium | psychrophilum | 2.23 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | reichenbachii | Flavobacterium | reichenbachii | 2 | Flavobacterium | reichenbachii | 1.9 | A | 1 | 1 | 0 | Flavobacterium reichenbachii | Flavobacterium reichenbachii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium | sp | Flavobacterium | saccharophilum | 2.04 | Flavobacterium | hibernum | 1.97 | A | 0 | 0 | 1 | Flavobacterium saccharophilum | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | sp-C2 | Flavobacterium | sp-C2 | 2.4 | Paraburkholderia | phenazinium | 1.57 | A | 1 | 1 | 0 | Flavobacterium sp-C2 | Paraburkholderia phenazinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-10323 | Flavobacterium | sp-CVUAS-10323 | 2.31 | Flavobacterium | chungbukense | 1.65 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-10323 | Flavobacterium chungbukense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium | sp-CVUAS-4146 | Flavobacterium | sp-CVUAS-4146 | 2.25 | Flavobacterium | branchiarum | 2.07 | B | 0 | 0 | 1 | Flavobacterium sp-CVUAS-4146 | Flavobacterium branchiarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4216,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-4216,2 | Flavobacterium | sp-CVUAS-4216,2 | 2.07 | Flavobacterium | sp-JQMS | 1.47 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-4216,2 | Flavobacterium sp-JQMS | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-5506 | Flavobacterium | sp-CVUAS-5506 | 2.48 | Flavobacterium | johnsoniae | 1.82 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-5506 | Flavobacterium johnsoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | sp-HMD2009 | Flavobacterium | sp-HMD2009 | 2.45 | Flavobacterium | succinicans | 1.69 | A | 1 | 1 | 0 | Flavobacterium sp-HMD2009 | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium | sp-JQMS | Flavobacterium | sp-JQMS | 2.59 | Flavobacterium | succinicans | 1.67 | A | 1 | 1 | 0 | Flavobacterium sp-JQMS | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium | succinicans | Flavobacterium | succinicans | 2.49 | Flavobacterium | tiangeerense | 1.78 | A | 1 | 1 | 0 | Flavobacterium succinicans | Flavobacterium tiangeerense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | tiangeerense | Flavobacterium | tiangeerense | 2.61 | Flavobacterium | succinicans | 2.14 | B | 0 | 0 | 1 | Flavobacterium tiangeerense | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2589,2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas | tenax | Epilithonimonas | tenax | 2.34 | Chryseobacterium | aquaticum | 1.4 | A | 1 | 1 | 0 | Epilithonimonas tenax | Chryseobacterium aquaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.27 | Ornithobacterium | rhinotracheale | 2.24 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.21 | Ornithobacterium | rhinotracheale | 2.04 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus | sp-CVUAS-11266 | Deinococcus | sp-CVUAS-11266 | 2.76 | Acidovorax | facilis | 1.23 | A | 1 | 1 | 0 | Deinococcus sp-CVUAS-11266 | Acidovorax facilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium | necrophorum | Fusobacterium | necrophorum | 2.48 | Fusobacterium | necrophorum | 2.41 | A | 1 | 1 | 0 | Fusobacterium necrophorum | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium | varium | Fusobacterium | varium | 2.09 | Fusobacterium | varium | 2.04 | A | 1 | 1 | 0 | Fusobacterium varium | Fusobacterium varium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter | abscessus | Caviibacter | abscessus | 2.32 | Caviibacter | abscessus | 2.18 | A | 1 | 1 | 0 | Caviibacter abscessus | Caviibacter abscessus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga | salmonicida | Oceanivirga | salmonicida | 2.26 | Oceanivirga | salmonicida | 2.03 | A | 1 | 1 | 0 | Oceanivirga salmonicida | Oceanivirga salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus | hongkongensis | Pseudostreptobacillus | hongkongensis | 2.21 | Pseudostreptobacillus | hongkongensis | 1.81 | A | 1 | 1 | 0 | Pseudostreptobacillus hongkongensis | Pseudostreptobacillus hongkongensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella | termitidis | Sebaldella | termitidis | 2.16 | Sebaldella | termitidis | 1.65 | A | 1 | 1 | 0 | Sebaldella termitidis | Sebaldella termitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella | termitidis | Sebaldella | termitidis | 2.47 | Sebaldella | termitidis | 1.41 | A | 1 | 1 | 0 | Sebaldella termitidis | Sebaldella termitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 38322 | 3; public strain collection | EFEx | 16.12.2015 | Sneathia | sanguinegens | Sneathia | sanguinegens | 2.76 | Sneathia | sanguinegens | 2.27 | A | 1 | 1 | 0 | Sneathia sanguinegens | Sneathia sanguinegens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 41628 | 3; type strain | EFEx | 2020-10-08 | Sneathia | sanguinegens | Sneathia | sanguinegens | 2.49 | Sneathia | sanguinegens | 2.45 | A | 1 | 1 | 0 | Sneathia sanguinegens | Sneathia sanguinegens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Sneathia | vaginalis | 2.61 | Sneathia | vaginalis | 2.6 | A | 1 | 1 | 0 | Sneathia vaginalis | Sneathia vaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Sneathia | vaginalis | 2.57 | Sneathia | vaginalis | 2.56 | A | 1 | 1 | 0 | Sneathia vaginalis | Sneathia vaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus | felis | Streptobacillus | felis | 2.52 | Streptobacillus | felis | 2.5 | A | 1 | 1 | 0 | Streptobacillus felis | Streptobacillus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.47 | Streptobacillus | moniliformis | 2.43 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.58 | Streptobacillus | moniliformis | 2.57 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus | notomytis | Streptobacillus | notomytis | 2.46 | Streptobacillus | notomytis | 2.11 | A | 1 | 1 | 0 | Streptobacillus notomytis | Streptobacillus notomytis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus | ratti | Streptobacillus | ratti | 2.4 | Pseudostreptobacillus | hongkongensis | 1.43 | A | 1 | 1 | 0 | Streptobacillus ratti | Pseudostreptobacillus hongkongensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2280,3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas | glaciistagni | Aureimonas | glaciistagni | 2.53 | Bacillus | thuringiensis | 1.31 | A | 1 | 1 | 0 | Aureimonas glaciistagni | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM373) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; D02 | Brucella | abortus | Brucella | microti | 2.24 | Brucella | sp. | 2.15 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM369) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; E01 | Brucella | abortus | Brucella | microti | 2.12 | Brucella | sp. | 2.01 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; I16 | Brucella | canis | Brucella | microti | 2.1 | Brucella | suis | 2.01 | B | 0 | 0 | 1 | Brucella microti | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.03.2014; E17 | Brucella | canis | Brucella | microti | 2.05 | Brucella | sp. | 1.95 | A | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM389) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; F04 | Brucella | ceti | Brucella | microti | 2.2 | Brucella | sp. | 2.03 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM403) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; K04 | Brucella | ceti | Brucella | microti | 2.21 | Brucella | sp. | 2.07 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM374) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I02 | Brucella | melitensis | Brucella | sp. | 2.24 | Brucella | microti | 2.15 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM407) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; J05 | Brucella | melitensis | Brucella | sp. | 2.22 | Brucella | microti | 2.21 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM408) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; H06 | Brucella | melitensis | Brucella | microti | 2.18 | Brucella | sp. | 2.1 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM410) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I06 | Brucella | melitensis | Brucella | microti | 2.31 | Brucella | sp. | 2.1 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 4915 | 3; type strain | DT | 2013-11-26 | Brucella | microti | Brucella | microti | 2.67 | Brucella | sp. | 2.39 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 4915 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.03.2015 C7 | Brucella | microti | Brucella | microti | 2.41 | Brucella | sp. | 2.26 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM421) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; F07 | Brucella | neotomae | Brucella | sp. | 2.19 | Brucella | microti | 2.13 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.03.2015; C15 | Brucella | neotomae | Brucella | sp. | 2.17 | Brucella | microti | 2.04 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM423) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; L07 | Brucella | ovis | Brucella | sp. | 2.15 | Brucella | microti | 2.04 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; F16 | Brucella | ovis | Brucella | microti | 1.93 | Brucella | suis | 1.86 | B | 0 | 0 | 1 | Brucella microti | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM405) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; E05 | Brucella | pinnipedialis | Brucella | microti | 2.23 | Brucella | sp. | 2.18 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; C03 | Brucella | pinnipedialis | Brucella | sp. | 2.19 | Brucella | microti | 1.97 | A | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM431) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; B10 | Brucella | suis | Brucella | microti | 2.26 | Brucella | sp. | 2.12 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM427) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; G08 | Brucella | suis | Brucella | sp. | 2.27 | Brucella | microti | 2.25 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella | suis | Brucella | suis | 2.62 | Brucella | microti | 1.84 | A | 1 | 1 | 0 | Brucella suis | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM428) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I08 | Brucella | suis | Brucella | sp. | 2.31 | Brucella | microti | 2.23 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | (BM430) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; I09 | Brucella | suis | Brucella | microti | 2.22 | Brucella | sp. | 2.14 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 15.15.2016 I7 | Brucella | vulpis | Brucella | sp. | 1.97 | Brucella | microti | 1.83 | B | 0 | 0 | 1 | Brucella sp. | Brucella microti | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum | ovis | Falsochrobactrum | ovis | 2.48 | Falsochrobactrum | ovis | 2.41 | A | 1 | 1 | 0 | Falsochrobactrum ovis | Falsochrobactrum ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_08 A | 2; 0 | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 02.04.2012 K8 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.42 | Ochrobactrum | anthropi | 2.39 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 03.06.2016 H12 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.54 | Ochrobactrum | anthropi | 2.42 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012806 | 2; 0 | DT | 2017-10-04 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.18 | Ochrobactrum | intermedium | 2.11 | A | 1 | 1 | 0 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.55 | Ochrobactrum | intermedium | 2.53 | A | 1 | 1 | 0 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum | sp | Ochrobactrum | anthropi | 2.46 | Ochrobactrum | anthropi | 2.41 | A | 0 | 0 | 1 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum | sp | Ochrobactrum | intermedium | 1.78 | Ochrobactrum | anthropi | 1.68 | B | 0 | 0 | 1 | Ochrobactrum intermedium | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012040 | 2; 0 | DT | 2017-09-14 | Ochrobactrum | sp | Ochrobactrum | tritici | 1.78 | Ochrobactrum | intermedium | 1.65 | B | 0 | 0 | 1 | Ochrobactrum tritici | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter | aganoensis | Aminobacter | aganoensis | 2.39 | Pseudomonas | mucidolens | 1.45 | A | 1 | 1 | 0 | Aminobacter aganoensis | Pseudomonas mucidolens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter | aminovorans | Aminobacter | aminovorans | 2.56 | Lactobacillus | curvatus | 1.5 | A | 1 | 1 | 0 | Aminobacter aminovorans | Lactobacillus curvatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2015 H22 | Paracoccus | yeei | Paracoccus | yeei | 2.45 | Paracoccus | yeei | 2.24 | A | 1 | 1 | 0 | Paracoccus yeei | Paracoccus yeei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas | mucosae | Roseomonas | mucosa | 2.53 | Roseomonas | mucosa | 2.49 | A | 0 | 0 | 1 | Roseomonas mucosa | Roseomonas mucosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 24.01.2007, H7, [0034] | Inquilinus | limosus | Inquilinus | limosus | 2.63 | Inquilinus | limosus | 2.62 | A | 1 | 1 | 0 | Inquilinus limosus | Inquilinus limosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter | denitrificans | Achromobacter | denitrificans | 2.46 | Achromobacter | denitrificans | 2.22 | A | 1 | 1 | 0 | Achromobacter denitrificans | Achromobacter denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter | sp-DSM-30031 | Achromobacter | sp-DSM-30031 | 2.22 | Achromobacter | piechaudii | 1.95 | A | 1 | 1 | 0 | Achromobacter sp-DSM-30031 | Achromobacter piechaudii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2008 E5 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.33 | Achromobacter | xylosoxidans | 2.16 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.02.2016 N7 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.26 | Achromobacter | xylosoxidans | 2.23 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea | sp-CVUAS-31213 | Basilea | sp-CVUAS-31213 | 2.18 | Streptomyces | badius | 1.35 | A | 1 | 1 | 0 | Basilea sp-CVUAS-31213 | Streptomyces badius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella | avium | Bordetella | avium | 2.39 | Bordetella | avium | 2.29 | A | 1 | 1 | 0 | Bordetella avium | Bordetella avium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32946 | 2; 0 | EFEx | 2021-02-18 | Kerstersia | gyiorum | Kerstersia | gyiorum | 2.48 | Kerstersia | gyiorum | 2.48 | A | 1 | 1 | 0 | Kerstersia gyiorum | Kerstersia gyiorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 2.56 | Oligella | urethralis | 2.43 | A | 1 | 1 | 0 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 2.51 | Oligella | urethralis | 2.39 | A | 1 | 1 | 0 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga | kullae | Pigmentiphaga | kullae | 2.24 | Pigmentiphaga | daeguensis | 1.57 | A | 1 | 1 | 0 | Pigmentiphaga kullae | Pigmentiphaga daeguensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.32 | Taylorella | asinigenitalis | 2.01 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella asinigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.12 | Taylorella | asinigenitalis | 2.09 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella asinigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.15 | Taylorella | equigenitalis | 2.1 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.35 | Taylorella | equigenitalis | 2.24 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007, K7 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.27 | Burkholderia | ambifaria | 2.24 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16087 LMG 19182 | 3; type starin | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K19 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.36 | Burkholderia | ambifaria | 2.15 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K21 | Burkholderia | anthina | Burkholderia | anthina | 2.35 | Burkholderia | anthina | 2.27 | A | 1 | 1 | 0 | Burkholderia anthina | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.08.2007, N10 | Burkholderia | anthina | Burkholderia | pyrrocinia | 2.17 | Burkholderia | anthina | 2.16 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2010, K12 | Burkholderia | arboris | Burkholderia | pyrrocinia | 2.34 | Burkholderia | cepacia | 2.23 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 3.4.2017; L6 | Burkholderia | arboris | Burkholderia | cenocepacia | 2.29 | Burkholderia | vietnamiensis | 2.22 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K11 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.48 | Burkholderia | cenocepacia | 2.36 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.07.2009; K21 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.42 | Burkholderia | cenocepacia | 2.17 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 4.5.2006; H15 | Burkholderia | cepacia | Burkholderia | cenocepacia | 2.46 | Burkholderia | cenocepacia | 2.15 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 9.5.2007; N9 | Burkholderia | cepacia | Burkholderia | cepacia | 2.4 | Burkholderia | pyrrocinia | 2.34 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.11.2006; H10 | Burkholderia | contaminans | Burkholderia | cepacia | 2.28 | Burkholderia | cepacia | 2.16 | A | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E10 | Burkholderia | contaminans | Burkholderia | cepacia | 2.42 | Burkholderia | cenocepacia | 2.17 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016 C13 | Burkholderia | diffusa | Burkholderia | ambifaria | 2.35 | Burkholderia | vietnamiensis | 2.32 | B | 0 | 0 | 1 | Burkholderia ambifaria | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L17 | Burkholderia | diffusa | Burkholderia | cenocepacia | 2.31 | Burkholderia | ambifaria | 2.21 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K16 | Burkholderia | dolosa | Burkholderia | dolosa | 2.29 | Burkholderia | multivorans | 2.15 | B | 0 | 0 | 1 | Burkholderia dolosa | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.11.2006; E09 | Burkholderia | dolosa | Burkholderia | dolosa | 2.08 | Burkholderia | cenocepacia | 2.07 | B | 0 | 0 | 1 | Burkholderia dolosa | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.01.2007; H22 | Burkholderia | gladioli | Burkholderia | gladioli | 2.45 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H15 | Burkholderia | gladioli | Burkholderia | gladioli | 2.42 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K03 | Burkholderia | glumae | Burkholderia | glumae | 2.33 | Burkholderia | plantarii | 2.06 | B | 0 | 0 | 1 | Burkholderia glumae | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.05.2012; K22 | Burkholderia | glumae | Burkholderia | glumae | 2.33 | Burkholderia | plantarii | 2.13 | B | 0 | 0 | 1 | Burkholderia glumae | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016; C10 | Burkholderia | lata | Burkholderia | cepacia | 2.49 | Burkholderia | lata | 2.41 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia lata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.08.2016; K19 | Burkholderia | lata | Burkholderia | cepacia | 2.47 | Burkholderia | lata | 2.38 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia lata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.05.2016; L06 | Burkholderia | latens | Burkholderia | latens | 2.32 | Burkholderia | vietnamiensis | 2.29 | B | 0 | 0 | 1 | Burkholderia latens | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.09.2009; K20 | Burkholderia | mallei | Burkholderia | thailandensis | 1.8 | Burkholderia | thailandensis | 1.65 | B | 0 | 0 | 1 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.08.2007; H13 | Burkholderia | mallei | Burkholderia | thailandensis | 1.91 | Burkholderia | thailandensis | 1.73 | B | 0 | 0 | 1 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2010; H14 | Burkholderia | metallica | Burkholderia | metallica | 2.4 | Burkholderia | cenocepacia | 2.27 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L21 | Burkholderia | metallica | Burkholderia | metallica | 2.31 | Burkholderia | cenocepacia | 2.23 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.12.2006; K19 | Burkholderia | multivorans | Burkholderia | multivorans | 2.22 | Burkholderia | multivorans | 2.21 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.11.2006; K04 | Burkholderia | multivorans | Burkholderia | multivorans | 2.13 | Burkholderia | multivorans | 2 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K17 | Burkholderia | oklahomensis | Burkholderia | multivorans | 1.77 | Burkholderia | ambifaria | 1.76 | B | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K18 | Burkholderia | oklahomensis | Burkholderia | thailandensis | 1.72 | Burkholderia | ambifaria | 1.65 | B | 0 | 0 | 1 | Burkholderia thailandensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K06 | Burkholderia | plantarii | Burkholderia | plantarii | 2.49 | Burkholderia | pyrrocinia | 1.89 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K05 | Burkholderia | plantarii | Burkholderia | plantarii | 2.5 | Burkholderia | pyrrocinia | 1.94 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012; K10 | Burkholderia | pseudomallei | Burkholderia | thailandensis | 2.03 | Burkholderia | thailandensis | 1.69 | A | 0 | 0 | 1 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; H21 | Burkholderia | pseudomallei | Burkholderia | thailandensis | 2.03 | Burkholderia | thailandensis | 1.73 | A | 0 | 0 | 1 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K07 | Burkholderia | pseudomultivorans | Burkholderia | multivorans | 2.1 | Burkholderia | multivorans | 2.07 | A | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K08 | Burkholderia | pseudomultivorans | Burkholderia | multivorans | 2.03 | Burkholderia | cenocepacia | 1.93 | A | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CAMPA 1043 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; K04 | Burkholderia | puraquae | Burkholderia | pyrrocinia | 2.27 | Burkholderia | cepacia_Group | 2.2 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | R-50214 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.03.2016; C04 | Burkholderia | puraquae | Burkholderia | pyrrocinia | 2.41 | Burkholderia | cepacia_Group | 2.32 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; N06 | Burkholderia | pyrrocinia | Burkholderia | pyrrocinia | 2.57 | Burkholderia | cepacia_Group | 2.33 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; I21 | Burkholderia | pyrrocinia | Burkholderia | cenocepacia | 2.13 | Burkholderia | pyrrocinia | 2.12 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H08 | Burkholderia | seminalis | Burkholderia | seminalis | 2.41 | Burkholderia | cepacia | 2.32 | B | 0 | 0 | 1 | Burkholderia seminalis | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.01.2007; K12 | Burkholderia | seminalis | Burkholderia | cepacia | 2.29 | Burkholderia | seminalis | 2.11 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia seminalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K13 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.56 | Burkholderia | stabilis | 2.44 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E08 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.51 | Burkholderia | stabilis | 2.28 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; C08 | Burkholderia | stagnalis | Burkholderia | anthina | 2.14 | Burkholderia | ambifaria | 2.11 | B | 0 | 0 | 1 | Burkholderia anthina | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F14 | Burkholderia | stagnalis | Burkholderia | vietnamiensis | 2.16 | Burkholderia | ambifaria | 2.01 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007; H19 | Burkholderia | taxonK | Burkholderia | cepacia | 2.45 | Burkholderia | cenocepacia | 2.16 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.06.2009; K13 | Burkholderia | taxonK | Burkholderia | cepacia | 2.23 | Burkholderia | cenocepacia | 2.13 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F16 | Burkholderia | territorii | Burkholderia | cenocepacia | 2.3 | Burkholderia | cenocepacia | 2.05 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; E14 | Burkholderia | territorii | Burkholderia | cenocepacia | 2.28 | Burkholderia | vietnamiensis | 2.21 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.04.2012, K6 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.38 | Burkholderia | thailandensis | 2.17 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13276 n | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 24.11.2010, K19 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.34 | Burkholderia | thailandensis | 2.26 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H11 | Burkholderia | ubonensis | Burkholderia | cenocepacia | 2.19 | Burkholderia | cepacia | 2.07 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; I04 | Burkholderia | ubonensis | Burkholderia | cepacia_Group | 2.1 | Burkholderia | plantarii | 2.09 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9631=LMG 6999 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.05.2007 N15 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.33 | Burkholderia | cenocepacia | 2.2 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10929 V | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.12.2006 N9 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.43 | Burkholderia | pyrrocinia | 2.14 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.04.2012, N07 | Caballeronia | glathei | Caballeronia | glathei | 2.55 | Caballeronia | turbans | 1.53 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H12 | Caballeronia | glathei | Caballeronia | glathei | 2.47 | Caballeronia | turbans | 1.56 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H15 | Pandoraea | apista | Pandoraea | apista | 2.36 | Pandoraea | apista | 2.28 | A | 1 | 1 | 0 | Pandoraea apista | Pandoraea apista | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H21 | Pandoraea | pnomenusa | Pandoraea | pnomenusa | 2.4 | Pandoraea | pnomenusa | 2.29 | A | 1 | 1 | 0 | Pandoraea pnomenusa | Pandoraea pnomenusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H16 | Pandoraea | pulmonicola | Pandoraea | pulmonicola | 2.49 | Pandoraea | pulmonicola | 2.25 | A | 1 | 1 | 0 | Pandoraea pulmonicola | Pandoraea pulmonicola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K7 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.52 | Paraburkholderia | caledonica | 2.45 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K8 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.46 | Paraburkholderia | caledonica | 2.37 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N9 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.25 | Paraburkholderia | azotifigens | 2.23 | B | 0 | 0 | 1 | Paraburkholderia caribensis | Paraburkholderia azotifigens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N10 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.23 | Paraburkholderia | azotifigens | 2.18 | B | 0 | 0 | 1 | Paraburkholderia caribensis | Paraburkholderia azotifigens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H03 | Paraburkholderia | hospita | Paraburkholderia | azotifigens | 1.81 | Lactobacillus | fructivorans | 1.71 | C | 0 | 0 | 1 | Paraburkholderia azotifigens | Lactobacillus fructivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H04 | Paraburkholderia | hospita | Paraburkholderia | hospita | 2.07 | Paraburkholderia | azotifigens | 1.9 | A | 1 | 1 | 0 | Paraburkholderia hospita | Paraburkholderia azotifigens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 1.12.2010; H15 | Paraburkholderia | kururiensis | Burkholderia | pyrrocinia | 1.64 | Burkholderia | cenocepacia | 1.54 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.11.2010; N19 | Paraburkholderia | phymatum | Paraburkholderia | azotifigens | 1.82 | Paraburkholderia | caribensis | 1.7 | B | 0 | 0 | 1 | Paraburkholderia azotifigens | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.11.2010; K17 | Paraburkholderia | phytofirmans | Paraburkholderia | phytofirmans | 2.44 | Paraburkholderia | phytofirmans | 2.42 | A | 1 | 1 | 0 | Paraburkholderia phytofirmans | Paraburkholderia phytofirmans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H16 | Paraburkholderia | tuberum | Paraburkholderia | tuberum | 2.36 | Paraburkholderia | xenovorans | 1.59 | A | 1 | 1 | 0 | Paraburkholderia tuberum | Paraburkholderia xenovorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.11.2010, N21 | Paraburkholderia | xenovorans | Paraburkholderia | xenovorans | 2.49 | Paraburkholderia | dipogonis | 2.33 | B | 0 | 0 | 1 | Paraburkholderia xenovorans | Paraburkholderia dipogonis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 E07 | Ralstonia | insidiosa | Ralstonia | insidiosa | 2.38 | Ralstonia | insidiosa | 2.35 | A | 1 | 1 | 0 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 C06 | Ralstonia | mannitolilytica | Ralstonia | mannitolilytica | 2.46 | Ralstonia | mannitolilytica | 2.22 | A | 1 | 1 | 0 | Ralstonia mannitolilytica | Ralstonia mannitolilytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27511 DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 22.01.2007 H03 | Ralstonia | pickettii | Ralstonia | pickettii | 2.15 | Ralstonia | pickettii | 2.12 | A | 1 | 1 | 0 | Ralstonia pickettii | Ralstonia pickettii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 07.05.2009 N15 | Ralstonia | pickettii | Ralstonia | insidiosa | 2.07 | Ralstonia | insidiosa | 2.07 | A | 0 | 0 | 1 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia | acidovorans | Delftia | acidovorans | 2.54 | Delftia | tsuruhatensis | 2.44 | B | 0 | 0 | 1 | Delftia acidovorans | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia | tsuruhatensis | Delftia | tsuruhatensis | 2.45 | Delftia | tsuruhatensis | 2.31 | A | 1 | 1 | 0 | Delftia tsuruhatensis | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga | intermedia | Hydrogenophaga | intermedia | 2.45 | Hydrogenophaga | intermedia | 2.4 | A | 1 | 1 | 0 | Hydrogenophaga intermedia | Hydrogenophaga intermedia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.07.2013 L13 | Xenophilus | sp | Aromatoleum | evansii | 1.43 | Ligilactobacillus | salivarius | 1.32 | B | 0 | 0 | 1 | Aromatoleum evansii | Ligilactobacillus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24713 extr | 3; type strain | Extr | 2016-02-02 | Deefgea | chitinilytica | Deefgea | chitinilytica | 2.72 | Mannheimia | haemolytica | 1.42 | A | 1 | 1 | 0 | Deefgea chitinilytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24713 | 3; type strain | eDT | 2015-09-22 | Deefgea | chitinilytica | Deefgea | chitinilytica | 2.11 | Pichia | cactophila | 1.24 | A | 1 | 1 | 0 | Deefgea chitinilytica | Pichia cactophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea | rivuli | Deefgea | rivuli | 2.6 | Arthrobacter | woluwensis | 1.45 | A | 1 | 1 | 0 | Deefgea rivuli | Arthrobacter woluwensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 08-F17.2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Deefgea | sp-CVUAS-2703.2 | 2.03 | Deefgea | chitinilytica | 1.7 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-2703.2 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Deefgea | sp-CVUAS-2703.2 | 2.08 | Deefgea | chitinilytica | 1.65 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-2703.2 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6995,3 | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea | sp-CVUAS-6995,3 | Deefgea | sp-CVUAS-6995.3 | 2.61 | Deefgea | sp-CVUAS-9920 | 1.56 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-6995.3 | Deefgea sp-CVUAS-9920 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-6995,3 | Deefgea | sp-CVUAS-6995.3 | 2.38 | Pseudomonas | orientalis | 1.48 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-6995.3 | Pseudomonas orientalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-9920 | Deefgea | sp-CVUAS-9920 | 2.42 | Deefgea | sp-CVUAS-2703.2 | 1.53 | A | 1 | 1 | 0 | Deefgea sp-CVUAS-9920 | Deefgea sp-CVUAS-2703.2 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea | sp-CVUAS-9920 | Deefgea | sp-CVUAS-9920 | 2.21 | Deefgea | chitinilytica | 1.37 | A | 1 | 1 | 0 | Deefgea sp-CVUAS-9920 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter | limnosediminis | Iodobacter | limnosediminis | 2.6 | Iodobacter | fluviatilis | 1.45 | A | 1 | 1 | 0 | Iodobacter limnosediminis | Iodobacter fluviatilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella | kingae | Kingella | kingae | 2.49 | Kingella | kingae | 2.41 | A | 1 | 1 | 0 | Kingella kingae | Kingella kingae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10067,2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria | animaloris | Neisseria | animaloris | 2.41 | Neisseria | animaloris | 2.17 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria | animaloris | Neisseria | animaloris | 2.65 | Neisseria | animaloris | 2.42 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-4931 | Neisseria | sp-CVUAS-4931 | 2.56 | Neisseria | animaloris | 1.42 | A | 1 | 1 | 0 | Neisseria sp-CVUAS-4931 | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-8997 | Neisseria | sp-CVUAS-8997 | 2.27 | Neisseria_flavescens_subflava | group | 1.48 | A | 1 | 1 | 0 | Neisseria sp-CVUAS-8997 | Neisseria_flavescens_subflava group | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-LHL-1296 | Neisseria | sp-LHL-1296 | 2.59 | Neisseria | animaloris | 1.61 | A | 1 | 1 | 0 | Neisseria sp-LHL-1296 | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5699,3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria | sp-N95-16 | Neisseria | sp-N95-16 | 2.73 | Bergeriella | denitrificans | 1.99 | C | 0 | 0 | 1 | Neisseria sp-N95-16 | Bergeriella denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae | sp-CVUAS-8636 | Neisseriaceae | sp-CVUAS-8636 | 2.14 | Pseudomonas | taetrolens | 1.43 | A | 1 | 1 | 0 | Neisseriaceae sp-CVUAS-8636 | Pseudomonas taetrolens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter | cryaerophilus | Arcobacter | cryaerophilus | 2.03 | Arcobacter | cryaerophilus | 2.01 | A | 0 | 0 | 1 | Arcobacter cryaerophilus | Arcobacter cryaerophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.06 | Aliarcobacter | butzleri | 1.94 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.24 | Aliarcobacter | butzleri | 2.21 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter | coli | Campylobacter | coli | 2.18 | Campylobacter | coli | 1.97 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4689 | 3; type strain | eDT | 2015-12-30 | Campylobacter | coli | Campylobacter | coli | 2.24 | Campylobacter | coli | 2.21 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter | fetus | Campylobacter | fetus | 2.34 | Campylobacter | fetus | 2.28 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.37 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5361 | 3; type strain | DT | 2013-09-24 | Campylobacter | fetus | Campylobacter | fetus | 2.51 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.19 | Campylobacter | hyointestinalis | 2.12 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6673,2 | 3; sequenced | DT | 2013-08-28 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.15 | Campylobacter | hyointestinalis | 2.06 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter | jejuni | Campylobacter | jejuni | 2.19 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter | jejuni | Campylobacter | jejuni | 2.39 | Campylobacter | jejuni | 2.32 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter | jejuni | Campylobacter | jejuni | 2.16 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | lari | Campylobacter | lari | 2.19 | Campylobacter | lari | 2.04 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-06556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.04 | Campylobacter | lari | 1.82 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.2 | Campylobacter | lari | 2.11 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 11375 | 3; type strain | eDT | 2015-12-30 | Campylobacter | lari | Campylobacter | lari | 2.42 | Campylobacter | lari | 2.19 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-07048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Burkholderia | plantarii | 1.42 | A | 1 | 1 | 0 | Campylobacter peloridis | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Campylobacter | sp-W1 | 1.48 | A | 1 | 1 | 0 | Campylobacter peloridis | Campylobacter sp-W1 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter | sputorum | Campylobacter | sputorum | 2.01 | Campylobacter | sputorum | 1.9 | A | 1 | 1 | 0 | Campylobacter sputorum | Campylobacter sputorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter | sp-W1 | Campylobacter | sp-W1 | 2.61 | Campylobacter | jejuni | 1.49 | A | 1 | 1 | 0 | Campylobacter sp-W1 | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter | upsaliensis | Campylobacter | upsaliensis | 2.05 | Campylobacter | upsaliensis | 1.96 | A | 1 | 1 | 0 | Campylobacter upsaliensis | Campylobacter upsaliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11185,3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter | canadensis | Helicobacter | canadensis | 2.72 | Clostridium | cochlearium | 1.42 | A | 1 | 1 | 0 | Helicobacter canadensis | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas | bestiarum | Aeromonas | bestiarum | 2.44 | Aeromonas | salmonicida | 2.42 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas | bivalvium | Aeromonas | veronii | 2.4 | Aeromonas | bivalvium | 2.18 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas bivalvium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas | caviae | Aeromonas | caviae | 2.29 | Aeromonas | hydrophila | 2.28 | B | 0 | 0 | 1 | Aeromonas caviae | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | hydrophila | 2.03 | Aeromonas | dhakensis | 2 | B | 0 | 0 | 1 | Aeromonas hydrophila | Aeromonas dhakensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | dhakensis | 2.42 | Aeromonas | dhakensis | 2.41 | A | 1 | 1 | 0 | Aeromonas dhakensis | Aeromonas dhakensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30187 RKI | 4; type strain, WGS sequenced (BfR) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E17 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.57 | Aeromonas | hydrophila | 2.46 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.27 | Aeromonas | hydrophila | 2.25 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas | jandaei | Aeromonas | jandaei | 2.34 | Aeromonas | jandaei | 2.05 | A | 1 | 1 | 0 | Aeromonas jandaei | Aeromonas jandaei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | media | Aeromonas | media | 2.22 | Aeromonas | media | 2.2 | A | 1 | 1 | 0 | Aeromonas media | Aeromonas media | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas | molluscorum | Aeromonas | molluscorum | 2.28 | Aeromonas | molluscorum | 2.15 | A | 1 | 1 | 0 | Aeromonas molluscorum | Aeromonas molluscorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | popoffii | Aeromonas | popoffii | 2.5 | Aeromonas | salmonicida | 2.26 | B | 0 | 0 | 1 | Aeromonas popoffii | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.22 | Aeromonas | salmonicida | 2.07 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.31 | Aeromonas | salmonicida | 2.25 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.55 | Aeromonas | salmonicida | 2.38 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.53 | Aeromonas | salmonicida | 2.22 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8400,2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.19 | Aeromonas | salmonicida | 2.07 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.38 | Aeromonas | salmonicida | 2.36 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.47 | Aeromonas | salmonicida | 2.3 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.18 | Aeromonas | salmonicida | 2.11 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas | simiae | Aeromonas | simiae | 2.31 | Aeromonas | simiae | 2.07 | A | 1 | 1 | 0 | Aeromonas simiae | Aeromonas simiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas | veronii | Aeromonas | veronii | 2.46 | Aeromonas | ichthiosmia | 2.18 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas ichthiosmia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7386 | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas | veronii | Aeromonas | veronii | 2.46 | Aeromonas | bivalvium | 2.26 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas bivalvium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas | pectinilytica | Pseudaeromonas | pectinilytica | 2.62 | Aeromonas | eucrenophila | 1.4 | A | 1 | 1 | 0 | Pseudaeromonas pectinilytica | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella | viscosa | Moritella | viscosa | 2.32 | Moritella | viscosa | 2.22 | A | 1 | 1 | 0 | Moritella viscosa | Moritella viscosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella | viscosa | Moritella | viscosa | 2.29 | Moritella | viscosa | 2.23 | A | 1 | 1 | 0 | Moritella viscosa | Moritella viscosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | sp-CVUAS-9366 | 2.64 | Shewanella | sp-CVUAS-9366 | 2.62 | A | 1 | 1 | 0 | Shewanella sp-CVUAS-9366 | Shewanella sp-CVUAS-9366 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | sp-CVUAS-9366 | 2.55 | Shewanella | sp-CVUAS-9366 | 2.5 | A | 1 | 1 | 0 | Shewanella sp-CVUAS-9366 | Shewanella sp-CVUAS-9366 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 81,2 | 3; sequenced | EFEx | 28:11.2013 | Shewanella | sp-RKVR22 | Shewanella | sp-RKVR22 | 2.4 | Shewanella | corallii | 1.59 | A | 1 | 1 | 0 | Shewanella sp-RKVR22 | Shewanella corallii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, H5 | Citrobacter | amalonaticus | Citrobacter | amalonaticus | 2.41 | Citrobacter | amalonaticus | 2.3 | A | 1 | 1 | 0 | Citrobacter amalonaticus | Citrobacter amalonaticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | freundii | 2.39 | Citrobacter | freundii | 2.37 | A | 1 | 1 | 0 | Citrobacter freundii | Citrobacter freundii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | braakii | 2.29 | Citrobacter | braakii | 2.23 | A | 0 | 0 | 1 | Citrobacter braakii | Citrobacter braakii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25408 | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 11.04.2008, N11, [0034]: Citrobacter diversus ATCC 25408\LMTG70N11_0\0_N11\1\1SLin\fid | Citrobacter | koseri | Citrobacter | koseri | 2.52 | Citrobacter | koseri | 2.47 | A | 1 | 1 | 0 | Citrobacter koseri | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter | condimenti | Cronobacter | condimenti | 2.49 | Cronobacter | malonaticus | 2.29 | B | 0 | 0 | 1 | Cronobacter condimenti | Cronobacter malonaticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter | dublinensis | Cronobacter | muytjensii | 2.22 | Cronobacter | sp | 2.2 | B | 0 | 0 | 1 | Cronobacter muytjensii | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18702 | 3; type strain | DT | 2019-03-05 | Cronobacter | malonaticus | Cronobacter | malonaticus | 2.56 | Cronobacter | sp | 2.46 | B | 0 | 0 | 1 | Cronobacter malonaticus | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-15 | Cronobacter | malonaticus | Cronobacter | sp | 2.29 | Cronobacter | malonaticus | 2.24 | B | 0 | 0 | 1 | Cronobacter sp | Cronobacter malonaticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32945 | 2; 0 | EFEx | 2021-02-18 | Cronobacter | sakazakii | Cronobacter | sp | 2.14 | Cronobacter | sp | 1.99 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter | sakazakii | Cronobacter | sp | 2.44 | Cronobacter | sp | 2.33 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter | turicensis | Cronobacter | turicensis | 2.52 | Cronobacter | sp | 2.36 | B | 0 | 0 | 1 | Cronobacter turicensis | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.11.2011 E6 | Enterobacter | cloacae | Enterobacter | hormaechei | 2.06 | Enterobacter | hormaechei | 2.04 | A | 0 | 0 | 1 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30054 2 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E18, [0034] | Enterobacter | cloacae | Enterobacter | cloacae | 2.34 | Enterobacter | cloacae | 2.08 | A | 1 | 1 | 0 | Enterobacter cloacae | Enterobacter cloacae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K4, [0034] | Enterobacter | sp-ATCC-27988 | Kosakonia | radicincitans | 1.94 | Raoultella | ornithinolytica | 1.68 | B | 0 | 0 | 1 | Kosakonia radicincitans | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae | sp-CVUAS-30072 | Enterobacteriaceae | sp-CVUAS-30072 | 2.52 | Kosakonia | cowanii | 1.45 | A | 1 | 1 | 0 | Enterobacteriaceae sp-CVUAS-30072 | Kosakonia cowanii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia | coli | Escherichia | coli | 2.32 | Escherichia | coli | 2.29 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0500,2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia | coli | Escherichia | coli | 2.41 | Escherichia | coli | 2.34 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30053 | 3; type strain | DT | 2013-09-17 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.53 | Klebsiella | aerogenes | 2.46 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NM 20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E20 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.48 | Klebsiella | aerogenes | 2.29 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella | grimontii | Klebsiella | oxytoca | 2.42 | Klebsiella | grimontii | 2.39 | B | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella grimontii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2486,3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella | grimontii | Klebsiella | grimontii | 2.66 | Klebsiella | michiganensis | 2.35 | B | 0 | 0 | 1 | Klebsiella grimontii | Klebsiella michiganensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11699 | 3; sequenced [BfR] | eDT | 2017-04-20 06:56 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.47 | Klebsiella | oxytoca | 2.38 | B | 0 | 0 | 1 | Klebsiella michiganensis | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2404,3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.71 | Klebsiella | grimontii | 2.46 | B | 0 | 0 | 1 | Klebsiella michiganensis | Klebsiella grimontii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.67 | Klebsiella | michiganensis | 2.37 | A | 1 | 1 | 0 | Klebsiella michiganensis | Klebsiella michiganensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.4 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10038,2 | 3; sequenced [BfR] | EFex | 2019-02-12 14:02 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.41 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11572 | 3; sequenced [BfR] | eDT | 2017-04-20 06:51 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.39 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.11.2013 E18 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.42 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1482,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:35 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.24 | Klebsiella | pneumoniae | 2.24 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1616,4 | 3; sequenced [BfR] | eDT | 2017-11-22 09:55 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.42 | Klebsiella | pneumoniae | 2.39 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1678 | 3; sequenced [BfR] | eDT | 2016-10-18 06:38 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.44 | Klebsiella | pneumoniae | 2.43 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1701 | 3; sequenced [BfR] | eDT | 2016-10-18 06:41 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.5 | Klebsiella | pneumoniae | 2.44 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30499 | 3; sequenced [BfR] | eDT | 2017-11-22 09:59 | Klebsiella | pneumoniae | Klebsiella | quasipneumoniae | 2.29 | Klebsiella | pneumoniae | 2.24 | B | 0 | 0 | 1 | Klebsiella quasipneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 753,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:34 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.44 | Klebsiella | pneumoniae | 2.32 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 11.04.2008 N21 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.33 | Klebsiella | pneumoniae | 2.33 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30104 | 3; type strain | EFEx | 2013-09-17 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.52 | Klebsiella | pneumoniae | 2.4 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10109 | 3; sequenced [BfR] | eDT | 2016-10-18 07:00 | Klebsiella | quasipneumoniae | Klebsiella | pneumoniae | 2.34 | Klebsiella | pneumoniae | 2.24 | A | 0 | 0 | 1 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | quasipneumoniae | 2.7 | Klebsiella | quasipneumoniae | 2.3 | A | 1 | 1 | 0 | Klebsiella quasipneumoniae | Klebsiella quasipneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | quasipneumoniae | 2.49 | Klebsiella | quasipneumoniae | 2.41 | A | 1 | 1 | 0 | Klebsiella quasipneumoniae | Klebsiella quasipneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella | variicola | Klebsiella | variicola | 2.12 | Klebsiella | variicola | 2.08 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11677 | 3; sequenced [BfR] | eDT | 2017-04-20 06:54 | Klebsiella | variicola | Klebsiella | variicola | 2.52 | Klebsiella | variicola | 2.42 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4888,2 | 3; sequenced [BfR] | EFex | 2017-07-19 11:36 | Klebsiella | variicola | Klebsiella | variicola | 2.54 | Klebsiella | variicola | 2.39 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | variicola | Klebsiella | variicola | 2.58 | Klebsiella | variicola | 2.52 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia | sacchari | Kosakonia | sacchari | 2.29 | Salmonella | sp | 1.72 | C | 0 | 0 | 1 | Kosakonia sacchari | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella | morganii | Morganella | morganii | 2.56 | Morganella | morganii | 2.55 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella | morganii | Morganella | morganii | 2.68 | Morganella | morganii | 2.61 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea | 0 | Pantoea | sp | 1.99 | Pantoea | agglomerans | 1.93 | B | 0 | 0 | 1 | Pantoea sp | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea | agglomerans | Pantoea | agglomerans | 2.52 | Pantoea | agglomerans | 2.47 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3493 | 3; type strain | eDT | 2015-07-09 | Pantoea | agglomerans | Pantoea | agglomerans | 2.35 | Pantoea | agglomerans | 2.31 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4263,2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea | sp-FF5 | Pantoea | sp-FF5 | 2.72 | Pantoea | septica | 1.84 | A | 1 | 1 | 0 | Pantoea sp-FF5 | Pantoea septica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.36 | Plesiomonas | shigelloides | 2.34 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.42 | Plesiomonas | shigelloides | 2.41 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N17, [0034] (Enterobacter gergoviae) | Pluralibacter | gergoviae | Pluralibacter | gergoviae | 2.32 | Pluralibacter | gergoviae | 2.24 | A | 1 | 1 | 0 | Pluralibacter gergoviae | Pluralibacter gergoviae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LT21 Ib 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H9 | Salmonella | enterica | Salmonella | sp | 2.36 | Salmonella | sp | 2.36 | A | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | II 06 06820 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H6 | Shigella | boydii | Escherichia | coli | 2.43 | Escherichia | coli | 2.39 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 3A 07 02503 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H7 | Shigella | flexneri | Escherichia | coli | 2.46 | Escherichia | coli | 2.37 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella | tarda | Edwardsiella | tarda | 2.47 | Edwardsiella | tarda | 2.47 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N13, [0034] | Edwardsiella | tarda | Edwardsiella | tarda | 2.49 | Edwardsiella | tarda | 2.37 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 18.04.2008, K9, [0034] | Hafnia | alvei | Hafnia | alvei | 2.42 | Hafnia | alvei | 2.32 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia | alvei | Hafnia | alvei | 2.36 | Hafnia | alvei | 2.29 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea | myxofaciens | Cosenzaea | myxofaciens | 2.56 | Cosenzaea | myxofaciens | 2.54 | A | 1 | 1 | 0 | Cosenzaea myxofaciens | Cosenzaea myxofaciens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus | cibarius | Proteus | cibarius | 2.61 | Proteus | terrae | 2.5 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | hauseri | Proteus | hauseri | 2.6 | Proteus | hauseri | 2.55 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus | hauseri | Proteus | hauseri | 2.47 | Proteus | hauseri | 2.45 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus | mirabilis | Proteus | mirabilis | 2.56 | Proteus | mirabilis | 2.45 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4479 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | mirabilis | Proteus | mirabilis | 2.52 | Proteus | mirabilis | 2.46 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | penneri | Proteus | penneri | 2.48 | Proteus | penneri | 2.43 | A | 1 | 1 | 0 | Proteus penneri | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | cibarius | 2.51 | Proteus | terrae | 2.33 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | vulgaris | 2.45 | Proteus | penneri | 2.41 | B | 0 | 0 | 1 | Proteus vulgaris | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | cibarius | 2.42 | Proteus | terrae | 2.4 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | terrae | Proteus | terrae | 2.66 | Proteus | cibarius | 2.44 | B | 0 | 0 | 1 | Proteus terrae | Proteus cibarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, E5 K16, [0034]: Proteus vulgaris ATCC 33420\LMTG72K16_0\0_K16\1\1SLin\fid | Proteus | vulgaris | Proteus | vulgaris | 2.3 | Proteus | vulgaris | 2.28 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2140 | 3; strain from public collection; sequenced (rpoB) | eDT | 30.11.2017 | Proteus | vulgaris | Proteus | vulgaris | 2.38 | Proteus | vulgaris | 2.33 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia | rettgeri | Providencia | rettgeri | 2.57 | Providencia | rettgeri | 2.52 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia | rettgeri | Providencia | rettgeri | 2.54 | Providencia | rettgeri | 2.31 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25827 | 3; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K15, [0034]: Proteus stuartii ATCC 25827\LMTG72K15_0\0_K15\1\1SLin\fid in vers 2016 | Providencia | stuartii | Providencia | stuartii | 2.32 | Providencia | stuartii | 2.31 | A | 1 | 1 | 0 | Providencia stuartii | Providencia stuartii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia | fonticola | Serratia | fonticola | 2.56 | Serratia | fonticola | 2.55 | A | 1 | 1 | 0 | Serratia fonticola | Serratia fonticola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.11.2008, E11, [0034] | Serratia | grimesii | Serratia | grimesii | 2.42 | Serratia | liquefaciens | 2.03 | B | 0 | 0 | 1 | Serratia grimesii | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.19 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.13 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia | quinivorans | Serratia | proteamaculans | 2.16 | Serratia | liquefaciens | 2.12 | B | 0 | 0 | 1 | Serratia proteamaculans | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6221 | 4; sequenced | DT | 2013-01-08 | Yersinia | aldovae | Yersinia | aldovae | 2.33 | Yersinia | aldovae | 2.32 | A | 1 | 1 | 0 | Yersinia aldovae | Yersinia aldovae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 09.03.2007, H12 | Yersinia | aleksiciae | Yersinia | aleksiciae | 2.28 | Yersinia | aleksiciae | 2.28 | A | 1 | 1 | 0 | Yersinia aleksiciae | Yersinia aleksiciae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia | bercovieri | Yersinia | bercovieri | 2.33 | Yersinia | bercovieri | 2.25 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.56 | Yersinia | bercovieri | 2.45 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2012 001153 | 3; sequenced | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.49 | Yersinia | bercovieri | 2.44 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4780 | 3; type strain | DT | 2013-05-08 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.49 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.48 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.43 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.34 | Yersinia | enterocolitica | 2.31 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.45 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.57 | Yersinia | enterocolitica | 2.53 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.ma_01\0_C13, [0034] | Yersinia | entomophaga | Yersinia | entomophaga | 2.8 | Yersinia | nurmii | 2.36 | B | 0 | 0 | 1 | Yersinia entomophaga | Yersinia nurmii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.24 | Yersinia | enterocolitica | 2.08 | B | 0 | 0 | 1 | Yersinia frederiksenii | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 21 | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 22.01.2007 H06 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.41 | Yersinia | frederiksenii | 2.36 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 21 | 3; well described strain | DT | 2013-01-30 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.57 | Yersinia | frederiksenii | 2.31 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1824 | 2; sequenced | DT | 2012-12-05 | Yersinia | intermedia | Yersinia | intermedia | 2.46 | Yersinia | intermedia | 2.35 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia | intermedia | Yersinia | intermedia | 2.51 | Yersinia | intermedia | 2.45 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.57 | Yersinia | kristensenii | 2.53 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 302 | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.46 | Yersinia | kristensenii | 2.39 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 25.11.2015, Y.ma_01\0_C11, [0034] | Yersinia | massiliensis | Yersinia | massiliensis | 2.71 | Yersinia | massiliensis | 2.33 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB 925699 | 3; sequenced | DT | 2013-03-19 | Yersinia | massiliensis | Yersinia | massiliensis | 2.33 | Yersinia | massiliensis | 1.85 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia | mollaretii | Yersinia | mollaretii | 2.32 | Yersinia | mollaretii | 2.25 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18520 | 3; type strain | DT | 16.01.2013 | Yersinia | mollaretii | Yersinia | mollaretii | 2.13 | Yersinia | mollaretii | 2.11 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 51 | 2; 0 | DT | 2013-01-16 | Yersinia | mollaretii | Yersinia | mollaretii | 2.42 | Yersinia | mollaretii | 2.4 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.nu_04\0_E10, [0034] | Yersinia | nurmii | Yersinia | nurmii | 2.75 | Yersinia | entomophaga | 2.33 | B | 0 | 0 | 1 | Yersinia nurmii | Yersinia entomophaga | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.pe_04\0_E8, [0034] | Yersinia | pekkanenii | Yersinia | pekkanenii | 2.56 | Yersinia | pekkanenii | 2.36 | A | 1 | 1 | 0 | Yersinia pekkanenii | Yersinia pekkanenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 2028 | 2; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H13 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.48 | Yersinia | pseudotuberculosis | 2.48 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H14 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.57 | Yersinia | pseudotuberculosis | 2.43 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0304,2 | 2; biochemistry | DT | 2013-01-31 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.45 | Yersinia | pseudotuberculosis | 2.44 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8992 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.04.2009, E12 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.69 | Yersinia | pseudotuberculosis | 2.67 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia | rohdei | Yersinia | rohdei | 2.46 | Yersinia | rohdei | 2.41 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 02.03.2007, E07 | Yersinia | rohdei | Yersinia | rohdei | 2.83 | Yersinia | rohdei | 2.31 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2772 | 3; sequenced | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.61 | Yersinia | ruckeri | 2.56 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.56 | Yersinia | ruckeri | 2.56 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia | similis | Yersinia | similis | 2.83 | Yersinia | similis | 2.51 | A | 1 | 1 | 0 | Yersinia similis | Yersinia similis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMB 4354 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 13.07.2007, E22 | Yersinia | similis | Yersinia | pseudotuberculosis | 2.52 | Yersinia | pseudotuberculosis | 2.49 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia | wautersii | Yersinia | wautersii | 2.29 | Yersinia | pseudotuberculosis | 2.23 | B | 0 | 0 | 1 | Yersinia wautersii | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.wa_04\0_E5, [0034] | Yersinia | wautersii | Yersinia | wautersii | 2.69 | Yersinia | pseudotuberculosis | 2.39 | B | 0 | 0 | 1 | Yersinia wautersii | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella | londiniensis | Legionella | londiniensis | 2.27 | Legionella | londiniensis | 1.59 | A | 1 | 1 | 0 | Legionella londiniensis | Legionella londiniensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella | pneumophila | Legionella | pneumophila | 2.21 | Legionella | pneumophila | 1.73 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4122 | 2; 0 | eDT | 20.01.2016 | Legionella | pneumophila | Legionella | pneumophila | 2.33 | Legionella | pneumophila | 2.29 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella | taurinensis | Legionella | taurinensis | 2.41 | Legionella | sp | 1.86 | A | 1 | 1 | 0 | Legionella taurinensis | Legionella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10873,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | stevensii | 2.33 | Halomonas | aquamarina | 1.54 | A | 1 | 1 | 0 | Halomonas stevensii | Halomonas aquamarina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10874,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | stevensii | 2.43 | Dermacoccus | sp[3] | 1.35 | A | 1 | 1 | 0 | Halomonas stevensii | Dermacoccus sp[3] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) | caballi | Pasteurella | caballi | 2.22 | Pasteurella | caballi | 2.15 | A | 0 | 0 | 1 | Pasteurella caballi | Pasteurella caballi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) | caballi | Pasteurella | caballi | 2.54 | (Pasteurella) | caballi | 2.53 | C | 0 | 0 | 1 | Pasteurella caballi | (Pasteurella) caballi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) | testudinis | (Pasteurella) | testudinis | 2.33 | Pasteurella | testudinis | 2.01 | C | 0 | 0 | 1 | (Pasteurella) testudinis | Pasteurella testudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 957 | 3; sequenced | eDT | 02.12.2015 | (Pasteurella) | testudinis | (Pasteurella) | testudinis | 2.11 | Pasteurella | testudinis | 1.87 | C | 0 | 0 | 1 | (Pasteurella) testudinis | Pasteurella testudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | anseriformium | Actinobacillus | anseriformium | 2.48 | Actinobacillus | equuli | 2.05 | B | 0 | 0 | 1 | Actinobacillus anseriformium | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15UGF 198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus | anseriformium | Actinobacillus | anseriformium | 2.53 | Actinobacillus | anseriformium | 2.22 | A | 1 | 1 | 0 | Actinobacillus anseriformium | Actinobacillus anseriformium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus | arthritidis | Actinobacillus | arthritidis | 2.33 | Actinobacillus | arthritidis_JRA_Pasturella-4 | 2 | B | 0 | 0 | 1 | Actinobacillus arthritidis | Actinobacillus arthritidis_JRA_Pasturella-4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.42 | Actinobacillus | capsulatus | 2.2 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus capsulatus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.52 | Actinobacillus | capsulatus | 2.42 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus capsulatus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.4 | Actinobacillus | equuli | 2.38 | A | 1 | 1 | 0 | Actinobacillus equuli | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.28 | Actinobacillus | suis | 2.22 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.36 | Actinobacillus | equuli | 2.35 | A | 1 | 1 | 0 | Actinobacillus equuli | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus | equuli | Actinobacillus | equuli | 2.52 | Actinobacillus | suis | 2.52 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus | lignieresii | Actinobacillus | pleuropneumoniae | 2.34 | Actinobacillus | lignieresii | 2.33 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9945,2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | lignieresii | Actinobacillus | pleuropneumoniae | 2.45 | Actinobacillus | lignieresii | 2.37 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus | rossii | Actinobacillus | rossii | 2.36 | Actinobacillus | rossii | 2.08 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus | rossii | Actinobacillus | rossii | 2.67 | Actinobacillus | rossii | 2.64 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32505,2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus | sp-CVUAS-32505,2 | Actinobacillus | sp-CVUAS-32505,2 | 2.54 | Actinobacillus | arthritidis_JRA_Pasturella-4 | 2.18 | B | 0 | 0 | 1 | Actinobacillus sp-CVUAS-32505,2 | Actinobacillus arthritidis_JRA_Pasturella-4 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-genomospecie- | Actinobacillus | sp-genomospecies-1 | 2.53 | Actinobacillus | arthritidis | 2.09 | B | 0 | 0 | 1 | Actinobacillus sp-genomospecies-1 | Actinobacillus arthritidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | sp-genomospecies-1 | 2.46 | Actinobacillus | pleuropneumoniae | 2.14 | B | 0 | 0 | 1 | Actinobacillus sp-genomospecies-1 | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus | sp-MK-2012 | Actinobacillus | sp-MK-2012 | 2.67 | Actinobacillus | sp-MK-2012 | 2.34 | A | 1 | 1 | 0 | Actinobacillus sp-MK-2012 | Actinobacillus sp-MK-2012 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-MK-2012 | Actinobacillus | sp-MK-2012 | 2.46 | Actinobacillus | sp-MK-2012 | 2.44 | A | 1 | 1 | 0 | Actinobacillus sp-MK-2012 | Actinobacillus sp-MK-2012 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus | suis | Actinobacillus | suis | 2.6 | Actinobacillus | suis | 2.55 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus | suis | Actinobacillus | suis | 2.5 | Actinobacillus | suis | 2.47 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium | gallinarum | Avibacterium | gallinarum | 2.19 | Avibacterium | gallinarum | 2.15 | A | 1 | 1 | 0 | Avibacterium gallinarum | Avibacterium gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium | gallinarum | Avibacterium | gallinarum | 2.72 | Avibacterium | gallinarum | 2.27 | A | 1 | 1 | 0 | Avibacterium gallinarum | Avibacterium gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1816,2 | 3; sequenced | DT | 2013-07-24 | Basfia | succiniciproducens | Basfia | succiniciproducens | 2.08 | Pasteurella | multocida | 1.75 | C | 0 | 0 | 1 | Basfia succiniciproducens | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.52 | Bibersteinia | trehalosi | 2.48 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.47 | Bibersteinia | trehalosi | 2.39 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter | oris | Chelonobacter | oris | 2.42 | Chelonobacter | oris | 2.41 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter | oris | Chelonobacter | oris | 2.11 | Chelonobacter | oris | 2.05 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3802,2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter | osteomyelitidis | Cricetibacter | osteomyelitidis | 2.54 | Haemophilus | haemolyticus | 1.43 | A | 1 | 1 | 0 | Cricetibacter osteomyelitidis | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia | canicola | Frederiksenia | canicola | 2.41 | Frederiksenia | canicola | 2.21 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia | canicola | Frederiksenia | canicola | 2.32 | Frederiksenia | canicola | 2.28 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium | anatis | Gallibacterium | anatis | 2.07 | Gallibacterium | anatis | 2.04 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium | anatis | Gallibacterium | anatis | 2.45 | Gallibacterium | anatis | 2.36 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium | sp-F151 | Gallibacterium | sp-F151 | 2.49 | Gallibacterium | anatis | 1.44 | A | 1 | 1 | 0 | Gallibacterium sp-F151 | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium | trehalosifermentans | Gallibacterium | trehalosifermentans | 2.65 | Gallibacterium | sp-F151 | 1.88 | A | 1 | 1 | 0 | Gallibacterium trehalosifermentans | Gallibacterium sp-F151 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus | haemoglobinophilus | Haemophilus | haemoglobinophilus | 2.29 | Haemophilus | sp-2269 | 1.91 | A | 1 | 1 | 0 | Haemophilus haemoglobinophilus | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | sp-2269 | 2.39 | Haemophilus | sp-2269 | 2.35 | A | 1 | 1 | 0 | Haemophilus sp-2269 | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | sp-2269 | 2.51 | Haemophilus | sp-2269 | 2.46 | A | 1 | 1 | 0 | Haemophilus sp-2269 | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P1066-2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus | somni | Histophilus | somni | 2.2 | Histophilus | somni | 1.82 | A | 1 | 1 | 0 | Histophilus somni | Histophilus somni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia | caviae | Mannheimia | caviae | 2.62 | Mannheimia | caviae | 2.22 | A | 1 | 1 | 0 | Mannheimia caviae | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1064,4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia | glucosida | Mannheimia | glucosida | 2.35 | Mannheimia | glucosida | 2.22 | A | 1 | 1 | 0 | Mannheimia glucosida | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia | glucosida | Mannheimia | glucosida | 2.59 | Mannheimia | glucosida | 2.47 | A | 1 | 1 | 0 | Mannheimia glucosida | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia | granulomatis | Mannheimia | sp-14UCF195 | 2.3 | Mannheimia | granulomatis | 2.05 | B | 0 | 0 | 1 | Mannheimia sp-14UCF195 | Mannheimia granulomatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia | granulomatis | Mannheimia | granulomatis | 2.48 | Mannheimia | sp-14UCF195 | 2.26 | B | 0 | 0 | 1 | Mannheimia granulomatis | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.53 | Mannheimia | haemolytica | 2.47 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6197,2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.57 | Mannheimia | haemolytica | 2.52 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia | pernigra | Mannheimia | pernigra | 2.51 | Mannheimia | pernigra | 2.11 | A | 1 | 1 | 0 | Mannheimia pernigra | Mannheimia pernigra | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia | pernigra | Mannheimia | pernigra | 2.32 | Mannheimia | pernigra | 2.32 | A | 1 | 1 | 0 | Mannheimia pernigra | Mannheimia pernigra | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.57 | Mannheimia | ruminalis | 2.3 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.3 | Mannheimia | ruminalis | 2.19 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10396,2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia | sp-10-12559-1 | Mannheimia | sp-10-12559-1 | 2.46 | Mannheimia | caviae | 2.03 | B | 0 | 0 | 1 | Mannheimia sp-10-12559-1 | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10183,3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia | sp-14UCF195 | Mannheimia | sp-14UCF195 | 2.64 | Mannheimia | sp-14UCF195 | 2.32 | A | 1 | 1 | 0 | Mannheimia sp-14UCF195 | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia | sp-14UCF195 | Mannheimia | sp-14UCF195 | 2.68 | Mannheimia | sp-14UCF195 | 2.29 | A | 1 | 1 | 0 | Mannheimia sp-14UCF195 | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia | sp-15USF15 | Mannheimia | sp-15USF15 | 2.68 | Mannheimia | sp-15USF15 | 2.11 | A | 1 | 1 | 0 | Mannheimia sp-15USF15 | Mannheimia sp-15USF15 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia | sp-15USF15 | Mannheimia | sp-15USF15 | 2.49 | Mannheimia | caviae | 2.13 | B | 0 | 0 | 1 | Mannheimia sp-15USF15 | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia | sp-CVUAS-11240 | Mannheimia | sp-CVUAS-11240 | 2.65 | Mannheimia | sp-15USF15 | 1.94 | A | 1 | 1 | 0 | Mannheimia sp-CVUAS-11240 | Mannheimia sp-15USF15 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia | sp-CVUAS-1496 | Mannheimia | sp-CVUAS-1496 | 2.47 | Mannheimia | glucosida | 2.37 | B | 0 | 0 | 1 | Mannheimia sp-CVUAS-1496 | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.5 | Necropsobacter | rosorum | 2.43 | A | 1 | 1 | 0 | Necropsobacter rosorum | Necropsobacter rosorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.36 | Necropsobacter | rosorum | 2.35 | A | 1 | 1 | 0 | Necropsobacter rosorum | Necropsobacter rosorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella | semolina | Nicoletella | semolina | 2.55 | Nicoletella | semolina | 2.43 | A | 1 | 1 | 0 | Nicoletella semolina | Nicoletella semolina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella | semolina | Nicoletella | semolina | 2.46 | Nicoletella | semolina | 2.46 | A | 1 | 1 | 0 | Nicoletella semolina | Nicoletella semolina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4060,3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella | aerogenes | Pasteurella | aerogenes | 2.47 | Pasteurella | aerogenes | 2.05 | A | 1 | 1 | 0 | Pasteurella aerogenes | Pasteurella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | canis | Pasteurella | canis | 2.07 | Pasteurella | canis | 2.04 | A | 1 | 1 | 0 | Pasteurella canis | Pasteurella canis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | mairii | Pasteurella | mairii | 2.11 | Actinobacillus | seminis | 2.06 | C | 0 | 0 | 1 | Pasteurella mairii | Actinobacillus seminis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella | multocida | Pasteurella | multocida | 2.48 | Pasteurella | multocida | 2.33 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32055 | 3; specific PCR | eDT | 2019-11-14 | Pasteurella | multocida | Pasteurella | multocida | 2.67 | Pasteurella | multocida | 2.59 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella | multocida | Pasteurella | multocida | 2.62 | Pasteurella | multocida | 2.52 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.57 | Pasteurella | multocida | 2.56 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.51 | Pasteurella | multocida | 2.5 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella | multocida | Pasteurella | multocida | 2.57 | Pasteurella | multocida | 2.49 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32506,2 | 3; sequenced | EFEx | 2020-11-19 | Pasteurellaceae | Bisgaard-Taxon-10 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.58 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.39 | A | 1 | 1 | 0 | Pasteurellaceae Bisgaard-Taxon-10 | Pasteurellaceae Bisgaard-Taxon-10 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Pasteurellaceae | Bisgaard-Taxon-10 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.36 | Pasteurellaceae | Bisgaard-Taxon-10 | 2.34 | A | 1 | 1 | 0 | Pasteurellaceae Bisgaard-Taxon-10 | Pasteurellaceae Bisgaard-Taxon-10 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | Bisgaard-Taxon-40 | Pasteurellaceae | Bisgaard-Taxon-40 | 2.45 | Haemophilus | parainfluenzae | 1.36 | A | 1 | 1 | 0 | Pasteurellaceae Bisgaard-Taxon-40 | Haemophilus parainfluenzae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9987,2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae | canine-oral-taxon-080 | Pasteurellaceae | canine-oral-taxon-080 | 2.56 | Frederiksenia | canicola | 1.83 | C | 0 | 0 | 1 | Pasteurellaceae canine-oral-taxon-080 | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 261,4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae | sp-CVUAS-261,4 | Pasteurellaceae | sp-CVUAS-261,4 | 2.64 | Klebsiella | oxytoca | 1.47 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-261,4 | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae | sp-CVUAS-31988 | Pasteurellaceae | sp-CVUAS-31988 | 2.61 | Haemophilus | haemolyticus | 1.65 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-31988 | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | sp-CVUAS-4868 | Pasteurellaceae | sp-CVUAS-4868 | 2.54 | Enterococcus | faecalis | 1.34 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-4868 | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 724,3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae | sp-CVUAS-724,3 | Pasteurellaceae | sp-CVUAS-724,3 | 2.6 | Rodentibacter | heylii | 1.47 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-724,3 | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae | sp-CVUAS-9305 | Pasteurellaceae | sp-CVUAS-9305 | 2.65 | Actinobacillus | equuli | 1.41 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-9305 | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9967,2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae | sp-CVUAS-9967,2 | Pasteurellaceae | sp-CVUAS-9967,2 | 2.45 | Haemophilus | haemolyticus | 1.74 | C | 0 | 0 | 1 | Pasteurellaceae sp-CVUAS-9967,2 | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae | sp-LHL-U-309 | Pasteurellaceae | sp-LHL-U-309 | 2.36 | Pasteurellaceae | sp-LHL-U-309 | 2.13 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-309 | Pasteurellaceae sp-LHL-U-309 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 443,2 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-309 | Pasteurellaceae | sp-LHL-U-309 | 2.46 | Pasteurellaceae | sp-LHL-U-309 | 2.22 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-309 | Pasteurellaceae sp-LHL-U-309 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 444,2 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-3444 | Pasteurellaceae | sp-LHL-U-3444 | 2.54 | Chelonobacter | oris | 1.54 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-3444 | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae | sp-MCCM-2539 | Pasteurellaceae | sp-MCCM-2539 | 2.54 | Haemophilus | pittmaniae | 1.53 | A | 1 | 1 | 0 | Pasteurellaceae sp-MCCM-2539 | Haemophilus pittmaniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.48 | Rodentibacter | heylii | 2.4 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.55 | Rodentibacter | heylii | 2.5 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium | mucosae | Spirabiliibacterium | mucosae | 2.39 | Spirabiliibacterium | mucosae | 2.29 | A | 1 | 1 | 0 | Spirabiliibacterium mucosae | Spirabiliibacterium mucosae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium | mucosae | Spirabiliibacterium | mucosae | 2.45 | Spirabiliibacterium | mucosae | 2.38 | A | 1 | 1 | 0 | Spirabiliibacterium mucosae | Spirabiliibacterium mucosae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8966 | 3; sequenced | DT | 23.07.2013 | Volucribacter | amazonae | Volucribacter | amazonae | 2.43 | Haemophilus | influenzae | 1.53 | A | 1 | 1 | 0 | Volucribacter amazonae | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 308,2 | 3; sequenced | DT | 2013-11-06 | Psittacicella | hinzii | Psittacicella | hinzii | 2.57 | Agromyces | cerinus | 1.43 | A | 1 | 1 | 0 | Psittacicella hinzii | Agromyces cerinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E22 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.51 | Acinetobacter | baumannii | 2.48 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2014 C05 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.46 | Acinetobacter | baumannii | 2.41 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter | defluvii | Acinetobacter | defluvii | 2.42 | Acinetobacter | defluvii | 2.16 | A | 1 | 1 | 0 | Acinetobacter defluvii | Acinetobacter defluvii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | guillouiae | Acinetobacter | guillouiae | 2.34 | Acinetobacter | guillouiae | 2.24 | A | 1 | 1 | 0 | Acinetobacter guillouiae | Acinetobacter guillouiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | johnsonii | Acinetobacter | johnsonii | 2.3 | Acinetobacter | johnsonii | 2.26 | A | 1 | 1 | 0 | Acinetobacter johnsonii | Acinetobacter johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | junii | Acinetobacter | junii | 2.33 | Acinetobacter | junii | 2.31 | A | 1 | 1 | 0 | Acinetobacter junii | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | lwoffii | Acinetobacter | lwoffii | 2.35 | Acinetobacter | lwoffii | 2.27 | A | 1 | 1 | 0 | Acinetobacter lwoffii | Acinetobacter lwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7379,3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.29 | Acinetobacter | pseudolwoffii | 2.26 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.08 | Acinetobacter | pseudolwoffii | 2.07 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | variabilis | Acinetobacter | variabilis | 2.35 | Acinetobacter | variabilis | 2.17 | A | 1 | 1 | 0 | Acinetobacter variabilis | Acinetobacter variabilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | ovis | 2.44 | Moraxella | bovoculi | 2.16 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | ovis | 2.09 | Moraxella | bovoculi | 2.06 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella | catarrhalis | Moraxella | catarrhalis | 2.53 | Moraxella | catarrhalis | 2.51 | A | 1 | 1 | 0 | Moraxella catarrhalis | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11674,3 | 3; sequenced | eDT | 2022-03-22 | Moraxella | cuniculi | Moraxella | cuniculi | 2.55 | Moraxella | catarrhalis | 1.63 | A | 1 | 1 | 0 | Moraxella cuniculi | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella | osloensis | Moraxella | osloensis | 2.53 | Moraxella | osloensis | 1.94 | A | 1 | 1 | 0 | Moraxella osloensis | Moraxella osloensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella | ovis | Moraxella | ovis | 2.39 | Moraxella | bovoculi | 2.32 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | frigidicola | Psychrobacter | frigidicola | 2.49 | Psychrobacter | pulmonis | 1.47 | A | 1 | 1 | 0 | Psychrobacter frigidicola | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter | pasteurii | Psychrobacter | pasteurii | 2.52 | Psychrobacter | pasteurii | 2.04 | A | 1 | 1 | 0 | Psychrobacter pasteurii | Psychrobacter pasteurii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.58 | Psychrobacter | pulmonis | 2.35 | A | 1 | 1 | 0 | Psychrobacter pulmonis | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.3 | Psychrobacter | faecalis | 1.83 | A | 1 | 1 | 0 | Psychrobacter pulmonis | Psychrobacter faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas | (alkylphenolica) | Pseudomonas | alkylphenolica | 2.4 | Pseudomonas | japonica | 1.59 | A | 0 | 0 | 1 | Pseudomonas alkylphenolica | Pseudomonas japonica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.33 | Pseudomonas | aeruginosa | 2.28 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007 H8 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.46 | Pseudomonas | aeruginosa | 2.38 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas | alcaligenes | Pseudomonas | alcaligenes | 2.44 | Pseudomonas | alcaligenes | 2.35 | A | 1 | 1 | 0 | Pseudomonas alcaligenes | Pseudomonas alcaligenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas | anguilliseptica | Pseudomonas | anguilliseptica | 2.02 | Pseudomonas | borbori | 1.95 | A | 1 | 1 | 0 | Pseudomonas anguilliseptica | Pseudomonas borbori | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas | baetica | Pseudomonas | baetica | 2.66 | Pseudomonas | koreensis | 2.06 | B | 0 | 0 | 1 | Pseudomonas baetica | Pseudomonas koreensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas | brenneri | Pseudomonas | brenneri | 2.14 | Pseudomonas | veronii | 1.84 | A | 1 | 1 | 0 | Pseudomonas brenneri | Pseudomonas veronii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.02.2008 N11 | Pseudomonas | chlororaphis | Pseudomonas | chlororaphis | 2.44 | Pseudomonas | chlororaphis | 2.3 | A | 1 | 1 | 0 | Pseudomonas chlororaphis | Pseudomonas chlororaphis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas | fulva | Pseudomonas | fulva | 2.29 | Pseudomonas | fulva | 1.9 | A | 1 | 1 | 0 | Pseudomonas fulva | Pseudomonas fulva | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas | koreensis | Pseudomonas | koreensis | 2.26 | Pseudomonas | corrugata | 1.89 | A | 1 | 1 | 0 | Pseudomonas koreensis | Pseudomonas corrugata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas | monteilii | Pseudomonas | monteilii | 2.01 | Pseudomonas | putida | 1.92 | A | 1 | 1 | 0 | Pseudomonas monteilii | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1_1 | 2; 0 | DT | 2017-08-31 | Pseudomonas | nitroreducens | Pseudomonas | nitroreducens | 2.07 | Pseudomonas | nitroreducens | 1.97 | A | 1 | 1 | 0 | Pseudomonas nitroreducens | Pseudomonas nitroreducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas | oryzihabitans | Pseudomonas | oryzihabitans | 2.12 | Pseudomonas | oryzihabitans | 2.09 | A | 1 | 1 | 0 | Pseudomonas oryzihabitans | Pseudomonas oryzihabitans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1_2 | 2; 0 | DT | 2017-08-31 | Pseudomonas | putida | Pseudomonas | putida | 2.49 | Pseudomonas | putida | 2.31 | A | 1 | 1 | 0 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas | taetrolens | Pseudomonas | taetrolens | 2 | Pseudomonas | lundensis | 1.95 | A | 1 | 1 | 0 | Pseudomonas taetrolens | Pseudomonas lundensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 12.12.2016, F19, [0096] | Allofrancisella | guangzhouensis | Allofrancisella | guangzhouensis | 2.69 | Paracoccus | denitrificans | 1.33 | A | 1 | 1 | 0 | Allofrancisella guangzhouensis | Paracoccus denitrificans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 18.11.2015, L16, [0096] | Francisella | hispaniensis | Francisella | hispaniensis | 2.75 | Francisella | tularensis | 1.99 | A | 1 | 1 | 0 | Francisella hispaniensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 25.11.2015, I3, [0096] | Francisella | naotunensis | Francisella | naotunensis | 2.79 | Francisella | philomiragia | 1.57 | A | 1 | 1 | 0 | Francisella naotunensis | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 02.04.2012, H21 | Francisella | tularensis | Francisella | tularensis | 2.55 | Francisella | tularensis | 2.07 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4680,2 | 3; FLI: specific PCR | DT | 2013-04-02 | Francisella | tularensis | Francisella | tularensis | 2.23 | Francisella | tularensis | 2.2 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 23.04.2013, F20, [0096] | Francisella | tularensis | Francisella | tularensis | 2.36 | Francisella | tularensis | 2.34 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 26.04.2013, I20, [0096] | Francisella | tularensis | Francisella | tularensis | 2.68 | Francisella | tularensis | 2.26 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 01.06.2017, F3, [0096] | Francisella | tularensis | Francisella | tularensis | 2.8 | Francisella | tularensis | 2.05 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas | sp-G9077 | Wohlfahrtiimonas | sp-G9077 | 2.42 | Wohlfahrtiimonas | chitiniclastica | 2.35 | B | 0 | 0 | 1 | Wohlfahrtiimonas sp-G9077 | Wohlfahrtiimonas chitiniclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33180 | 3; sequenced | DT | 2021-08-04 | Photobacterium | damselae | Photobacterium | damselae | 2.22 | Photobacterium | damselae | 2.02 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium | damselae | Photobacterium | damselae | 2.14 | Photobacterium | damselae | 2.07 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1266,2 | 3; sequenced | EFEx | 25.09.1019 | Photobacterium | piscicola | Photobacterium | piscicola | 2.6 | Photobacterium | iliopiscarium | 1.57 | A | 1 | 1 | 0 | Photobacterium piscicola | Photobacterium iliopiscarium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 842,2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium | swingsii | Photobacterium | swingsii | 2.63 | Providencia | alcalifaciens | 1.37 | A | 1 | 1 | 0 | Photobacterium swingsii | Providencia alcalifaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio | aestuarianus | Vibrio | aestuarianus | 2.52 | Vibrio | aestuarianus | 2.5 | A | 1 | 1 | 0 | Vibrio aestuarianus | Vibrio aestuarianus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7079 | 3; spezific PCR (CVFR) | eDT | 2013-07-17 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.18 | Vibrio | alginolyticus | 2.06 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2171 | 3; type strain | DT | 2016-10-12 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.3 | Vibrio | alginolyticus | 2.29 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio | anguillarum | Vibrio | anguillarum | 2.34 | Vibrio | anguillarum | 2.3 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | A8-12 | 3; gyrB sequence | eDT | 2016-06-16 | Vibrio | anguillarum | Vibrio | anguillarum | 2.22 | Vibrio | anguillarum | 2.2 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio | campbelli | Vibrio | campbellii | 2.32 | Vibrio | harveyi | 2.12 | B | 0 | 0 | 1 | Vibrio campbellii | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio | campbelli | Vibrio | campbellii | 2.37 | Vibrio | rotiferianus | 2.2 | B | 0 | 0 | 1 | Vibrio campbellii | Vibrio rotiferianus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16-VB00067-0_BfR 1_eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio | cholerae | Vibrio | mimicus | 1.93 | Vibrio | mimicus | 1.93 | B | 0 | 0 | 1 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10227 | 3; sequenced | eDT | 2022-06-23 | Vibrio | cholerae | Vibrio | mimicus | 2.04 | Vibrio | albensis | 1.95 | A | 0 | 0 | 1 | Vibrio mimicus | Vibrio albensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 45VB066-10 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 1.52 | Vibrio | europaeus | 1.43 | B | 0 | 0 | 1 | Vibrio cincinnatiensis | Vibrio europaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00004 LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 2.14 | Vibrio | europaeus | 1.52 | A | 1 | 1 | 0 | Vibrio cincinnatiensis | Vibrio europaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.08 | Vibrio | diazotrophicus | 1.98 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8995 | 3; sequenced | eDT | 2013-04-10 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.14 | Vibrio | diazotrophicus | 2.14 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio | fluvialis | Vibrio | fluvialis | 2.42 | Vibrio | fluvialis | 2.16 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00036 LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio | fluvialis | Vibrio | fluvialis | 2.22 | Vibrio | fluvialis | 2.18 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 14383 | 3; public strain collection, sequenced | eDT | 2013-04-16 | Vibrio | furnissi | Vibrio | furnissii | 2.21 | Vibrio | furnissii | 2.06 | A | 0 | 0 | 1 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1754 | 3; sequenced | eDT | 2013-04-16 | Vibrio | furnissii | Vibrio | furnissii | 2.2 | Vibrio | furnissii | 2.05 | A | 1 | 1 | 0 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 16.08.2012, K12, [0034] | Vibrio | harveyi | Vibrio | harveyi | 2.28 | Vibrio | harveyi | 2.25 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio | harveyi | Vibrio | harveyi | 2.39 | Vibrio | harveyi | 2.36 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | P340-7 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio | jasicida | Vibrio | harveyi | 2.49 | Vibrio | jasicida | 2.46 | B | 0 | 0 | 1 | Vibrio harveyi | Vibrio jasicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | P340-9 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio | jasicida | Vibrio | jasicida | 2.36 | Vibrio | jasicida | 2.29 | A | 1 | 1 | 0 | Vibrio jasicida | Vibrio jasicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 12-VBH0498 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metoecus | Vibrio | mimicus | 2.48 | Vibrio | mimicus | 2.38 | A | 0 | 0 | 1 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00016 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | metoecus | Vibrio | mimicus | 2.14 | Vibrio | mimicus | 2.06 | A | 0 | 0 | 1 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metschnikovii | Vibrio | metschnikovii | 2.15 | Vibrio | metschnikovii | 2.05 | A | 1 | 1 | 0 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 619 | 2; 0 | eDT | 2013-08-14 | Vibrio | metschnikovii | Vibrio | metschnikovii | 2.35 | Vibrio | metschnikovii | 2.29 | A | 1 | 1 | 0 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7872 | 3; sequenced | eDT | 2013-04-10 | Vibrio | mimicus | Vibrio | mimicus | 2.41 | Vibrio | mimicus | 2.35 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19130 | 4; type strain; WGS sequenced | DT | 2021-02-24 | Vibrio | mimicus | Vibrio | mimicus | 2.31 | Vibrio | mimicus | 2.28 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | navarrensis | Vibrio | navarrensis | 2.16 | Vibrio | navarrensis | 1.99 | A | 1 | 1 | 0 | Vibrio navarrensis | Vibrio navarrensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio | owensii | Vibrio | owensii | 2.32 | Vibrio | jasicida | 2.27 | B | 0 | 0 | 1 | Vibrio owensii | Vibrio jasicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00254-0_BfR 18 | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.28 | Vibrio | parahaemolyticus | 2.26 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10027 | 3; type strain | EFEx | 2013-05-22 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.21 | Vibrio | parahaemolyticus | 2.16 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 871,2 | 3; sequenced | eDT | 2014-08-21 | Vibrio | sp-CVUAS-871,2 | Vibrio | sp-CVUAS-871,2 | 2.28 | Vibrio | mimicus | 1.45 | A | 1 | 1 | 0 | Vibrio sp-CVUAS-871,2 | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2778,2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio | tubiashii | Vibrio | tubiashii | 2.17 | Vibrio | tubiashii | 2.05 | A | 1 | 1 | 0 | Vibrio tubiashii | Vibrio tubiashii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2778,2 | 3; sequenced | eDT | 2019-10-24 | Vibrio | tubiashii | Vibrio | tubiashii | 2.51 | Vibrio | tubiashii | 2.2 | A | 1 | 1 | 0 | Vibrio tubiashii | Vibrio tubiashii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio | vulnificus | Vibrio | vulnificus | 2.32 | Vibrio | vulnificus | 2.3 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LVM05 LVU05 | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-15 | Vibrio | vulnificus | Vibrio | vulnificus | 2.46 | Vibrio | vulnificus | 2.31 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria | indica | Ignatzschineria | indica | 2.43 | Ignatzschineria | indica | 2.18 | A | 1 | 1 | 0 | Ignatzschineria indica | Ignatzschineria indica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria | ureiclastica | Ignatzschineria | ureiclastica | 2.58 | Ignatzschineria | ureiclastica | 2.42 | A | 1 | 1 | 0 | Ignatzschineria ureiclastica | Ignatzschineria ureiclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas | chelatiphaga | Stenotrophomonas | chelatiphaga | 2.14 | Stenotrophomonas | maltophilia | 1.61 | A | 1 | 1 | 0 | Stenotrophomonas chelatiphaga | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2009 K17 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.26 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.05.2009 K22 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.22 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.42 | Stenotrophomonas | rhizophila | 1.79 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas rhizophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.05.2015 C17 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.4 | Stenotrophomonas | rhizophila | 1.69 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas rhizophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4239,2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas | tumulicola | Stenotrophomonas | tumulicola | 2.47 | Stenotrophomonas | chelatiphaga | 1.69 | A | 1 | 1 | 0 | Stenotrophomonas tumulicola | Stenotrophomonas chelatiphaga | 0 | 0 | 0 | 0

## Test for spectral proximity
Parameter: | Bacillus | cytotoxicus | Isolate ID | Validity | preparation method (if given) | Date | Organism | Score | correct negative | false match
Test auf spektral proximity |  |  | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces bovis | 0.56 | 1 | 0
Number of isolates: | 1408 |  | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces denticolens | 0.53 | 1 | 0
|  |  | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces sp-CVUAS-31303 | 0.6 | 1 | 0
correct-negative: | 1404 |  | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces viscosus | 0.69 | 1 | 0
|  |  | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces weissii | 0.02 | 1 | 0
false-positive: | 4 |  | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium bovis | 0.28 | 1 | 0
|  |  | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium canis | 0.5 | 1 | 0
| 0.2841% |  | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium haemolyticum | 0.6 | 1 | 0
|  |  | 17-6203 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium hippocoleae | 0.15 | 1 | 0
|  |  | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium hippocoleae | 0.83 | 1 | 0
score limit | 2 | Species | DSM 10002 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium phocae | 0.69 | 1 | 0
settings | 1.7 | Genus | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium phocae | 1.03 | 1 | 0
It is checked against a database consisting only of the parameter. |  |  | 2698T | 3; type strain | DT | 2013-08-07 | Arcanobacterium phocisimile | 0.94 | 1 | 0
|  |  | 14418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium pinnipediorum | 0.74 | 1 | 0
Score levels smaller than 0.1 are displayed as 0.1. |  |  | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium pinnipediorum | 0.62 | 1 | 0
|  |  | DSM 13483 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium pluranimalium | 0.81 | 1 | 0
|  |  | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 1.03 | 1 | 0
|  |  | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 0.72 | 1 | 0
|  |  | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium sp-2701 | 0.54 | 1 | 0
|  |  | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium sp-2701 | 0.52 | 1 | 0
|  |  | CVUAS 1060,2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium wilhelmae | 0.1 | 1 | 0
|  |  | CVUAS 647 | 3; type strain | DT | 2014-09-04 | Arcanobacterium wilhelmae | 0.31 | 1 | 0
|  |  | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia sp-CVUAS-30561 | 0.86 | 1 | 0
|  |  | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia canis | 0.48 | 1 | 0
|  |  | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia cardiffensis | 0.29 | 1 | 0
|  |  | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia hyovaginalis | 0.8 | 1 | 0
|  |  | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia hyovaginalis | 0.94 | 1 | 0
|  |  | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia sp-CVUAS-32183 | 0.57 | 1 | 0
|  |  | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia sp-CVUAS-8688 | 0.7 | 1 | 0
|  |  | CVUAS 31845,2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia sp-CVUAS-8688 | 0.67 | 1 | 0
|  |  | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella abortisuis | 0.39 | 1 | 0
|  |  | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella abortisuis | 0.76 | 1 | 0
|  |  | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella bernardiae | 0.53 | 1 | 0
|  |  | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella bernardiae | 0.5 | 1 | 0
|  |  | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella bialowiezensis | 0.55 | 1 | 0
|  |  | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella bonasi | 0.6 | 1 | 0
|  |  | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.26 | 1 | 0
|  |  | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.5 | 1 | 0
|  |  | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella pyogenes | 0.79 | 1 | 0
|  |  | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella pyogenes | 0.56 | 1 | 0
|  |  | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium amycolatum | 0.72 | 1 | 0
|  |  | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2014 E8 | Corynebacterium amycolatum | 0.74 | 1 | 0
|  |  | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium aurimucosum | 0.54 | 1 | 0
|  |  | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium belfantii | 0.25 | 1 | 0
|  |  | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium belfantii | 0.15 | 1 | 0
|  |  | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium bovis | 0.25 | 1 | 0
|  |  | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium camporealensis | 0.04 | 1 | 0
|  |  | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium casei | 0.42 | 1 | 0
|  |  | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium confusum | 0.34 | 1 | 0
|  |  | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium confusum | 0.57 | 1 | 0
|  |  | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium cystitidis | 0.12 | 1 | 0
|  |  | KL0163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.21 | 1 | 0
|  |  | KL0167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.41 | 1 | 0
|  |  | KL0173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.6 | 1 | 0
|  |  | KL0956 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 0.45 | 1 | 0
|  |  | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium frankenforstense | 0.79 | 1 | 0
|  |  | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium frankenforstense | 0.64 | 1 | 0
|  |  | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium freiburgense | 0.17 | 1 | 0
|  |  | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium hansenii | 0.96 | 1 | 0
|  |  | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium hansenii | 0.1 | 1 | 0
|  |  | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium humireducens | 0.07 | 1 | 0
|  |  | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium kutscheri | 0.6 | 1 | 0
|  |  | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium lactis | 0.71 | 1 | 0
|  |  | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium lactis | 0.53 | 1 | 0
|  |  | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium phoceense | 0.52 | 1 | 0
|  |  | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium pilosum | 0.5 | 1 | 0
|  |  | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium pseudotuberculosis | 0.2 | 1 | 0
|  |  | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium pseudotuberculosis | 0.23 | 1 | 0
|  |  | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium pseudotuberculosis | 0.24 | 1 | 0
|  |  | DSM 20689 | 3; type strain | DT | 2012-12-18 | Corynebacterium pseudotuberculosis | 0.46 | 1 | 0
|  |  | C138 | 3; sequenced (rpoB) UGent | EFEx | 2020-07-30 | Corynebacterium rouxii | 0.57 | 1 | 0
|  |  | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium rouxii | 0.21 | 1 | 0
|  |  | CVUAS 3559,2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium rouxii | 0.45 | 1 | 0
|  |  | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium rouxii | 0.16 | 1 | 0
|  |  | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium silvaticum | 0.7 | 1 | 0
|  |  | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium silvaticum | 0.4 | 1 | 0
|  |  | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium sp-208-U2-6 | 0.68 | 1 | 0
|  |  | CVUAS 32165,2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium sp-2300500 | 0.71 | 1 | 0
|  |  | CVUAS 5775,2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium sp-2300500 | 0.98 | 1 | 0
|  |  | CVUAS 2554,2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium sp-2396 | 0.06 | 1 | 0
|  |  | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium sp-243-z-3 | 0.35 | 1 | 0
|  |  | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium sp-257-y-4 | 0.94 | 1 | 0
|  |  | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium sp-284z-2 | 0.72 | 1 | 0
|  |  | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium sp-CVUAS-11664 | 0.7 | 1 | 0
|  |  | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium sp-CVUAS-1888 | 0.46 | 1 | 0
|  |  | CVUAS 1721,3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium sp-CVUAS-2484.4 | 0.73 | 1 | 0
|  |  | CVUAS 2484,4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388 = 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium sp-CVUAS-2484.4 | 0.78 | 1 | 0
|  |  | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium sp-CVUAS-30404 | 0.55 | 1 | 0
|  |  | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-CVUAS-30404 | 0.56 | 1 | 0
|  |  | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium sp-CVUAS-31533 | 0.63 | 1 | 0
|  |  | CVUAS 32278,2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium sp-CVUAS-32278.2 | 0.68 | 1 | 0
|  |  | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium sp-CVUAS-874,4 | 0.83 | 1 | 0
|  |  | CVUAS 874,4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium sp-CVUAS-874,4 | 0.6 | 1 | 0
|  |  | LALLF 21TRD0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-LALLF-21TRB0553 | 0.7 | 1 | 0
|  |  | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium sp-LHL-0497 | 0.15 | 1 | 0
|  |  | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium sp-LHL-0558 | 0.37 | 1 | 0
|  |  | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium sp-LHL-1322 | 0.36 | 1 | 0
|  |  | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium stationis | 0.45 | 1 | 0
|  |  | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium testudinoris | 0.28 | 1 | 0
|  |  | CCM 2823 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 0.39 | 1 | 0
|  |  | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium ulcerans | 0.19 | 1 | 0
|  |  | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium ulcerans | 0.47 | 1 | 0
|  |  | JB362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium ulcerans | 0.92 | 1 | 0
|  |  | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 0.26 | 1 | 0
|  |  | CVUAS 6522,4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium ulcerans | 0.59 | 1 | 0
|  |  | CVUAS 5437,3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium variabile | 0.47 | 1 | 0
|  |  | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium xerosis | 0.65 | 1 | 0
|  |  | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium xerosis-group | 0.95 | 1 | 0
|  |  | CVUAS 2379,2 | 3; sequenced | EFex | 2021-04-08 | Gordonia paraffinivorans | 0.76 | 1 | 0
|  |  | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia paraffinivorans | 0.68 | 1 | 0
|  |  | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia rubripertincta | 0.41 | 1 | 0
|  |  | CVUAS 10185,3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium smegmatis | 0.82 | 1 | 0
|  |  | CVUAS 31296,3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia cyriacigeorgica | 0.74 | 1 | 0
|  |  | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia farcinica | 0.43 | 1 | 0
|  |  | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus equi | 0.82 | 1 | 0
|  |  | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 1.03 | 1 | 0
|  |  | CVUAS 1447,3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella sp | 0.84 | 1 | 0
|  |  | CVUAS 2895,3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium luteolum | 0.36 | 1 | 0
|  |  | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas flavigena | 0.76 | 1 | 0
|  |  | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium sp-234z-2 | 0.77 | 1 | 0
|  |  | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium sp-CVUAS-30560 | 0.07 | 1 | 0
|  |  | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus sedentarius | 0.54 | 1 | 0
|  |  | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus congolensis | 0.56 | 1 | 0
|  |  | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus congolensis | 0.79 | 1 | 0
|  |  | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium flaccumfaciens | 0.81 | 1 | 0
|  |  | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia sp-MUVA-A-508948 | 0.49 | 1 | 0
|  |  | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium sp | 0.67 | 1 | 0
|  |  | IMET 10353 DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.09.2011 L8 | Glutamicibacter nicotianae | 0.65 | 1 | 0
|  |  | IMET 10432 DSM 20127 | 3; type strain | unknown | 08.09.2011, L13, [0034] | Sinomonas atrocyanea | 0.44 | 1 | 0
|  |  | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter protophormiae | 1.06 | 1 | 0
|  |  | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria atrinae | 0.75 | 1 | 0
|  |  | 56a-5 | 3; sequenced | OSextr | 2020-05-20 | Kocuria gwangalensis | 0.58 | 1 | 0
|  |  | DSM 20032 | 3; type strain | eDT | 2015-07-09 | Kocuria kristinae | 0.63 | 1 | 0
|  |  | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2014 E22 | Kocuria palustris | 0.1 | 1 | 0
|  |  | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015 K12 | Micrococcus luteus | 0.64 | 1 | 0
|  |  | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium cellulans | 0.73 | 1 | 0
|  |  | CVUAS 749,2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter inulinus | 0.47 | 1 | 0
|  |  | A10 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E13 | Bacillus anthracis | 1.84 | 1 | 0
|  |  | A12 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E15 | Bacillus anthracis | 1.75 | 1 | 0
|  |  | A18 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.04.2006; H12 | Bacillus anthracis | 1.72 | 1 | 0
|  |  | A25 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.04.2006; H17 | Bacillus anthracis | 1.25 | 1 | 0
|  |  | A30 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.04.2006; H21 | Bacillus anthracis | 1.39 | 1 | 0
|  |  | A4 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E07 | Bacillus anthracis | 1.44 | 1 | 0
|  |  | A5 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E08 | Bacillus anthracis | 1.84 | 1 | 0
|  |  | A7 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E10 | Bacillus anthracis | 1.74 | 1 | 0
|  |  | A9 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E12 | Bacillus anthracis | 1.76 | 1 | 0
|  |  | ATCC 14185 A2 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2006; E05 | Bacillus anthracis | 1.71 | 1 | 0
|  |  | 161004439-1 | 2; 0 | DT | 2016-04-12 | Bacillus cereus | 1.69 | 1 | 0
|  |  | 161004439-2 | 2; 0 | DT | 2016-04-12 | Bacillus cereus | 1.81 | 1 | 0
|  |  | 161005557-1 | 2; 0 | DT | 2016-08-11 | Bacillus cereus | 1.85 | 1 | 0
|  |  | 171012589-5 | 2; 0 | DT | 2017-09-19 | Bacillus cereus | 1.84 | 1 | 0
|  |  | 181015909-7 | 2; 0 | DT | 2018-11-02 | Bacillus cereus | 1.72 | 1 | 0
|  |  | ATCC 11778 LHL | 3; public strain collection | DT | 2015-06-23 | Bacillus cereus | 1.68 | 1 | 0
|  |  | ATCC 14579 IGV | 3; type strain | DT | 2016-04-06 | Bacillus cereus | 1.96 | 1 | 0
|  |  | DSM 4312 IGV | 3; public strain collection | DT | 2016-04-13 | Bacillus cereus | 1.76 | 1 | 0
|  |  | DSM 4313 IGV | 3; public strain collection | DT | 2016-04-06 | Bacillus cereus | 1.89 | 1 | 0
|  |  | 141014442-3 | 3; VETQAS | DT | 2014-11-07 | Bacillus licheniformis | 0.83 | 1 | 0
|  |  | 141015909-3 | 2; 0 | DT | 2014-12-12 | Bacillus licheniformis | 0.57 | 1 | 0
|  |  | 161004445 | 2; 0 | DT | 2016-04-12 | Bacillus licheniformis | 0.28 | 1 | 0
|  |  | 161007249-3 | 2; 0 | DT | 2016-06-14 | Bacillus licheniformis | 0.73 | 1 | 0
|  |  | ATCC 14580 | 3; type strain | DT | 2015-06-23 | Bacillus licheniformis | 0.42 | 1 | 0
|  |  | DSM 6022 | 3; public strain collection | DT | 2016-04-06 | Bacillus thuringiensis | 1.64 | 1 | 0
|  |  | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus amylolyticus | 0.74 | 1 | 0
|  |  | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus amylolyticus | 0.42 | 1 | 0
|  |  | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus barengoltzii | 0.35 | 1 | 0
|  |  | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus barengoltzii | 0.68 | 1 | 0
|  |  | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus glucanolyticus | 0.29 | 1 | 0
|  |  | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus glucanolyticus | 0.28 | 1 | 0
|  |  | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus larvae | 0.64 | 1 | 0
|  |  | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus larvae | 0.18 | 1 | 0
|  |  | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 0.54 | 1 | 0
|  |  | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 0.88 | 1 | 0
|  |  | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus macerans | 0.59 | 1 | 0
|  |  | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus pabuli | 0.56 | 1 | 0
|  |  | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus pabuli | 0.74 | 1 | 0
|  |  | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus pasadenensis | 0.7 | 1 | 0
|  |  | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus validus | 0.54 | 1 | 0
|  |  | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium barati | 0.85 | 1 | 0
|  |  | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium butyricum | 0.46 | 1 | 0
|  |  | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium chauvoei | 0.12 | 1 | 0
|  |  | 171002729 | 2; 0 | DT | 2017-02-28 | Clostridium colchlearium | 0.76 | 1 | 0
|  |  | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium colicanis | 0.41 | 1 | 0
|  |  | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium novyi | 0.07 | 1 | 0
|  |  | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium septicum | 0.43 | 1 | 0
|  |  | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium sporogenes | 0.41 | 1 | 0
|  |  | 141015223 aerob | 2; 0 | DT | 2014-11-27 | Clostridium tertium | 0.26 | 1 | 0
|  |  | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium tetani | 0.68 | 1 | 0
|  |  | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya limosa | 0.1 | 1 | 0
|  |  | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium sordellii | 0.91 | 1 | 0
|  |  | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium bifermentans | 0.55 | 1 | 0
|  |  | IGV 015 | 3; sequenced | DT | 2016-11-30 | Bacillus amyloliquefaciens | 0.66 | 1 | 0
|  |  | IGV 016 | 3; sequenced | DT | 2016-11-30 | Bacillus amyloliquefaciens | 0.1 | 1 | 0
|  |  | B 44 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N11 | Bacillus badius | 0.56 | 1 | 0
|  |  | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.03.2017 I18 | Bacillus ectoiniformans | 0.87 | 1 | 0
|  |  | B 62 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 H21 | Bacillus licheniformis | 0.75 | 1 | 0
|  |  | DSM 299 | 3; public strain collection | DT | 2016-04-08 | Bacillus mycoides | 1.8 | 1 | 0
|  |  | B 17 | 3; sequenced (CVUAS) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F7 | Bacillus pumilus | 0.92 | 1 | 0
|  |  | B 020 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2.12.2014, H11 | Bacillus subtilis | 0.69 | 1 | 0
|  |  | B 032 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014, K18 | Bacillus subtilis | 0.44 | 1 | 0
|  |  | IGV 013 | 3; sequenced | DT | 2016-11-09 | Bacillus subtilis | 0.65 | 1 | 0
|  |  | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F22 | Cytobacillus firmus | 1.11 | 1 | 0
|  |  | B 48 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.01.2015 H18 | Cytobacillus firmus | 0.95 | 1 | 0
|  |  | B 70 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.02.2015 N3 | Cytobacillus firmus | 1.16 | 1 | 0
|  |  | B 88 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.02.2015 L16 | Cytobacillus firmus | 0.94 | 1 | 0
|  |  | B 46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N14 | Niallia circulans | 0.63 | 1 | 0
|  |  | B 2 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2014 E5 | Priestia megaterium | 0.54 | 1 | 0
|  |  | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 N5 | Virgibacillus pantothenticus | 0.8 | 1 | 0
|  |  | A281_1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2012 N7 | Alkalihalobacillus clausii | 0.13 | 1 | 0
|  |  | CVUAS 31071,2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus clausii | 0.08 | 1 | 0
|  |  | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.02.2015 N6 | Alkalihalobacillus clausii | 0.14 | 1 | 0
|  |  | Oman 24 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2012 N9 | Alkalihalobacillus clausii | 0.53 | 1 | 0
|  |  | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 13.02.2015, K10, [0034] | Bacillus alveayuensis | 0.6 | 1 | 0
|  |  | DSM 7 B159 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 01.03.2016 E10 | Bacillus amyloliquefaciens | 0.59 | 1 | 0
|  |  | DSM 7 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCA/TA2 (A) | 16.02.2015, K16, [0034] | Bacillus amyloliquefaciens | 0.36 | 1 | 0
|  |  | DSM 7 | 3; type strain | DT | 2014-01-29 | Bacillus amyloliquefaciens | 0.56 | 1 | 0
|  |  | A269-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.04.2012; H15 | Bacillus anthracis | 1.2 | 1 | 0
|  |  | A1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.05.2015; I20 | Bacillus anthracis | 1.72 | 1 | 0
|  |  | A100 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; E16 | Bacillus anthracis | 1.69 | 1 | 0
|  |  | A101 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E08 | Bacillus anthracis | 1.57 | 1 | 0
|  |  | A102 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E09 | Bacillus anthracis | 1.63 | 1 | 0
|  |  | A103 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E10 | Bacillus anthracis | 1.56 | 1 | 0
|  |  | A106 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2016; E17 | Bacillus anthracis | 1.77 | 1 | 0
|  |  | A107 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2016; C19 | Bacillus anthracis | 1.67 | 1 | 0
|  |  | A108 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E13 | Bacillus anthracis | 1.5 | 1 | 0
|  |  | A11 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.05.2015; C08 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A110 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.11.2015; F11 | Bacillus anthracis | 1.5 | 1 | 0
|  |  | A111 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E15 | Bacillus anthracis | 1.73 | 1 | 0
|  |  | A112 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.02.2016; F09 | Bacillus anthracis | 1.43 | 1 | 0
|  |  | A112-1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.02.2013; C19 | Bacillus anthracis | 1.23 | 1 | 0
|  |  | A113 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E04 | Bacillus anthracis | 1.58 | 1 | 0
|  |  | A118 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H13 | Bacillus anthracis | 1.73 | 1 | 0
|  |  | A120 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H15 | Bacillus anthracis | 1.6 | 1 | 0
|  |  | A122 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H17 | Bacillus anthracis | 1.67 | 1 | 0
|  |  | A123 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H18 | Bacillus anthracis | 1.55 | 1 | 0
|  |  | A124 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H19 | Bacillus anthracis | 1.74 | 1 | 0
|  |  | A125 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2006; H20 | Bacillus anthracis | 1.65 | 1 | 0
|  |  | A126 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; H21 | Bacillus anthracis | 1.83 | 1 | 0
|  |  | A127 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.11.2015; F22 | Bacillus anthracis | 1.68 | 1 | 0
|  |  | A128 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.02.2016; H13 | Bacillus anthracis | 1.56 | 1 | 0
|  |  | A130 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.11.2015; K08 | Bacillus anthracis | 1.43 | 1 | 0
|  |  | A131 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H13 | Bacillus anthracis | 1.32 | 1 | 0
|  |  | A132 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2006; K21 | Bacillus anthracis | 1.82 | 1 | 0
|  |  | A133 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.02.2016; H20 | Bacillus anthracis | 1.6 | 1 | 0
|  |  | A135 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H14 | Bacillus anthracis | 1.35 | 1 | 0
|  |  | A136 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2006; N07 | Bacillus anthracis | 1.24 | 1 | 0
|  |  | A137 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.11.2015; L07 | Bacillus anthracis | 1.55 | 1 | 0
|  |  | A138-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.03.2013; F14 | Bacillus anthracis | 1.32 | 1 | 0
|  |  | A15 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.06.2015; H04 | Bacillus anthracis | 1.73 | 1 | 0
|  |  | A16 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.06.2015; H06 | Bacillus anthracis | 1.44 | 1 | 0
|  |  | A173-2 (Ringversuch A938-2) | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.05.2017; C04 | Bacillus anthracis | 1.55 | 1 | 0
|  |  | A19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; L09 | Bacillus anthracis | 1.57 | 1 | 0
|  |  | A22 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I07 | Bacillus anthracis | 1.6 | 1 | 0
|  |  | A23 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I09 | Bacillus anthracis | 1.5 | 1 | 0
|  |  | A24 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I11 | Bacillus anthracis | 1.46 | 1 | 0
|  |  | A27 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; K16 | Bacillus anthracis | 1.74 | 1 | 0
|  |  | A28 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; K18 | Bacillus anthracis | 1.7 | 1 | 0
|  |  | A29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; I20 | Bacillus anthracis | 1.64 | 1 | 0
|  |  | A294 1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; H10 | Bacillus anthracis | 1.23 | 1 | 0
|  |  | A3 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.05.2015; H08 | Bacillus anthracis | 1.45 | 1 | 0
|  |  | A300-1a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.03.2013; N20 | Bacillus anthracis | 1.22 | 1 | 0
|  |  | A300-1b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.03.2013; I04 | Bacillus anthracis | 1.46 | 1 | 0
|  |  | A300-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.10.2012; E16 | Bacillus anthracis | 1.36 | 1 | 0
|  |  | A300-2b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.03.2013; F10 | Bacillus anthracis | 1.47 | 1 | 0
|  |  | A302-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.03.2013; C13 | Bacillus anthracis | 1.33 | 1 | 0
|  |  | A306-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.03.2013; E19 | Bacillus anthracis | 1.22 | 1 | 0
|  |  | A315-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.03.2013; F05 | Bacillus anthracis | 1.23 | 1 | 0
|  |  | A315-2 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.03.2013; E16 | Bacillus anthracis | 1.23 | 1 | 0
|  |  | A32 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.06.2015; K22 | Bacillus anthracis | 1.51 | 1 | 0
|  |  | A33 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; L13 | Bacillus anthracis | 1.44 | 1 | 0
|  |  | A34 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L06 | Bacillus anthracis | 1.56 | 1 | 0
|  |  | A35 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.04.2006; K19 | Bacillus anthracis | 1.69 | 1 | 0
|  |  | A36 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.04.2006; K20 | Bacillus anthracis | 1.62 | 1 | 0
|  |  | A37 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L10 | Bacillus anthracis | 1.48 | 1 | 0
|  |  | A38 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; L15 | Bacillus anthracis | 1.49 | 1 | 0
|  |  | A39 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L14 | Bacillus anthracis | 1.62 | 1 | 0
|  |  | A40 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.04.2006; N06 | Bacillus anthracis | 1.59 | 1 | 0
|  |  | A41 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; N18 | Bacillus anthracis | 1.59 | 1 | 0
|  |  | A42 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L20 | Bacillus anthracis | 1.49 | 1 | 0
|  |  | A43 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.06.2015; L21 | Bacillus anthracis | 1.54 | 1 | 0
|  |  | A44 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; E03 | Bacillus anthracis | 1.42 | 1 | 0
|  |  | A45 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; E06 | Bacillus anthracis | 1.35 | 1 | 0
|  |  | A46 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; E07 | Bacillus anthracis | 1.53 | 1 | 0
|  |  | A49 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2015; C12 | Bacillus anthracis | 1.5 | 1 | 0
|  |  | A58 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; E13 | Bacillus anthracis | 1.54 | 1 | 0
|  |  | A59 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; E15 | Bacillus anthracis | 1.57 | 1 | 0
|  |  | A6 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.05.2015; F13 | Bacillus anthracis | 1.57 | 1 | 0
|  |  | A60 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; C17 | Bacillus anthracis | 1.38 | 1 | 0
|  |  | A61 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; C20 | Bacillus anthracis | 1.61 | 1 | 0
|  |  | A62 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; C21 | Bacillus anthracis | 1.55 | 1 | 0
|  |  | A63 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.04.2006; E14 | Bacillus anthracis | 1.81 | 1 | 0
|  |  | A64 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; H05 | Bacillus anthracis | 1.56 | 1 | 0
|  |  | A65 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; F07 | Bacillus anthracis | 1.2 | 1 | 0
|  |  | A66 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; F09 | Bacillus anthracis | 1.59 | 1 | 0
|  |  | A67 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2015; H12 | Bacillus anthracis | 1.59 | 1 | 0
|  |  | A68 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2006; H13 | Bacillus anthracis | 1.84 | 1 | 0
|  |  | A69 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; F16 | Bacillus anthracis | 1.68 | 1 | 0
|  |  | A72 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2006; H17 | Bacillus anthracis | 1.65 | 1 | 0
|  |  | A73 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2006; H18 | Bacillus anthracis | 1.82 | 1 | 0
|  |  | A74 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; I07 | Bacillus anthracis | 1.43 | 1 | 0
|  |  | A8 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.05.2015; H18 | Bacillus anthracis | 1.67 | 1 | 0
|  |  | A80 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; K09 | Bacillus anthracis | 1.39 | 1 | 0
|  |  | A81 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; K12 | Bacillus anthracis | 1.62 | 1 | 0
|  |  | A82 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; K13 | Bacillus anthracis | 1.58 | 1 | 0
|  |  | A83 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; K17 | Bacillus anthracis | 1.7 | 1 | 0
|  |  | A84 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.02.2016; N20 | Bacillus anthracis | 1.71 | 1 | 0
|  |  | A85 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; K19 | Bacillus anthracis | 1.01 | 1 | 0
|  |  | A86 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2016; E09 | Bacillus anthracis | 1.59 | 1 | 0
|  |  | A87 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; K21 | Bacillus anthracis | 1.54 | 1 | 0
|  |  | A88 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; N05 | Bacillus anthracis | 1.79 | 1 | 0
|  |  | A88 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; N04 | Bacillus anthracis | 1.84 | 1 | 0
|  |  | A89 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2006; N07 | Bacillus anthracis | 1.37 | 1 | 0
|  |  | A89 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; C10 | Bacillus anthracis | 1.57 | 1 | 0
|  |  | A90 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; C09 | Bacillus anthracis | 1.39 | 1 | 0
|  |  | A90 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.02.2016; E12 | Bacillus anthracis | 1.51 | 1 | 0
|  |  | A91 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.04.2006; E04 | Bacillus anthracis | 1.34 | 1 | 0
|  |  | A92 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; C13 | Bacillus anthracis | 1.55 | 1 | 0
|  |  | A93 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.11.2015; E15 | Bacillus anthracis | 1.59 | 1 | 0
|  |  | AF 28 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H07 | Bacillus anthracis | 1.44 | 1 | 0
|  |  | AF 29 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H08 | Bacillus anthracis | 1.42 | 1 | 0
|  |  | AF 31 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H09 | Bacillus anthracis | 1.45 | 1 | 0
|  |  | AF 71 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H10 | Bacillus anthracis | 1.4 | 1 | 0
|  |  | AF 96 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H11 | Bacillus anthracis | 1.39 | 1 | 0
|  |  | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E06 | Bacillus anthracis | 1.47 | 1 | 0
|  |  | ATCC 14578 A70 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.08.2015; H18 | Bacillus anthracis | 1.58 | 1 | 0
|  |  | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; E04 | Bacillus anthracis | 1.92 | 1 | 0
|  |  | SA 20 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.06.2009; H12 | Bacillus anthracis | 1.18 | 1 | 0
|  |  | Sterne | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015; E03 | Bacillus anthracis | 1.66 | 1 | 0
|  |  | vollum | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2007; E22 | Bacillus anthracis | 1.19 | 1 | 0
|  |  | Wirt | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015; C05 | Bacillus anthracis | 1.52 | 1 | 0
|  |  | ATCC 9372 WIS-St. Nr. 39 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007 E10 | Bacillus atrophaeus | 0.34 | 1 | 0
|  |  | B 004 | 2; 0 | DT | 2014-01-22 | Bacillus atrophaeus | 0.79 | 1 | 0
|  |  | B 4 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 E7 | Bacillus atrophaeus | 0.73 | 1 | 0
|  |  | B 93 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 N3 | Bacillus atrophaeus | 0.64 | 1 | 0
|  |  | DSM 2277 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E13, [0034] | Bacillus atrophaeus | 0.9 | 1 | 0
|  |  | DSM 675 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 13.04.2007, H17, [0034] | Bacillus atrophaeus | 0.46 | 1 | 0
|  |  | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 L8 | Bacillus atrophaeus | 0.8 | 1 | 0
|  |  | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H20 | Bacillus badius | 1.01 | 1 | 0
|  |  | ATCC 11778 IGV | 3; public strain collection | DT | 2016-02-03 | Bacillus cereus | 1.9 | 1 | 0
|  |  | AaCol | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2009 H3 | Bacillus cereus | 1.47 | 1 | 0
|  |  | SK1b | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2009 D21 | Bacillus cereus | 1.7 | 1 | 0
|  |  | B194 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H10 | Bacillus cereus | 1.62 | 1 | 0
|  |  | B195 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H11 | Bacillus cereus | 1.51 | 1 | 0
|  |  | B196 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H12 | Bacillus cereus | 1.31 | 1 | 0
|  |  | CVUAS 34551 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus cereus | 1.78 | 1 | 0
|  |  | CVUAS 34552 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus cereus | 1.55 | 1 | 0
|  |  | 301101RA0432 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2007; E20 | Bacillus cereus | 1.17 | 1 | 0
|  |  | 43 Paciflor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; E16 | Bacillus cereus | 1.36 | 1 | 0
|  |  | 8 Toyocerin | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.06.2006 E14 | Bacillus cereus | 1.88 | 1 | 0
|  |  | ATCC 10876 B270 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I09 | Bacillus cereus | 1.48 | 1 | 0
|  |  | ATCC 10987 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 23.10.2015, E6, [0034] | Bacillus cereus | 1.53 | 1 | 0
|  |  | ATCC 11778 B23 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.08.2006 E21 | Bacillus cereus | 1.59 | 1 | 0
|  |  | DSM 31 B69 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 06.02.2015 I21 | Bacillus cereus | 1.7 | 1 | 0
|  |  | ATCC 14579 RKI | 3; type strain | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E7, [0034] | Bacillus cereus | 1.63 | 1 | 0
|  |  | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus cereus | 1.99 | 1 | 0
|  |  | ATCC 33018 B245 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2016; H22 | Bacillus cereus | 1.5 | 1 | 0
|  |  | ATCC 33019 B244 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.04.2015; H15 | Bacillus cereus | 1.53 | 1 | 0
|  |  | ATCC 6464 B331 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; K14 | Bacillus cereus | 1.58 | 1 | 0
|  |  | ATCC 7004 B284 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; L06 | Bacillus cereus | 1.66 | 1 | 0
|  |  | B10 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 1.12.2014 C18 | Bacillus cereus | 1.52 | 1 | 0
|  |  | B108 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 15.08.2006 N6 | Bacillus cereus | 2.03 | 0 | 1
|  |  | B109 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.03.2015 N20 | Bacillus cereus | 1.24 | 1 | 0
|  |  | B110 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 15.08.2006 N8 | Bacillus cereus | 1.3 | 1 | 0
|  |  | B146 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 27.03.2015 I11 | Bacillus cereus | 1.78 | 1 | 0
|  |  | B16 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.12.2014 F6 | Bacillus cereus | 1.53 | 1 | 0
|  |  | B197 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.04.2015; H22 | Bacillus cereus | 1.35 | 1 | 0
|  |  | B198 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2015; K09 | Bacillus cereus | 1.53 | 1 | 0
|  |  | B199 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H15 | Bacillus cereus | 1.72 | 1 | 0
|  |  | B200 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.08.2006; H16 | Bacillus cereus | 1.44 | 1 | 0
|  |  | B232 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.08.2006; K10 | Bacillus cereus | 1.42 | 1 | 0
|  |  | B24 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H08 | Bacillus cereus | 1.62 | 1 | 0
|  |  | B252 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; K03 | Bacillus cereus | 1.51 | 1 | 0
|  |  | B253 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H16 | Bacillus cereus | 1.66 | 1 | 0
|  |  | B254 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H17 | Bacillus cereus | 1.71 | 1 | 0
|  |  | B255 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; F20 | Bacillus cereus | 1.59 | 1 | 0
|  |  | B256 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; K06 | Bacillus cereus | 1.65 | 1 | 0
|  |  | B257 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; K05 | Bacillus cereus | 1.65 | 1 | 0
|  |  | B258 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; I08 | Bacillus cereus | 1.62 | 1 | 0
|  |  | B259 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; K09 | Bacillus cereus | 1.5 | 1 | 0
|  |  | B260 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; I07 | Bacillus cereus | 1.76 | 1 | 0
|  |  | B261 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; I14 | Bacillus cereus | 1.26 | 1 | 0
|  |  | B262 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.09.2006; H06 | Bacillus cereus | 1.33 | 1 | 0
|  |  | B272 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I13 | Bacillus cereus | 1.56 | 1 | 0
|  |  | B273 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I15 | Bacillus cereus | 1.59 | 1 | 0
|  |  | B274 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; K17 | Bacillus cereus | 1.74 | 1 | 0
|  |  | B275 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; K19 | Bacillus cereus | 1.12 | 1 | 0
|  |  | B277 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; I20 | Bacillus cereus | 1.39 | 1 | 0
|  |  | B278 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; L04 | Bacillus cereus | 1.76 | 1 | 0
|  |  | B282 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 29.04.2015; N07 | Bacillus cereus | 1.56 | 1 | 0
|  |  | B293 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.09.2006; K06 | Bacillus cereus | 1.49 | 1 | 0
|  |  | B294 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; L14 | Bacillus cereus | 1.74 | 1 | 0
|  |  | B295 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; L16 | Bacillus cereus | 1.46 | 1 | 0
|  |  | B296 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; N18 | Bacillus cereus | 1.64 | 1 | 0
|  |  | B297 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C07 | Bacillus cereus | 1.52 | 1 | 0
|  |  | B298 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; E09 | Bacillus cereus | 1.54 | 1 | 0
|  |  | B299 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.09.2006; K12 | Bacillus cereus | 1.43 | 1 | 0
|  |  | B3 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.08.2006 E6 | Bacillus cereus | 1.53 | 1 | 0
|  |  | B300 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C14 | Bacillus cereus | 1.66 | 1 | 0
|  |  | B302 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C18 | Bacillus cereus | 1.38 | 1 | 0
|  |  | B303 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; C20 | Bacillus cereus | 1.58 | 1 | 0
|  |  | B304 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; E21 | Bacillus cereus | 1.56 | 1 | 0
|  |  | B305 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; H04 | Bacillus cereus | 1.63 | 1 | 0
|  |  | B306 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; H06 | Bacillus cereus | 1.67 | 1 | 0
|  |  | B308 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.05.2015; H10 | Bacillus cereus | 1.59 | 1 | 0
|  |  | B309 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.09.2006; E10 | Bacillus cereus | 1.73 | 1 | 0
|  |  | B310 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; F13 | Bacillus cereus | 1.6 | 1 | 0
|  |  | B311 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; H16 | Bacillus cereus | 1.64 | 1 | 0
|  |  | B316 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; H22 | Bacillus cereus | 1.58 | 1 | 0
|  |  | B33 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.12.2014 I20 | Bacillus cereus | 1.54 | 1 | 0
|  |  | B332 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; I16 | Bacillus cereus | 1.76 | 1 | 0
|  |  | B74 CCM 88 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 14.08.2006 K9 | Bacillus cereus | 1.48 | 1 | 0
|  |  | BW-B | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015; E07 | Bacillus cereus | 1.36 | 1 | 0
|  |  | CCM 99 B249 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; H10 | Bacillus cereus | 1.81 | 1 | 0
|  |  | CVUAS 1107 | 2; 0 | eDT | 27.01.2016 | Bacillus cereus | 1.92 | 1 | 0
|  |  | CVUAS 3901 | 3; 0 | eDT | 26.04.2016 | Bacillus cereus | 1.81 | 1 | 0
|  |  | DSM 1274 B247 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2015; F06 | Bacillus cereus | 1.55 | 1 | 0
|  |  | DSM 2299 B190 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.08.2006; H08 | Bacillus cereus | 1.33 | 1 | 0
|  |  | DSM 2302 B192 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.04.2015; F13 | Bacillus cereus | 1.84 | 1 | 0
|  |  | DSM 31 B181 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 30.03.2015 E21 | Bacillus cereus | 1.92 | 1 | 0
|  |  | DSM 31 | 3; type strain | eDT | 2016-01-27 | Bacillus cereus | 1.85 | 1 | 0
|  |  | DSM 31 RKI | 3; type strain | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 21.09.2015, E13, [0034] | Bacillus cereus | 1.79 | 1 | 0
|  |  | DSM 3101 B237 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2015; C20 | Bacillus cereus | 1.58 | 1 | 0
|  |  | DSM 318 B233 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2015; C13 | Bacillus cereus | 1.38 | 1 | 0
|  |  | DSM 336 B234 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.08.2006; K12 | Bacillus cereus | 1.35 | 1 | 0
|  |  | DSM 345 B170 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 30.03.2015 C11 | Bacillus cereus | 1.94 | 1 | 0
|  |  | DSM 345 | 3; public strain collection | eDT | 2016-01-27 | Bacillus cereus | 1.79 | 1 | 0
|  |  | DSM 351 B235 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.08.2006; K13 | Bacillus cereus | 1.74 | 1 | 0
|  |  | DSM 4218 B238 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2015; C21 | Bacillus cereus | 1.54 | 1 | 0
|  |  | DSM 4312 B269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.09.2006; H11 | Bacillus cereus | 1.47 | 1 | 0
|  |  | B329 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.05.2015; K11 | Bacillus cereus | 1.54 | 1 | 0
|  |  | DSM 4312 | 3; public strain collection | eDT | 27.01.2016 | Bacillus cereus | 1.88 | 1 | 0
|  |  | DSM 4312 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E5, [0034] | Bacillus cereus | 1.1 | 1 | 0
|  |  | DSM 4313 | 3; public strain collection | eDT | 2016-01-27 | Bacillus cereus | 1.77 | 1 | 0
|  |  | DSM 4490 | 3; public strain collection | in 80% TFA, Zentrifugation und Sterilfiltration, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 21.09.2015, E9, [0034] | Bacillus cereus | 1.87 | 1 | 0
|  |  | DSM 487 B169 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006 E15 | Bacillus cereus | 1.59 | 1 | 0
|  |  | DSM 487 B185 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006 H4 | Bacillus cereus | 1.59 | 1 | 0
|  |  | DSM 508 B186 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 7.04.2015 F6 | Bacillus cereus | 1.91 | 1 | 0
|  |  | DSM 609 | 3; strain from public collection | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 21.09.2015, C17, [0034] | Bacillus cereus | 1.66 | 1 | 0
|  |  | DSM 626 B187 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.04.2015; F07 | Bacillus cereus | 1.62 | 1 | 0
|  |  | DSM 679 B242 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.09.2006; E06 | Bacillus cereus | 1.64 | 1 | 0
|  |  | DSM 689 B243 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.09.2006; E07 | Bacillus cereus | 1.65 | 1 | 0
|  |  | Hohenheim | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2007; K04 | Bacillus cereus | 1.35 | 1 | 0
|  |  | ATCC 9139 B289 | 3; reference strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.04.2015; N07 | Bacillus cereus | 1.71 | 1 | 0
|  |  | NCIMB 11925 | 3; reference strain collection | DT | 2015-07-22 | Bacillus cereus | 1.67 | 1 | 0
|  |  | NCTC 10320 B148 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006 E5 | Bacillus cereus | 1.64 | 1 | 0
|  |  | NCTC 10989 B268 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; N04 | Bacillus cereus | 1.71 | 1 | 0
|  |  | NCTC 6474 B265 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; K20 | Bacillus cereus | 1.55 | 1 | 0
|  |  | NCTC 9939 B267 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.04.2015; I22 | Bacillus cereus | 1.47 | 1 | 0
|  |  | NiMe_1 | 3; 16S, ces-pos | eDT | 2016-01-27 | Bacillus cereus | 1.86 | 1 | 0
|  |  | Paciflor | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2007; K04 | Bacillus cereus | 1.5 | 1 | 0
|  |  | SL2009001164 | 3; 0 | eDT | 2016-01-27 | Bacillus cereus | 1.79 | 1 | 0
|  |  | SL2009005808 | 2; 0 | eDT | 2019-09-23 | Bacillus cereus | 2.03 | 0 | 1
|  |  | SL2009007190 | 2; 0 | eDT | 2019-09-25 | Bacillus cereus | 2.12 | 0 | 1
|  |  | SL2014002835 | 2; CytK2 pos, Krist neg | eDT | 2019-09-25 | Bacillus cereus | 2.01 | 0 | 1
|  |  | Toyocerin | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; H07 | Bacillus cereus | 1.72 | 1 | 0
|  |  | Wasserprobe | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; H06 | Bacillus cereus | 1.77 | 1 | 0
|  |  | WIS-St. Nr. 29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007; E09 | Bacillus cereus | 1.34 | 1 | 0
|  |  | WIS-St. Nr. 311 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007; E12 | Bacillus cereus | 1.59 | 1 | 0
|  |  | WSBC 10028 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H07 | Bacillus cereus | 1.79 | 1 | 0
|  |  | WSBC 10030 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H08 | Bacillus cereus | 1.75 | 1 | 0
|  |  | WSBC 10286 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.06.2006; E09 | Bacillus cereus | 1.53 | 1 | 0
|  |  | WSBC 10288 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H10 | Bacillus cereus | 1.98 | 1 | 0
|  |  | WSBC 10441 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006; H11 | Bacillus cereus | 1.61 | 1 | 0
|  |  | WSBC 10925 | 3; well defined isolat, ces-pos | eDT | 27.01.2016 | Bacillus cereus | 1.83 | 1 | 0
|  |  | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.05.2015; E09 | Bacillus tropicus | 1.84 | 1 | 0
|  |  | CCEB 627 | 3; type strain | DT | 2014-01-22 | Bacillus firmus | 0.58 | 1 | 0
|  |  | CCM 37 | 3; reference strain collection | DT | 2014-01-22 | Bacillus firmus | 1.35 | 1 | 0
|  |  | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.11.2017 F14 | Bacillus halotolerans | 1.13 | 1 | 0
|  |  | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.12.2017 I8 | Bacillus halotolerans | 0.74 | 1 | 0
|  |  | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.12.2016 [H3] | Bacillus inaquosorum | 0.92 | 1 | 0
|  |  | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.02.2016 C20 | Bacillus infantis | 0.63 | 1 | 0
|  |  | ATCC 12759 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007 E10 | Bacillus licheniformis | 0.33 | 1 | 0
|  |  | B 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015, F12 | Bacillus licheniformis | 0.59 | 1 | 0
|  |  | B 118 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015, F18 | Bacillus licheniformis | 0.53 | 1 | 0
|  |  | B 72 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.08.2006 K8 | Bacillus licheniformis | 0.64 | 1 | 0
|  |  | B084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus licheniformis | 0.48 | 1 | 0
|  |  | B116 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.03.2015 F13 | Bacillus licheniformis | 0.71 | 1 | 0
|  |  | BRLK08_04 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.04.2015 I16 | Bacillus licheniformis | 0.48 | 1 | 0
|  |  | DSM 13 B178 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 16.08.2006; E20 | Bacillus licheniformis | 0.74 | 1 | 0
|  |  | DSM 13 | 3; type strain | eDT | 2016-01-27 | Bacillus licheniformis | 0.48 | 1 | 0
|  |  | DSM 13 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.04.2007 E15 | Bacillus licheniformis | 0.39 | 1 | 0
|  |  | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K15 | Bacillus mojavensis | 0.84 | 1 | 0
|  |  | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F4 | Bacillus mojavensis | 0.79 | 1 | 0
|  |  | ATCC 12826 WIS-St. Nr. 283 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007 E11 | Bacillus mycoides | 1.56 | 1 | 0
|  |  | B103 | 2; 0 | DT | 2015-07-22 | Bacillus mycoides | 1.71 | 1 | 0
|  |  | DSM 104119 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.05.2018 I15 | Bacillus mycoides | 1.67 | 1 | 0
|  |  | DSM 11821 | 3; public strain collection | eDT | 2016-01-27 | Bacillus mycoides | 1.44 | 1 | 0
|  |  | DSM 11821 RKI | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 03.05.2007, E4, [0034] | Bacillus mycoides | 1.28 | 1 | 0
|  |  | DSM 2048 | 3; typestrain | eDT | 09.12.2015 | Bacillus mycoides | 1.56 | 1 | 0
|  |  | DSM 2048 RKI | 3; typestrain | EFex [BRU03] | 2022-01-13 | Bacillus mycoides | 1.63 | 1 | 0
|  |  | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus mycoides | 1.67 | 1 | 0
|  |  | WSBC 10204 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K5, [0034] | Bacillus mycoides | 1.53 | 1 | 0
|  |  | WSBC 10207 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K6, [0034] | Bacillus mycoides | 1.65 | 1 | 0
|  |  | WSBC 10277 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H14 | Bacillus mycoides | 1.6 | 1 | 0
|  |  | WSBC 10279 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H15 | Bacillus mycoides | 1.27 | 1 | 0
|  |  | WSBC 10291 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H16 | Bacillus mycoides | 1.56 | 1 | 0
|  |  | WSBC 10295 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K7, [0034] | Bacillus mycoides | 1.69 | 1 | 0
|  |  | WSBC 10360 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.11.2006 H17 | Bacillus mycoides | 1.51 | 1 | 0
|  |  | WSBC 10363 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K8, [0034] | Bacillus mycoides | 1.56 | 1 | 0
|  |  | WSBC 10387 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 11.07.2006, K9, [0034] | Bacillus mycoides | 1.64 | 1 | 0
|  |  | ATCC 11778 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E8, [0034] | Bacillus pacificus | 1.58 | 1 | 0
|  |  | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus pseudomycoides | 1.98 | 1 | 0
|  |  | DSM 12442 | 3; type strain | eDT | 2016-01-27 | Bacillus pseudomycoides | 1.58 | 1 | 0
|  |  | DSM 12442 RKI | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 03.05.2007, N9, [0034] | Bacillus pseudomycoides | 1.84 | 1 | 0
|  |  | DSM 12443 | 3; public strain collection | eDT | 09.12.2015 | Bacillus pseudomycoides | 1.69 | 1 | 0
|  |  | B 031 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2006, H6 | Bacillus pumilus | 0.75 | 1 | 0
|  |  | B 91 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.08.2006 K15 | Bacillus pumilus | 1 | 1 | 0
|  |  | B017 | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus pumilus | 1.14 | 1 | 0
|  |  | CCM 340 | 3; reference strain collection | DT | 2014-01-22 | Bacillus pumilus | 0.88 | 1 | 0
|  |  | DSM 13835 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2007, N7 | Bacillus pumilus | 0.38 | 1 | 0
|  |  | DSM 27 B165 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E7 | Bacillus pumilus | 0.82 | 1 | 0
|  |  | DSM 27 Neukauf2017 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 , K22 | Bacillus pumilus | 1.07 | 1 | 0
|  |  | DSM 27 | 3; type strain | DT | 2014-01-22 | Bacillus pumilus | 1.14 | 1 | 0
|  |  | DSM 355 B161 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; N10 | Bacillus pumilus | 0.88 | 1 | 0
|  |  | DSM 361 B162 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E6 | Bacillus pumilus | 0.88 | 1 | 0
|  |  | DSM 492 B160 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.08.2006, E11 | Bacillus pumilus | 0.89 | 1 | 0
|  |  | MM 171 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007 H6 | Bacillus pumilus | 0.55 | 1 | 0
|  |  | NVH-819-05 | 3; well described isolate [0017] | DT | 2015-04-21 | Bacillus pumilus | 0.29 | 1 | 0
|  |  | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.05.2018, N12 | Bacillus siamensis | 0.82 | 1 | 0
|  |  | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2016, K18 | Bacillus sonorensis | 0.56 | 1 | 0
|  |  | DSM 12369 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.04.2007, K15 | Bacillus sp-DSM-12369 | 0.66 | 1 | 0
|  |  | ATCC 6633 B13 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2014; E22 | Bacillus spizizenii | 1.01 | 1 | 0
|  |  | ATCC 6633 | 3; public strain collection | DT | 2016-02-11 | Bacillus spizizenii | 0.57 | 1 | 0
|  |  | ATCC 6633 DSM 347 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007 E9 | Bacillus spizizenii | 0.83 | 1 | 0
|  |  | DSM 1087 B156 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; K19 | Bacillus spizizenii | 0.9 | 1 | 0
|  |  | DSM 1087 | 3; public strain collection | DT | 2014-01-22 | Bacillus spizizenii | 0.88 | 1 | 0
|  |  | DSM 1087 RKI | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 18.12.2017; N5 | Bacillus spizizenii | 0.69 | 1 | 0
|  |  | DSM 347 | 3; public strain collection | DT | 2014-01-10 | Bacillus spizizenii | 0.83 | 1 | 0
|  |  | DSM 618 | 4; strain copy from public strain collection | eDT | 27.01.2016 | Bacillus spizizenii | 0.74 | 1 | 0
|  |  | B 130 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015, I10 | Bacillus subtilis | 0.84 | 1 | 0
|  |  | DSM 402 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 28.11.2017 F4 | Bacillus subtilis | 0.72 | 1 | 0
|  |  | ATCC 6051 B158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; N03 | Bacillus subtilis | 0.68 | 1 | 0
|  |  | B 061 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 07.01.2015; F20 | Bacillus subtilis | 0.77 | 1 | 0
|  |  | B rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2007; N14 | Bacillus subtilis | 0.83 | 1 | 0
|  |  | B129 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016; N19 | Bacillus subtilis | 0.83 | 1 | 0
|  |  | B14 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.08.2006; E14 | Bacillus subtilis | 0.67 | 1 | 0
|  |  | B154 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2015; I18 | Bacillus subtilis | 0.74 | 1 | 0
|  |  | B19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2014; H09 | Bacillus subtilis | 0.75 | 1 | 0
|  |  | BRLK08_02 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015; I15 | Bacillus subtilis | 0.47 | 1 | 0
|  |  | CCM 110 B068 | 3; reference strain collection | DT | 2015-07-22 | Bacillus subtilis | 0.78 | 1 | 0
|  |  | DSM 10 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.03.2016 C20 | Bacillus subtilis | 0.5 | 1 | 0
|  |  | DSM 1092 | 3; reference strain collection | DT | 2014-01-22 | Bacillus subtilis | 0.74 | 1 | 0
|  |  | E rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.04.2007; N15 | Bacillus subtilis | 0.59 | 1 | 0
|  |  | I rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.04.2007; E09 | Bacillus subtilis | 0.49 | 1 | 0
|  |  | LVM03 LVU 03 | 3; sequenced (LGL) | eDT | 2015-07-09 | Bacillus subtilis | 0.66 | 1 | 0
|  |  | MM 508 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.05.2007; H11 | Bacillus subtilis | 0.9 | 1 | 0
|  |  | RKI-w2gr | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2008; H16 | Bacillus subtilis | 1.06 | 1 | 0
|  |  | WIS-St. Nr. 491 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.01.2007; E14 | Bacillus subtilis | 0.58 | 1 | 0
|  |  | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 1.84 | 1 | 0
|  |  | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 1.88 | 1 | 0
|  |  | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus thuringiensis | 1.79 | 1 | 0
|  |  | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 1.82 | 1 | 0
|  |  | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 1.73 | 1 | 0
|  |  | B 8 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 1.12.2014 E15 | Bacillus thuringiensis | 1.66 | 1 | 0
|  |  | CVUAS 10203 | 4; Cry-pos, Cyt-K2, Krist. Pos | eDT | 27.01.2016 | Bacillus thuringiensis | 1.77 | 1 | 0
|  |  | CVUAS 5760 | 4; Cry-pos, Krist. Pos | eDT | 27.01.2016 | Bacillus thuringiensis | 1.91 | 1 | 0
|  |  | CVUAS 5761 | 4; Cry-pos, Krist. Pos | eDT | 27.01.2016 | Bacillus thuringiensis | 1.74 | 1 | 0
|  |  | DSM 2046 | 3; type strain | DT | 2014-01-22 | Bacillus thuringiensis | 1.61 | 1 | 0
|  |  | DSM 2046 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015 F5 | Bacillus thuringiensis | 1.45 | 1 | 0
|  |  | DSM 350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.02.2015 K6 | Bacillus thuringiensis | 1.77 | 1 | 0
|  |  | DSM 5815 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.09.2015 C21 | Bacillus thuringiensis | 1.42 | 1 | 0
|  |  | DSM 6073 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.12.2015 E18 | Bacillus thuringiensis | 1.78 | 1 | 0
|  |  | DSM 6890 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2015, L4 | Bacillus thuringiensis | 1.66 | 1 | 0
|  |  | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 07.11.2006 K4 | Bacillus thuringiensis | 1.47 | 1 | 0
|  |  | CVUAS 2786,2 | 3; sequenced | DT | 2013-07-31 | Bacillus timorensis | 0.56 | 1 | 0
|  |  | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.07.2017, H17 | Bacillus vallismortis | 0.87 | 1 | 0
|  |  | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.02.2016 K19 | Cytobacillus horneckiae | 1.17 | 1 | 0
|  |  | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.04.2014 H5 | Cytobacillus horneckiae | 0.99 | 1 | 0
|  |  | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus kochii | 1.08 | 1 | 0
|  |  | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 19.02.2015, E18 | Domibacillus aminovorans | 1.27 | 1 | 0
|  |  | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 15.02.2015, E8 | Domibacillus aminovorans | 1.06 | 1 | 0
|  |  | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.02.2015 H5 | Evansella clarkii | 1.03 | 1 | 0
|  |  | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F7 | Evansella clarkii | 0.66 | 1 | 0
|  |  | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus stearothermophilus | 0.75 | 1 | 0
|  |  | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 I10 | Halalkalihalobacillus halodurans | 0.61 | 1 | 0
|  |  | DSM 6939 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 C18 | Halalkalihalobacillus halodurans | 0.81 | 1 | 0
|  |  | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 E19 | Halalkalihalobacillus halodurans | 0.62 | 1 | 0
|  |  | B041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia sporothermodurans | 0.14 | 1 | 0
|  |  | B106 | 3; sequenced | DT | 2015-07-22 | Heyndrickxia sporothermodurans | 0.71 | 1 | 0
|  |  | DSM 10599 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015, H9 | Heyndrickxia sporothermodurans | 0.78 | 1 | 0
|  |  | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017; N12 | Lederbergia lenta | 0.77 | 1 | 0
|  |  | DSM 9 | 3; type strain | DT | 2014-01-22 | Lederbergia lenta | 0.67 | 1 | 0
|  |  | DSM 9 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2007 H7 | Lederbergia lenta | 0.81 | 1 | 0
|  |  | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus bataviensis | 1.16 | 1 | 0
|  |  | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017 N13 | Neobacillus cucumis | 1.26 | 1 | 0
|  |  | B 216 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 L4 | Niallia circulans | 0.76 | 1 | 0
|  |  | DSM 11 | 3; type strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 13.04.2007, H11, [0034] | Niallia circulans | 0.55 | 1 | 0
|  |  | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E14, [0034] | Niallia circulans | 0.64 | 1 | 0
|  |  | CVUAS 2582,2 | 3; sequenced | eDT | 2015-09-23 | Niallia nealsonii | 0.93 | 1 | 0
|  |  | CVUAS 33253 | 2; sequenced | EFex | 2023-01-12 | Niallia sp | 0.75 | 1 | 0
|  |  | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K20 | Peribacillus muralis | 0.96 | 1 | 0
|  |  | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016; K20 | Peribacillus psychrosaccharolyticus | 0.6 | 1 | 0
|  |  | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus simplex | 0.64 | 1 | 0
|  |  | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2016 H13 | Peribacillus simplex | 0.63 | 1 | 0
|  |  | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 15.12.2016, K15, [0096] | Peribacillus simplex | 0.63 | 1 | 0
|  |  | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016 K15 | Priestia flexa | 0.79 | 1 | 0
|  |  | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H18 | Priestia flexa | 0.55 | 1 | 0
|  |  | ATCC 14581 | 3; type strain | DT | 2014-01-29 | Priestia megaterium | 0.74 | 1 | 0
|  |  | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015 H21 | Priestia megaterium | 0.3 | 1 | 0
|  |  | B 333 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.05.2015 K18 | Priestia megaterium | 0.53 | 1 | 0
|  |  | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia megaterium | 0.28 | 1 | 0
|  |  | DSM 32 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.04.2007 K17 | Priestia megaterium | 0.39 | 1 | 0
|  |  | DSM 333 B175 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015 C16 | Priestia megaterium | 0.66 | 1 | 0
|  |  | DSM 509 B172 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015 E14 | Priestia megaterium | 0.7 | 1 | 0
|  |  | DSM 90 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2007 H10 | Priestia megaterium | 0.71 | 1 | 0
|  |  | RRLK06_01 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.05.2015 K18 | Priestia megaterium | 0.68 | 1 | 0
|  |  | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E3 | Salibacillus salarius | 0.84 | 1 | 0
|  |  | DSM 6307 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 H11 | Sutcliffiella cohnii | 0.88 | 1 | 0
|  |  | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F13 | Sutcliffiella cohnii | 0.85 | 1 | 0
|  |  | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 K8 | Sutcliffiella cohnii | 0.4 | 1 | 0
|  |  | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus saccharophilus | 0.6 | 1 | 0
|  |  | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus proomii | 0.5 | 1 | 0
|  |  | CVUAS 3616,2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus proomii | 0.67 | 1 | 0
|  |  | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H15 | Weizmannia coagulans | 0.78 | 1 | 0
|  |  | CVUAS 71,2 | 3; sequenced | DT | 2013-12-13 | Weizmannia ginsengihumi | 0.49 | 1 | 0
|  |  | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.55 | 1 | 0
|  |  | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.86 | 1 | 0
|  |  | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia zopfii | 1.15 | 1 | 0
|  |  | B092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus fusiformis | 0.8 | 1 | 0
|  |  | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus fusiformis | 0.45 | 1 | 0
|  |  | CVUAS 7192,2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus macroides | 1.03 | 1 | 0
|  |  | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus sphaericus | 0.71 | 1 | 0
|  |  | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus sphaericus | 0.92 | 1 | 0
|  |  | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus saudimassiliensis | 0.86 | 1 | 0
|  |  | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus odysseyi | 0.67 | 1 | 0
|  |  | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium acetylicum | 0.52 | 1 | 0
|  |  | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium enclense | 0.57 | 1 | 0
|  |  | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium profundum | 0.64 | 1 | 0
|  |  | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix thermosphacta | 0.42 | 1 | 0
|  |  | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix thermosphacta | 0.92 | 1 | 0
|  |  | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria aquatica | 0.74 | 1 | 0
|  |  | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria aquatica | 0.25 | 1 | 0
|  |  | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria fleischmannii | 0.36 | 1 | 0
|  |  | LU2006 3 | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria fleischmannii | 0.62 | 1 | 0
|  |  | CIP 105447 | 3; type strain | EFEx | 2017-02-15 | Listeria grayi | 0.68 | 1 | 0
|  |  | DSM 20596 | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria grayi | 0.46 | 1 | 0
|  |  | CVUAS 10132,2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria innocua | 0.85 | 1 | 0
|  |  | CVUAS 8569 | 3; 16S Sequenz | DT | 19.02.2013 | Listeria innocua | 0.54 | 1 | 0
|  |  | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria ivanovii | 0.64 | 1 | 0
|  |  | CVUAS 32100 | 2; sequenced | EFex | 2019-12-13 | Listeria ivanovii | 0.88 | 1 | 0
|  |  | DSM 23813 | 3; type strain | eDT | 2015-02-18 | Listeria marthii | 0.1 | 1 | 0
|  |  | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.41 | 1 | 0
|  |  | CVUAS 873,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.33 | 1 | 0
|  |  | CVUAS 240,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.13 | 1 | 0
|  |  | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.14 | 1 | 0
|  |  | 10BfR L0206 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.41 | 1 | 0
|  |  | 10BfR L0066 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.55 | 1 | 0
|  |  | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.35 | 1 | 0
|  |  | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.93 | 1 | 0
|  |  | CVUAS 7366,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.5 | 1 | 0
|  |  | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria newyorkensis | 0.27 | 1 | 0
|  |  | DSM 22097 | 3; type strain | eDT | 2015-11-20 | Listeria rocourtiae | 0.32 | 1 | 0
|  |  | CVUAS 2763 | 3; 16S Sequenz | DT | 19.02.2013 | Listeria seeligeri | 0.31 | 1 | 0
|  |  | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria seeligeri | 0.51 | 1 | 0
|  |  | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria sp | 0.36 | 1 | 0
|  |  | CVUAS 2412,3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria welshimeri | 0.53 | 1 | 0
|  |  | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria welshimeri | 0.66 | 1 | 0
|  |  | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus laterosporus | 0.32 | 1 | 0
|  |  | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus laterosporus | 0.24 | 1 | 0
|  |  | B210 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 K19 | Paenibacillus alvei | 0.79 | 1 | 0
|  |  | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus chibensis | 0.74 | 1 | 0
|  |  | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus nuruki | 0.36 | 1 | 0
|  |  | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.09.2013 C8 | Paenibacillus polymyxa | 0.99 | 1 | 0
|  |  | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.11.2017 C18 | Paenibacillus polymyxa | 1.07 | 1 | 0
|  |  | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus sp-CVUAS-6369 | 0.65 | 1 | 0
|  |  | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus sp-E8a | 0.64 | 1 | 0
|  |  | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus taichungensis | 0.4 | 1 | 0
|  |  | FLI003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus taichungensis | 0.46 | 1 | 0
|  |  | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus tundrae | 0.63 | 1 | 0
|  |  | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus halophilus | 0.36 | 1 | 0
|  |  | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus huakuii | 0.58 | 1 | 0
|  |  | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus psychrophilus | 0.32 | 1 | 0
|  |  | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus carouselicus | 0.55 | 1 | 0
|  |  | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 14.03.2013 | Macrococcus caseolyticus | 0.84 | 1 | 0
|  |  | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 0.97 | 1 | 0
|  |  | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus lentus | 0.42 | 1 | 0
|  |  | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus sciuri | 0.27 | 1 | 0
|  |  | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus sciuri | 0.13 | 1 | 0
|  |  | CVUAS 9937,2 | 3; 16S rDNA-sequencing | DT | 01.12.2015 | Mammaliicoccus vitulinus | 0.25 | 1 | 0
|  |  | CVUAS 343,2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus vitulinus | 0.6 | 1 | 0
|  |  | CVUAS 5687 | 3; sequenced | DT | 2013-04-16 | Staphylococcus agnetis | 0.87 | 1 | 0
|  |  | TUM 7385 | 3; sequenced | DT | 02.10.2013 | Staphylococcus agnetis | 1.07 | 1 | 0
|  |  | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus argenteus | 0.98 | 1 | 0
|  |  | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus argenteus | 1.06 | 1 | 0
|  |  | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus arlettae | 0.43 | 1 | 0
|  |  | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus arlettae | 0.68 | 1 | 0
|  |  | ATCC 29213 | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus aureus | 0.89 | 1 | 0
|  |  | CVUAS 9914,2 | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus aureus | 0.56 | 1 | 0
|  |  | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus auricularis | 1 | 1 | 0
|  |  | DSM 20326 (T) | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus capitis | 0.75 | 1 | 0
|  |  | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus capitis | 0.46 | 1 | 0
|  |  | DSM 20608 (T) | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus caprae | 0.79 | 1 | 0
|  |  | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus caprae | 0.69 | 1 | 0
|  |  | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus carnosus | 1.02 | 1 | 0
|  |  | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus carnosus | 0.34 | 1 | 0
|  |  | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus chromogenes | 0.91 | 1 | 0
|  |  | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus chromogenes | 0.98 | 1 | 0
|  |  | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus cohnii | 0.8 | 1 | 0
|  |  | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus cohnii | 0.3 | 1 | 0
|  |  | x | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 15.11.2011 E20 | Staphylococcus delphini | 0.87 | 1 | 0
|  |  | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus delphini | 0.82 | 1 | 0
|  |  | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus devriesei | 0.66 | 1 | 0
|  |  | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 0.43 | 1 | 0
|  |  | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus epidermidis | 1.07 | 1 | 0
|  |  | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus equorum | 0.29 | 1 | 0
|  |  | DSM 20674 T | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus equorum | 0.84 | 1 | 0
|  |  | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus felis | 0.72 | 1 | 0
|  |  | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus felis | 0.41 | 1 | 0
|  |  | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.08.2011 H9 | Staphylococcus haemolyticus | 0.7 | 1 | 0
|  |  | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus haemolyticus | 0.65 | 1 | 0
|  |  | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N6 [0034] | Staphylococcus hominis | 0.83 | 1 | 0
|  |  | CVUAS 789,2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus hominis | 0.71 | 1 | 0
|  |  | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus hyicus | 1.22 | 1 | 0
|  |  | CVUAS 3767 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus hyicus | 0.95 | 1 | 0
|  |  | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, N8, [0034] | Staphylococcus intermedius | 0.76 | 1 | 0
|  |  | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N7, [0034] | Staphylococcus intermedius | 0.75 | 1 | 0
|  |  | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 05.04.2012, N5, [0034] | Staphylococcus lugdunensis | 0.96 | 1 | 0
|  |  | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, H16 | Staphylococcus lugdunensis | 0.43 | 1 | 0
|  |  | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus microti | 0.57 | 1 | 0
|  |  | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus pasteuri | 0.5 | 1 | 0
|  |  | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus pasteuri | 0.98 | 1 | 0
|  |  | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.03.2016, H05, [0034] | Staphylococcus pseudintermedius | 0.94 | 1 | 0
|  |  | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus pseudintermedius | 0.7 | 1 | 0
|  |  | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus saprophyticus | 0.02 | 1 | 0
|  |  | CVUAS 503,2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus saprophyticus | 0.71 | 1 | 0
|  |  | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus schleiferi | 0.69 | 1 | 0
|  |  | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus schweitzeri | 0.57 | 1 | 0
|  |  | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus simulans | 0.55 | 1 | 0
|  |  | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus simulans | 0.78 | 1 | 0
|  |  | CVUAS 2439,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 0.91 | 1 | 0
|  |  | CVUAS 761,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 1.09 | 1 | 0
|  |  | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus warneri | 0.41 | 1 | 0
|  |  | DSM 20316 | 3; type strain | DT | 2013-10-02 | Staphylococcus warneri | 0.53 | 1 | 0
|  |  | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.24 | 1 | 0
|  |  | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.72 | 1 | 0
|  |  | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus sp-185-U1-2 | 0.8 | 1 | 0
|  |  | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus suis | 0.52 | 1 | 0
|  |  | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus urinaeequi | 0.11 | 1 | 0
|  |  | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 27.09.2016 | Aerococcus viridans | 0.6 | 1 | 0
|  |  | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter phocae | 0.71 | 1 | 0
|  |  | CVUAS 2442,2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium divergens | 0.31 | 1 | 0
|  |  | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium divergens | 0.9 | 1 | 0
|  |  | CVUAS 6273 | 3; sequenced | eDT | 13.01.2016 | Carnobacterium maltaromaticum | 0.44 | 1 | 0
|  |  | CVUAS 1519,3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella adiacens | 0.44 | 1 | 0
|  |  | CVUAS 2616 | 3; sequenced | DT | 25.09.2013 | Marinilactibacillus sp-CVUAS-2616 | 0.84 | 1 | 0
|  |  | CVUAS 8980,3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus sp-CVUAS-2616 | 0.52 | 1 | 0
|  |  | 16 | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011 H20 | Enterococcus casseliflavus | 0.32 | 1 | 0
|  |  | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H12 | Enterococcus durans | 0.63 | 1 | 0
|  |  | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H13 | Enterococcus durans | 0.47 | 1 | 0
|  |  | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.07.2011 N8 | Enterococcus faecium | 0.72 | 1 | 0
|  |  | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N13 | Enterococcus gallinaru | 0.44 | 1 | 0
|  |  | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N3 | Enterococcus hirae | 0.5 | 1 | 0
|  |  | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus casseliflavus | 0.78 | 1 | 0
|  |  | CVUAS 5030 | 3; sequenced | eDT | 23.07.2014 | Enterococcus cecorum | 0.67 | 1 | 0
|  |  | CVUAS 8897 | 3; sequenced | eDT | 23.07.2014 | Enterococcus cecorum | 0.09 | 1 | 0
|  |  | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus faecalis | 0.39 | 1 | 0
|  |  | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus faecalis | 0.65 | 1 | 0
|  |  | MR11-207 LVU04 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus faecium | 0.82 | 1 | 0
|  |  | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus gallinarum | 0.76 | 1 | 0
|  |  | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus solitarius | 0.67 | 1 | 0
|  |  | CVUAS 1549,2 | 3; sequenced | DT | 2016-09-27 | Vagococcus salmoninarum | 0.1 | 1 | 0
|  |  | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus acidophilus | 0.9 | 1 | 0
|  |  | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus johnsonii | 0.78 | 1 | 0
|  |  | CVUAS 32277,2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus sp-CVUAS-32277,2 | 0.92 | 1 | 0
|  |  | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus claussenii | 1.06 | 1 | 0
|  |  | LTH 1213 | 2; 0 | DT | 2015-07-09 | Leuconostoc mesenteroides | 0.76 | 1 | 0
|  |  | LTH 1146 | 2; 0 | DT | 2015-07-09 | Lactococcus lactis | 0.37 | 1 | 0
|  |  | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus agalactiae | 0.68 | 1 | 0
|  |  | DSM 6784 | 3; public strain collection | eDT | 2015-03-11 | Streptococcus agalactiae | 0.37 | 1 | 0
|  |  | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus canis | 0.77 | 1 | 0
|  |  | LHL Mi-845_17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus canis | 0.15 | 1 | 0
|  |  | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus castoreus | 0.86 | 1 | 0
|  |  | DSM 17536 | 3; type strain | EFEx | 2018-05-03 | Streptococcus castoreus | 0.86 | 1 | 0
|  |  | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus catagoni | 0.55 | 1 | 0
|  |  | 99-1-2017 | 3; typestrain | EFEx | 2018-05-03 | Streptococcus catagoni | 0.23 | 1 | 0
|  |  | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus dysgalactiae | 0.83 | 1 | 0
|  |  | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0.56 | 1 | 0
|  |  | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus dysgalactiae | 1.04 | 1 | 0
|  |  | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0.73 | 1 | 0
|  |  | CVUAS 32462,2 | 3; specific PCR | eDT | 2022-05-30 | Streptococcus equi | 0.56 | 1 | 0
|  |  | CVUAS 32770 | 3; specific PCR | eDT | 2022-05-30 | Streptococcus equi | 0.5 | 1 | 0
|  |  | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus equi | 0.56 | 1 | 0
|  |  | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus equi | 0.75 | 1 | 0
|  |  | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus iniae | 0.93 | 1 | 0
|  |  | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus orisasini | 0.62 | 1 | 0
|  |  | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus parauberis | 0.52 | 1 | 0
|  |  | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus phocae | 0.78 | 1 | 0
|  |  | 12-13862-001-01 LVU11 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus pluranimalium | 0.48 | 1 | 0
|  |  | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.61 | 1 | 0
|  |  | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.47 | 1 | 0
|  |  | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.49 | 1 | 0
|  |  | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus pneumoniae | 0.17 | 1 | 0
|  |  | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus porcinus | 0.11 | 1 | 0
|  |  | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.04.2008, K11, [0034] | Streptococcus pyogenes | 0.61 | 1 | 0
|  |  | CVUAS 10145,3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus salivarius | 0.41 | 1 | 0
|  |  | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus sp-CVUAS-32349 | 0.1 | 1 | 0
|  |  | CVUAS 2260,4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus troglodytidis | 1.13 | 1 | 0
|  |  | CVUAS 5218,4 | 3; sequenced; 16S; NRZ Strep Uni AA | DT | 2015-06-11 | Streptococcus troglodytidis | 0.79 | 1 | 0
|  |  | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0.67 | 1 | 0
|  |  | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0.1 | 1 | 0
|  |  | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.61 | 1 | 0
|  |  | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.49 | 1 | 0
|  |  | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.23 | 1 | 0
|  |  | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.3 | 1 | 0
|  |  | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.29 | 1 | 0
|  |  | C bot B L677 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.27 | 1 | 0
|  |  | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.26 | 1 | 0
|  |  | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.35 | 1 | 0
|  |  | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 1.02 | 1 | 0
|  |  | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium botulinum | 1.1 | 1 | 0
|  |  | CVUAS 32971 | 3; sequenced | EFEx | 2021-02-25 | Clostridium botulinum | 0.68 | 1 | 0
|  |  | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium chauvoei | 0.6 | 1 | 0
|  |  | CVUAS 849,2 | 3; sequenced | DT | 2014-12-17 | Clostridium isatidis | 0.07 | 1 | 0
|  |  | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium nigeriense | 0.45 | 1 | 0
|  |  | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium perfringens | 0.65 | 1 | 0
|  |  | B | 3; case rerort | DT | 2013-08-06 | Clostridium perfringens | 0.57 | 1 | 0
|  |  | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium septicum | 0.67 | 1 | 0
|  |  | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium sp-DSM-1985 | 0.56 | 1 | 0
|  |  | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium sporogenes | 0.62 | 1 | 0
|  |  | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium tertium | 0.56 | 1 | 0
|  |  | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium tetani | 0.42 | 1 | 0
|  |  | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix anatis | 0.78 | 1 | 0
|  |  | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix anatis | 0.71 | 1 | 0
|  |  | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix aquatica | 0.56 | 1 | 0
|  |  | VA-92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix aquatica | 0.55 | 1 | 0
|  |  | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix larvae | 0.5 | 1 | 0
|  |  | DSM 110099 | 3; type strain | DT | 2021-10-28 | Erysipelothrix piscisicarius | 0.54 | 1 | 0
|  |  | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix rhusiopathiae | 0.31 | 1 | 0
|  |  | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix rhusiopathiae | 0.3 | 1 | 0
|  |  | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix tonsillarum | 0.54 | 1 | 0
|  |  | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-27 | Erysipelothrix tonsillarum | 0.39 | 1 | 0
|  |  | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus sp-CVUAS-10169 | 0.39 | 1 | 0
|  |  | CVUAS 32948 | 2; 0 | EFEx | 2021-02-18 | Finegoldia magna | 0.64 | 1 | 0
|  |  | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus kunzii | 0.39 | 1 | 0
|  |  | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus kunzii | 0.46 | 1 | 0
|  |  | CVUAS 433,2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus ovis | 0.49 | 1 | 0
|  |  | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus ovis | 0.43 | 1 | 0
|  |  | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus sp-CVUAS-32321 | 0.26 | 1 | 0
|  |  | CVUAS 692,2 | 3; sequenced | DT | 2013-02-19 | Parvimonas sp-CVUAS-692,2 | 0.48 | 1 | 0
|  |  | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides finegoldii | 0.65 | 1 | 0
|  |  | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides fragilis | 0.63 | 1 | 0
|  |  | 181001478 | 2; 0 | DT | 2018-05-02 | Bacteroides pyogenes | 0.57 | 1 | 0
|  |  | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella heparinolytica | 0.86 | 1 | 0
|  |  | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium gleum | 0.18 | 1 | 0
|  |  | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium indologenes | 0.49 | 1 | 0
|  |  | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium indologenes | 0.35 | 1 | 0
|  |  | CVUAS 32947 | 2; 0 | EFEx | 2021-02-18 | Bacteroides fragilis | 0.27 | 1 | 0
|  |  | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium hibernum | 0.76 | 1 | 0
|  |  | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium sp | 0.54 | 1 | 0
|  |  | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium chaponense | 0.12 | 1 | 0
|  |  | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium hominis | 0.4 | 1 | 0
|  |  | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium manosquense | 0.15 | 1 | 0
|  |  | 185-U1-6 | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium manosquense | 0.57 | 1 | 0
|  |  | CVUAS 11450,4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium sp-I-7 | 0.34 | 1 | 0
|  |  | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium aquidurense | 0.66 | 1 | 0
|  |  | CVUAS 5322,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium branchiarum | 0.27 | 1 | 0
|  |  | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium chungangense | 0.69 | 1 | 0
|  |  | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium columnare | 0.92 | 1 | 0
|  |  | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium frigidimaris | 0.42 | 1 | 0
|  |  | CVUAS 4217,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium frigidimaris | 0.29 | 1 | 0
|  |  | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium piscis | 0.65 | 1 | 0
|  |  | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium plurextorum | 0.61 | 1 | 0
|  |  | CVUAS 5323,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium psychoterrae | 0.72 | 1 | 0
|  |  | CVUAS 11550,4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium psychrophilum | 1.05 | 1 | 0
|  |  | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium psychrophilum | 0.8 | 1 | 0
|  |  | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium reichenbachii | 0.88 | 1 | 0
|  |  | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium sp-C2 | 0.71 | 1 | 0
|  |  | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-10323 | 0.69 | 1 | 0
|  |  | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium sp-CVUAS-4146 | 0.1 | 1 | 0
|  |  | CVUAS 4216,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-4216,2 | 0.22 | 1 | 0
|  |  | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-5506 | 1.02 | 1 | 0
|  |  | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium sp-HMD2009 | 0.62 | 1 | 0
|  |  | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium sp-JQMS | 0.3 | 1 | 0
|  |  | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium succinicans | 0.14 | 1 | 0
|  |  | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium tiangeerense | 0.09 | 1 | 0
|  |  | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella columbina | 0.1 | 1 | 0
|  |  | CVUAS 2589,2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas tenax | 0.49 | 1 | 0
|  |  | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium rhinotracheale | 0.79 | 1 | 0
|  |  | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium rhinotracheale | 0.1 | 1 | 0
|  |  | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus sp-CVUAS-11266 | 0.25 | 1 | 0
|  |  | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium necrophorum | 0.56 | 1 | 0
|  |  | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium varium | 0.76 | 1 | 0
|  |  | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus moniliformis | 0.52 | 1 | 0
|  |  | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter abscessus | 0.36 | 1 | 0
|  |  | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga salmonicida | 0.61 | 1 | 0
|  |  | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus hongkongensis | 0.57 | 1 | 0
|  |  | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella termitidis | 0.27 | 1 | 0
|  |  | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella termitidis | 0.49 | 1 | 0
|  |  | CCUG 38322 | 3; public strain collection | EFEx | 16.12.2015 | Sneathia sanguinegens | 0.01 | 1 | 0
|  |  | CCUG 41628 | 3; type strain | EFEx | 2020-10-08 | Sneathia sanguinegens | 0.1 | 1 | 0
|  |  | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.54 | 1 | 0
|  |  | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.33 | 1 | 0
|  |  | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus felis | 0.2 | 1 | 0
|  |  | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus moniliformis | 0.59 | 1 | 0
|  |  | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus notomytis | 0.44 | 1 | 0
|  |  | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus ratti | 0.52 | 1 | 0
|  |  | 171012806 | 2; 0 | DT | 2017-10-04 | Ochrobactrum intermedium | 0.21 | 1 | 0
|  |  | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum intermedium | 0.1 | 1 | 0
|  |  | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum sp | 0.23 | 1 | 0
|  |  | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum sp | 0.1 | 1 | 0
|  |  | 171012040 | 2; 0 | DT | 2017-09-14 | Ochrobactrum sp | 0.17 | 1 | 0
|  |  | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia coli | 0.64 | 1 | 0
|  |  | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella oxytoca | 0.5 | 1 | 0
|  |  | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella variicola | 1.14 | 1 | 0
|  |  | CVUAS 1816,2 | 3; sequenced | DT | 2013-07-24 | Basfia succiniciproducens | 0.56 | 1 | 0
|  |  | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas alcaligenes | 0.38 | 1 | 0
|  |  | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas anguilliseptica | 0.1 | 1 | 0
|  |  | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas brenneri | 0.83 | 1 | 0
|  |  | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas fulva | 0.33 | 1 | 0
|  |  | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas koreensis | 0.35 | 1 | 0
|  |  | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas monteilii | 0.67 | 1 | 0
|  |  | 171011722-1_1 | 2; 0 | DT | 2017-08-31 | Pseudomonas nitroreducens | 0.51 | 1 | 0
|  |  | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas oryzihabitans | 0.72 | 1 | 0
|  |  | 171011722-1_2 | 2; 0 | DT | 2017-08-31 | Pseudomonas putida | 0.52 | 1 | 0
|  |  | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas taetrolens | 0.29 | 1 | 0
|  |  | CVUAS 2280,3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas glaciistagni | 0.35 | 1 | 0
|  |  | A-269_08 A | 2; 0 | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 02.04.2012 K8 | Ochrobactrum anthropi | 0.73 | 1 | 0
|  |  | (BM373) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; D02 | Brucella abortus | 0.38 | 1 | 0
|  |  | (BM369) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; E01 | Brucella abortus | 0.59 | 1 | 0
|  |  | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; I16 | Brucella canis | 0.76 | 1 | 0
|  |  | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.03.2014; E17 | Brucella canis | 0.65 | 1 | 0
|  |  | (BM389) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; F04 | Brucella ceti | 0.33 | 1 | 0
|  |  | (BM403) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; K04 | Brucella ceti | 0.38 | 1 | 0
|  |  | (BM374) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I02 | Brucella melitensis | 0.5 | 1 | 0
|  |  | (BM407) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; J05 | Brucella melitensis | 0.51 | 1 | 0
|  |  | (BM408) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; H06 | Brucella melitensis | 0.51 | 1 | 0
|  |  | (BM410) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I06 | Brucella melitensis | 0.24 | 1 | 0
|  |  | CCM 4915 | 3; type strain | DT | 2013-11-26 | Brucella microti | 0.18 | 1 | 0
|  |  | CCM 4915 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.03.2015 C7 | Brucella microti | 0.97 | 1 | 0
|  |  | (BM421) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; F07 | Brucella neotomae | 0.9 | 1 | 0
|  |  | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.03.2015; C15 | Brucella neotomae | 0.71 | 1 | 0
|  |  | (BM423) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; L07 | Brucella ovis | 0.61 | 1 | 0
|  |  | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; F16 | Brucella ovis | 0.94 | 1 | 0
|  |  | (BM405) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; E05 | Brucella pinnipedialis | 0.41 | 1 | 0
|  |  | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; C03 | Brucella pinnipedialis | 0.32 | 1 | 0
|  |  | (BM431) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; B10 | Brucella suis | 0.7 | 1 | 0
|  |  | (BM427) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; G08 | Brucella suis | 0.04 | 1 | 0
|  |  | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella suis | 0.45 | 1 | 0
|  |  | (BM428) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I08 | Brucella suis | 0.62 | 1 | 0
|  |  | (BM430) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; I09 | Brucella suis | 0.52 | 1 | 0
|  |  | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 15.15.2016 I7 | Brucella vulpis | 0.67 | 1 | 0
|  |  | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum ovis | 0.42 | 1 | 0
|  |  | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 03.06.2016 H12 | Ochrobactrum anthropi | 0.37 | 1 | 0
|  |  | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter aganoensis | 0.21 | 1 | 0
|  |  | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter aminovorans | 0.48 | 1 | 0
|  |  | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2015 H22 | Paracoccus yeei | 0.54 | 1 | 0
|  |  | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas mucosae | 0.71 | 1 | 0
|  |  | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 24.01.2007, H7, [0034] | Inquilinus limosus | 0.68 | 1 | 0
|  |  | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella asinigenitalis | 0.72 | 1 | 0
|  |  | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter denitrificans | 0.72 | 1 | 0
|  |  | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter sp-DSM-30031 | 0.63 | 1 | 0
|  |  | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2008 E5 | Achromobacter xylosoxidans | 0.45 | 1 | 0
|  |  | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.02.2016 N7 | Achromobacter xylosoxidans | 0.44 | 1 | 0
|  |  | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea sp-CVUAS-31213 | 0.47 | 1 | 0
|  |  | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella avium | 0.47 | 1 | 0
|  |  | CVUAS 32946 | 2; 0 | EFEx | 2021-02-18 | Kerstersia gyiorum | 0.65 | 1 | 0
|  |  | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.64 | 1 | 0
|  |  | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.4 | 1 | 0
|  |  | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga kullae | 0.49 | 1 | 0
|  |  | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella asinigenitalis | 0.26 | 1 | 0
|  |  | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella equigenitalis | 0.05 | 1 | 0
|  |  | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella equigenitalis | 0.1 | 1 | 0
|  |  | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012; K10 | Burkholderia pseudomallei | 0.67 | 1 | 0
|  |  | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; H21 | Burkholderia pseudomallei | 0.55 | 1 | 0
|  |  | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.04.2012, K6 | Burkholderia thailandensis | 0.46 | 1 | 0
|  |  | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007, K7 | Burkholderia ambifaria | 0.69 | 1 | 0
|  |  | DSM 16087 LMG 19182 | 3; type starin | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K19 | Burkholderia ambifaria | 0.78 | 1 | 0
|  |  | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K21 | Burkholderia anthina | 0.75 | 1 | 0
|  |  | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.08.2007, N10 | Burkholderia anthina | 0.63 | 1 | 0
|  |  | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2010, K12 | Burkholderia arboris | 0.92 | 1 | 0
|  |  | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 3.4.2017; L6 | Burkholderia arboris | 0.57 | 1 | 0
|  |  | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K11 | Burkholderia cenocepacia | 0.82 | 1 | 0
|  |  | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.07.2009; K21 | Burkholderia cenocepacia | 0.74 | 1 | 0
|  |  | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 4.5.2006; H15 | Burkholderia cepacia | 1.06 | 1 | 0
|  |  | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 9.5.2007; N9 | Burkholderia cepacia | 1.01 | 1 | 0
|  |  | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.11.2006; H10 | Burkholderia contaminans | 0.62 | 1 | 0
|  |  | BC 01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E10 | Burkholderia contaminans | 0.92 | 1 | 0
|  |  | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016 C13 | Burkholderia diffusa | 0.42 | 1 | 0
|  |  | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L17 | Burkholderia diffusa | 0.67 | 1 | 0
|  |  | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K16 | Burkholderia dolosa | 0.65 | 1 | 0
|  |  | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.11.2006; E09 | Burkholderia dolosa | 0.9 | 1 | 0
|  |  | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.01.2007; H22 | Burkholderia gladioli | 0.57 | 1 | 0
|  |  | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H15 | Burkholderia gladioli | 0.69 | 1 | 0
|  |  | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K03 | Burkholderia glumae | 0.96 | 1 | 0
|  |  | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.05.2012; K22 | Burkholderia glumae | 0.48 | 1 | 0
|  |  | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016; C10 | Burkholderia lata | 0.85 | 1 | 0
|  |  | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.08.2016; K19 | Burkholderia lata | 0.84 | 1 | 0
|  |  | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.05.2016; L06 | Burkholderia latens | 0.77 | 1 | 0
|  |  | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.09.2009; K20 | Burkholderia mallei | 0.68 | 1 | 0
|  |  | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.08.2007; H13 | Burkholderia mallei | 0.76 | 1 | 0
|  |  | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2010; H14 | Burkholderia metallica | 0.92 | 1 | 0
|  |  | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L21 | Burkholderia metallica | 0.66 | 1 | 0
|  |  | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.12.2006; K19 | Burkholderia multivorans | 0.49 | 1 | 0
|  |  | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.11.2006; K04 | Burkholderia multivorans | 0.75 | 1 | 0
|  |  | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K17 | Burkholderia oklahomensis | 0.28 | 1 | 0
|  |  | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K18 | Burkholderia oklahomensis | 0.48 | 1 | 0
|  |  | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K06 | Burkholderia plantarii | 0.78 | 1 | 0
|  |  | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K05 | Burkholderia plantarii | 0.86 | 1 | 0
|  |  | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K07 | Burkholderia pseudomultivorans | 0.74 | 1 | 0
|  |  | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K08 | Burkholderia pseudomultivorans | 0.7 | 1 | 0
|  |  | CAMPA 1043 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; K04 | Burkholderia puraquae | 0.67 | 1 | 0
|  |  | R-50214 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.03.2016; C04 | Burkholderia puraquae | 0.87 | 1 | 0
|  |  | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; N06 | Burkholderia pyrrocinia | 1.05 | 1 | 0
|  |  | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; I21 | Burkholderia pyrrocinia | 0.1 | 1 | 0
|  |  | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H08 | Burkholderia seminalis | 1.02 | 1 | 0
|  |  | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.01.2007; K12 | Burkholderia seminalis | 0.77 | 1 | 0
|  |  | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K13 | Burkholderia stabilis | 0.68 | 1 | 0
|  |  | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E08 | Burkholderia stabilis | 0.72 | 1 | 0
|  |  | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; C08 | Burkholderia stagnalis | 0.41 | 1 | 0
|  |  | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F14 | Burkholderia stagnalis | 0.26 | 1 | 0
|  |  | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007; H19 | Burkholderia taxonK | 0.83 | 1 | 0
|  |  | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.06.2009; K13 | Burkholderia taxonK | 0.77 | 1 | 0
|  |  | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F16 | Burkholderia territorii | 0.6 | 1 | 0
|  |  | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; E14 | Burkholderia territorii | 0.52 | 1 | 0
|  |  | DSM 13276 n | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 24.11.2010, K19 | Burkholderia thailandensis | 0.51 | 1 | 0
|  |  | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H11 | Burkholderia ubonensis | 0.53 | 1 | 0
|  |  | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; I04 | Burkholderia ubonensis | 0.14 | 1 | 0
|  |  | CCUG 9631=LMG 6999 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.05.2007 N15 | Burkholderia vietnamiensis | 0.68 | 1 | 0
|  |  | LMG 10929 V | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.12.2006 N9 | Burkholderia vietnamiensis | 0.8 | 1 | 0
|  |  | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.04.2012, N07 | Caballeronia glathei | 0.23 | 1 | 0
|  |  | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H12 | Caballeronia glathei | 0.02 | 1 | 0
|  |  | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H15 | Pandoraea apista | 0.13 | 1 | 0
|  |  | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H21 | Pandoraea pnomenusa | 0.47 | 1 | 0
|  |  | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H16 | Pandoraea pulmonicola | 0.47 | 1 | 0
|  |  | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K7 | Paraburkholderia caledonica | 0.47 | 1 | 0
|  |  | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K8 | Paraburkholderia caledonica | 0.74 | 1 | 0
|  |  | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N9 | Paraburkholderia caribensis | 0.62 | 1 | 0
|  |  | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N10 | Paraburkholderia caribensis | 0.57 | 1 | 0
|  |  | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H03 | Paraburkholderia hospita | 0.85 | 1 | 0
|  |  | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H04 | Paraburkholderia hospita | 0.85 | 1 | 0
|  |  | DSM 13464 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 1.12.2010; H15 | Paraburkholderia kururiensis | 0.71 | 1 | 0
|  |  | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.11.2010; N19 | Paraburkholderia phymatum | 0.91 | 1 | 0
|  |  | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.11.2010; K17 | Paraburkholderia phytofirmans | 0.46 | 1 | 0
|  |  | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H16 | Paraburkholderia tuberum | 0.73 | 1 | 0
|  |  | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.11.2010, N21 | Paraburkholderia xenovorans | 0.67 | 1 | 0
|  |  | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 E07 | Ralstonia insidiosa | 0.1 | 1 | 0
|  |  | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 C06 | Ralstonia mannitolilytica | 0.42 | 1 | 0
|  |  | ATCC 27511 DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 22.01.2007 H03 | Ralstonia pickettii | 0.51 | 1 | 0
|  |  | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 07.05.2009 N15 | Ralstonia pickettii | 0.37 | 1 | 0
|  |  | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia acidovorans | 0.46 | 1 | 0
|  |  | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia tsuruhatensis | 0.56 | 1 | 0
|  |  | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga intermedia | 0.43 | 1 | 0
|  |  | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.07.2013 L13 | Xenophilus sp | 0.47 | 1 | 0
|  |  | DSM 24713 extr | 3; type strain | Extr | 2016-02-02 | Deefgea chitinilytica | 0.29 | 1 | 0
|  |  | DSM 24713 | 3; type strain | eDT | 2015-09-22 | Deefgea chitinilytica | 0.24 | 1 | 0
|  |  | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea rivuli | 0.58 | 1 | 0
|  |  | 08-F17.2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.48 | 1 | 0
|  |  | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.02 | 1 | 0
|  |  | CVUAS 6995,3 | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea sp-CVUAS-6995,3 | 0.7 | 1 | 0
|  |  | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-6995,3 | 0.47 | 1 | 0
|  |  | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-9920 | 0.27 | 1 | 0
|  |  | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea sp-CVUAS-9920 | 0.28 | 1 | 0
|  |  | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter limnosediminis | 0.32 | 1 | 0
|  |  | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella kingae | 0.58 | 1 | 0
|  |  | CVUAS 10067,2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria animaloris | 0.28 | 1 | 0
|  |  | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria animaloris | 0.07 | 1 | 0
|  |  | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-4931 | 1.01 | 1 | 0
|  |  | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-8997 | 0.52 | 1 | 0
|  |  | CVUAS 8040 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-LHL-1296 | 0.42 | 1 | 0
|  |  | CVUAS 5699,3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria sp-N95-16 | 0.92 | 1 | 0
|  |  | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae sp-CVUAS-8636 | 0.25 | 1 | 0
|  |  | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter fetus | 0.33 | 1 | 0
|  |  | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter fetus | 0.69 | 1 | 0
|  |  | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter fetus | 0.9 | 1 | 0
|  |  | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter hyointestinalis | 0.69 | 1 | 0
|  |  | BfR-CA-07048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter peloridis | 0.63 | 1 | 0
|  |  | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter peloridis | 0.78 | 1 | 0
|  |  | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter cryaerophilus | 0.62 | 1 | 0
|  |  | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter butzleri | 0.51 | 1 | 0
|  |  | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter butzleri | 0.71 | 1 | 0
|  |  | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter coli | 0.49 | 1 | 0
|  |  | DSM 4689 | 3; type strain | eDT | 2015-12-30 | Campylobacter coli | 0.62 | 1 | 0
|  |  | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter jejuni | 0.17 | 1 | 0
|  |  | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter jejuni | 0.28 | 1 | 0
|  |  | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter jejuni | 0.28 | 1 | 0
|  |  | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter lari | 0.52 | 1 | 0
|  |  | BfR-CA-06556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.7 | 1 | 0
|  |  | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.54 | 1 | 0
|  |  | DSM 11375 | 3; type strain | eDT | 2015-12-30 | Campylobacter lari | 0.65 | 1 | 0
|  |  | DSM 5361 | 3; type strain | DT | 2013-09-24 | Campylobacter fetus | 0.94 | 1 | 0
|  |  | CVUAS 6673,2 | 3; sequenced | DT | 2013-08-28 | Campylobacter hyointestinalis | 0.52 | 1 | 0
|  |  | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter sputorum | 0.9 | 1 | 0
|  |  | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter sp-W1 | 0.77 | 1 | 0
|  |  | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter upsaliensis | 0.17 | 1 | 0
|  |  | CVUAS 11185,3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter canadensis | 0.96 | 1 | 0
|  |  | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas bestiarum | 0.92 | 1 | 0
|  |  | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas bivalvium | 0.07 | 1 | 0
|  |  | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas caviae | 0.45 | 1 | 0
|  |  | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.44 | 1 | 0
|  |  | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.54 | 1 | 0
|  |  | DSM 30187 RKI | 4; type strain, WGS sequenced (BfR) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E17 | Aeromonas hydrophila | 0.04 | 1 | 0
|  |  | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas hydrophila | 0.23 | 1 | 0
|  |  | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas jandaei | 0.24 | 1 | 0
|  |  | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas media | 0.06 | 1 | 0
|  |  | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas molluscorum | 0.46 | 1 | 0
|  |  | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas popoffii | 0.32 | 1 | 0
|  |  | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas salmonicida | 0.59 | 1 | 0
|  |  | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas salmonicida | 0.73 | 1 | 0
|  |  | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas salmonicida | 0.24 | 1 | 0
|  |  | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas salmonicida | 0.29 | 1 | 0
|  |  | CVUAS 8400,2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas salmonicida | 0.27 | 1 | 0
|  |  | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas salmonicida | 0.24 | 1 | 0
|  |  | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas salmonicida | 0.56 | 1 | 0
|  |  | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas salmonicida | 0.05 | 1 | 0
|  |  | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas simiae | 0.4 | 1 | 0
|  |  | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas veronii | 0.47 | 1 | 0
|  |  | DSM 7386 | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas veronii | 0.56 | 1 | 0
|  |  | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas pectinilytica | 0.48 | 1 | 0
|  |  | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella viscosa | 0.22 | 1 | 0
|  |  | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella viscosa | 0.1 | 1 | 0
|  |  | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.25 | 1 | 0
|  |  | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.54 | 1 | 0
|  |  | CVUAS 81,2 | 3; sequenced | EFEx | 28:11.2013 | Shewanella sp-RKVR22 | 0.25 | 1 | 0
|  |  | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter condimenti | 1.15 | 1 | 0
|  |  | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter dublinensis | 0.29 | 1 | 0
|  |  | DSM 18702 | 3; type strain | DT | 2019-03-05 | Cronobacter malonaticus | 1.13 | 1 | 0
|  |  | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter turicensis | 0.95 | 1 | 0
|  |  | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.11.2011 E6 | Enterobacter cloacae | 0.51 | 1 | 0
|  |  | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, H5 | Citrobacter amalonaticus | 0.68 | 1 | 0
|  |  | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter freundii | 0.34 | 1 | 0
|  |  | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter freundii | 0.09 | 1 | 0
|  |  | ATCC 25408 | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 11.04.2008, N11, [0034]: Citrobacter diversus ATCC 25408\LMTG70N11_0\0_N11\1\1SLin\fid | Citrobacter koseri | 0.41 | 1 | 0
|  |  | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-15 | Cronobacter malonaticus | 0.36 | 1 | 0
|  |  | CVUAS 32945 | 2; 0 | EFEx | 2021-02-18 | Cronobacter sakazakii | 0.53 | 1 | 0
|  |  | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter sakazakii | 0.89 | 1 | 0
|  |  | DSM 30054 2 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E18, [0034] | Enterobacter cloacae | 0.37 | 1 | 0
|  |  | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K4, [0034] | Enterobacter sp-ATCC-27988 | 0.7 | 1 | 0
|  |  | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae sp-CVUAS-30072 | 0.28 | 1 | 0
|  |  | CVUAS 0500,2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia coli | 0.51 | 1 | 0
|  |  | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella grimontii | 1.01 | 1 | 0
|  |  | CVUAS 2486,3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella grimontii | 0.68 | 1 | 0
|  |  | CVUAS 11699 | 3; sequenced [BfR] | eDT | 2017-04-20 06:56 | Klebsiella michiganensis | 0.89 | 1 | 0
|  |  | CVUAS 2404,3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella michiganensis | 0.41 | 1 | 0
|  |  | CVUAS 10038,2 | 3; sequenced [BfR] | EFex | 2019-02-12 14:02 | Klebsiella oxytoca | 0.75 | 1 | 0
|  |  | CVUAS 11572 | 3; sequenced [BfR] | eDT | 2017-04-20 06:51 | Klebsiella oxytoca | 0.73 | 1 | 0
|  |  | CVUAS 1482,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:35 | Klebsiella pneumoniae | 0.58 | 1 | 0
|  |  | CVUAS 1616,4 | 3; sequenced [BfR] | eDT | 2017-11-22 09:55 | Klebsiella pneumoniae | 0.5 | 1 | 0
|  |  | CVUAS 1678 | 3; sequenced [BfR] | eDT | 2016-10-18 06:38 | Klebsiella pneumoniae | 0.67 | 1 | 0
|  |  | CVUAS 1701 | 3; sequenced [BfR] | eDT | 2016-10-18 06:41 | Klebsiella pneumoniae | 0.51 | 1 | 0
|  |  | CVUAS 30499 | 3; sequenced [BfR] | eDT | 2017-11-22 09:59 | Klebsiella pneumoniae | 0.76 | 1 | 0
|  |  | CVUAS 753,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:34 | Klebsiella pneumoniae | 0.74 | 1 | 0
|  |  | CVUAS 10109 | 3; sequenced [BfR] | eDT | 2016-10-18 07:00 | Klebsiella quasipneumoniae | 0.8 | 1 | 0
|  |  | CVUAS 11677 | 3; sequenced [BfR] | eDT | 2017-04-20 06:54 | Klebsiella variicola | 0.72 | 1 | 0
|  |  | CVUAS 4888,2 | 3; sequenced [BfR] | EFex | 2017-07-19 11:36 | Klebsiella variicola | 0.81 | 1 | 0
|  |  | DSM 30053 | 3; type strain | DT | 2013-09-17 | Klebsiella aerogenes | 0.1 | 1 | 0
|  |  | NM 20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E20 | Klebsiella aerogenes | 0.37 | 1 | 0
|  |  | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella michiganensis | 0.35 | 1 | 0
|  |  | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.11.2013 E18 | Klebsiella oxytoca | 0.67 | 1 | 0
|  |  | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 11.04.2008 N21 | Klebsiella pneumoniae | 0.84 | 1 | 0
|  |  | DSM 30104 | 3; type strain | EFEx | 2013-09-17 | Klebsiella pneumoniae | 0.78 | 1 | 0
|  |  | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.1 | 1 | 0
|  |  | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.74 | 1 | 0
|  |  | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella variicola | 0.7 | 1 | 0
|  |  | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia sacchari | 0.65 | 1 | 0
|  |  | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.67 | 1 | 0
|  |  | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.59 | 1 | 0
|  |  | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N17, [0034] (Enterobacter gergoviae) | Pluralibacter gergoviae | 0.16 | 1 | 0
|  |  | LT21 Ib 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H9 | Salmonella enterica | 0.72 | 1 | 0
|  |  | II 06 06820 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H6 | Shigella boydii | 0.65 | 1 | 0
|  |  | 3A 07 02503 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H7 | Shigella flexneri | 0.54 | 1 | 0
|  |  | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea agglomerans | 0.59 | 1 | 0
|  |  | DSM 3493 | 3; type strain | eDT | 2015-07-09 | Pantoea agglomerans | 0.47 | 1 | 0
|  |  | CVUAS 4263,2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea sp-FF5 | 0.68 | 1 | 0
|  |  | CVUAS 8773 | 2; sequenced | DT | 2013-03-20 | Pantoea | 0.6 | 1 | 0
|  |  | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella tarda | 0.96 | 1 | 0
|  |  | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N13, [0034] | Edwardsiella tarda | 0.81 | 1 | 0
|  |  | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 18.04.2008, K9, [0034] | Hafnia alvei | 0.51 | 1 | 0
|  |  | LHL1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia alvei | 0.62 | 1 | 0
|  |  | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea myxofaciens | 0.07 | 1 | 0
|  |  | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella morganii | 0.37 | 1 | 0
|  |  | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella morganii | 0.24 | 1 | 0
|  |  | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus cibarius | 0.63 | 1 | 0
|  |  | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus hauseri | 0.79 | 1 | 0
|  |  | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus hauseri | 0.74 | 1 | 0
|  |  | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus mirabilis | 0.64 | 1 | 0
|  |  | DSM 4479 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus mirabilis | 0.82 | 1 | 0
|  |  | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus penneri | 0.34 | 1 | 0
|  |  | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.25 | 1 | 0
|  |  | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.03 | 1 | 0
|  |  | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.31 | 1 | 0
|  |  | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus terrae | 0.31 | 1 | 0
|  |  | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, E5 K16, [0034]: Proteus vulgaris ATCC 33420\LMTG72K16_0\0_K16\1\1SLin\fid | Proteus vulgaris | 0.56 | 1 | 0
|  |  | DSM 2140 | 3; strain from public collection; sequenced (rpoB) | eDT | 30.11.2017 | Proteus vulgaris | 0.63 | 1 | 0
|  |  | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia rettgeri | 0.47 | 1 | 0
|  |  | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia rettgeri | 0.66 | 1 | 0
|  |  | ATCC 25827 | 3; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K15, [0034]: Proteus stuartii ATCC 25827\LMTG72K15_0\0_K15\1\1SLin\fid in vers 2016 | Providencia stuartii | 0.69 | 1 | 0
|  |  | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia frederiksenii | 0.68 | 1 | 0
|  |  | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia mollaretii | 0.56 | 1 | 0
|  |  | NCTC 2028 | 2; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H13 | Yersinia pestis | 0.8 | 1 | 0
|  |  | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H14 | Yersinia pestis | 0.8 | 1 | 0
|  |  | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia wautersii | 0.57 | 1 | 0
|  |  | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia fonticola | 0.26 | 1 | 0
|  |  | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.11.2008, E11, [0034] | Serratia grimesii | 0.63 | 1 | 0
|  |  | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia marcescens | 0.21 | 1 | 0
|  |  | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia marcescens | 0.79 | 1 | 0
|  |  | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia quinivorans | 0.26 | 1 | 0
|  |  | CVUAS 6221 | 4; sequenced | DT | 2013-01-08 | Yersinia aldovae | 0.66 | 1 | 0
|  |  | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 09.03.2007, H12 | Yersinia aleksiciae | 0.48 | 1 | 0
|  |  | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia bercovieri | 0.6 | 1 | 0
|  |  | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia bercovieri | 0.51 | 1 | 0
|  |  | SL2012 001153 | 3; sequenced | DT | 2013-01-30 | Yersinia bercovieri | 0.64 | 1 | 0
|  |  | DSM 4780 | 3; type strain | DT | 2013-05-08 | Yersinia enterocolitica | 0.78 | 1 | 0
|  |  | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.61 | 1 | 0
|  |  | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia enterocolitica | 0.54 | 1 | 0
|  |  | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.58 | 1 | 0
|  |  | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia enterocolitica | 0.49 | 1 | 0
|  |  | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia enterocolitica | 0.42 | 1 | 0
|  |  | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.ma_01\0_C13, [0034] | Yersinia entomophaga | 1.24 | 1 | 0
|  |  | 21 | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 22.01.2007 H06 | Yersinia frederiksenii | 0.47 | 1 | 0
|  |  | RKI 21 | 3; well described strain | DT | 2013-01-30 | Yersinia frederiksenii | 0.66 | 1 | 0
|  |  | CVUAS 1824 | 2; sequenced | DT | 2012-12-05 | Yersinia intermedia | 0.65 | 1 | 0
|  |  | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia intermedia | 0.83 | 1 | 0
|  |  | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia kristensenii | 0.51 | 1 | 0
|  |  | CVUAS 302 | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia kristensenii | 0.41 | 1 | 0
|  |  | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 25.11.2015, Y.ma_01\0_C11, [0034] | Yersinia massiliensis | 0.8 | 1 | 0
|  |  | VB 925699 | 3; sequenced | DT | 2013-03-19 | Yersinia massiliensis | 0.67 | 1 | 0
|  |  | DSM 18520 | 3; type strain | DT | 16.01.2013 | Yersinia mollaretii | 0.29 | 1 | 0
|  |  | RKI 51 | 2; 0 | DT | 2013-01-16 | Yersinia mollaretii | 0.84 | 1 | 0
|  |  | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.nu_04\0_E10, [0034] | Yersinia nurmii | 0.22 | 1 | 0
|  |  | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.pe_04\0_E8, [0034] | Yersinia pekkanenii | 0.24 | 1 | 0
|  |  | CVUAS 0304,2 | 2; biochemistry | DT | 2013-01-31 | Yersinia pseudotuberculosis | 0.57 | 1 | 0
|  |  | DSM 8992 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.04.2009, E12 | Yersinia pseudotuberculosis | 0.78 | 1 | 0
|  |  | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia rohdei | 0.52 | 1 | 0
|  |  | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 02.03.2007, E07 | Yersinia rohdei | 0.59 | 1 | 0
|  |  | CVUAS 2772 | 3; sequenced | DT | 2013-01-30 | Yersinia ruckeri | 0.56 | 1 | 0
|  |  | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia ruckeri | 0.27 | 1 | 0
|  |  | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia similis | 0.16 | 1 | 0
|  |  | IMB 4354 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 13.07.2007, E22 | Yersinia similis | 0.75 | 1 | 0
|  |  | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.wa_04\0_E5, [0034] | Yersinia wautersii | 0.37 | 1 | 0
|  |  | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella londiniensis | 0.2 | 1 | 0
|  |  | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella pneumophila | 0.64 | 1 | 0
|  |  | CVUAS 4122 | 2; 0 | eDT | 20.01.2016 | Legionella pneumophila | 0.67 | 1 | 0
|  |  | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella taurinensis | 0.79 | 1 | 0
|  |  | CVUAS 10873,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0.63 | 1 | 0
|  |  | CVUAS 10874,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0.78 | 1 | 0
|  |  | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) caballi | 0.63 | 1 | 0
|  |  | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) caballi | 0.68 | 1 | 0
|  |  | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) testudinis | 0.28 | 1 | 0
|  |  | LHL 957 | 3; sequenced | eDT | 02.12.2015 | (Pasteurella) testudinis | 0.11 | 1 | 0
|  |  | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus anseriformium | 0.51 | 1 | 0
|  |  | 15UGF 198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus anseriformium | 0.4 | 1 | 0
|  |  | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus arthritidis | 0.39 | 1 | 0
|  |  | 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus capsulatus | 0.38 | 1 | 0
|  |  | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus capsulatus | 0.48 | 1 | 0
|  |  | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.49 | 1 | 0
|  |  | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.57 | 1 | 0
|  |  | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.19 | 1 | 0
|  |  | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus equuli | 0.37 | 1 | 0
|  |  | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus lignieresii | 0.36 | 1 | 0
|  |  | CVUAS 9945,2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus lignieresii | 0.28 | 1 | 0
|  |  | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus rossii | 0.56 | 1 | 0
|  |  | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus rossii | 0.7 | 1 | 0
|  |  | CVUAS 32505,2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus sp-CVUAS-32505,2 | 0.23 | 1 | 0
|  |  | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-genomospecie- | 0.71 | 1 | 0
|  |  | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus sp-genomospecies-1 | 0.66 | 1 | 0
|  |  | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus sp-MK-2012 | 0.61 | 1 | 0
|  |  | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-MK-2012 | 0.81 | 1 | 0
|  |  | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus suis | 0.56 | 1 | 0
|  |  | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus suis | 0.69 | 1 | 0
|  |  | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium gallinarum | 0.71 | 1 | 0
|  |  | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium gallinarum | 0.32 | 1 | 0
|  |  | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 0.44 | 1 | 0
|  |  | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 0.87 | 1 | 0
|  |  | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter oris | 0.5 | 1 | 0
|  |  | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter oris | 0.39 | 1 | 0
|  |  | CVUAS 3802,2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter osteomyelitidis | 0.6 | 1 | 0
|  |  | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia canicola | 0.43 | 1 | 0
|  |  | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia canicola | 0.26 | 1 | 0
|  |  | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium anatis | 0.44 | 1 | 0
|  |  | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium anatis | 0.47 | 1 | 0
|  |  | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium sp-F151 | 0.25 | 1 | 0
|  |  | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium trehalosifermentans | 0.49 | 1 | 0
|  |  | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus haemoglobinophilus | 0.58 | 1 | 0
|  |  | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus sp-2269 | 0.65 | 1 | 0
|  |  | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus sp-2269 | 0.8 | 1 | 0
|  |  | P1066-2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus somni | 0.54 | 1 | 0
|  |  | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia caviae | 0.52 | 1 | 0
|  |  | CVUAS 1064,4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia glucosida | 0.21 | 1 | 0
|  |  | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia glucosida | 0.48 | 1 | 0
|  |  | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia granulomatis | 0.17 | 1 | 0
|  |  | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia granulomatis | 0.1 | 1 | 0
|  |  | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia haemolytica | 0.35 | 1 | 0
|  |  | CVUAS 6197,2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia haemolytica | 0.5 | 1 | 0
|  |  | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia pernigra | 0.07 | 1 | 0
|  |  | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia pernigra | 0.46 | 1 | 0
|  |  | 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia ruminalis | 0.1 | 1 | 0
|  |  | 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia ruminalis | 0.12 | 1 | 0
|  |  | CVUAS 10396,2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia sp-10-12559-1 | 0.81 | 1 | 0
|  |  | CVUAS 10183,3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia sp-14UCF195 | 0.48 | 1 | 0
|  |  | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia sp-14UCF195 | 0.09 | 1 | 0
|  |  | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia sp-15USF15 | 0.47 | 1 | 0
|  |  | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia sp-15USF15 | 0.5 | 1 | 0
|  |  | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia sp-CVUAS-11240 | 0.5 | 1 | 0
|  |  | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia sp-CVUAS-1496 | 0.34 | 1 | 0
|  |  | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter rosorum | 0.61 | 1 | 0
|  |  | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter rosorum | 0.75 | 1 | 0
|  |  | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella semolina | 0.18 | 1 | 0
|  |  | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella semolina | 0.57 | 1 | 0
|  |  | CVUAS 4060,3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella aerogenes | 0.53 | 1 | 0
|  |  | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella canis | 0.28 | 1 | 0
|  |  | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | Pasteurella mairii | 0.97 | 1 | 0
|  |  | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella multocida | 0.49 | 1 | 0
|  |  | CVUAS 32055 | 3; specific PCR | eDT | 2019-11-14 | Pasteurella multocida | 0.52 | 1 | 0
|  |  | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella multocida | 0.39 | 1 | 0
|  |  | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella multocida | 0.47 | 1 | 0
|  |  | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella multocida | 0.29 | 1 | 0
|  |  | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella multocida | 0.5 | 1 | 0
|  |  | CVUAS 32506,2 | 3; sequenced | EFEx | 2020-11-19 | Pasteurellaceae Bisgaard-Taxon-10 | 0.57 | 1 | 0
|  |  | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Pasteurellaceae Bisgaard-Taxon-10 | 0.3 | 1 | 0
|  |  | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae Bisgaard-Taxon-40 | 0.52 | 1 | 0
|  |  | CVUAS 9987,2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae canine-oral-taxon-080 | 0.41 | 1 | 0
|  |  | CVUAS 261,4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae sp-CVUAS-261,4 | 0.4 | 1 | 0
|  |  | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae sp-CVUAS-31988 | 0.1 | 1 | 0
|  |  | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae sp-CVUAS-4868 | 0.1 | 1 | 0
|  |  | CVUAS 724,3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae sp-CVUAS-724,3 | 0.36 | 1 | 0
|  |  | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae sp-CVUAS-9305 | 0.67 | 1 | 0
|  |  | CVUAS 9967,2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae sp-CVUAS-9967,2 | 0.14 | 1 | 0
|  |  | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae sp-LHL-U-309 | 0.5 | 1 | 0
|  |  | CVUAS 443,2 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-309 | 0.58 | 1 | 0
|  |  | CVUAS 444,2 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-3444 | 0.72 | 1 | 0
|  |  | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae sp-MCCM-2539 | 0.34 | 1 | 0
|  |  | CVUAS 308,2 | 3; sequenced | DT | 2013-11-06 | Psittacicella hinzii | 0.1 | 1 | 0
|  |  | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.65 | 1 | 0
|  |  | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.74 | 1 | 0
|  |  | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium mucosae | 0.63 | 1 | 0
|  |  | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium mucosae | 0.39 | 1 | 0
|  |  | CVUAS 8966 | 3; sequenced | DT | 23.07.2013 | Volucribacter amazonae | 0.74 | 1 | 0
|  |  | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E22 | Acinetobacter baumannii | 0.22 | 1 | 0
|  |  | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2014 C05 | Acinetobacter baumannii | 0.54 | 1 | 0
|  |  | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter defluvii | 0.1 | 1 | 0
|  |  | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter guillouiae | 0.12 | 1 | 0
|  |  | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter johnsonii | 0.58 | 1 | 0
|  |  | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter junii | 0.66 | 1 | 0
|  |  | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter lwoffii | 0.72 | 1 | 0
|  |  | CVUAS 7379,3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter pseudolwoffii | 0.32 | 1 | 0
|  |  | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter pseudolwoffii | 0.86 | 1 | 0
|  |  | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter variabilis | 0.51 | 1 | 0
|  |  | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 0.42 | 1 | 0
|  |  | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 0.1 | 1 | 0
|  |  | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella catarrhalis | 0.25 | 1 | 0
|  |  | CVUAS 11674,3 | 3; sequenced | eDT | 2022-03-22 | Moraxella cuniculi | 0.45 | 1 | 0
|  |  | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella osloensis | 0.84 | 1 | 0
|  |  | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella ovis | 0.63 | 1 | 0
|  |  | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter frigidicola | 0.42 | 1 | 0
|  |  | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter pasteurii | 0.5 | 1 | 0
|  |  | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.3 | 1 | 0
|  |  | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.49 | 1 | 0
|  |  | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas (alkylphenolica) | 0.71 | 1 | 0
|  |  | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas aeruginosa | 0.09 | 1 | 0
|  |  | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007 H8 | Pseudomonas aeruginosa | 0.73 | 1 | 0
|  |  | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas baetica | 0.85 | 1 | 0
|  |  | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.02.2008 N11 | Pseudomonas chlororaphis | 0.46 | 1 | 0
|  |  | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 02.04.2012, H21 | Francisella tularensis | 0.77 | 1 | 0
|  |  | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 12.12.2016, F19, [0096] | Allofrancisella guangzhouensis | 0.46 | 1 | 0
|  |  | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 18.11.2015, L16, [0096] | Francisella hispaniensis | 1.01 | 1 | 0
|  |  | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 25.11.2015, I3, [0096] | Francisella naotunensis | 0.52 | 1 | 0
|  |  | CVUAS 4680,2 | 3; FLI: specific PCR | DT | 2013-04-02 | Francisella tularensis | 0.1 | 1 | 0
|  |  | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 23.04.2013, F20, [0096] | Francisella tularensis | 0.4 | 1 | 0
|  |  | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 26.04.2013, I20, [0096] | Francisella tularensis | 0.2 | 1 | 0
|  |  | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 01.06.2017, F3, [0096] | Francisella tularensis | 0.58 | 1 | 0
|  |  | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio aestuarianus | 0.68 | 1 | 0
|  |  | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio anguillarum | 0.39 | 1 | 0
|  |  | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio campbelli | 0.63 | 1 | 0
|  |  | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio campbelli | 0.66 | 1 | 0
|  |  | 16-VB00067-0_BfR 1_eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio cholerae | 0.69 | 1 | 0
|  |  | 10-VBH0202 DT-48h | 4; WGS sequenced | DT | 2021-02-02 | Vibrio cincinnatiensis | 0.58 | 1 | 0
|  |  | BfR-VB-00004 LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio cincinnatiensis | 0.67 | 1 | 0
|  |  | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio diazotrophicus | 0.6 | 1 | 0
|  |  | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio fluvialis | 0.62 | 1 | 0
|  |  | BfR-VB-00036 LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio fluvialis | 0.88 | 1 | 0
|  |  | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 16.08.2012, K12, [0034] | Vibrio harveyi | 0.82 | 1 | 0
|  |  | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio harveyi | 0.3 | 1 | 0
|  |  | 20-VB00047 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio jasicida | 0.92 | 1 | 0
|  |  | 20-VB00053 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio jasicida | 0.75 | 1 | 0
|  |  | 12-VBH0498 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio metoecus | 0.62 | 1 | 0
|  |  | 20-VB00016 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio metoecus | 0.83 | 1 | 0
|  |  | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio metschnikovii | 0.28 | 1 | 0
|  |  | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio navarrensis | 0.6 | 1 | 0
|  |  | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio owensii | 0.88 | 1 | 0
|  |  | 17-VB00254-0_BfR 18 | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio parahaemolyticus | 0.39 | 1 | 0
|  |  | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio vulnificus | 0.1 | 1 | 0
|  |  | CVUAS 33180 | 3; sequenced | DT | 2021-08-04 | Photobacterium damselae | 0.51 | 1 | 0
|  |  | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium damselae | 0.33 | 1 | 0
|  |  | CVUAS 1266,2 | 3; sequenced | EFEx | 25.09.1019 | Photobacterium piscicola | 0.32 | 1 | 0
|  |  | CVUAS 842,2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium swingsii | 0.53 | 1 | 0
|  |  | CVUAS 7079 | 3; spezific PCR (CVFR) | eDT | 2013-07-17 | Vibrio alginolyticus | 0.6 | 1 | 0
|  |  | DSM 2171 | 3; type strain | DT | 2016-10-12 | Vibrio alginolyticus | 0.27 | 1 | 0
|  |  | A8-12 | 3; gyrB sequence | eDT | 2016-06-16 | Vibrio anguillarum | 0.1 | 1 | 0
|  |  | CVUAS 10227 | 3; sequenced | eDT | 2022-06-23 | Vibrio cholerae | 0.49 | 1 | 0
|  |  | CVUAS 8995 | 3; sequenced | eDT | 2013-04-10 | Vibrio diazotrophicus | 0.5 | 1 | 0
|  |  | DSM 14383 | 3; public strain collection, sequenced | eDT | 2013-04-16 | Vibrio furnissi | 0.41 | 1 | 0
|  |  | CVUAS 1754 | 3; sequenced | eDT | 2013-04-16 | Vibrio furnissii | 0.68 | 1 | 0
|  |  | CVUAS 619 | 2; 0 | eDT | 2013-08-14 | Vibrio metschnikovii | 0.55 | 1 | 0
|  |  | CVUAS 7872 | 3; sequenced | eDT | 2013-04-10 | Vibrio mimicus | 0.59 | 1 | 0
|  |  | DSM 19130 | 4; type strain; WGS sequenced | DT | 2021-02-24 | Vibrio mimicus | 0.39 | 1 | 0
|  |  | DSM 10027 | 3; type strain | EFEx | 2013-05-22 | Vibrio parahaemolyticus | 0.62 | 1 | 0
|  |  | CVUAS 871,2 | 3; sequenced | eDT | 2014-08-21 | Vibrio sp-CVUAS-871,2 | 0.27 | 1 | 0
|  |  | CVUAS 2778,2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio tubiashii | 0.74 | 1 | 0
|  |  | CVUAS 2778,2 | 3; sequenced | eDT | 2019-10-24 | Vibrio tubiashii | 0.74 | 1 | 0
|  |  | LVM05 LVU05 | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-15 | Vibrio vulnificus | 0.26 | 1 | 0
|  |  | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas sp-G9077 | 0.96 | 1 | 0
|  |  | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria indica | 0.2 | 1 | 0
|  |  | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria ureiclastica | 0.64 | 1 | 0
|  |  | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas chelatiphaga | 0.98 | 1 | 0
|  |  | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2009 K17 | Stenotrophomonas maltophilia | 0.46 | 1 | 0
|  |  | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.05.2009 K22 | Stenotrophomonas maltophilia | 0.26 | 1 | 0
|  |  | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas rhizophila | 0.73 | 1 | 0
|  |  | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.05.2015 C17 | Stenotrophomonas rhizophila | 0.1 | 1 | 0
|  |  | CVUAS 4239,2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas tumulicola | 1.09 | 1 | 0

## Report (Species)
Targeted Identification
following the Guidelines for Validating Species Identifications using MALDI-TOF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022
Parameter: Bacillus cytotoxicus
Version: | 3
generated by | Dyk | checked by | Rau
generated on | 2023-04-05 | checked on | 2023-04-12
Application | isolates from food & veterinary diagnostics
Environment | Food microbiology, veterinary diagnostics
Applied database / software version |  | MBT(K) / BRU 11897 BT 4.0 + MUP 1157 | BT 4.0
Validation set
Number of target parameter isolates: 183 |  | number of non-target parameter isolates: 1408
cultivation conditions: SBA, 30°/37° C, 24h, aerob/anaerob/mikroaerophil
preparation acc. to manufacturer instructions: DT, eDT, EFex; RKI-specific preparation TFA-method
Only identified samples are used for determination of the inclusivity and exclusivity.
True-positive: 1. Hit = targeted parameter shows classifcation score ≥ 2 and score 2. Hit < 1,7 or score 2. Hit ≥ 1,7 & genus is the same as 1. Hit or score 2. Hit ≥ 2 & species is the same as 1. Hit |  | True-negative: 1. Hit without assignment to the target parameter with score ≥ 2
False-negative: true-positive criteria are not fulfilled |  | False-positive: true-negative criteria are not fullfilled
Validation result | Target parameter: | Bacillus cytotoxicus |  | target | fullfilled
Total number of spectra | 183
of which with ID result | 182 | rate of identified spectra | 99.45% | 20 | ja | 1
Total number of own spectra | 93 |  |  | 36.4 | ja | 1
With respect to identified spectra: |  | Inclusivity
positive accordance (PA) | 182 | True positive rate | 100% | 95% | ja | 1
negative deviation (ND) | 0 | False negative rate | 0% | 1% | ja | 1
Nicht-Ziel-Parameter
Total number of spectra | 1408
of which with ID result | 916 | rate of identified spectra | 65.06% | 30 | ja | 1
Total number of own spectra | 619 |  |  | 183.2 | ja | 1
With respect to identified spectra: |  | Exclusivity
negative accordance (NA) | 916 | True negative rate | 100% | 99% | ja | 1
positive deviation (PD) | 0 | False-Positive Rate | 0% | 1% | ja | 1
Test for spectral proximity | 1408 | Proportion of closely related spectra | 0.2841%
Using the complete database, out of 182 identified target parameter spectra 182 (=100%) were identified correctly (inclusivity). 0 (=0%) of identified spectra of the parameter were assigned to a wrong species.
Using the complete database, out of 916 identified non-target parameter spectra 100% were counted correctly as negative (exclusivity). Out of these 916 spectra 0 (=0%) were allocated incorrectly to the parameter.
Using a database containing only the parameter, of all 1408 non-target parameter spectra for 4 (=0,3%) spectra a score > 2 was shown.
Criteria for the | targeted identification | are thus | fulfilled
comment:
The spectra of the RKI (Zenodo) are integrated into the validation in vers. 3 (as of 2018).
This validation is described in Rau et al., 2024, Food Control
Released for site | for department | on (date) | by (signature)
CVUA Stuttgart / Germany | T, MT, D | 2023-04-12 | Rau
The procedure is suitable for the intended purpose and is thereby released for usage at site.
Analytical parameter added to LIMS:
A signed version of this validation report is kept within the generating lab.
Other microbes that may be present in this matrix.
Other gems growing on the used agar
Non parameter isolates/materials that can provide similar spectra.
yes
no