## Settings
Settings for the Validation
Targeted Identification
following the Guidelines for Validating Species Identifications using MALDI-TOF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022
Institution / Labor | CVUA Stuttgart / Germany
Parameter | Genus |  | Species
|  | Corynebacterium | pseudotuberculosis
Differentiation limits
| Score for differentiation on genus level | 1.7 | Score for differentiation on species level | 2

## Target parameter (Species)
Parameter: | Corynebacterium | pseudotuberculosis | own spectrum | own ID | Isolate ID | Validity | preparation method (if given) | date | Genus | Species | Score | Genus | Species | Score | Consistence | identified | not assigned | correct | false negative | wrong species identified | second spectra of isolate: | 1
|  |  | 1 | 1 | VB997349 | 2; 0 | EFEx | 2024-08-28 | Corynebacterium | pseudotuberculosis | 2.14 | Corynebacterium | pseudotuberculosis | 1.94 | A | 1 | 0 | 1 | 0 | 0
Total number of spectra | 80 |  | 1 | 1 | 992 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.16 | Corynebacterium | pseudotuberculosis | 2.1 | A | 1 | 0 | 1 | 0 | 0
of which with ID result | 80 |  | 1 | 1 | 993 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.23 | A | 1 | 0 | 1 | 0 | 0
not identified | 0 |  | 1 | 1 | CVUAS 3337 | 4; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.31 | Corynebacterium | pseudotuberculosis | 2.29 | A | 1 | 0 | 1 | 0 | 0
with respect to identified spectra: |  |  | 1 | 1 | CVUAS 3338 | 4; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.36 | Corynebacterium | pseudotuberculosis | 2.31 | A | 1 | 0 | 1 | 0 | 0
positive accordance (PA) | 80 |  | 1 | 1 | CVUAS 3339 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.34 | Corynebacterium | pseudotuberculosis | 2.29 | A | 1 | 0 | 1 | 0 | 0
negative deviation (ND) | 0 |  | 1 | 1 | CVUAS 3354 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | pseudotuberculosis | 2.44 | Corynebacterium | pseudotuberculosis | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3357 | 4; sequenced | EFEx | 2018-07-10 | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3361 | 4; sequenced | DT | 2012-12-12 | Corynebacterium | pseudotuberculosis | 2.52 | Corynebacterium | pseudotuberculosis | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3362 | 3; sequenced | DT | 2013-01-09 | Corynebacterium | pseudotuberculosis | 2.32 | Corynebacterium | pseudotuberculosis | 2.25 | A | 1 | 0 | 1 | 0 | 0
score limit | 2 | Species | 1 | 1 | CVUAS 3364 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.27 | Corynebacterium | pseudotuberculosis | 2.26 | A | 1 | 0 | 1 | 0 | 0
settings | 1.7 | Genus | 1 | 1 | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0
The complete database to be checked is used. |  |  | 1 | 1 | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium | pseudotuberculosis | 2.48 | Corynebacterium | pseudotuberculosis | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL0269 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.08 | Corynebacterium | pseudotuberculosis | 2.06 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | DSM 20689 CVUAS | 3; type strain | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0 |  | 1
number of individual isolates | 79 |  | 0 | 0 | DSM 20689 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-05-27 | Corynebacterium | pseudotuberculosis | 2.55 | Corynebacterium | pseudotuberculosis | 2.49 | A | 1 | 0 | 1 | 0 | 0
mean score first hit | 2.359 |  | 1 | 1 | CVUAS 32656 | 4; WGS sequenced | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.39 | A | 1 | 0 | 1 | 0 | 0
Standard derivation of mean score first hit | 0.1163 |  | 1 | 1 | CVUAS 32689 | 4; WGS sequenced | eDT | 2020-10-15 | Corynebacterium | pseudotuberculosis | 2.38 | Corynebacterium | pseudotuberculosis | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32746 | 4; WGS sequenced | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.45 | Corynebacterium | pseudotuberculosis | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32834,2 | 4; WGS sequenced | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.4 | Corynebacterium | pseudotuberculosis | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32842 | 4; WGS sequenced | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.47 | Corynebacterium | pseudotuberculosis | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 33314 | 3; WGS sequenced | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.37 | A | 1 | 0 | 1 | 0 | 0
of which with ID result | 80 |  | 1 | 1 | CVUAS 33501 | 3; WGS sequenced | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.35 | Corynebacterium | pseudotuberculosis | 2.27 | A | 1 | 0 | 1 | 0 | 0
positive accordance (PA) | 80 |  | 1 | 1 | CVUAS 34900 | 4; WGS sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.3 | Corynebacterium | pseudotuberculosis | 2.28 | A | 1 | 0 | 1 | 0 | 0
negative deviation (ND) | 0 |  | 1 | 1 | CVUAS 34905 | 4; WGS sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.34 | Corynebacterium | pseudotuberculosis | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 36894 | 4; WGS sequenced | EFEx | 2024-08-28 | Corynebacterium | pseudotuberculosis | 2.32 | Corynebacterium | pseudotuberculosis | 2.26 | A | 1 | 0 | 1 | 0 | 0
not identified in % | 0% |  | 1 | 1 | CVUAS 4258,2 | 4; WGS sequenced | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.4 | Corynebacterium | pseudotuberculosis | 2.3 | A | 1 | 0 | 1 | 0 | 0
mismatched in % | 0% |  | 1 | 1 | CVUAS 5583,2 | 4; WGS sequenced | eDT | 2015-08-11 | Corynebacterium | pseudotuberculosis | 2.37 | Corynebacterium | pseudotuberculosis | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | IHIT 39383 | 2; 0 | EFEx | 2024-10-09 | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.41 | A | 1 | 0 | 1 | 0 | 0
thereof |  |  | 1 | 1 | IHIT 39989 | 2; 0 | EFEx | 2024-10-09 | Corynebacterium | pseudotuberculosis | 2.46 | Corynebacterium | pseudotuberculosis | 2.31 | A | 1 | 0 | 1 | 0 | 0
own spectra (CVUAS) | 74 |  | 1 | 1 | KL0293 | 4; WGS; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.22 | Corynebacterium | pseudotuberculosis | 2.12 | A | 1 | 0 | 1 | 0 | 0
own spectra [%] | 92.5% |  | 1 | 1 | KL0302 | 4; WGS; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.18 | Corynebacterium | pseudotuberculosis | 2.14 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL0359 | 4; WGS; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.34 | Corynebacterium | pseudotuberculosis | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2997 | 4; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.46 | Corynebacterium | pseudotuberculosis | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32354 | 3; WGS sequenced | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4988 | 4; WGS sequenced | DT | 2012-10-08 | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | IHIT 38290 | 2; 0 | EFEx | 2024-10-09 | Corynebacterium | pseudotuberculosis | 2.28 | Corynebacterium | pseudotuberculosis | 2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | IHIT 41887 | 2; 0 | EFEx | 2024-10-09 | Corynebacterium | pseudotuberculosis | 2.37 | Corynebacterium | pseudotuberculosis | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | IHIT 42954 | 2; 0 | EFEx | 2024-10-09 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | IHIT 45389 | 2; 0 | EFEx | 2024-10-15 | Corynebacterium | pseudotuberculosis | 2.25 | Corynebacterium | pseudotuberculosis | 1.92 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | IHIT 54308 | 2; 0 | EFEx | 2024-10-09 | Corynebacterium | pseudotuberculosis | 2.37 | Corynebacterium | pseudotuberculosis | 2.05 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | 23UOF7 | 3; sequenced | eDT | 2023-02-03 | Corynebacterium | pseudotuberculosis | 2.27 | Corynebacterium | pseudotuberculosis | 1.9 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 31248,2 | 3; sequenced | eDT | 2018-11-09 | Corynebacterium | pseudotuberculosis | 2.49 | Corynebacterium | pseudotuberculosis | 2.48 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32166 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3221 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.29 | Corynebacterium | pseudotuberculosis | 2.25 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32632 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.42 | Corynebacterium | pseudotuberculosis | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32654 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.38 | Corynebacterium | pseudotuberculosis | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32745 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.36 | Corynebacterium | pseudotuberculosis | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32773 | 3; sequenced (CVUAS) | eDT | 2020-11-17 | Corynebacterium | pseudotuberculosis | 2.31 | Corynebacterium | pseudotuberculosis | 2.06 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32774 | 3; sequenced (CVUAS) | eDT | 2020-11-18 | Corynebacterium | pseudotuberculosis | 2.38 | Corynebacterium | pseudotuberculosis | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32830 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.5 | Corynebacterium | pseudotuberculosis | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3353 | 3; sequenced | EFEx | 2018-07-10 | Corynebacterium | pseudotuberculosis | 2.45 | Corynebacterium | pseudotuberculosis | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3355 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.4 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3365 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.35 | Corynebacterium | pseudotuberculosis | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3366 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.09 | Corynebacterium | pseudotuberculosis | 2.09 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34899 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34901 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.47 | Corynebacterium | pseudotuberculosis | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34902 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.37 | Corynebacterium | pseudotuberculosis | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34903 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.37 | Corynebacterium | pseudotuberculosis | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34904 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.35 | Corynebacterium | pseudotuberculosis | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34906 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34907 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.13 | Corynebacterium | pseudotuberculosis | 2.1 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34908 | 3; sequenced | eDT | 2023-02-17 | Corynebacterium | pseudotuberculosis | 2.11 | Corynebacterium | pseudotuberculosis | 2.06 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4418 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.38 | Corynebacterium | pseudotuberculosis | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4621,2 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.22 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4987 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | 2.23 | Corynebacterium | pseudotuberculosis | 2.21 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5188 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | 2.43 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 9749,2 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | 2.28 | Corynebacterium | pseudotuberculosis | 2.25 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL0160 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.13 | Corynebacterium | pseudotuberculosis | 2.04 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL0177 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.08 | Corynebacterium | pseudotuberculosis | 2.03 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | KL0184 | 3; KL Diphtheriae, LGL | eDT | 2016-04-12 | Corynebacterium | pseudotuberculosis | 2.24 | Corynebacterium | pseudotuberculosis | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 10301,2 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 8472,2 | 3; 0 | eDT | 2022-05-22 | Corynebacterium | pseudotuberculosis | 2.4 | Corynebacterium | pseudotuberculosis | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | KL0262 | 4; WGS sequenced (LGL) | 0 | 2011-05-06 | Corynebacterium | pseudotuberculosis | 2.6 | Corynebacterium | pseudotuberculosis | 2.51 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | KL0263 | 4; WGS sequenced (LGL) | 0 | 2011-05-06 | Corynebacterium | pseudotuberculosis | 2.59 | Corynebacterium | pseudotuberculosis | 2.52 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | KL0264 | 4; WGS sequenced (LGL) | 0 | 2011-05-09 | Corynebacterium | pseudotuberculosis | 2.51 | Corynebacterium | pseudotuberculosis | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | KL0265 | 4; WGS sequenced (LGL) | 0 | 2011-05-09 | Corynebacterium | pseudotuberculosis | 2.57 | Corynebacterium | pseudotuberculosis | 2.56 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | KL0266 | 4; WGS sequenced (LGL) | 0 | 2011-05-10 | Corynebacterium | pseudotuberculosis | 2.42 | Corynebacterium | pseudotuberculosis | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | LALLF 21TRD1263 | 3; sequenced | EFEx | 2022-05-09 | Corynebacterium | pseudotuberculosis | 2.17 | Corynebacterium | pseudotuberculosis | 2.17 | A | 1 | 0 | 1 | 0 | 0

## Non-target parameter (Species)
Parameter: | Corynebacterium pseudotuberculosis |  | own spectrum | own ID | Isolate ID | Validity | preparation method (if given) | date | Genus | Species | Genus | Species | Score | Genus | Species | Score | Consistence | identified | correct ID | not assigned on sepcies level | 1. Hit | 2. Hit | 1. Hit is met with the parameter | 2. Hit is met with the parameter | wrong identified as parameter | parameter in the first or the second hit
|  |  | 1 | 0 | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.3 | Corynebacterium | diphtheriae | 2.26 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
total number of spectra | 1529 |  | 1 | 0 | KL0171 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.16 | Corynebacterium | diphtheriae | 2.13 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
of which with ID result | 723 |  | 1 | 0 | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.28 | Corynebacterium | diphtheriae | 2.28 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
with respect to identified spectra: |  |  | 0 | 0 | KL2649 | 4; WGS sequenced (LGL) | 0 | 2023-04-24 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.05 | Corynebacterium | diphtheriae | 2.01 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
negative accordance (NA) (correct identified not as the parameter) | 723 |  | 0 | 0 | KL2656 | 4; WGS sequenced (LGL) | 0 | 2023-04-27 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.28 | Corynebacterium | diphtheriae | 2.24 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
wrong identified as parameter (PD) | 0 |  | 0 | 0 | KL3178 | 4; WGS sequenced (LGL) | 0 | 2024-02-07 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.12 | Corynebacterium | diphtheriae | 2.11 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL3290 | 4; WGS sequenced (LGL) | 0 | 2024-05-07 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.15 | Corynebacterium | diphtheriae | 2.13 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 44123 CVUAS | 3; type strain | eDT | 2024-01-26 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.53 | Corynebacterium | diphtheriae | 2.45 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 44123 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-05-27 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.39 | Corynebacterium | diphtheriae | 2.39 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.37 | Corynebacterium | diphtheriae | 2.36 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
Parameter is assigned as: |  |  | 1 | 1 | KL0167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.45 | Corynebacterium | diphtheriae | 2.37 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0187 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.34 | Corynebacterium | diphtheriae | 2.32 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
1. Hit is met with the parameter | 0 |  | 1 | 1 | KL0232 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.45 | Corynebacterium | diphtheriae | 2.32 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
2. Hit is met with the parameter | 0 |  | 1 | 1 | KL0237 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.48 | Corynebacterium | diphtheriae | 2.47 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
PM is among the first 2 hits | 0 |  | 1 | 1 | KL0240 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-13 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.52 | Corynebacterium | diphtheriae | 2.32 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
in % | 0% |  | 0 | 0 | KL0932 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.51 | Corynebacterium | diphtheriae | 2.45 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0933 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.38 | Corynebacterium | diphtheriae | 2.38 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0935 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.4 | Corynebacterium | diphtheriae | 2.36 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
score limit | 2 | Species | 0 | 0 | KL0936 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.34 | Corynebacterium | diphtheriae | 2.3 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
settings | 1.7 | Genus | 0 | 0 | KL0950 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.44 | Corynebacterium | diphtheriae | 2.38 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
The complete database to be checked is used. |  |  | 0 | 0 | KL0952 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.37 | Corynebacterium | diphtheriae | 2.32 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.5 | Corynebacterium | diphtheriae | 2.41 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
of which with ID result | 723 |  | 1 | 1 | KL0179 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.41 | Corynebacterium | diphtheriae | 2.36 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
positive accordance (PA) | 723 |  | 1 | 1 | KL0235 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.51 | Corynebacterium | diphtheriae | 2.41 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
negative deviation (ND) | 0 |  | 0 | 0 | KL0940 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.53 | Corynebacterium | diphtheriae | 2.45 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0954 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.25 | Corynebacterium | diphtheriae | 2.19 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
not identified in % | 52.71% |  | 0 | 0 | KL0956 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.04 | Corynebacterium | diphtheriae | 2 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
mismatched in % | 0% |  | 1 | 1 | ATCC 13812 | 3; public strain collection | eDT | 2024-09-26 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.26 | Corynebacterium | diphtheriae | 2.24 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 43988 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-05-27 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.25 | Corynebacterium | diphtheriae | 2.21 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
own spectra (CVUAS) | 891 |  | 0 | 0 | KL3381 | 4; WGS sequenced (LGL) | 0 | 2024-07-17 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.48 | Corynebacterium | diphtheriae | 2.44 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
own spectra [%] | 53.12% |  | 0 | 0 | KL3391 | 4; WGS sequenced (LGL) | 0 | 2024-07-22 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.43 | Corynebacterium | diphtheriae | 2.4 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL3445 | 4; WGS sequenced (LGL) | 0 | 2024-08-28 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.4 | Corynebacterium | diphtheriae | 2.32 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
own spectra identified | 352 |  | 0 | 0 | KL3446 | 4; WGS sequenced (LGL) | 0 | 2024-08-28 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.39 | Corynebacterium | diphtheriae | 2.33 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
own spectra identified [%] | 55.87% |  | 0 | 0 | KL3482 | 4; WGS sequenced (LGL) | 0 | 2024-09-20 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.55 | Corynebacterium | diphtheriae | 2.54 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL3485 | 4; WGS sequenced (LGL) | 0 | 2024-09-24 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.49 | Corynebacterium | diphtheriae | 2.45 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
external spectra (CVUAS) | 660 |  | 0 | 0 | KL3488 | 4; WGS sequenced (LGL) | 0 | 2024-09-25 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.36 | Corynebacterium | diphtheriae | 2.33 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
external spectra [%] | 46.88% |  | 0 | 0 | KL3506 | 4; WGS sequenced (LGL) | 0 | 2024-10-10 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.47 | Corynebacterium | diphtheriae | 2.39 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB2024000664 | 3; KL-LGL | EFEx | 2024-09-24 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.43 | Corynebacterium | diphtheriae | 2.34 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C86 | 4; sequenced (rpoB) UGent; NGS (CVUAS) | eDT | 2020-06-18 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.34 | Corynebacterium | ulcerans | 2.2 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 2823 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.31 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34267 | 4; KL Diphteria; NGS | eDT | 2022-09-02 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.23 | Corynebacterium | ulcerans | 2.13 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C104 | 4; sequenced (rpoB) UGent; NGS (CVUAS) | DT | 2020-06-18 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.33 | Corynebacterium | ulcerans | 2.22 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C139 | 4; sequenced (rpoB) UGent; NGS (CVUAS) | eDT | 2020-06-18 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.25 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C167 | 4; sequenced (rpoB) UGent; NGS (CVUAS) | eDT | 2020-06-18 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.25 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 112226 eDT | 3; type strain | eDT | 2024-08-13 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.35 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 112226 EFex | 3; type strain | EFEx | 2024-08-13 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.37 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL3299A | 4; WGS sequenced (LGL) | 0 | 2024-05-15 | Corynebacterium | ramonii | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.28 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 211002017-1 | 3; 0 | DT | 2021-02-11 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.12 | Corynebacterium | diphtheriae | 2.01 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 211002017-2 | 3; 0 | DT | 2021-02-11 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2 | Corynebacterium | diphtheriae | 1.99 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.09 | Corynebacterium | diphtheriae | 2.01 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C138 | 4; NGS (CVUAS) | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.32 | Corynebacterium | diphtheriae | 2.29 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3559,2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.16 | Corynebacterium | diphtheriae | 2.15 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.24 | Corynebacterium | diphtheriae | 2.07 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | KL1306 | 3; sequenced / variant upgraded to species level Coynebacterium rouxii [0137] here: NGS data (LGA) C. rouxi | EFEx | 2020-10-20 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.04 | Corynebacterium | diphtheriae | 1.98 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | KL1663 | 3; sequenced / variant upgraded to species level Coynebacterium rouxii [0137] here: NGS data (LGA) C. rouxi | EFEx | 2020-12-15 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 1.98 | Corynebacterium | diphtheriae | 1.97 | B | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141000461-2 | 1; 0 | 0 | 2014-01-20 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.06 | Corynebacterium | ulcerans | 1.89 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141016278 | 3; Confirmation by KL Diphtheriae (LGL); WBC: FT-IR | 0 | 2014-12-22 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2 | Corynebacterium | pseudotuberculosis | 1.84 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-1 | 1; 0 | 0 | 2016-03-15 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.05 | Corynebacterium | ulcerans | 1.79 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-2 | 1; 0 | 0 | 2016-03-15 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.88 | Corynebacterium | ulcerans | 1.79 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-3 | 3; Confirmation by KL Diphtheriae (LGL) | 0 | 2016-03-16 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.06 | Corynebacterium | pseudotuberculosis | 1.91 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-4 | 1; 0 | 0 | 2016-03-14 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.13 | Corynebacterium | ulcerans | 1.93 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-5 | 1; 0 | 0 | 2016-03-14 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 2.01 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161003120-6 | 1; 0 | 0 | 2016-03-14 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.04 | Corynebacterium | ulcerans | 1.95 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10-7-D-00025 | 3; sequenced | DT | 2013-01-08 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 1.97 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11USF28 | 3; sequenced | DT | 2012-12-11 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.1 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11USF53 | 3; sequenced | DT | 2013-01-16 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.09 | Corynebacterium | ulcerans | 1.94 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11USF78 | 3; sequenced | DT | 2012-12-11 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.11 | Corynebacterium | ulcerans | 2.03 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11USF79 | 3; sequenced | DT | 2012-12-11 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.1 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 121017479 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.03 | Corynebacterium | ulcerans | 1.84 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 131000349 | 3; sequenced | DT | 2013-01-18 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.22 | Corynebacterium | ulcerans | 2.02 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.28 | Corynebacterium | ulcerans | 1.94 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.04 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 057 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.18 | Corynebacterium | ulcerans | 2.08 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 076 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.97 | Corynebacterium | ulcerans | 1.83 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 085 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.18 | Corynebacterium | ulcerans | 1.94 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 091 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.16 | Corynebacterium | ulcerans | 2.14 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 098 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.08 | Corynebacterium | ulcerans | 2.02 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 111 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.11 | Corynebacterium | ulcerans | 2.02 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 123 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.16 | Corynebacterium | ulcerans | 2.08 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 128 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.21 | Corynebacterium | ulcerans | 2.06 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 130 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.13 | Corynebacterium | ulcerans | 2 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 139 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.23 | Corynebacterium | ulcerans | 2.06 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 141 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 1.88 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 145 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.13 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 147 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.03 | Corynebacterium | ulcerans | 1.97 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 149 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.06 | Corynebacterium | ulcerans | 1.77 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 156 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.1 | Corynebacterium | ulcerans | 1.88 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 166 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 168 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.13 | Corynebacterium | ulcerans | 1.92 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 182 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.08 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 183 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.2 | Corynebacterium | ulcerans | 1.99 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 188 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 2.03 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 233 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.04 | Corynebacterium | ulcerans | 2.01 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 257 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 2.02 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | AGES WS 332 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.22 | Corynebacterium | ulcerans | 2 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 344 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.07 | Corynebacterium | ulcerans | 1.95 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 351 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.04 | Corynebacterium | ulcerans | 1.88 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AGES WS 353 | 2; 0 | 0 | 2019-02-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.15 | Corynebacterium | ulcerans | 2.01 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BU125-97 | 3; sequenced | DT | 2012-10-19 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 2.01 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10179,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.14 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10701 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.07 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11363 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.86 | Corynebacterium | ulcerans | 1.83 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11488 | 3; FT-IR (NTTB-D) | eDT | 2017-01-11 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.94 | Corynebacterium | ulcerans | 1.88 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1400,3 | 3; FT-IR (NTTB-D) | eDT | 2018-03-02 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.01 | Corynebacterium | ulcerans | 1.92 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1717,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.66 | Corynebacterium | ulcerans | 1.66 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2040,2 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.1 | Corynebacterium | ulcerans | 1.97 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2738,2 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.09 | Corynebacterium | ulcerans | 1.97 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4292 | 3; type strain [0127] | DT | 2012-12-08 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.13 | Corynebacterium | ulcerans | 2.02 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5160 | 3; sequenced | DT | 2012-10-08 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.08 | Corynebacterium | ulcerans | 1.94 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5243,2 | 3; FT-IR (NTTB-D) | DT | 2013-11-18 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.96 | Corynebacterium | ulcerans | 1.88 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6455 | 3; sequenced, strain of the species desciption [0127] | DT | 2013-01-08 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.99 | Corynebacterium | ulcerans | 1.89 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6678,3 | 3; FT-IR (NTTB-D) | eDT | 2016-11-23 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.03 | Corynebacterium | ulcerans | 1.98 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 704,2 | 2; FT-IR (NTTB-D) | eDT | 2013-12-09 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2 | Corynebacterium | ulcerans | 1.99 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8972,3 | 3; FT-IR (NTTB-D) | eDT | 2015-12-02 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 2.03 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 97,2 | 2; FT-IR (NTTB-D) | eDT | 2013-11-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2 | Corynebacterium | ulcerans | 1.9 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1281 | 4; WGS sequenced (LGL) | 0 | 2018-12-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.29 | Corynebacterium | ulcerans | 2.2 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1848 | 4; WGS sequenced (LGL) | 0 | 2011-05-09 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.3 | Corynebacterium | ulcerans | 2.15 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-1014-16 | 3; FT-IR (NTTB-D) | eDT | 2017-01-11 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.07 | Corynebacterium | ulcerans | 2.06 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-1111-3-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-1408-1-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.07 | Corynebacterium | ulcerans | 1.99 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-1449-2-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 1.92 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-1514-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.1 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-1591-2-17 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.15 | Corynebacterium | ulcerans | 2 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-1627-5-12 | 3; sequenced | DT | 2012-10-08 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.05 | Corynebacterium | ulcerans | 2 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-241-6-18 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 1.9 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-241-7-18 | 3; FT-IR (NTTB-D) | eDT | 2018-03-15 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 2.02 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-28-3-13 | 3; sequenced | DT | 2013-01-18 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.01 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-308-2-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.21 | Corynebacterium | ulcerans | 2.05 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-309-1-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.18 | Corynebacterium | pseudotuberculosis | 1.91 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-310-2-18 | 3; FT-IR (NTTB-D) | EFEx | 2018-03-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 1.92 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-349-18 | 3; FT-IR (NTTB-D) | eDT | 2018-04-05 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.91 | Corynebacterium | ulcerans | 1.82 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-349-3-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.04 | Corynebacterium | ulcerans | 1.93 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-427-3-15 | 3; FT-IR (NTTB-D) | eDT | 2015-03-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2 | Corynebacterium | ulcerans | 1.94 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-434-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.98 | Corynebacterium | ulcerans | 1.87 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-443-17 | 3; FT-IR (NTTB-D) | eDT | 2017-09-27 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.09 | Corynebacterium | ulcerans | 2.01 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-450-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.12 | Corynebacterium | ulcerans | 2.02 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-468-1-18 | 3; FT-IR (NTTB-D) | eDT | 2018-07-04 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.06 | Corynebacterium | ulcerans | 2 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-619-1-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 1.98 | Corynebacterium | ulcerans | 1.94 | B | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-793-3-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.11 | Corynebacterium | ulcerans | 1.96 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-803-14 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.11 | Corynebacterium | ulcerans | 1.85 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S-884-18 | 3; FT-IR (NTTB-D) | eDT | 2018-08-21 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.21 | Corynebacterium | ulcerans | 2.05 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | W30-B05227 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.04 | Corynebacterium | ulcerans | 1.85 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | W67-B16966 | 3; FT-IR (NTTB-D) | eDT | 2018-07-09 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.02 | Corynebacterium | ulcerans | 1.9 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | C57 UGhent | 3; sequenced (rpoB) UGent | DT | 2020-06-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.36 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141005571-1 | 1; 0 | 0 | 2014-05-09 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.31 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141013039-2 | 1; 0 | 0 | 2014-05-09 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.34 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151004530 | 1; 0 | 0 | 2015-04-15 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.37 | Corynebacterium | ulcerans | 2.32 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151012433-2-2 | 1; 0 | 0 | 2015-09-28 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161007221 | 1; 0 | 0 | 2016-06-13 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.29 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131010012-1 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.33 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131010012-2 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.28 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131011719 | 3; sequenced | eDT | 2018-03-29 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.18 | Corynebacterium | ulcerans | 2.17 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131013415-1 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.2 | Corynebacterium | ulcerans | 2.14 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131013415-2 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.26 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131015432-10 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.34 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131015432-13 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.18 | Corynebacterium | ulcerans | 2.12 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 131015432-14 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.27 | Corynebacterium | ulcerans | 2.21 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 141001018 | 3; sequenced | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.26 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 141001548 | 3; sequenced | eDT | 2018-03-29 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.19 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C133 | 3; sequenced (rpoB) UGent; confirmed NGS (CVUAS) | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.08 | Corynebacterium | ulcerans | 2.03 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C15 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.28 | Corynebacterium | ulcerans | 2.19 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C44 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.37 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C55 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.22 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C64 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.51 | Corynebacterium | ulcerans | 2.49 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C78 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.28 | Corynebacterium | ulcerans | 2.19 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.49 | Corynebacterium | ulcerans | 2.36 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 46325 CVUAS | 3; type strain | DT | 2013-01-09 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.32 | Corynebacterium | ulcerans | 2.27 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 46325 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.05.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Corynebacterium ulcerans DSM 46325, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-05-27 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.6 | Corynebacterium | ulcerans | 2.42 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0199 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.36 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0853 | 3; MLST | eDT | 2016-10-27 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0941 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.47 | Corynebacterium | ulcerans | 2.4 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1017 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.4 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB 921348 | 2; 0 | eDT | 2015-01-30 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.11 | Corynebacterium | ulcerans | 2.04 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 171008197 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.33 | Corynebacterium | ulcerans | 2.2 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C156 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.32 | Corynebacterium | ulcerans | 2.19 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C164 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C166 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.29 | Corynebacterium | ulcerans | 2.23 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C169 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.24 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C174 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.55 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C25 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.17 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C92 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.29 | Corynebacterium | ulcerans | 2.24 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C95 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.52 | Corynebacterium | ulcerans | 2.32 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.34 | Corynebacterium | ulcerans | 2.29 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.25 | Corynebacterium | ulcerans | 2.2 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0188 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.37 | Corynebacterium | ulcerans | 2.24 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0195 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.29 | Corynebacterium | ulcerans | 2.14 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0818 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.32 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0867 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0870 | 3; 0 | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.36 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0885 | 3; FT-IR no NTTB-D, MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1015 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL1025 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 151012433-2 | 3; KL Diphtheria | eDT | 2021-06-22 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.22 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 181011209 001 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.2 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 181011209 002 | 3; 0 | EFEx | 2018-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.22 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C1 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C11 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.26 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C112 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C12 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.35 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C123 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.37 | Corynebacterium | ulcerans | 2.26 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C152 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.29 | Corynebacterium | ulcerans | 2.18 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C17 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 2.16 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C60 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.54 | Corynebacterium | ulcerans | 2.44 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C61 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C63 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.52 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C65 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.5 | Corynebacterium | ulcerans | 2.33 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C70 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.29 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C71 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C72 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.49 | Corynebacterium | ulcerans | 2.34 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C73 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.55 | Corynebacterium | ulcerans | 2.46 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C76 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.5 | Corynebacterium | ulcerans | 2.29 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C80 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C81 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.26 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C82 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.27 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C83 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.46 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C90 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.5 | Corynebacterium | ulcerans | 2.35 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C93 | 3; sequenced (rpoB) UGent | DT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.28 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1750 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.38 | Corynebacterium | ulcerans | 2.35 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.28 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10306 | 3; sequenced | EFex | 2018-04-24 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.3 | Corynebacterium | ulcerans | 2.19 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33950 | 3; KL Diphteria | eDT | 2022-06-13 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.11 | Corynebacterium | ulcerans | 2.11 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34094 | 3; KL Diphteria | eDT | 2022-08-01 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.34 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6522,4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.23 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 360523 | 3; 0 | eDT | 2013-01-30 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.21 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0190 | 3; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.41 | Corynebacterium | ulcerans | 2.29 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0246 | 3; WGS data LGL [0085]; Strain for comparison, send from the KL Diphtheria LGL | DT | 2012-12-12 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0796 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0825 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0832 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.36 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0840 | 3; MLST | eDT | 0 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.44 | Corynebacterium | ulcerans | 2.37 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S-477-6-21 | 3; KL Diphtheria | eDT | 2021-04-21 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.2 | Corynebacterium | ulcerans | 2.09 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | WFL-1217-18 | 3; 0 | eDT | 2018-04-17 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.27 | Corynebacterium | ulcerans | 2.23 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C117 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.39 | Corynebacterium | ulcerans | 2.34 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C146 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.45 | Corynebacterium | ulcerans | 2.4 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C162 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.33 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C18 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.3 | Corynebacterium | ulcerans | 2.23 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C20 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.23 | Corynebacterium | ulcerans | 2.13 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C28 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.4 | Corynebacterium | ulcerans | 2.35 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C31 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.3 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C32 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.15 | Corynebacterium | ulcerans | 2.14 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C39 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.52 | Corynebacterium | ulcerans | 2.39 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C41 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.36 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C51 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.24 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C52 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.3 | Corynebacterium | ulcerans | 2.23 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C53 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.4 | Corynebacterium | ulcerans | 2.25 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C57 CVUAS | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.17 | Corynebacterium | ulcerans | 2.11 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C67 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.36 | Corynebacterium | ulcerans | 2.32 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C98 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.31 | Corynebacterium | ulcerans | 2.22 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | bovis | Actinomyces | bovis | 1.92 | Actinomyces | bovis | 1.68 | B | 0 | 0 | 1 | Actinomyces bovis | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces | denticolens | Actinomyces | denticolens | 2.07 | Actinomyces | denticolens | 2.05 | A | 1 | 1 | 0 | Actinomyces denticolens | Actinomyces denticolens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces | sp-CVUAS-31303 | Campylobacter | jejuni | 1.26 | Staphylococcus | carnosus | 1.22 | B | 0 | 0 | 1 | Campylobacter jejuni | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces | viscosus | Actinomyces | viscosus | 2.04 | Actinomyces | viscosus | 1.87 | A | 1 | 1 | 0 | Actinomyces viscosus | Actinomyces viscosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | weissii | Actinomyces | weissii | 2.35 | Actinomyces | weissii | 2.35 | A | 1 | 1 | 0 | Actinomyces weissii | Actinomyces weissii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium | bovis | Arcanobacterium | pluranimalium | 1.77 | Arcanobacterium | pluranimalium | 1.52 | B | 0 | 0 | 1 | Arcanobacterium pluranimalium | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium | canis | Arcanobacterium | canis | 2.19 | Legionella | adelaidensis | 1.26 | A | 1 | 1 | 0 | Arcanobacterium canis | Legionella adelaidensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium | haemolyticum | Arcanobacterium | haemolyticum | 2.3 | Arcanobacterium | haemolyticum | 2.18 | A | 1 | 1 | 0 | Arcanobacterium haemolyticum | Arcanobacterium haemolyticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-2056 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.17 | Paucilactobacillus | oligofermentans | 1.34 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Paucilactobacillus oligofermentans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.1 | Arcanobacterium | hippocoleae | 1.49 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10002 | 3; type strain | EFex | 2017-10-04 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.5 | Arcanobacterium | phocae | 2.4 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.4 | Arcanobacterium | phocae | 2.29 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2698 | 3; type strain | DT | 2013-08-07 | Arcanobacterium | phocisimile | Arcanobacterium | phocisimile | 2.15 | Arcanobacterium | phocisimile | 1.84 | A | 1 | 1 | 0 | Arcanobacterium phocisimile | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.12 | Arcanobacterium | phocisimile | 1.79 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.34 | Arcanobacterium | phocisimile | 1.92 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13483 | 3; type strain | EFex | 2017-10-04 | Arcanobacterium | pluranimalium | Arcanobacterium | pluranimalium | 2.1 | Arcanobacterium | pluranimalium | 1.57 | A | 1 | 1 | 0 | Arcanobacterium pluranimalium | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | hippocoleae | 2.2 | Arcanobacterium | hippocoleae | 2.17 | A | 0 | 0 | 1 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | hippocoleae | 2.28 | Arcanobacterium | hippocoleae | 2.25 | A | 0 | 0 | 1 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium | sp-2701 | Arcanobacterium | phocae | 2.49 | Arcanobacterium | phocae | 2.36 | A | 0 | 0 | 1 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium | sp-2701 | Arcanobacterium | phocae | 2.46 | Arcanobacterium | phocae | 2.37 | A | 0 | 0 | 1 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1060,2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium | wilhelmae | Staphylococcus | schleiferi | 1.33 | Staphylococcus | schleiferi | 1.3 | B | 0 | 0 | 1 | Staphylococcus schleiferi | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 647 | 4; type strain | DT | 2014-09-04 | Arcanobacterium | wilhelmae | Lactobacillus | amylovorus | 1.34 | Torulaspora | delbrueckii | 1.25 | B | 0 | 0 | 1 | Lactobacillus amylovorus | Torulaspora delbrueckii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia | sp-CVUAS-30561 | Streptomyces | polychromogenes | 1.39 | Streptomyces | racemochromogenes | 1.38 | B | 0 | 0 | 1 | Streptomyces polychromogenes | Streptomyces racemochromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia | canis | Schaalia | canis | 2.26 | Schaalia | canis | 2.24 | A | 1 | 1 | 0 | Schaalia canis | Schaalia canis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia | cardiffensis | Schaalia | cardiffensis | 1.4 | Veillonella | atypica | 1.29 | B | 0 | 0 | 1 | Schaalia cardiffensis | Veillonella atypica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.12 | Schaalia | hyovaginalis | 2.06 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.4 | Schaalia | hyovaginalis | 2.29 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia | sp-CVUAS-32183 | Schaalia | funkei | 1.46 | Schaalia | funkei | 1.42 | B | 0 | 0 | 1 | Schaalia funkei | Schaalia funkei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia | sp-CVUAS-8688 | Schaalia | odontolytica | 1.39 | Schaalia | odontolytica | 1.3 | B | 0 | 0 | 1 | Schaalia odontolytica | Schaalia odontolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31845.2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia | sp-CVUAS-8688 | Schaalia | odontolytica | 1.44 | Schaalia | odontolytica | 1.37 | B | 0 | 0 | 1 | Schaalia odontolytica | Schaalia odontolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella | abortisuis | Trueperella | abortisuis | 2.27 | Trueperella | abortisuis | 2.2 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | abortisuis | Trueperella | abortisuis | 1.78 | Trueperella | abortisuis | 1.75 | B | 0 | 0 | 1 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella | bernardiae | Trueperella | bernardiae | 2.11 | Trueperella | bernardiae | 2.03 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella | bernardiae | Trueperella | bernardiae | 2.23 | Trueperella | bernardiae | 2.15 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella | bialowiezensis | Trueperella | bialowiezensis | 2.02 | Oceanobacillus | picturae | 1.3 | A | 1 | 1 | 0 | Trueperella bialowiezensis | Oceanobacillus picturae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella | bonasi | Trueperella | bonasi | 2.47 | Trueperella | bonasi | 2.32 | A | 1 | 1 | 0 | Trueperella bonasi | Trueperella bonasi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pyogenes | 1.6 | Trueperella | pyogenes | 1.5 | B | 0 | 0 | 1 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Bergeyella | zoohelcum | 1.44 | Levilactobacillus | brevis | 1.34 | B | 0 | 0 | 1 | Bergeyella zoohelcum | Levilactobacillus brevis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.42 | Trueperella | pyogenes | 2.39 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.32 | Trueperella | pyogenes | 2.3 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.19 | Corynebacterium | amycolatum | 2.15 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-17 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.3 | Corynebacterium | amycolatum | 2.29 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium | aurimucosum | Corynebacterium | aurimucosum | 2.07 | Corynebacterium | aurimucosum | 2.05 | A | 1 | 1 | 0 | Corynebacterium aurimucosum | Corynebacterium aurimucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium | bovis | Corynebacterium | bovis | 2.35 | Corynebacterium | bovis | 2.33 | A | 1 | 1 | 0 | Corynebacterium bovis | Corynebacterium bovis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium | camporealensis | Corynebacterium | camporealensis | 1.96 | Corynebacterium | flavescens | 1.55 | B | 0 | 0 | 1 | Corynebacterium camporealensis | Corynebacterium flavescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium | casei | Corynebacterium | casei | 1.98 | Corynebacterium | stationis | 1.64 | B | 0 | 0 | 1 | Corynebacterium casei | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium | confusum | Corynebacterium | confusum | 1.78 | Corynebacterium | confusum | 1.68 | B | 0 | 0 | 1 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | confusum | Corynebacterium | confusum | 1.64 | Corynebacterium | confusum | 1.51 | B | 0 | 0 | 1 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium | cystitidis | Corynebacterium | cystitidis | 2.36 | Pseudomonas | abietaniphila | 1.37 | A | 1 | 1 | 0 | Corynebacterium cystitidis | Pseudomonas abietaniphila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 1.97 | Candida | tropicalis | 1.57 | B | 0 | 0 | 1 | Corynebacterium frankenforstense | Candida tropicalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.48 | Streptococcus | pneumoniae | 1.29 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium | freiburgense | Pseudomonas | gessardii | 1.27 | Inquilinus | limosus | 1.24 | B | 0 | 0 | 1 | Pseudomonas gessardii | Inquilinus limosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | hansenii | Corynebacterium | xerosis | 1.42 | Corynebacterium | amycolatum | 1.35 | B | 0 | 0 | 1 | Corynebacterium xerosis | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | hansenii | Corynebacterium | xerosis | 1.58 | Corynebacterium | hansenii | 1.53 | B | 0 | 0 | 1 | Corynebacterium xerosis | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | humireducens | Corynebacterium | glutamicum | 1.44 | Gluconobacter | oxydans | 1.3 | B | 0 | 0 | 1 | Corynebacterium glutamicum | Gluconobacter oxydans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | kutscheri | Corynebacterium | kutscheri | 2.56 | Corynebacterium | kutscheri | 2.37 | A | 1 | 1 | 0 | Corynebacterium kutscheri | Corynebacterium kutscheri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | lactis | Corynebacterium | amycolatum | 1.46 | Tsukamurella | spumae | 1.3 | B | 0 | 0 | 1 | Corynebacterium amycolatum | Tsukamurella spumae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | lactis | Phocaeicola | vulgatus | 1.33 | Pseudomonas | veronii | 1.31 | B | 0 | 0 | 1 | Phocaeicola vulgatus | Pseudomonas veronii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium | phoceense | Corynebacterium | phoceense | 2.19 | Corynebacterium | phoceense | 1.89 | A | 1 | 1 | 0 | Corynebacterium phoceense | Corynebacterium phoceense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | pilosum | Corynebacterium | pilosum | 1.78 | Burkholderia | cenocepacia | 1.32 | B | 0 | 0 | 1 | Corynebacterium pilosum | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium | sp-208-U2-6 | Sphingomonas | yabuuchiae | 1.39 | Acidiphilium | acidophilum | 1.38 | B | 0 | 0 | 1 | Sphingomonas yabuuchiae | Acidiphilium acidophilum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32165.2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium | sp-2300500 | Acinetobacter | baumannii | 1.26 | Lysinibacillus | endophyticus | 1.22 | B | 0 | 0 | 1 | Acinetobacter baumannii | Lysinibacillus endophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5775,2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium | sp-2300500 | Acinetobacter | baumannii | 1.29 | Agromyces | fucosus | 1.28 | B | 0 | 0 | 1 | Acinetobacter baumannii | Agromyces fucosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2554,2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium | sp-2396 | Corynebacterium | sp | 1.92 | Arthrobacter | pyridinolis | 1.35 | B | 0 | 0 | 1 | Corynebacterium sp | Arthrobacter pyridinolis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium | sp-243-z-3 | Corynebacterium | resistens | 1.78 | Corynebacterium | resistens | 1.76 | B | 0 | 0 | 1 | Corynebacterium resistens | Corynebacterium resistens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | sp-257-y-4 | Proteus | mirabilis | 1.36 | Ligilactobacillus | agilis | 1.27 | B | 0 | 0 | 1 | Proteus mirabilis | Ligilactobacillus agilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | sp-284z-2 | Proteus | mirabilis | 1.3 | Loigolactobacillus | bifermentans | 1.3 | B | 0 | 0 | 1 | Proteus mirabilis | Loigolactobacillus bifermentans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium | sp-CVUAS-11664 | Pseudomonas | aeruginosa | 1.32 | Liquorilactobacillus | satsumensis | 1.24 | B | 0 | 0 | 1 | Pseudomonas aeruginosa | Liquorilactobacillus satsumensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | sp-CVUAS-1888 | Corynebacterium | striatum | 1.3 | Pseudomonas | aeruginosa | 1.28 | B | 0 | 0 | 1 | Corynebacterium striatum | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1721,3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium | sp-CVUAS-2484.4 | Streptococcus | gallolyticus | 1.5 | Corynebacterium | felinum | 1.41 | B | 0 | 0 | 1 | Streptococcus gallolyticus | Corynebacterium felinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2484.4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388; 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium | sp-CVUAS-2484.4 | Pseudomonas | citronellolis | 1.29 | Lacticaseibacillus | paracasei | 1.27 | B | 0 | 0 | 1 | Pseudomonas citronellolis | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | falsenii | 1.42 | Corynebacterium | sp | 1.32 | B | 0 | 0 | 1 | Corynebacterium falsenii | Corynebacterium sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 20TRD0812 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp | 1.4 | Veillonella | atypica | 1.32 | B | 0 | 0 | 1 | Corynebacterium sp | Veillonella atypica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp | 1.57 | Arthrobacter | psychrolactophilus | 1.36 | B | 0 | 0 | 1 | Corynebacterium sp | Arthrobacter psychrolactophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium | sp-CVUAS-31533 | Corynebacterium | amycolatum | 1.45 | Paeniglutamicibacter | sulfureus | 1.38 | B | 0 | 0 | 1 | Corynebacterium amycolatum | Paeniglutamicibacter sulfureus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32278.2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium | sp-CVUAS-32278.2 | Agromyces | cerinus | 1.2 | Listeria | grayi | 1.17 | B | 0 | 0 | 1 | Agromyces cerinus | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium | sp-CVUAS-874,4 | Clostridium | beijerinckii | 1.38 | Staphylococcus | carnosus | 1.34 | B | 0 | 0 | 1 | Clostridium beijerinckii | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 874,4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium | sp-CVUAS-874,4 | Neobacillus | bataviensis | 1.3 | Paucilactobacillus | suebicus | 1.29 | B | 0 | 0 | 1 | Neobacillus bataviensis | Paucilactobacillus suebicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 21TRD0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-LALLF-21TRD0553 | Cryptococcus | neoformans | 1.38 | Corynebacterium | flavescens | 1.37 | B | 0 | 0 | 1 | Cryptococcus neoformans | Corynebacterium flavescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium | sp-LHL-0497 | Haemophilus | influenzae | 1.26 | Staphylococcus | epidermidis | 1.17 | B | 0 | 0 | 1 | Haemophilus influenzae | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | sp-LHL-0558 | Nocardia | cyriacigeorgica | 1.32 | Glutamicibacter | protophormiae | 1.3 | B | 0 | 0 | 1 | Nocardia cyriacigeorgica | Glutamicibacter protophormiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium | sp-LHL-1322 | Agromyces | humatus | 1.41 | Corynebacterium | xerosis | 1.29 | B | 0 | 0 | 1 | Agromyces humatus | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | stationis | Corynebacterium | stationis | 2.27 | Corynebacterium | stationis | 2.21 | A | 1 | 1 | 0 | Corynebacterium stationis | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium | testudinoris | Bacteroides | fragilis | 1.29 | Citrobacter | koseri | 1.28 | B | 0 | 0 | 1 | Bacteroides fragilis | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5437,3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium | variabile | Corynebacterium | variabile | 1.62 | Blastomonas | ursincola | 1.34 | B | 0 | 0 | 1 | Corynebacterium variabile | Blastomonas ursincola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium | xerosis | Corynebacterium | xerosis | 1.57 | Levilactobacillus | brevis | 1.35 | B | 0 | 0 | 1 | Corynebacterium xerosis | Levilactobacillus brevis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium | xerosis-group | Corynebacterium | xerosis | 1.58 | Aromatoleum | bremensis | 1.45 | B | 0 | 0 | 1 | Corynebacterium xerosis | Aromatoleum bremensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2379,2 | 3; sequenced | EFex | 2021-04-08 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.04 | Gordonia | paraffinivorans | 2 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.14 | Gordonia | paraffinivorans | 1.92 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6272 | 2; API Coryne profile 0150006 | EFex | 2021-03-25 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 1.98 | Gordonia | paraffinivorans | 1.93 | B | 0 | 0 | 1 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7903,3 eDT | 3; sequenced | eDT | 2017-07-27 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.06 | Gordonia | paraffinivorans | 1.74 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7903,3 EFex | 3; sequenced | EFex | 2017-07-27 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.09 | Gordonia | paraffinivorans | 1.97 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia | rubripertincta | Gordonia | rubripertincta | 2.25 | Gordonia | rubripertincta | 2.17 | A | 1 | 1 | 0 | Gordonia rubripertincta | Gordonia rubripertincta | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10185.3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium | smegmatis | Mycobacterium | smegmatis | 1.48 | Mycobacterium | smegmatis | 1.42 | B | 0 | 0 | 1 | Mycobacterium smegmatis | Mycobacterium smegmatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31296,3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia | cyriacigeorgica | Nocardia | cyriacigeorgica | 2.45 | Nocardia | cyriacigeorgica | 2.28 | A | 1 | 1 | 0 | Nocardia cyriacigeorgica | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia | farcinica | Nocardia | farcinica | 2.36 | Nocardia | kroppenstedtii | 2.26 | B | 0 | 0 | 1 | Nocardia farcinica | Nocardia kroppenstedtii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.25 | Rhodococcus | hoagii | 2.22 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2755,3 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.29 | Rhodococcus | hoagii | 2.18 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2998 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.27 | Rhodococcus | hoagii | 2.13 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33761 sir | 2; 0 | EFEx | 2022-04-27 | Rhodococcus | equi | Rhodococcus | hoagii | 2.35 | Rhodococcus | hoagii | 2.35 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33761 | 2; 0 | EFEx | 2022-04-27 | Rhodococcus | equi | Rhodococcus | hoagii | 2.38 | Rhodococcus | hoagii | 2.38 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4057 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.23 | Rhodococcus | hoagii | 2.18 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5384,2 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.21 | Rhodococcus | hoagii | 2.09 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.19 | Rhodococcus | hoagii | 2.04 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1447,3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella | sp | Tsukamurella | tyrosinosolvens | 1.54 | Tsukamurella | paurometabola | 1.54 | B | 0 | 0 | 1 | Tsukamurella tyrosinosolvens | Tsukamurella paurometabola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2895.3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium | luteolum | Brevibacterium | luteolum | 1.92 | Brevibacterium | luteolum | 1.71 | B | 0 | 0 | 1 | Brevibacterium luteolum | Brevibacterium luteolum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas | flavigena | Bacteroides | fragilis | 1.43 | Neisseria | meningitidis | 1.35 | B | 0 | 0 | 1 | Bacteroides fragilis | Neisseria meningitidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium | sp-234z-2 | Glutamicibacter | protophormiae | 1.31 | Phocaeicola | vulgatus | 1.29 | B | 0 | 0 | 1 | Glutamicibacter protophormiae | Phocaeicola vulgatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium | sp-CVUAS-30560 | Raoultella | ornithinolytica | 1.39 | Glutamicibacter | ardleyensis | 1.31 | B | 0 | 0 | 1 | Raoultella ornithinolytica | Glutamicibacter ardleyensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus | sedentarius | Kytococcus | sedentarius | 2.1 | Kytococcus | sedentarius | 1.84 | A | 1 | 1 | 0 | Kytococcus sedentarius | Kytococcus sedentarius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus | congolensis | Dermatophilus | congolensis | 1.62 | Dermatophilus | congolensis | 1.4 | B | 0 | 0 | 1 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus | congolensis | Dermatophilus | congolensis | 1.58 | Chryseobacterium | arthrosphaerae | 1.39 | B | 0 | 0 | 1 | Dermatophilus congolensis | Chryseobacterium arthrosphaerae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium | flaccumfaciens | Curtobacterium | flaccumfaciens | 2.42 | Curtobacterium | flaccumfaciens | 2.31 | A | 1 | 1 | 0 | Curtobacterium flaccumfaciens | Curtobacterium flaccumfaciens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia | sp-MUVA-A-508948 | Staphylococcus | aureus | 1.46 | Clostridium | chauvoei | 1.45 | B | 0 | 0 | 1 | Staphylococcus aureus | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium | sp | Microbacterium | phyllosphaerae | 1.84 | Microbacterium | foliorum | 1.84 | B | 0 | 0 | 1 | Microbacterium phyllosphaerae | Microbacterium foliorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34837 | 3; sequenced | EFEx | 2023-01-19 | Microbacterium | sp-CVUAS-34837 | Microbacterium | lacticum | 1.39 | Lactobacillus | sakei | 1.3 | B | 0 | 0 | 1 | Microbacterium lacticum | Lactobacillus sakei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34838 | 3; sequenced | EFEx | 2023-01-19 | Microbacterium | sp-G5-3 | Microbacterium | sorbitolivorans | 1.48 | Lactobacillus | amylovorus | 1.44 | B | 0 | 0 | 1 | Microbacterium sorbitolivorans | Lactobacillus amylovorus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2011-09-01 | Glutamicibacter | nicotianae | Lactobacillus | pentosus | 1.33 | Burkholderia | cepacia | 1.26 | B | 0 | 0 | 1 | Lactobacillus pentosus | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10432 DSM 20127 | 3; type strain | unknown | 2011-09-08 | Sinomonas | atrocyanea | Sinomonas | atrocyanea | 2.52 | Sinomonas | atrocyanea | 2.24 | A | 1 | 1 | 0 | Sinomonas atrocyanea | Sinomonas atrocyanea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter | protophormiae | Lactobacillus | sakei | 1.37 | Loigolactobacillus | coryniformis | 1.31 | B | 0 | 0 | 1 | Lactobacillus sakei | Loigolactobacillus coryniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria | atrinae | Kocuria | atrinae | 2 | Kocuria | carniphila | 1.83 | A | 1 | 1 | 0 | Kocuria atrinae | Kocuria carniphila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 56a-5 | 3; sequenced | EFEx | 2020-05-20 | Kocuria | gwangalensis | Kocuria | carniphila | 1.64 | Kocuria | carniphila | 1.59 | B | 0 | 0 | 1 | Kocuria carniphila | Kocuria carniphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-14 | Kocuria | palustris | Kocuria | palustris | 2.38 | Kocuria | palustris | 2.33 | A | 1 | 1 | 0 | Kocuria palustris | Kocuria palustris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20030 | 3; type strain | #CON=3-5 loops in 100 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 2:2 HCCA¦TA2 (A) | 2011-08-25 | Micrococcus | luteus | Micrococcus | luteus | 1.98 | Micrococcus | luteus | 1.91 | B | 0 | 0 | 1 | Micrococcus luteus | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Micrococcus | luteus | Micrococcus | luteus | 2.37 | Micrococcus | luteus | 2.24 | A | 1 | 1 | 0 | Micrococcus luteus | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20032 | 3; type strain | eDT | 2015-07-09 | Rothia | kristinae | Rothia | kristinae | 2.26 | Rothia | kristinae | 2.25 | A | 1 | 1 | 0 | Rothia kristinae | Rothia kristinae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium | cellulans | Cellulosimicrobium | cellulans | 2.05 | Cellulosimicrobium | cellulans | 1.97 | A | 1 | 1 | 0 | Cellulosimicrobium cellulans | Cellulosimicrobium cellulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 749,2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter | inulinus | Sanguibacter | inulinus | 2.46 | Sanguibacter | keddieii | 1.76 | A | 1 | 1 | 0 | Sanguibacter inulinus | Sanguibacter keddieii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108415 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Cutibacterium acnes DSM 108415: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Cutibacterium | acnes | Cutibacterium | acnes | 2.24 | Cutibacterium | acnes | 2.17 | A | 1 | 1 | 0 | Cutibacterium acnes | Cutibacterium acnes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1079 | 3; public strain collection | EFEx | 2017-01-20 | Penicillium | roqueforti | Penicillium | camemberti | 1.55 | Staphylococcus | warneri | 1.49 | B | 0 | 0 | 1 | Penicillium camemberti | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10354.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | tonsurans | 1.38 | Limosilactobacillus | fermentum | 1.25 | B | 0 | 0 | 1 | Trichophyton tonsurans | Limosilactobacillus fermentum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10357.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | interdigitale | 1.38 | Trichophyton | mentagrophytes | 1.33 | B | 0 | 0 | 1 | Trichophyton interdigitale | Trichophyton mentagrophytes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1386 | 3; public strain collection | Fextr | 2016-03-10 | Candida | albicans | Candida | albicans | 2.2 | Candida | albicans | 2.14 | A | 1 | 1 | 0 | Candida albicans | Candida albicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31652 | 3; standard identification | EFEx | 2019-03-25 | Saccharomyces | cerevisiae | Saccharomyces | cerevisiae | 2.18 | Saccharomyces | cerevisiae | 1.92 | A | 1 | 1 | 0 | Saccharomyces cerevisiae | Saccharomyces cerevisiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 011 | 3; sequenced | EFEx | 2016-12-08 | Sarocladium | strictum | Sarocladium | strictum | 1.54 | Sarocladium | strictum | 1.41 | B | 0 | 0 | 1 | Sarocladium strictum | Sarocladium strictum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2520.4 | 3; sequenced (ITS) | F-Extr | 2018-07-26 | Fusarium | sambucinum | Kodamaea | ohmeri | 1.23 | Lentilactobacillus | parabuchneri | 1.2 | B | 0 | 0 | 1 | Kodamaea ohmeri | Lentilactobacillus parabuchneri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus | amylolyticus | Paenibacillus | amylolyticus | 2.26 | Paenibacillus | amylolyticus | 2.25 | A | 1 | 1 | 0 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus | amylolyticus | Paenibacillus | amylolyticus | 2.17 | Paenibacillus | amylolyticus | 2.12 | A | 1 | 1 | 0 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.53 | Paenibacillus | barengoltzii | 2.5 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.43 | Paenibacillus | barengoltzii | 2.4 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.18 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.25 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus | larvae | Paenibacillus | larvae | 2.32 | Paenibacillus | larvae | 2.28 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.33 | Paenibacillus | larvae | 2.24 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.33 | Paenibacillus | larvae | 2.24 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.18 | Paenibacillus | larvae | 2.12 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus | macerans | Paenibacillus | macerans | 2.28 | Paenibacillus | macerans | 2.07 | A | 1 | 1 | 0 | Paenibacillus macerans | Paenibacillus macerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.12 | Paenibacillus | pabuli | 2.02 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.22 | Paenibacillus | pabuli | 2.15 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus | pasadenensis | Paenibacillus | pasadenensis | 2.52 | Paenibacillus | pasadenensis | 2.05 | A | 1 | 1 | 0 | Paenibacillus pasadenensis | Paenibacillus pasadenensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus | validus | Paenibacillus | validus | 2.26 | Paenibacillus | validus | 2.16 | A | 1 | 1 | 0 | Paenibacillus validus | Paenibacillus validus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium | barati | Clostridium | budayi | 2.31 | Clostridium | baratii | 2.28 | B | 0 | 0 | 1 | Clostridium budayi | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium | butyricum | Clostridium | butyricum | 2.61 | Clostridium | butyricum | 2.48 | A | 1 | 1 | 0 | Clostridium butyricum | Clostridium butyricum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium | chauvoei | Clostridium | chauvoei | 2.13 | Clostridium | chauvoei | 2.01 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002729 | 2; 0 | DT | 2017-02-28 | Clostridium | cochlearium | Clostridium | cochlearium | 2.51 | Clostridium | cochlearium | 2.5 | A | 1 | 1 | 0 | Clostridium cochlearium | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium | colicanis | Clostridium | colicanis | 2.29 | Bifidobacterium | boum | 1.42 | A | 1 | 1 | 0 | Clostridium colicanis | Bifidobacterium boum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium | novyi | Clostridium | novyi | 2.03 | Clostridium | novyi | 1.96 | A | 1 | 1 | 0 | Clostridium novyi | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium | septicum | Clostridium | septicum | 2.44 | Clostridium | septicum | 2.26 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium | sporogenes | Clostridium | sporogenes | 2.34 | Clostridium | sporogenes | 2.21 | A | 1 | 1 | 0 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141015223 aerob | 2; 0 | DT | 2014-11-27 | Clostridium | tertium | Clostridium | tertium | 2.44 | Clostridium | tertium | 2.22 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium | tetani | Clostridium | tetani | 2.23 | Clostridium | tetani | 2.16 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya | limosa | Hathewaya | limosa | 2.29 | Hathewaya | limosa | 2.25 | A | 1 | 1 | 0 | Hathewaya limosa | Hathewaya limosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium | sordellii | Paeniclostridium | sordellii | 2.32 | Paeniclostridium | sordellii | 2.27 | A | 1 | 1 | 0 | Paeniclostridium sordellii | Paeniclostridium sordellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium | bifermentans | Paraclostridium | bifermentans | 2.02 | Paraclostridium | bifermentans | 1.92 | A | 1 | 1 | 0 | Paraclostridium bifermentans | Paraclostridium bifermentans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCEB 623 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-10 | Bacillus | badius | Bacillus | badius | 2.06 | Bacillus | badius | 1.99 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-31 | Bacillus | ectoiniformans | Streptococcus | dysgalactiae | 1.49 | Bacillus | licheniformis | 1.28 | B | 0 | 0 | 1 | Streptococcus dysgalactiae | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Cytobacillus | firmus | Cytobacillus | firmus | 2.45 | Cytobacillus | firmus | 2.24 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 37 B 88 | 3; reference strain collection | 0 | 2015-02-09 | Cytobacillus | firmus | Cytobacillus | firmus | 2.06 | Cytobacillus | firmus | 1.98 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-10 | Niallia | circulans | Niallia | circulans | 1.98 | Niallia | circulans | 1.88 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Virgibacillus | pantothenticus | Virgibacillus | pantothenticus | 2.01 | Virgibacillus | pantothenticus | 1.84 | A | 1 | 1 | 0 | Virgibacillus pantothenticus | Virgibacillus pantothenticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31071,2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 1.69 | Alkalihalobacillus | clausii | 1.58 | B | 0 | 0 | 1 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 1.71 | Alkalihalobacillus | clausii | 1.7 | B | 0 | 0 | 1 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13666 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Alkalihalobacillus okuhidensis DSM 13666: preparation and measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-19 | Alkalihalobacillus | okuhidensis | Alkalihalobacillus | okuhidensis | 2.23 | Alkalihalobacillus | halodurans | 1.96 | A | 1 | 1 | 0 | Alkalihalobacillus okuhidensis | Alkalihalobacillus halodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28875 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-01-25 | Bacillus | albus | Bacillus | cereus | 2.08 | Bacillus | cereus | 2.07 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21631 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Strain obtained from DSMZ strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-05-09 | Bacillus | altitudinis | Bacillus | altitudinis | 1.95 | Bacillus | altitudinis | 1.9 | B | 0 | 0 | 1 | Bacillus altitudinis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26896 | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-05 | Bacillus | altitudinis | Bacillus | altitudinis | 2.14 | Bacillus | altitudinis | 2.13 | A | 1 | 1 | 0 | Bacillus altitudinis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 2015-02-13 | Bacillus | alveayuensis | Bacillus | alveayuensis | 2.52 | Peribacillus | psychrosaccharolyticus | 1.35 | A | 1 | 1 | 0 | Bacillus alveayuensis | Peribacillus psychrosaccharolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2016-03-01 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.24 | Bacillus | amyloliquefaciens | 1.99 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7 CVUAS | 3; type strain | DT | 2014-01-29 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.15 | Bacillus | amyloliquefaciens | 2.15 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-09 | Bacillus | anthracis | Bacillus | cereus | 2.01 | Bacillus | cereus | 1.96 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-26 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 004 | 2; 0 | DT | 2014-01-22 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.14 | Bacillus | atrophaeus | 2.14 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-15 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.43 | Bacillus | atrophaeus | 2.36 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DGL6 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 07.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-19 | Bacillus | axarquiensis | Bacillus | mojavensis | 1.89 | Bacillus | mojavensis | 1.73 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNXM37 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-16 | Bacillus | axarquiensis | Bacillus | mojavensis | 2.09 | Bacillus | subtilis | 1.77 | A | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Bacillus | badius | Bacillus | badius | 2.45 | Bacillus | badius | 2.26 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25427 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 05.10.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus bingmayongensis DSM 25427, preparation and measurements: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-11-05 | Bacillus | bingmayongensis | Bacillus | cereus | 1.99 | Bacillus | pseudomycoides | 1.88 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus | cereus | Bacillus | cereus | 2.05 | Bacillus | thuringiensis | 2.02 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1107 | 2; 0 | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | cereus | 2.41 | Bacillus | cereus | 2.38 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2833 | 4; 16S, draft genome [0009] | eDT | 2016-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.48 | Bacillus | cytotoxicus | 2.45 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NVH 391-98 | 3; type strain | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.26 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29500 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Strain obtained from DSMZ strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-09-05 | Bacillus | gobiensis | Stenotrophomonas | pictorum | 1.52 | Metabacillus | indicus | 1.46 | B | 0 | 0 | 1 | Stenotrophomonas pictorum | Metabacillus indicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-30 | Bacillus | halotolerans | Bacillus | mojavensis | 1.96 | Bacillus | mojavensis | 1.83 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-12 | Bacillus | halotolerans | Bacillus | mojavensis | 2.23 | Bacillus | subtilis | 2.1 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21200 | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-28 | Bacillus | inaquosorum | Bacillus | subtilis | 2.12 | Bacillus | subtilis | 2.09 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-12-02 | Bacillus | inaquosorum | Bacillus | subtilis | 2.27 | Bacillus | subtilis | 2.12 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-12 | Bacillus | infantis | Bacillus | infantis | 2.46 | Bacillus | infantis | 2.39 | A | 1 | 1 | 0 | Bacillus infantis | Bacillus infantis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus | licheniformis | Bacillus | licheniformis | 2.17 | Bacillus | licheniformis | 2.15 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13 | 3; type strain | eDT | 2016-01-27 | Bacillus | licheniformis | Bacillus | licheniformis | 2.27 | Bacillus | licheniformis | 1.81 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28872 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-04 | Bacillus | luti | Bacillus | cereus | 2.36 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26473 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-07-25 | Bacillus | manliponensis | Bacillus | manliponensis | 2.42 | Bacillus | manliponensis | 2.14 | A | 1 | 1 | 0 | Bacillus manliponensis | Bacillus manliponensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28877 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-04 | Bacillus | mobilis | Bacillus | cereus | 1.91 | Bacillus | cereus | 1.84 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-10 | Bacillus | mojavensis | Bacillus | subtilis | 2.13 | Bacillus | subtilis | 2.08 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Bacillus | mojavensis | Bacillus | mojavensis | 2.33 | Bacillus | subtilis | 2.27 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2048 | 3; type strain | eDT | 09.12.2015 | Bacillus | mycoides | Bacillus | mycoides | 2.22 | Bacillus | mycoides | 2.08 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | mycoides | Bacillus | mycoides | 2.04 | Bacillus | mycoides | 2.01 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28871 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-01-25 | Bacillus | nitratireducens | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.07 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10987 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-10-23 | Bacillus | pacificus | Bacillus | cereus | 1.92 | Bacillus | mycoides | 1.62 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SN4-1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-04-11 | Bacillus | pacificus | Bacillus | cereus | 2.32 | Bacillus | cereus | 2.25 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 12759 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus | paralicheniformis | Staphylococcus | epidermidis | 1.45 | Staphylococcus | carnosus | 1.45 | B | 0 | 0 | 1 | Staphylococcus epidermidis | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ01-02 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-02-24 | Bacillus | paralicheniformis | Bacillus | licheniformis | 1.7 | Bacillus | licheniformis | 1.66 | B | 0 | 0 | 1 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28876 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 3.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Bacillus paramycoides LMG 28876, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-02-04 | Bacillus | paramycoides | Bacillus | cereus | 2.23 | Bacillus | mycoides | 2.02 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28873 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-11 | Bacillus | paranthracis | Bacillus | cereus | 2.14 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28870 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-01 | Bacillus | proteolyticus | Bacillus | cereus | 2.23 | Bacillus | cereus | 2.11 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus | pseudomycoides | Bacillus | cereus | 2.43 | Bacillus | cereus | 2.4 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 12443 | 3; public strain collection | eDT | 09.12.2015 | Bacillus | pseudomycoides | Bacillus | mycoides | 2.13 | Bacillus | cereus | 2.13 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B017 | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus | pumilus | Bacillus | pumilus | 2.15 | Bacillus | pumilus | 1.96 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 340 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | pumilus | Bacillus | pumilus | 2.04 | Bacillus | pumilus | 1.95 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 103081 | 3; public strain collection | see [0185] | 2023-03-02 | Bacillus | safensis | Bacillus | altitudinis | 1.44 | Bacillus | safensis | 1.4 | B | 0 | 0 | 1 | Bacillus altitudinis | Bacillus safensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19292 | 3; type strain | see [0185] | 2023-03-02 | Bacillus | safensis | Bacillus | safensis | 1.92 | Bacillus | pumilus | 1.85 | B | 0 | 0 | 1 | Bacillus safensis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-05-24 | Bacillus | siamensis | Bacillus | subtilis | 1.84 | Bacillus | mojavensis | 1.8 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-15 | Bacillus | sonorensis | Bacillus | sonorensis | 1.95 | Listeria | grayi | 1.56 | B | 0 | 0 | 1 | Bacillus sonorensis | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12369 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Bacillus | sp-DSM-12369 | Bacillus | sonorensis | 1.53 | Thauera | mechernichensis | 1.39 | B | 0 | 0 | 1 | Bacillus sonorensis | Thauera mechernichensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 6633 | 3; public strain collection | eDT | 2015-07-22 | Bacillus | spizizenii | Bacillus | subtilis | 2.08 | Bacillus | subtilis | 1.96 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 618 | 4; public strain collection | eDT | 2016-01-27 | Bacillus | spizizenii | Bacillus | subtilis | 2.09 | Bacillus | subtilis | 2.04 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 110 B068 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.33 | Bacillus | subtilis | 2.15 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1092 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.23 | Bacillus | subtilis | 2.2 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 2.17 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.04 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus | thuringiensis | Bacillus | cereus | 2.03 | Bacillus | cereus | 1.97 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 2.24 | Bacillus | cereus | 2.19 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 1.99 | Bacillus | cereus | 1.95 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6890 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-12-02 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.18 | Bacillus | cereus | 2.04 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-11-07 | Bacillus | thuringiensis | Bacillus | cereus | 1.75 | Bacillus | cereus | 1.7 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2786,2 | 3; sequenced | DT | 2013-07-31 | Bacillus | timorensis | Bacillus | timonensis | 1.49 | Bacillus | smithii | 1.38 | B | 0 | 0 | 1 | Bacillus timonensis | Bacillus smithii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-08 | Bacillus | tropicus | Bacillus | cereus | 2.28 | Bacillus | cereus | 2.19 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-07-20 | Bacillus | vallismortis | Bacillus | vallismortis | 2.37 | Bacillus | subtilis | 1.89 | A | 1 | 1 | 0 | Bacillus vallismortis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102050 | 3; public strain collection | see [0185] | 2018-05-23 | Bacillus | wiedmannii | Bacillus | cereus | 2.2 | Bacillus | mycoides | 2.15 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSC29 (01) | 2; strain collection of the RKI | see [0185] | 2018-05-17 | Bacillus | wiedmannii | Bacillus | mycoides | 2.11 | Bacillus | thuringiensis | 2.1 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29883 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2018-05-24 | Bacillus | xiamenensis | Bacillus | pumilus | 1.74 | Bacillus | pumilus | 1.56 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LLTC93 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-07-18 | Bacillus | zhangzhouensis | Bacillus | pumilus | 1.81 | Bacillus | pumilus | 1.76 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCEB 627 DSM 12 | 3; type strain | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.19 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 37 | 3; reference strain collection | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.33 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-15 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.11 | Cytobacillus | horneckiae | 2.01 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-15 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.32 | Cytobacillus | horneckiae | 2.18 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus | kochii | Cytobacillus | kochii | 1.86 | Cytobacillus | kochii | 1.83 | B | 0 | 0 | 1 | Cytobacillus kochii | Cytobacillus kochii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-02-19 | Domibacillus | aminovorans | Pseudomonas | aeruginosa | 1.37 | Micrococcus | luteus | 1.35 | B | 0 | 0 | 1 | Pseudomonas aeruginosa | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-02-15 | Domibacillus | aminovorans | Staphylococcus | lugdunensis | 1.46 | Lodderomyces | elongisporus | 1.45 | B | 0 | 0 | 1 | Staphylococcus lugdunensis | Lodderomyces elongisporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-19 | Evansella | clarkii | Bacillus | clarkii | 2.35 | Bacillus | mycoides | 1.45 | A | 0 | 0 | 1 | Bacillus clarkii | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Evansella | clarkii | Bacillus | clarkii | 2.04 | Alkalihalobacillus | krulwichiae | 1.46 | A | 0 | 0 | 1 | Bacillus clarkii | Alkalihalobacillus krulwichiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus | stearothermophilus | Geobacillus | stearothermophilus | 2.39 | Geobacillus | stearothermophilus | 2.29 | A | 1 | 1 | 0 | Geobacillus stearothermophilus | Geobacillus stearothermophilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-25 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.19 | Alkalihalobacillus | okuhidensis | 1.98 | A | 0 | 0 | 1 | Alkalihalobacillus halodurans | Alkalihalobacillus okuhidensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | okuhidensis | 2.14 | Alkalihalobacillus | halodurans | 1.92 | A | 0 | 0 | 1 | Alkalihalobacillus okuhidensis | Alkalihalobacillus halodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Heyndrickxia | coagulans | Weizmannia | coagulans | 2.45 | Weizmannia | coagulans | 2.12 | A | 0 | 0 | 1 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30760 | 2; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Weizmannia coagulans DSM 2308: Preparation & measurement: N. Fischer#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-02 | Heyndrickxia | coagulans | Weizmannia | coagulans | 2.15 | Weizmannia | coagulans | 2.02 | A | 0 | 0 | 1 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia | sporothermodurans | Heyndrickxia | sporothermodurans | 1.34 | Acidovorax | citrulli | 1.3 | B | 0 | 0 | 1 | Heyndrickxia sporothermodurans | Acidovorax citrulli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 10599 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Heyndrickxia | sporothermodurans | Heyndrickxia | sporothermodurans | 2.35 | Paracoccus | versutus | 1.32 | A | 1 | 1 | 0 | Heyndrickxia sporothermodurans | Paracoccus versutus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-20 | Lederbergia | lenta | Lederbergia | lenta | 2.43 | Lederbergia | lenta | 2.33 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9 | 3; type strain | DT | 2014-01-22 | Lederbergia | lenta | Lederbergia | lenta | 2.22 | Lederbergia | lenta | 2.08 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus | bataviensis | Neobacillus | bataviensis | 2.29 | Staphylococcus | cohnii | 1.41 | A | 1 | 1 | 0 | Neobacillus bataviensis | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-20 | Neobacillus | cucumis | Neobacillus | bataviensis | 2.02 | Bacillus | cereus | 1.54 | A | 0 | 0 | 1 | Neobacillus bataviensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-03-05 | Niallia | circulans | Niallia | circulans | 1.9 | Niallia | circulans | 1.65 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2582,2 | 3; sequenced | eDT | 2015-09-23 | Niallia | nealsonii | Niallia | taxi | 2.04 | Citrobacter | freundii | 1.42 | A | 0 | 0 | 1 | Niallia taxi | Citrobacter freundii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-10 | Peribacillus | muralis | Peribacillus | muralis | 2.38 | Peribacillus | simplex | 1.81 | A | 1 | 1 | 0 | Peribacillus muralis | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-18 | Peribacillus | psychrosaccharolyticus | Peribacillus | psychrosaccharolyticus | 2.21 | Staphylococcus | simiae | 1.44 | A | 1 | 1 | 0 | Peribacillus psychrosaccharolyticus | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus | simplex | Peribacillus | simplex | 1.67 | Schleiferilactobacillus | perolens | 1.45 | B | 0 | 0 | 1 | Peribacillus simplex | Schleiferilactobacillus perolens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-09-03 | Peribacillus | simplex | Peribacillus | simplex | 1.78 | Peribacillus | simplex | 1.76 | B | 0 | 0 | 1 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-12-15 | Peribacillus | simplex | Peribacillus | simplex | 2.19 | Peribacillus | simplex | 1.64 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-18 | Priestia | flexa | Priestia | flexa | 2.54 | Priestia | flexa | 2.17 | A | 1 | 1 | 0 | Priestia flexa | Priestia flexa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Priestia | flexa | Priestia | megaterium | 2.01 | Priestia | megaterium | 1.8 | A | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-26 | Priestia | megaterium | Priestia | megaterium | 1.92 | Priestia | megaterium | 1.61 | B | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia | megaterium | Priestia | megaterium | 2.06 | Priestia | megaterium | 2.06 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Salibacillus | salarius | Salibacterium | salarium | 2.05 | Aeromonas | salmonicida | 1.35 | A | 0 | 0 | 1 | Salibacterium salarium | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.13 | Sutcliffiella | cohnii | 1.86 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-25 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.03 | Sutcliffiella | cohnii | 1.79 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus | saccharophilus | Lodderomyces | elongisporus | 1.45 | Liquorilactobacillus | nagelii | 1.34 | B | 0 | 0 | 1 | Lodderomyces elongisporus | Liquorilactobacillus nagelii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus | proomii | Virgibacillus | proomii | 2.02 | Virgibacillus | proomii | 1.99 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3616,2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus | proomii | Virgibacillus | proomii | 1.82 | Virgibacillus | proomii | 1.71 | B | 0 | 0 | 1 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 71,2 | 3; sequenced | DT | 2013-12-13 | Weizmannia | ginsengihumi | Weizmannia | ginsengihumi | 2.19 | Weizmannia | ginsengihumi | 2.06 | A | 1 | 1 | 0 | Weizmannia ginsengihumi | Weizmannia ginsengihumi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.47 | Kurthia | gibsonii | 2.39 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.27 | Kurthia | gibsonii | 2.26 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia | zopfii | Kurthia | zopfii | 1.66 | Kurthia | zopfii | 1.61 | B | 0 | 0 | 1 | Kurthia zopfii | Kurthia zopfii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 100506 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 03.11.2021¦ 1:1 HCCA | 2021-11-03 | Lysinibacillus | endophyticus | Lysinibacillus | endophyticus | 2.45 | Enterococcus | faecium | 1.64 | A | 1 | 1 | 0 | Lysinibacillus endophyticus | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.34 | Lysinibacillus | fusiformis | 2.23 | A | 1 | 1 | 0 | Lysinibacillus fusiformis | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus | fusiformis | Lysinibacillus | sphaericus | 2.36 | Lysinibacillus | fusiformis | 2.09 | B | 0 | 0 | 1 | Lysinibacillus sphaericus | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7192,2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus | macroides | Lysinibacillus | xylanilyticus | 2.23 | Lysinibacillus | boronitolerans | 2.18 | B | 0 | 0 | 1 | Lysinibacillus xylanilyticus | Lysinibacillus boronitolerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 54 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-03 | Lysinibacillus | macroides | Lysinibacillus | macroides | 1.8 | Bacillus | mycoides | 1.44 | B | 0 | 0 | 1 | Lysinibacillus macroides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24784 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-15 | Lysinibacillus | pakistanensis | Lysinibacillus | pakistanensis | 2.3 | Lysinibacillus | sp[2] | 1.98 | A | 1 | 1 | 0 | Lysinibacillus pakistanensis | Lysinibacillus sp[2] | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.34 | Lysinibacillus | sphaericus | 2.31 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.11 | Lysinibacillus | sphaericus | 1.96 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23493 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-15 | Lysinibacillus | xylanilyticus | Lysinibacillus | sphaericus | 1.6 | Lysinibacillus | sp[2] | 1.44 | B | 0 | 0 | 1 | Lysinibacillus sphaericus | Lysinibacillus sp[2] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus | saudimassiliensis | Lysinibacillus | fusiformis | 1.48 | Ureibacillus | thermosphaericus | 1.37 | B | 0 | 0 | 1 | Lysinibacillus fusiformis | Ureibacillus thermosphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus | odysseyi | Ureibacillus | sinduriensis | 1.47 | Lysinibacillus | sp[2] | 1.4 | B | 0 | 0 | 1 | Ureibacillus sinduriensis | Lysinibacillus sp[2] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium | acetylicum | Exiguobacterium | indicum | 1.82 | Exiguobacterium | soli | 1.6 | B | 0 | 0 | 1 | Exiguobacterium indicum | Exiguobacterium soli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium | enclense | Exiguobacterium | soli | 1.68 | Exiguobacterium | sp[2] | 1.33 | B | 0 | 0 | 1 | Exiguobacterium soli | Exiguobacterium sp[2] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium | profundum | Exiguobacterium | sp[3] | 2.41 | Exiguobacterium | sp[3] | 2.36 | A | 0 | 0 | 1 | Exiguobacterium sp[3] | Exiguobacterium sp[3] | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.32 | Brochothrix | thermosphacta | 2.24 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-14 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.39 | Brochothrix | thermosphacta | 2.37 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria | aquatica | Listeria | aquatica | 2.26 | Listeria | grayi | 1.48 | A | 1 | 1 | 0 | Listeria aquatica | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria | aquatica | Listeria | aquatica | 2.29 | Listeria | grayi | 1.44 | A | 1 | 1 | 0 | Listeria aquatica | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.11 | Listeria | fleischmannii | 2.1 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LU2006-3 | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.03 | Listeria | fleischmannii | 1.77 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105447 | 3; type strain | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.32 | Listeria | grayi | 2.32 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20596 | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.25 | Listeria | grayi | 2.12 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10132,2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | innocua | Listeria | innocua | 2.45 | Listeria | innocua | 2.44 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8569 | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | innocua | Listeria | innocua | 2.48 | Listeria | innocua | 2.46 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria | ivanovii | Listeria | ivanovii | 2.33 | Listeria | ivanovii | 2.32 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32100 | 2; sequenced | EFex | 2019-12-13 | Listeria | ivanovii | Listeria | ivanovii | 2.41 | Listeria | ivanovii | 2.35 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 23813 | 3; type strain | EFEx | 2017-02-15 | Listeria | marthii | Listeria | monocytogenes | 2.28 | Listeria | monocytogenes | 2.19 | A | 0 | 0 | 1 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.42 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 873,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.43 | Listeria | monocytogenes | 2.41 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 240,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.29 | Listeria | monocytogenes | 2.24 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CLIP 70676 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SLCC 2479 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.19 | Listeria | monocytogenes | 2.19 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.22 | Listeria | monocytogenes | 2.21 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.37 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7366,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.54 | Listeria | monocytogenes | 2.51 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria | newyorkensis | Listeria | newyorkensis | 2.07 | Listeria | cornellensis | 2.06 | B | 0 | 0 | 1 | Listeria newyorkensis | Listeria cornellensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 36192 | 3; sequenced | EFex | 2024-03-06 | Listeria | newyorkensis | Listeria | newyorkensis | 2.12 | Listeria | cornellensis | 2.04 | B | 0 | 0 | 1 | Listeria newyorkensis | Listeria cornellensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 22097 | 3; type strain | EFEx | 2017-02-15 | Listeria | rocourtiae | Listeria | rocourtiae | 2.1 | Listeria | cornellensis | 1.94 | A | 1 | 1 | 0 | Listeria rocourtiae | Listeria cornellensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2763 | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | seeligeri | Listeria | seeligeri | 2.45 | Listeria | seeligeri | 2.45 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | seeligeri | Listeria | seeligeri | 2.36 | Listeria | seeligeri | 2.3 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria | sp | Listeria | newyorkensis | 2.13 | Listeria | cornellensis | 1.89 | A | 0 | 0 | 1 | Listeria newyorkensis | Listeria cornellensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2412,3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | welshimeri | Listeria | welshimeri | 2.38 | Listeria | welshimeri | 2.28 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria | welshimeri | Listeria | welshimeri | 2.44 | Listeria | welshimeri | 2.41 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 1.95 | Brevibacillus | laterosporus | 1.91 | B | 0 | 0 | 1 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.42 | Brevibacillus | laterosporus | 2.25 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD1.4A | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-08 | Brevibacillus | parabrevis | Brevibacillus | parabrevis | 1.98 | Brevibacillus | brevis | 1.61 | B | 0 | 0 | 1 | Brevibacillus parabrevis | Brevibacillus brevis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.3A | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Brevibacillus | parabrevis | Brevibacillus | parabrevis | 2.1 | Brevibacillus | parabrevis | 1.84 | A | 1 | 1 | 0 | Brevibacillus parabrevis | Brevibacillus parabrevis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HB1.1 | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Brevibacillus | porteri | Brevibacillus | choshinensis | 1.69 | Brevibacillus | formosus | 1.63 | B | 0 | 0 | 1 | Brevibacillus choshinensis | Brevibacillus formosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HB1.2 | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Brevibacillus | porteri | Brevibacillus | choshinensis | 1.81 | Brevibacillus | formosus | 1.56 | B | 0 | 0 | 1 | Brevibacillus choshinensis | Brevibacillus formosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 210 | 3; eudo | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Paenibacillus | alvei | Paenibacillus | alvei | 1.82 | Paenibacillus | curdlanolyticus | 1.45 | B | 0 | 0 | 1 | Paenibacillus alvei | Paenibacillus curdlanolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus | chibensis | Paenibacillus | chibensis | 2.28 | Paenibacillus | chibensis | 2.15 | A | 1 | 1 | 0 | Paenibacillus chibensis | Paenibacillus chibensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus | nuruki | Ochrobactrum | anthropi | 1.29 | Paucilactobacillus | suebicus | 1.23 | B | 0 | 0 | 1 | Ochrobactrum anthropi | Paucilactobacillus suebicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-09-12 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 1.9 | Paenibacillus | polymyxa | 1.78 | B | 0 | 0 | 1 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-24 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 2.1 | Paenibacillus | polymyxa | 2.02 | A | 1 | 1 | 0 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus | sp-CVUAS-6369 | Lentilactobacillus | parabuchneri | 1.35 | Lactobacillus | amylovorus | 1.3 | B | 0 | 0 | 1 | Lentilactobacillus parabuchneri | Lactobacillus amylovorus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus | sp-E8a | Paenibacillus | amylolyticus | 2.24 | Paenibacillus | amylolyticus | 2.17 | A | 0 | 0 | 1 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus | taichungensis | Paenibacillus | illinoisensis | 1.77 | Paenibacillus | tylopili | 1.76 | B | 0 | 0 | 1 | Paenibacillus illinoisensis | Paenibacillus tylopili | 0 | 0 | 0 | 0
|  |  | 1 | 0 | FLI 003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus | taichungensis | Paenibacillus | pabuli | 1.83 | Paenibacillus | illinoisensis | 1.75 | B | 0 | 0 | 1 | Paenibacillus pabuli | Paenibacillus illinoisensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus | tundrae | Paenibacillus | pabuli | 1.78 | Paenibacillus | tylopili | 1.77 | B | 0 | 0 | 1 | Paenibacillus pabuli | Paenibacillus tylopili | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | halophilus | Listeria | innocua | 1.68 | Listeria | innocua | 1.48 | B | 0 | 0 | 1 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | huakuii | Macrococcus | canis | 1.37 | Neisseria | meningitidis | 1.36 | B | 0 | 0 | 1 | Macrococcus canis | Neisseria meningitidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | psychrophilus | Jeotgalicoccus | psychrophilus | 1.84 | Legionella | geestiana | 1.41 | B | 0 | 0 | 1 | Jeotgalicoccus psychrophilus | Legionella geestiana | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 2013-03-14 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.08 | Macrococcus | caseolyticus | 1.67 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 1.87 | Macrococcus | caseolyticus | 1.77 | B | 0 | 0 | 1 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus | carouselicus | Staphylococcus | simulans | 1.4 | Fredinandcohnia | humi | 1.39 | B | 0 | 0 | 1 | Staphylococcus simulans | Fredinandcohnia humi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus | lentus | Staphylococcus | lentus | 1.8 | Staphylococcus | lentus | 1.51 | B | 0 | 0 | 1 | Staphylococcus lentus | Staphylococcus lentus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus | sciuri | Staphylococcus | sciuri | 2.37 | Staphylococcus | sciuri | 2.27 | A | 0 | 0 | 1 | Staphylococcus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus | sciuri | Staphylococcus | sciuri | 1.99 | Staphylococcus | sciuri | 1.68 | B | 0 | 0 | 1 | Staphylococcus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9937.2 | 3; 16S rDNA-sequencing | EFex | 2016-04-06 | Mammaliicoccus | vitulinus | Staphylococcus | vitulinus | 1.87 | Staphylococcus | vitulinus | 1.81 | B | 0 | 0 | 1 | Staphylococcus vitulinus | Staphylococcus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 343.2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus | vitulinus | Staphylococcus | vitulinus | 2.12 | Staphylococcus | vitulinus | 2.05 | A | 0 | 0 | 1 | Staphylococcus vitulinus | Staphylococcus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5687 | 3; sequenced | DT | 2013-04-16 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | hyicus | 1.91 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TUM 7385 | 3; sequenced | DT | 02.10.2013 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | hyicus | 1.96 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.39 | Staphylococcus | argenteus | 2.38 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.28 | Staphylococcus | argenteus | 2.2 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 254-U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus | arlettae | Staphylococcus | arlettae | 1.72 | Staphylococcus | arlettae | 1.51 | B | 0 | 0 | 1 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus | arlettae | Staphylococcus | arlettae | 1.6 | Staphylococcus | arlettae | 1.58 | B | 0 | 0 | 1 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 29213 | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9914.2 | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus | auricularis | Staphylococcus | auricularis | 2.14 | Staphylococcus | auricularis | 1.85 | A | 1 | 1 | 0 | Staphylococcus auricularis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20326 | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.25 | Staphylococcus | capitis | 2.13 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.2 | Staphylococcus | capitis | 2.19 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20608 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.16 | Staphylococcus | caprae | 2.06 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.06 | Staphylococcus | caprae | 2.03 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.18 | Staphylococcus | carnosus | 1.55 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.24 | Staphylococcus | carnosus | 1.46 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.26 | Staphylococcus | hyicus | 1.85 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.28 | Staphylococcus | chromogenes | 1.91 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | cohnii | Staphylococcus | xylosus | 1.36 | Staphylococcus | schleiferi | 1.36 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus | cohnii | Staphylococcus | cohnii | 1.88 | Staphylococcus | ureilyticus | 1.84 | B | 0 | 0 | 1 | Staphylococcus cohnii | Staphylococcus ureilyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | x | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-04 | Staphylococcus | delphini | Staphylococcus | delphini | 2.32 | Staphylococcus | delphini | 2.01 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus | delphini | Staphylococcus | delphini | 1.9 | Staphylococcus | intermedius | 1.83 | B | 0 | 0 | 1 | Staphylococcus delphini | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.09 | Staphylococcus | epidermidis | 2.08 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.31 | Staphylococcus | epidermidis | 1.99 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | equorum | Staphylococcus | equorum | 2.16 | Staphylococcus | equorum | 1.8 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20674 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | equorum | Staphylococcus | equorum | 2.11 | Staphylococcus | equorum | 1.88 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | felis | Staphylococcus | felis | 2.13 | Staphylococcus | felis | 1.96 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | felis | Staphylococcus | felis | 2.14 | Staphylococcus | felis | 2 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2011-08-04 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.46 | Staphylococcus | haemolyticus | 2.43 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.12 | Staphylococcus | haemolyticus | 2.04 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus | hominis | Staphylococcus | hominis | 2.27 | Staphylococcus | hominis | 2.23 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 789,2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus | hominis | Staphylococcus | hominis | 2.13 | Staphylococcus | hominis | 1.96 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.2 | Staphylococcus | hyicus | 2.14 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3767 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.38 | Staphylococcus | hyicus | 2.12 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.15 | Staphylococcus | intermedius | 2.02 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.07 | Staphylococcus | intermedius | 1.97 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.4 | Staphylococcus | lugdunensis | 2.4 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.5 | Staphylococcus | lugdunensis | 2.43 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | microti | Staphylococcus | microti | 1.85 | Staphylococcus | microti | 1.85 | B | 0 | 0 | 1 | Staphylococcus microti | Staphylococcus microti | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | pasteuri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 2.17 | A | 0 | 0 | 1 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus | pasteuri | Staphylococcus | pasteuri | 2 | Staphylococcus | warneri | 1.81 | A | 1 | 1 | 0 | Staphylococcus pasteuri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-03-02 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.05 | Staphylococcus | pseudintermedius | 1.97 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 1.84 | Staphylococcus | pseudintermedius | 1.82 | B | 0 | 0 | 1 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | saprophyticus | Staphylococcus | kloosii | 1.43 | Staphylococcus | cohnii | 1.26 | B | 0 | 0 | 1 | Staphylococcus kloosii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 503,2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.18 | Staphylococcus | saprophyticus | 2.06 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus | schleiferi | Staphylococcus | schleiferi | 2.15 | Staphylococcus | schleiferi | 2.02 | A | 1 | 1 | 0 | Staphylococcus schleiferi | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus | schweitzeri | Staphylococcus | schweitzeri | 2.17 | Staphylococcus | schweitzeri | 2.17 | A | 1 | 1 | 0 | Staphylococcus schweitzeri | Staphylococcus schweitzeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | simulans | Staphylococcus | simulans | 2.13 | Staphylococcus | simulans | 2.11 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | simulans | Staphylococcus | simulans | 2.17 | Staphylococcus | simulans | 2.04 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2439.2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.26 | Staphylococcus | succinus | 2.23 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 761,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.28 | Staphylococcus | succinus | 2.21 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus | warneri | Staphylococcus | warneri | 2.07 | Staphylococcus | warneri | 2.02 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20316 | 3; type strain | DT | 2013-10-02 | Staphylococcus | warneri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 1.96 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.04 | Staphylococcus | xylosus | 2.02 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.88 | Staphylococcus | xylosus | 1.76 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus | devriesei | Staphylococcus | capitis | 1.46 | Staphylococcus | devriesei | 1.45 | B | 0 | 0 | 1 | Staphylococcus capitis | Staphylococcus devriesei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus | sp-185-U1-2 | Xanthomonas | hortorum | 1.38 | Xanthomonas | hortorum | 1.37 | B | 0 | 0 | 1 | Xanthomonas hortorum | Xanthomonas hortorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus | suis | Aerococcus | suis | 2.3 | Levilactobacillus | brevis | 1.42 | A | 1 | 1 | 0 | Aerococcus suis | Levilactobacillus brevis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus | urinaeequi | Aerococcus | viridans | 1.72 | Aerococcus | viridans | 1.65 | B | 0 | 0 | 1 | Aerococcus viridans | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 2016-09-27 | Aerococcus | viridans | Aerococcus | viridans | 2.04 | Aerococcus | viridans | 1.94 | A | 1 | 1 | 0 | Aerococcus viridans | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter | phocae | Atopobacter | phocae | 2.33 | Atopobacter | phocae | 2.31 | A | 1 | 1 | 0 | Atopobacter phocae | Atopobacter phocae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2442,2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium | divergens | Carnobacterium | divergens | 1.99 | Carnobacterium | divergens | 1.96 | B | 0 | 0 | 1 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium | divergens | Carnobacterium | divergens | 2.37 | Carnobacterium | divergens | 2.27 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6273 | 3; sequenced | eDT | 2016-01-13 | Carnobacterium | maltaromaticum | Carnobacterium | maltaromaticum | 2.55 | Carnobacterium | maltaromaticum | 2.54 | A | 1 | 1 | 0 | Carnobacterium maltaromaticum | Carnobacterium maltaromaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1519,3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella | adiacens | Granulicatella | adiacens | 2.3 | Granulicatella | adiacens | 2.24 | A | 1 | 1 | 0 | Granulicatella adiacens | Granulicatella adiacens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2616 | 3; sequenced | DT | 2013-09-25 | Marinilactibacillus | sp-CVUAS-2616 | Lacticaseibacillus | sharpeae | 1.42 | Staphylococcus | simulans | 1.34 | B | 0 | 0 | 1 | Lacticaseibacillus sharpeae | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8980,3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus | sp-CVUAS-2616 | Enterococcus | cecorum | 1.53 | Ligilactobacillus | murinus | 1.5 | B | 0 | 0 | 1 | Enterococcus cecorum | Ligilactobacillus murinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16 | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.3 | Enterococcus | casseliflavus | 2.2 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus | durans | Enterococcus | hirae | 2.09 | Enterococcus | hirae | 1.8 | A | 0 | 0 | 1 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus | durans | Enterococcus | durans | 2 | Enterococcus | durans | 1.94 | A | 1 | 1 | 0 | Enterococcus durans | Enterococcus durans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2011-07-11 | Enterococcus | faecium | Enterococcus | faecium | 2.09 | Enterococcus | faecium | 2.05 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-23 | Enterococcus | gallinarum | Enterococcus | gallinarum | 2.04 | Enterococcus | gallinarum | 1.98 | A | 1 | 1 | 0 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-23 | Enterococcus | hirae | Enterococcus | hirae | 2.33 | Enterococcus | hirae | 2.16 | A | 1 | 1 | 0 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.33 | Enterococcus | casseliflavus | 2.32 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5030 | 3; sequenced | eDT | 2014-07-23 | Enterococcus | cecorum | Enterococcus | cecorum | 2.34 | Enterococcus | cecorum | 2.29 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8897 | 3; sequenced | eDT | 2014-07-23 | Enterococcus | cecorum | Enterococcus | cecorum | 1.41 | Enterococcus | cecorum | 1.38 | B | 0 | 0 | 1 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus | faecalis | Enterococcus | faecalis | 2.51 | Enterococcus | faecalis | 2.41 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus | faecalis | Enterococcus | faecalis | 2.39 | Enterococcus | faecalis | 2.38 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MR11-207 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus | faecium | Enterococcus | faecium | 2.42 | Enterococcus | faecium | 2.1 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus | gallinarum | Enterococcus | gallinarum | 2.42 | Enterococcus | gallinarum | 2.38 | A | 1 | 1 | 0 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34954 | 3; sequenced | eDT | 2023-11-30 | Enterococcus | hirae | Enterococcus | hirae | 2.39 | Enterococcus | hirae | 2.37 | A | 1 | 1 | 0 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34219 | 3; sequenced | eDT | 2023-11-30 | Enterococcus | mundtii | Enterococcus | mundtii | 2.39 | Enterococcus | mundtii | 2.37 | A | 1 | 1 | 0 | Enterococcus mundtii | Enterococcus mundtii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11490 | 3; sequenced | eDT | 2023-11-30 | Enterococcus | phoeniculicola | Enterococcus | phoeniculicola | 2.13 | Brevundimonas | vesicularis | 1.45 | A | 1 | 1 | 0 | Enterococcus phoeniculicola | Brevundimonas vesicularis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus | solitarius | Tetragenococcus | solitarius | 2.14 | Halomonas | elongata | 1.38 | A | 1 | 1 | 0 | Tetragenococcus solitarius | Halomonas elongata | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1549,2 | 3; sequenced | DT | 2016-09-27 | Vagococcus | salmoninarum | Vagococcus | salmoninarum | 2.09 | Limosilactobacillus | oris | 1.43 | A | 1 | 1 | 0 | Vagococcus salmoninarum | Limosilactobacillus oris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus | acidophilus | Lactobacillus | acidophilus | 2.32 | Lactobacillus | acidophilus | 2.24 | A | 1 | 1 | 0 | Lactobacillus acidophilus | Lactobacillus acidophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2023000188 | 2; corresponds to expected result of PT | EFEx | 2023-09-07 | Lacticaseibacillus | rhamnosus | Lacticaseibacillus | rhamnosus | 2.6 | Lacticaseibacillus | rhamnosus | 2.52 | A | 1 | 1 | 0 | Lacticaseibacillus rhamnosus | Lacticaseibacillus rhamnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus | johnsonii | Lactobacillus | johnsonii | 2.42 | Lactobacillus | johnsonii | 2.35 | A | 1 | 1 | 0 | Lactobacillus johnsonii | Lactobacillus johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32277,2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus | sp-CVUAS-32277,2 | Limosilactobacillus | gorillae | 1.25 | Secundilactobacillus | malefermentans | 1.24 | B | 0 | 0 | 1 | Limosilactobacillus gorillae | Secundilactobacillus malefermentans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus | claussenii | Pediococcus | claussenii | 2.25 | Pediococcus | claussenii | 2.04 | A | 1 | 1 | 0 | Pediococcus claussenii | Pediococcus claussenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1213 | 2; 0 | DT | 2013-07-09 | Leuconostoc | mesenteroides | Leuconostoc | mesenteroides | 2.28 | Leuconostoc | mesenteroides | 2.09 | A | 1 | 1 | 0 | Leuconostoc mesenteroides | Leuconostoc mesenteroides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20622 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | acidominimus | Streptococcus | acidominimus | 1.89 | Streptococcus | suis | 1.78 | B | 0 | 0 | 1 | Streptococcus acidominimus | Streptococcus suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20728 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | alactolyticus | Streptococcus | alactolyticus | 2.04 | Streptococcus | alactolyticus | 2.04 | A | 1 | 1 | 0 | Streptococcus alactolyticus | Streptococcus alactolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20563 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | anginosus | Streptococcus | anginosus | 2.47 | Streptococcus | anginosus | 2.21 | A | 1 | 1 | 0 | Streptococcus anginosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25818 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | anginosus | Streptococcus | anginosus | 2.13 | Streptococcus | anginosus | 2.08 | A | 1 | 1 | 0 | Streptococcus anginosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15627 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | australis | Streptococcus | australis | 2.09 | Streptococcus | parasanguinis | 1.88 | A | 1 | 1 | 0 | Streptococcus australis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 103678 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | azizii | Streptococcus | pneumoniae | 1.64 | Streptococcus | cristatus | 1.6 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 69277 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | bovimastitidis | Streptococcus | penaeicida | 1.89 | Streptococcus | uberis | 1.69 | B | 0 | 0 | 1 | Streptococcus penaeicida | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19004 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | caballi | Streptococcus | caballi | 2.44 | Streptococcus | mitis_oralis | 1.53 | A | 1 | 1 | 0 | Streptococcus caballi | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 73951 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | caledonicus | Streptococcus | porci | 1.58 | Streptococcus | ratti | 1.53 | B | 0 | 0 | 1 | Streptococcus porci | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105751 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | cameli | Streptococcus | pneumoniae | 1.53 | Streptococcus | pneumoniae | 1.52 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 67170 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | caprae | Streptococcus | pyogenes | 1.51 | Streptococcus | alactolyticus | 1.5 | B | 0 | 0 | 1 | Streptococcus pyogenes | Streptococcus alactolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102819 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | caviae | Streptococcus | devriesei | 2.11 | Streptococcus | devriesei | 2.07 | A | 0 | 0 | 1 | Streptococcus devriesei | Streptococcus devriesei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 106182 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | chenjunshii | Streptococcus | canis | 1.59 | Streptococcus | ferus | 1.55 | B | 0 | 0 | 1 | Streptococcus canis | Streptococcus ferus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14507 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | constellatus | Streptococcus | constellatus | 2.31 | Streptococcus | constellatus | 2.24 | A | 1 | 1 | 0 | Streptococcus constellatus | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17475 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | constellatus | Streptococcus | constellatus | 2.52 | Streptococcus | constellatus | 1.99 | A | 1 | 1 | 0 | Streptococcus constellatus | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25819 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | constellatus | Streptococcus | constellatus | 2.32 | Streptococcus | constellatus | 2.1 | A | 1 | 1 | 0 | Streptococcus constellatus | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20562 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | criceti | Streptococcus | criceti | 2.51 | Streptococcus | criceti | 2.35 | A | 1 | 1 | 0 | Streptococcus criceti | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8249 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | cristatus | Streptococcus | cristatus | 2.36 | Streptococcus | cristatus | 2.35 | A | 1 | 1 | 0 | Streptococcus cristatus | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 65085 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | cuniculi | Streptococcus | suis | 1.7 | Streptococcus | pneumoniae | 1.6 | B | 0 | 0 | 1 | Streptococcus suis | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 66496 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | cuniculipharyngis | Pseudomonas | mucidolens | 1.54 | Pseudomonas | migulae | 1.51 | B | 0 | 0 | 1 | Pseudomonas mucidolens | Pseudomonas migulae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22233 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | danieliae | Streptococcus | pneumoniae | 1.53 | Streptococcus | pneumoniae | 1.49 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21999 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | dentapri | Streptococcus | canis | 1.61 | Streptococcus | equi_ssp_zooepidemicus | 1.55 | B | 0 | 0 | 1 | Streptococcus canis | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27381 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | dentiloxodontae | Streptococcus | equi_ssp_zooepidemicus | 1.6 | Streptococcus | salivarius | 1.59 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18963 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | dentirousetti | Streptococcus | dentirousetti | 2.15 | Streptococcus | downei | 2.14 | B | 0 | 0 | 1 | Streptococcus dentirousetti | Streptococcus downei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25137 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | dentisani | Streptococcus | downei | 1.76 | Streptococcus | criceti | 1.63 | B | 0 | 0 | 1 | Streptococcus downei | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 47155 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | devriesei | Streptococcus | devriesei | 2.59 | Streptococcus | devriesei | 2.41 | A | 1 | 1 | 0 | Streptococcus devriesei | Streptococcus devriesei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15616 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | didelphis | Streptococcus | didelphis | 2.28 | Streptococcus | pyogenes | 1.59 | A | 1 | 1 | 0 | Streptococcus didelphis | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5635 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | downei | Streptococcus | downei | 2.4 | Streptococcus | dentirousetti | 1.9 | A | 1 | 1 | 0 | Streptococcus downei | Streptococcus dentirousetti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 73139 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | downii | Streptococcus | mitis_oralis | 2.46 | Streptococcus | mitis_oralis | 2.38 | A | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14446 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | entericus | Streptococcus | entericus | 2.36 | Streptococcus | pyogenes | 1.42 | A | 1 | 1 | 0 | Streptococcus entericus | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20558 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | equinus | Streptococcus | equinus | 2.31 | Streptococcus | equinus | 2.21 | A | 1 | 1 | 0 | Streptococcus equinus | Streptococcus equinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20646 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ferus | Streptococcus | ferus | 2.45 | Streptococcus | ferus | 2.24 | A | 1 | 1 | 0 | Streptococcus ferus | Streptococcus ferus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15349 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | gallinaceus | Streptococcus | gallinaceus | 2.39 | Streptococcus | mitis_oralis | 1.51 | A | 1 | 1 | 0 | Streptococcus gallinaceus | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16831 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | gallolyticus | Streptococcus | gallolyticus | 2.39 | Streptococcus | gallolyticus | 2.25 | A | 1 | 1 | 0 | Streptococcus gallolyticus | Streptococcus gallolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15879 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | gallolyticus | Streptococcus | gallolyticus | 2.27 | Streptococcus | gallolyticus | 2.1 | A | 1 | 1 | 0 | Streptococcus gallolyticus | Streptococcus gallolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15351 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | gallolyticus | Streptococcus | gallolyticus | 2.42 | Streptococcus | gallolyticus | 2.12 | A | 1 | 1 | 0 | Streptococcus gallolyticus | Streptococcus gallolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6777 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | gordonii | Streptococcus | gordonii | 2.22 | Streptococcus | gordonii | 2.22 | A | 1 | 1 | 0 | Streptococcus gordonii | Streptococcus gordonii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17028 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | halichoeri | Streptococcus | halichoeri | 2.4 | Streptococcus | equi_ssp_zooepidemicus | 1.75 | A | 1 | 1 | 0 | Streptococcus halichoeri | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101996 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | halotolerans | Streptococcus | pluranimalium | 1.69 | Streptococcus | pluranimalium | 1.67 | B | 0 | 0 | 1 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19005 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | henryi | Streptococcus | henryi | 2.43 | Streptococcus | dysgalactiae | 1.53 | A | 1 | 1 | 0 | Streptococcus henryi | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 107951 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | hillyeri | Streptococcus | agalactiae | 1.53 | Streptococcus | dysgalactiae | 1.49 | B | 0 | 0 | 1 | Streptococcus agalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101997 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | himalayensis | Streptococcus | sanguinis | 1.62 | Pseudomonas | luteola | 1.6 | B | 0 | 0 | 1 | Streptococcus sanguinis | Pseudomonas luteola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26014 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | hongkongensis | Streptococcus | equi_ssp_zooepidemicus | 1.84 | Streptococcus | agalactiae | 1.84 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20770 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | hyointestinalis | Streptococcus | hyointestinalis | 2.56 | Streptococcus | hyointestinalis | 2.3 | A | 1 | 1 | 0 | Streptococcus hyointestinalis | Streptococcus hyointestinalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12219 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | hyovaginalis | Streptococcus | hyovaginalis | 2.43 | Streptococcus | hyovaginalis | 2.3 | A | 1 | 1 | 0 | Streptococcus hyovaginalis | Streptococcus hyovaginalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 52536 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ictaluri | Streptococcus | canis | 1.71 | Streptococcus | castoreus | 1.7 | B | 0 | 0 | 1 | Streptococcus canis | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22960 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | infantarius | Streptococcus | infantarius | 2.03 | Streptococcus | infantarius | 1.87 | A | 1 | 1 | 0 | Streptococcus infantarius | Streptococcus infantarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 39817 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | infantis | Streptococcus | infantis | 2.3 | Streptococcus | infantis | 1.92 | A | 1 | 1 | 0 | Streptococcus infantis | Streptococcus infantis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20576 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | iniae | Streptococcus | iniae | 2.27 | Streptococcus | iniae | 2.19 | A | 1 | 1 | 0 | Streptococcus iniae | Streptococcus iniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20573 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | intermedius | Streptococcus | intermedius | 2.14 | Streptococcus | intermedius | 2.05 | A | 1 | 1 | 0 | Streptococcus intermedius | Streptococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23027 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | lactarius | Streptococcus | parasanguinis | 2.07 | Streptococcus | parasanguinis | 1.94 | A | 0 | 0 | 1 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27382 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | loxodontisalivarius | Streptococcus | hyointestinalis | 1.65 | Streptococcus | infantarius | 1.64 | B | 0 | 0 | 1 | Streptococcus hyointestinalis | Streptococcus infantarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15350 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | lutetiensis | Streptococcus | lutetiensis | 2.39 | Streptococcus | infantarius | 2.13 | B | 0 | 0 | 1 | Streptococcus lutetiensis | Streptococcus infantarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20724 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | macacae | Streptococcus | macacae | 2.25 | Streptococcus | ratti | 1.51 | A | 1 | 1 | 0 | Streptococcus macacae | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18627 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | marimammalium | Streptococcus | marimammalium | 2.03 | Streptococcus | pneumoniae | 1.55 | A | 1 | 1 | 0 | Streptococcus marimammalium | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101995 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | marmotae | Streptococcus | suis | 1.76 | Streptococcus | suis | 1.74 | B | 0 | 0 | 1 | Streptococcus suis | Streptococcus suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 49690 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | massiliensis | Streptococcus | massiliensis | 2.28 | Streptococcus | massiliensis | 2.17 | A | 1 | 1 | 0 | Streptococcus massiliensis | Streptococcus massiliensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19192 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | merionis | Streptococcus | sp | 2.55 | Streptococcus | merionis | 2.43 | B | 0 | 0 | 1 | Streptococcus sp | Streptococcus merionis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17118 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | minor | Streptococcus | minor | 2.35 | Streptococcus | minor | 2.27 | A | 1 | 1 | 0 | Streptococcus minor | Streptococcus minor | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12643 EFEx | 3; type strain | EFEx | 2023-09-05 | Streptococcus | mitis | Streptococcus | mitis_oralis | 2.25 | Streptococcus | pneumoniae | 2.23 | B | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105143 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | moroccensis | Streptococcus | moroccensis | 2.13 | Streptococcus | sanguinis | 1.43 | A | 1 | 1 | 0 | Streptococcus moroccensis | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20523 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | mutans | Streptococcus | mutans | 2.27 | Streptococcus | mutans | 2.14 | A | 1 | 1 | 0 | Streptococcus mutans | Streptococcus mutans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27088 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oralis | Streptococcus | mitis_oralis | 2.38 | Streptococcus | mitis_oralis | 2.35 | A | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20627 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | oralis | Streptococcus | mitis_oralis | 2.24 | Streptococcus | mitis_oralis | 2.23 | A | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24864 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oralis | Streptococcus | pneumoniae | 2.34 | Streptococcus | mitis_oralis | 2.32 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 100101 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oricebi | Streptococcus | sanguinis | 1.64 | Streptococcus | constellatus | 1.62 | B | 0 | 0 | 1 | Streptococcus sanguinis | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27377 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oriloxodontae | Streptococcus | criceti | 2.36 | Streptococcus | criceti | 2.28 | A | 0 | 0 | 1 | Streptococcus criceti | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25193 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | orisasini | Streptococcus | ratti | 1.97 | Streptococcus | ratti | 1.96 | B | 0 | 0 | 1 | Streptococcus ratti | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 43577 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | orisratti | Streptococcus | orisratti | 2.32 | Streptococcus | equi_ssp_zooepidemicus | 1.78 | A | 1 | 1 | 0 | Streptococcus orisratti | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18307 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | orisuis | Streptococcus | criceti | 2.45 | Streptococcus | criceti | 2.43 | A | 0 | 0 | 1 | Streptococcus criceti | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 39485 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | ovis | Streptococcus | ovis | 1.86 | Streptococcus | ovis | 1.72 | B | 0 | 0 | 1 | Streptococcus ovis | Streptococcus ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 69612 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | ovuberis | Streptococcus | gallolyticus | 1.55 | Streptococcus | cristatus | 1.48 | B | 0 | 0 | 1 | Streptococcus gallolyticus | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 111148 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | pacificus | Paucilactobacillus | oligofermentans | 1.45 | Streptococcus | gallolyticus | 1.36 | B | 0 | 0 | 1 | Paucilactobacillus oligofermentans | Streptococcus gallolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29921 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | panodentis | Streptococcus | pneumoniae | 1.62 | Streptococcus | sanguinis | 1.58 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102135 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | pantholopis | Streptococcus | constellatus | 1.49 | Streptococcus | cristatus | 1.48 | B | 0 | 0 | 1 | Streptococcus constellatus | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6778 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | parasanguinis | Streptococcus | parasanguinis | 2.27 | Streptococcus | parasanguinis | 2.25 | A | 1 | 1 | 0 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29126 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | parasuis | Streptococcus | suis | 1.93 | Streptococcus | suis | 1.73 | B | 0 | 0 | 1 | Streptococcus suis | Streptococcus suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6631 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | parauberis | Streptococcus | parauberis | 2.33 | Streptococcus | parauberis | 2.01 | A | 1 | 1 | 0 | Streptococcus parauberis | Streptococcus parauberis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26545 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | penaeicida | Streptococcus | penaeicida | 2.41 | Streptococcus | canis | 1.42 | A | 1 | 1 | 0 | Streptococcus penaeicida | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 39814 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | peroris | Streptococcus | mitis_oralis | 1.78 | Streptococcus | mitis_oralis | 1.74 | B | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 16735 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | phocae | Streptococcus | phocae | 2.15 | Streptococcus | phocae | 2.03 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24768 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | phocae | Streptococcus | phocae | 2.28 | Streptococcus | phocae | 2.2 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14177 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | pluranimalium | Streptococcus | pluranimalium | 2.15 | Streptococcus | pluranimalium | 2.05 | A | 1 | 1 | 0 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22810 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | plurextorum | Streptococcus | plurextorum | 2.4 | Streptococcus | merionis | 1.54 | A | 1 | 1 | 0 | Streptococcus plurextorum | Streptococcus merionis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23759 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | porci | Streptococcus | porci | 2.33 | Streptococcus | porci | 1.92 | A | 1 | 1 | 0 | Streptococcus porci | Streptococcus porci | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20725 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | porcinus | Streptococcus | porcinus | 2.51 | Streptococcus | porcinus | 2.47 | A | 1 | 1 | 0 | Streptococcus porcinus | Streptococcus porcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28302 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | porcorum | Streptococcus | porcorum | 2.29 | Streptococcus | pyogenes | 1.66 | A | 1 | 1 | 0 | Streptococcus porcorum | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 49455 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | pseudopneumoniae | Streptococcus | pseudopneumoniae | 2.26 | Streptococcus | pseudopneumoniae | 2.23 | A | 1 | 1 | 0 | Streptococcus pseudopneumoniae | Streptococcus pseudopneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18513 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | pseudoporcinus | Streptococcus | pseudoporcinus | 2.15 | Streptococcus | pseudoporcinus | 2.08 | A | 1 | 1 | 0 | Streptococcus pseudoporcinus | Streptococcus pseudoporcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20564 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | ratti | Streptococcus | ratti | 2.54 | Streptococcus | ratti | 2.36 | A | 1 | 1 | 0 | Streptococcus ratti | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101998 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | respiraculi | Streptococcus | suis | 1.78 | Streptococcus | suis | 1.75 | B | 0 | 0 | 1 | Streptococcus suis | Streptococcus suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105174 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | rifensis | Streptococcus | moroccensis | 1.83 | Streptococcus | downei | 1.43 | B | 0 | 0 | 1 | Streptococcus moroccensis | Streptococcus downei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26920 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | rubneri | Streptococcus | parasanguinis | 2.13 | Streptococcus | parasanguinis | 1.95 | A | 0 | 0 | 1 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 104980 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ruminantium | Streptococcus | suis | 1.64 | Streptococcus | mitis_oralis | 1.63 | B | 0 | 0 | 1 | Streptococcus suis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28303 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | rupicaprae | Streptococcus | gallolyticus | 1.4 | Weissella | minor | 1.31 | B | 0 | 0 | 1 | Streptococcus gallolyticus | Weissella minor | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20560 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | salivarius | Streptococcus | salivarius | 2.4 | Streptococcus | salivarius | 2.19 | A | 1 | 1 | 0 | Streptococcus salivarius | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20617 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | salivarius | Streptococcus | thermophilus | 2.43 | Streptococcus | thermophilus | 2.35 | A | 0 | 0 | 1 | Streptococcus thermophilus | Streptococcus thermophilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27513 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | saliviloxodontae | Streptococcus | equi_ssp_zooepidemicus | 1.88 | Streptococcus | equi_ssp_zooepidemicus | 1.68 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20567 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | sanguinis | Streptococcus | sanguinis | 2.31 | Streptococcus | sanguinis | 2.22 | A | 1 | 1 | 0 | Streptococcus sanguinis | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14990 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | sinensis | Streptococcus | sinensis | 2.14 | Streptococcus | cristatus | 2.11 | B | 0 | 0 | 1 | Streptococcus sinensis | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33478 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | sobrinus | Streptococcus | sobrinus | 2.45 | Streptococcus | sobrinus | 2.41 | A | 1 | 1 | 0 | Streptococcus sobrinus | Streptococcus sobrinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9682 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | suis | Streptococcus | suis | 2.15 | Streptococcus | suis | 2.13 | A | 1 | 1 | 0 | Streptococcus suis | Streptococcus suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105173 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | tangierensis | Streptococcus | equi_ssp_zooepidemicus | 1.46 | Loigolactobacillus | coryniformis | 1.45 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Loigolactobacillus coryniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 32906 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | thoraltensis | Streptococcus | thoraltensis | 2.35 | Streptococcus | thoraltensis | 2.29 | A | 1 | 1 | 0 | Streptococcus thoraltensis | Streptococcus thoraltensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25324 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | troglodytae | Streptococcus | mutans | 2.18 | Streptococcus | mutans | 2.14 | A | 0 | 0 | 1 | Streptococcus mutans | Streptococcus mutans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KCTC 33006 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | troglodytidis | Streptococcus | constellatus | 1.73 | Streptococcus | anginosus | 1.66 | B | 0 | 0 | 1 | Streptococcus constellatus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 41590 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | urinalis | Streptococcus | urinalis | 2.43 | Streptococcus | urinalis | 2.23 | A | 1 | 1 | 0 | Streptococcus urinalis | Streptococcus urinalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22768 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ursoris | Streptococcus | ratti | 2.17 | Streptococcus | ratti | 2.1 | A | 0 | 0 | 1 | Streptococcus ratti | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5636 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | vestibularis | Streptococcus | vestibularis | 2.39 | Streptococcus | vestibularis | 2.08 | A | 1 | 1 | 0 | Streptococcus vestibularis | Streptococcus vestibularis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 110741 | 3; type strain; WGS (PMID: 33999792) | EFEx | 2023-06-20 | Streptococcus | vicugnae | Streptococcus | equinus | 1.87 | Streptococcus | lutetiensis | 1.84 | B | 0 | 0 | 1 | Streptococcus equinus | Streptococcus lutetiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 110742 | 3; type strain, WGS (PMID: 33999792) | EFEx | 2023-06-20 | Streptococcus | zalophi | Lacticaseibacillus | paracasei | 1.42 | Streptococcus | entericus | 1.38 | B | 0 | 0 | 1 | Lacticaseibacillus paracasei | Streptococcus entericus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1146 | 2; 0 | DT | 2013-07-09 | Lactococcus | lactis | Lactococcus | lactis | 2.42 | Lactococcus | lactis | 2.4 | A | 1 | 1 | 0 | Lactococcus lactis | Lactococcus lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.45 | Streptococcus | agalactiae | 2.44 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6784 | 3; public strain collection | eDT | 2019-07-09 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.49 | Streptococcus | agalactiae | 2.48 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus | canis | Streptococcus | canis | 2.53 | Streptococcus | canis | 2.35 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL Mi-845-17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus | canis | Streptococcus | canis | 2.47 | Streptococcus | canis | 2.45 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.19 | Streptococcus | dysgalactiae | 2.04 | B | 0 | 0 | 1 | Streptococcus castoreus | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17536 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | castoreus | Streptococcus | castoreus | 2.35 | Streptococcus | pyogenes | 1.96 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus | catagoni | Streptococcus | uberis | 1.74 | Streptococcus | phocae | 1.69 | B | 0 | 0 | 1 | Streptococcus uberis | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 99-1-2017 | 3; type strain | EFEx | 2018-05-03 | Streptococcus | catagoni | Streptococcus | uberis | 1.81 | Streptococcus | uberis | 1.71 | B | 0 | 0 | 1 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.14 | Streptococcus | canis | 2.07 | B | 0 | 0 | 1 | Streptococcus dysgalactiae | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.26 | Streptococcus | dysgalactiae | 1.98 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus | dysgalactiae | Streptococcus | canis | 2.12 | Streptococcus | dysgalactiae | 2.07 | B | 0 | 0 | 1 | Streptococcus canis | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.11 | Streptococcus | dysgalactiae | 2.1 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32462,2 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi_ssp_equi | 2.12 | Streptococcus | equi_ssp_equi | 2.06 | A | 0 | 0 | 1 | Streptococcus equi_ssp_equi | Streptococcus equi_ssp_equi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32770 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi_ssp_equi | 2.29 | Streptococcus | equi_ssp_equi | 2.24 | A | 0 | 0 | 1 | Streptococcus equi_ssp_equi | Streptococcus equi_ssp_equi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus | equi | Streptococcus | equi_ssp_zooepidemicus | 2.34 | Streptococcus | equi_ssp_zooepidemicus | 2.34 | A | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus | equi | Streptococcus | equi_ssp_zooepidemicus | 2.21 | Streptococcus | equi_ssp_zooepidemicus | 2.18 | A | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus | iniae | Streptococcus | iniae | 1.89 | Streptococcus | iniae | 1.74 | B | 0 | 0 | 1 | Streptococcus iniae | Streptococcus iniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12643 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 11.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus agalactiae DSM 2134, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-02-18 | Streptococcus | mitis | Streptococcus | pneumoniae | 2.14 | Streptococcus | pneumoniae | 2.06 | A | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus | orisasini | Streptococcus | ratti | 2.02 | Streptococcus | ratti | 1.97 | A | 0 | 0 | 1 | Streptococcus ratti | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 9056 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 11.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus agalactiae DSM 2134, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-09-01 | Streptococcus | parasanguinis | Streptococcus | parasanguinis | 2.3 | Streptococcus | parasanguinis | 2.26 | A | 1 | 1 | 0 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | parauberis | Streptococcus | parauberis | 2.1 | Streptococcus | parauberis | 1.9 | A | 1 | 1 | 0 | Streptococcus parauberis | Streptococcus parauberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | phocae | Streptococcus | phocae | 2.18 | Streptococcus | phocae | 1.97 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 12-13862-001-01 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus | pluranimalium | Streptococcus | pluranimalium | 2.2 | Streptococcus | pluranimalium | 2.18 | A | 1 | 1 | 0 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.48 | Streptococcus | pneumoniae | 2.44 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.16 | Streptococcus | pneumoniae | 2.16 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.29 | Streptococcus | pneumoniae | 2.23 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.49 | Streptococcus | pneumoniae | 2.46 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus | porcinus | Streptococcus | porcinus | 2.1 | Streptococcus | porcinus | 2.09 | A | 1 | 1 | 0 | Streptococcus porcinus | Streptococcus porcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 1567 | 2; strain collection of the RKI | #CON=approx. 3 loops per Eppend. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, Inactivation date: 19.08.2021¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus pneumoniae B2046, preparation: C. Hartl, measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-09-02 | Streptococcus | pyogenes | Streptococcus | pyogenes | 2.5 | Streptococcus | pyogenes | 2.5 | A | 1 | 1 | 0 | Streptococcus pyogenes | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-04 | Streptococcus | pyogenes | Streptococcus | pyogenes | 2.43 | Streptococcus | pyogenes | 2.4 | A | 1 | 1 | 0 | Streptococcus pyogenes | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10145,3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus | salivarius | Streptococcus | salivarius | 2.13 | Streptococcus | salivarius | 2.13 | A | 1 | 1 | 0 | Streptococcus salivarius | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20068 | 3; public strain collection | TFAextr | 2022-02-18 | Streptococcus | sanguinis | Streptococcus | sanguinis | 1.96 | Streptococcus | sanguinis | 1.75 | B | 0 | 0 | 1 | Streptococcus sanguinis | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus | sp-CVUAS-32349 | Streptococcus | dysgalactiae | 1.52 | Streptococcus | hyovaginalis | 1.48 | B | 0 | 0 | 1 | Streptococcus dysgalactiae | Streptococcus hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2260,4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus | troglodytidis | Streptococcus | anginosus | 1.57 | Streptococcus | anginosus | 1.51 | B | 0 | 0 | 1 | Streptococcus anginosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.03 | Streptococcus | uberis | 1.81 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.23 | Streptococcus | uberis | 2.18 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 1.94 | Clostridioides | difficile | 1.71 | B | 0 | 0 | 1 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.17 | Clostridioides | difficile | 2.06 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 2.05 | Clostridium | sporogenes | 1.9 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 1.88 | Clostridium | sporogenes | 1.78 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 2.01 | Clostridium | sporogenes | 1.93 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 1.59 | Clostridium | sporogenes | 1.51 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 1.67 | Clostridium | sporogenes | 1.59 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Arthrobacter | ramosus | 1.51 | Aeromonas | caviae | 1.36 | B | 0 | 0 | 1 | Arthrobacter ramosus | Aeromonas caviae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium | botulinum | Bacillus | subtilis | 1.49 | Clostridium | cochlearium | 1.42 | B | 0 | 0 | 1 | Bacillus subtilis | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium | chauvoei | Clostridium | chauvoei | 2.38 | Clostridium | chauvoei | 2.17 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 849,2 | 3; sequenced | DT | 2014-12-17 | Clostridium | isatidis | Clostridium | isatidis | 1.6 | Pseudomonas | boreopolis | 1.34 | B | 0 | 0 | 1 | Clostridium isatidis | Pseudomonas boreopolis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108256 | 3; Public strain collection | EFEx | 0 | Clostridium | leptum | Clostridium | paraputrificum | 1.43 | Staphylococcus | simulans | 1.4 | B | 0 | 0 | 1 | Clostridium paraputrificum | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.23.22 | 3; sequenced [0207] | EFEx | 0 | Clostridium | nexile | Aeromonas | caviae | 1.24 | Pseudomonas | agarici | 1.24 | B | 0 | 0 | 1 | Aeromonas caviae | Pseudomonas agarici | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium | nigeriense | Clostridium | tertium | 1.75 | Clostridium | tertium | 1.68 | B | 0 | 0 | 1 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium | perfringens | Clostridium | perfringens | 2.09 | Clostridium | perfringens | 1.9 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B | 3; case rerort | DT | 2013-08-06 | Clostridium | perfringens | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.4 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 10419 | 3; Public strain collection | eDT | 2021-10-26 | Clostridium | scindens | Clostridium | scindens | 2.23 | Clostridium | scindens | 2.15 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.1.26 | 3; sequenced [0207] | eDT | 0 | Clostridium | scindens | Clostridium | scindens | 2.2 | Clostridium | scindens | 2.06 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium | septicum | Clostridium | septicum | 2.27 | Clostridium | septicum | 2.23 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34091 | 3; sequenced | EFex | 2022-08-05 | Clostridium | sp-CVUAS-34091 | Clostridium | sartagoforme | 1.73 | Clostridium | tertium | 1.62 | B | 0 | 0 | 1 | Clostridium sartagoforme | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium | sp-DSM-1985 | Campylobacter | jejuni | 1.33 | Campylobacter | jejuni | 1.32 | B | 0 | 0 | 1 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium | sporogenes | Clostridium | sporogenes | 1.95 | Clostridium | sporogenes | 1.79 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.1.9 | 3; sequenced [0207] | eDT | 0 | Clostridium | symbiosum | Clostridium | symbiosum | 2.43 | Clostridium | symbiosum | 1.96 | A | 1 | 1 | 0 | Clostridium symbiosum | Clostridium symbiosum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WAL-14163 | 3; sequenced [0207] | eDT | 0 | Clostridium | symbiosum | Clostridium | symbiosum | 2.45 | Clostridium | symbiosum | 2.07 | A | 1 | 1 | 0 | Clostridium symbiosum | Clostridium symbiosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium | tertium | Clostridium | tertium | 2.33 | Clostridium | tertium | 2.28 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium | tetani | Clostridium | tetani | 2.55 | Clostridium | tetani | 2.47 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16296 | 3; Public strain collection | EFEx | 0 | Eubacterium | ramulus | Lacticaseibacillus | paracasei | 1.38 | Terrimonas | ferruginea | 1.33 | B | 0 | 0 | 1 | Lacticaseibacillus paracasei | Terrimonas ferruginea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H258 | 3; sequenced [0207] | eDT | 0 | Agathobacter | rectalis | Pseudomonas | putida | 1.45 | Pseudomonas | monteilii | 1.38 | B | 0 | 0 | 1 | Pseudomonas putida | Pseudomonas monteilii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.9.13 | 3; sequenced [0207] | eDT | 0 | Agathobacter | rectalis | Pseudomonas | putida | 1.39 | Lactococcus | plantarum | 1.31 | B | 0 | 0 | 1 | Pseudomonas putida | Lactococcus plantarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Clinical isolate | 3; sequenced [0207] | EFEx | 0 | Anaerobutyricum | hallii | Staphylococcus | simulans | 1.2 | Bacillus | subtilis | 1.16 | B | 0 | 0 | 1 | Staphylococcus simulans | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H53 | 3; sequenced [0207] | EFEx | 0 | Anaerostipes | hadrus | Arthrobacter | parietis | 1.22 | Delftia | tsuruhatensis | 1.22 | B | 0 | 0 | 1 | Arthrobacter parietis | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29492 | 3; type strain | EFEx | 0 | Blautia | caecimuris | Rhizorhabdus | wittichii | 1.27 | Enterococcus | faecium | 1.27 | B | 0 | 0 | 1 | Rhizorhabdus wittichii | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.18.51 | 3; sequenced [0207] | EFEx | 0 | Blautia | faecis | Terrimonas | ferruginea | 1.29 | Limosilactobacillus | antri | 1.21 | B | 0 | 0 | 1 | Terrimonas ferruginea | Limosilactobacillus antri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22028 | 3; type strain | EFEx | 0 | Blautia | glucerasea | Neisseria | oralis | 1.29 | Parvimonas | micra | 1.28 | B | 0 | 0 | 1 | Neisseria oralis | Parvimonas micra | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20583 | 3; type strain | EFEx | 0 | Blautia | hansenii | Aeromonas | molluscorum | 1.4 | Staphylococcus | simiae | 1.34 | B | 0 | 0 | 1 | Aeromonas molluscorum | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.19.60 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Halococcus | morrhuae | 1.29 | Acinetobacter | johnsonii | 1.28 | B | 0 | 0 | 1 | Halococcus morrhuae | Acinetobacter johnsonii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.9.19 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Weissella | viridescens | 1.44 | Aeromonas | eucrenophila | 1.22 | B | 0 | 0 | 1 | Weissella viridescens | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.20.66 | 3; sequenced [0207] | eDT | 0 | Blautia | obeum | Streptomyces | phaeochromogenes | 1.46 | Aeromonas | eucrenophila | 1.43 | B | 0 | 0 | 1 | Streptomyces phaeochromogenes | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9990 | 3; type strain | eDT | 0 | Blautia | producta | Blautia | coccoides | 2.03 | Lactobacillus | paraplantarum | 1.29 | A | 0 | 0 | 1 | Blautia coccoides | Lactobacillus paraplantarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.4.13 | 2; data not listed [0207] | eDT | 0 | Blautia | producta | Clostridium | haemolyticum | 1.43 | Clostridium | baratii | 1.4 | B | 0 | 0 | 1 | Clostridium haemolyticum | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26115 | 3; Public strain collection | EFEx | 0 | Blautia | pseudococcoides | Blautia | coccoides | 1.61 | Acinetobacter | tjernbergiae | 1.28 | B | 0 | 0 | 1 | Blautia coccoides | Acinetobacter tjernbergiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.6.27 | 3; sequenced [0207] | EFEx | 2021-10-20 | Blautia | schinkii | Flavobacterium | pectinovorum | 1.28 | Staphylococcus | capitis | 1.27 | B | 0 | 0 | 1 | Flavobacterium pectinovorum | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.6.32 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia | schinkii | Paracoccus | denitrificans | 1.34 | Staphylococcus | epidermidis | 1.31 | B | 0 | 0 | 1 | Paracoccus denitrificans | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.18.16 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia | wexlerae | Staphylococcus | sciuri | 1.29 | Achromobacter | spanius | 1.28 | B | 0 | 0 | 1 | Staphylococcus sciuri | Achromobacter spanius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31264 | 3; Public strain collection | eDT | 2021-10-20 | Coprococcus | comes | Clostridium | baratii | 1.28 | Staphylococcus | epidermidis | 1.25 | B | 0 | 0 | 1 | Clostridium baratii | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.17.21 | 3; sequenced [0207] | eDT | 2021-10-20 | Coprococcus | comes | Clostridium | baratii | 1.38 | Listeria | ivanovii | 1.35 | B | 0 | 0 | 1 | Clostridium baratii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.10.6 | 3; sequenced [0207] | EFEx | 2021-10-20 | Coprococcus | eutactus | Staphylococcus | schleiferi | 1.4 | Pseudomonas | thermotolerans | 1.33 | B | 0 | 0 | 1 | Staphylococcus schleiferi | Pseudomonas thermotolerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31256 | 3; Public strain collection | eDT | 0 | Dorea | formicigenerans | Dorea | formicigenerans | 1.56 | Veillonella | atypica | 1.25 | B | 0 | 0 | 1 | Dorea formicigenerans | Veillonella atypica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.19.39 | 3; sequenced [0207] | eDT | 0 | Dorea | formicigenerans | Dorea | formicigenerans | 1.52 | Lacticaseibacillus | paracasei | 1.28 | B | 0 | 0 | 1 | Dorea formicigenerans | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.18.49 | 3; sequenced [0207] | eDT | 0 | Dorea | longicatena | Lactobacillus | helveticus | 1.37 | Actinomyces | naeslundii | 1.36 | B | 0 | 0 | 1 | Lactobacillus helveticus | Actinomyces naeslundii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108191 | 3; Public strain collection | EFEx | 0 | Enterocloster | clostridioformis | Enterocloster | clostridioformis | 2.34 | Enterocloster | clostridioformis | 2.32 | A | 1 | 1 | 0 | Enterocloster clostridioformis | Enterocloster clostridioformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.20.55 | 3; sequenced [0207] | eDT | 0 | Faecalicatena | fissicatena | Stenotrophomonas | maltophilia | 1.35 | Clostridium | novyi | 1.29 | B | 0 | 0 | 1 | Stenotrophomonas maltophilia | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.9.3 | 3; sequenced [0207] | eDT | 0 | Faecalicatena | fissicatena | Metabacillus | indicus | 1.25 | Paracoccus | denitrificans | 1.24 | B | 0 | 0 | 1 | Metabacillus indicus | Paracoccus denitrificans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.14.54 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter | saccharivorans | Serratia | rubidaea | 1.35 | Terrimonas | ferruginea | 1.34 | B | 0 | 0 | 1 | Serratia rubidaea | Terrimonas ferruginea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.15.44 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter | saccharivorans | Pseudomonas | antarctica | 1.38 | Terrimonas | ferruginea | 1.37 | B | 0 | 0 | 1 | Pseudomonas antarctica | Terrimonas ferruginea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H260 | 3; sequenced [0207] | EFEx | 0 | Lachnospira | eligens | Bordetella | pertussis | 1.27 | Staphylococcus | auricularis | 1.21 | B | 0 | 0 | 1 | Bordetella pertussis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105336 | 3; Public strain collection | EFEx | 0 | Lacrimispora | celerecrescens | Lacrimispora | celerecrescens | 2.33 | Lacrimispora | sphenoides | 2.18 | B | 0 | 0 | 1 | Lacrimispora celerecrescens | Lacrimispora sphenoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H164 | 3; sequenced [0207] | EFEx | 0 | Roseburia | faecis | Staphylococcus | ureilyticus | 1.42 | Staphylococcus | ureilyticus | 1.33 | B | 0 | 0 | 1 | Staphylococcus ureilyticus | Staphylococcus ureilyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16839 | 3; type strain | EFEx | 0 | Roseburia | hominis | Limosilactobacillus | mucosae | 1.4 | Neisseria | gonorrhoeae | 1.37 | B | 0 | 0 | 1 | Limosilactobacillus mucosae | Neisseria gonorrhoeae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108070 | 3; Public strain collection | EFEx | 0 | Roseburia | inulinivorans | Bacillus | pumilus | 1.49 | Enterocloster | clostridioformis | 1.4 | B | 0 | 0 | 1 | Bacillus pumilus | Enterocloster clostridioformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.1.25 | 3; sequenced [0207] | eDT | 0 | Sellimonas | intestinalis | Magnusiomyces | capitatus | 1.4 | Hungatella | hathewayi | 1.37 | B | 0 | 0 | 1 | Magnusiomyces capitatus | Hungatella hathewayi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31255 | 3; Public strain collection | eDT | 0 | Anaerotruncus | colihominis | Anaerotruncus | colihominis | 2.38 | Anaerotruncus | colihominis | 2.24 | A | 1 | 1 | 0 | Anaerotruncus colihominis | Anaerotruncus colihominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H281 | 3; sequenced [0207] | EFEx | 0 | Faecalibacterium | prausnitzii | Sphingomonas | aurantiaca | 1.28 | Agromyces | mediolanus | 1.26 | B | 0 | 0 | 1 | Sphingomonas aurantiaca | Agromyces mediolanus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H271 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | bromii | Clostridium | budayi | 1.31 | Ralstonia | sp | 1.29 | B | 0 | 0 | 1 | Clostridium budayi | Ralstonia sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.22.53 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | gnavus | Ruminococcus | gnavus | 2.13 | Ruminococcus | gnavus | 1.97 | A | 1 | 1 | 0 | Ruminococcus gnavus | Ruminococcus gnavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.23.60 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | gnavus | Ruminococcus | gnavus | 2.13 | Ruminococcus | gnavus | 2.07 | A | 1 | 1 | 0 | Ruminococcus gnavus | Ruminococcus gnavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H279 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus | lactaris | Clostridioides | difficile | 1.32 | Alteromonas | stellipolaris | 1.32 | B | 0 | 0 | 1 | Clostridioides difficile | Alteromonas stellipolaris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H261 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus | torques | Staphylococcus | arlettae | 1.46 | Moraxella | catarrhalis | 1.27 | B | 0 | 0 | 1 | Staphylococcus arlettae | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix | anatis | Lactobacillus | amylovorus | 1.43 | Erysipelothrix | tonsillarum | 1.41 | B | 0 | 0 | 1 | Lactobacillus amylovorus | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix | anatis | Staphylococcus | carnosus | 1.36 | Erysipelothrix | rhusiopathiae | 1.26 | B | 0 | 0 | 1 | Staphylococcus carnosus | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix | aquatica | Erysipelothrix | rhusiopathiae | 1.42 | Erysipelothrix | rhusiopathiae | 1.37 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | VA92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | aquatica | Neisseria_flavescens_subflava | group | 1.35 | Fusobacterium | necrophorum | 1.31 | B | 0 | 0 | 1 | Neisseria_flavescens_subflava group | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | larvae | Loigolactobacillus | bifermentans | 1.46 | Fusobacterium | necrophorum | 1.44 | B | 0 | 0 | 1 | Loigolactobacillus bifermentans | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 110099 | 3; type strain | EFex [BRU03] | 2021-10-18 | Erysipelothrix | piscisicarius | Erysipelothrix | rhusiopathiae | 1.75 | Erysipelothrix | rhusiopathiae | 1.63 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.28 | Erysipelothrix | rhusiopathiae | 2.16 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.03 | Erysipelothrix | rhusiopathiae | 2.02 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix | tonsillarum | Erysipelothrix | rhusiopathiae | 1.92 | Erysipelothrix | rhusiopathiae | 1.83 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-23 | Erysipelothrix | tonsillarum | Erysipelothrix | rhusiopathiae | 1.88 | Erysipelothrix | rhusiopathiae | 1.79 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus | sp-CVUAS-10169 | Pseudomonas | putida | 1.34 | Pseudomonas | putida | 1.21 | B | 0 | 0 | 1 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32948 | 3; confirmed proficency test | EFEx | 2021-02-18 | Finegoldia | magna | Finegoldia | magna | 2.26 | Finegoldia | magna | 2.08 | A | 1 | 1 | 0 | Finegoldia magna | Finegoldia magna | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.25 | Helcococcus | kunzii | 2.25 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.45 | Helcococcus | kunzii | 2.4 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 433,2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus | ovis | Helcococcus | ovis | 2.39 | Helcococcus | ovis | 2.3 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus | ovis | Helcococcus | ovis | 2.28 | Helcococcus | ovis | 2.17 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus | sp-CVUAS-32321 | Helcococcus | sueciensis | 1.77 | Helcococcus | sueciensis | 1.73 | B | 0 | 0 | 1 | Helcococcus sueciensis | Helcococcus sueciensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 692,2 | 3; sequenced | DT | 2013-02-19 | Parvimonas | sp-CVUAS-692,2 | Actinomyces | bovis | 1.44 | Actinomyces | bovis | 1.31 | B | 0 | 0 | 1 | Actinomyces bovis | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides | finegoldii | Bacteroides | finegoldii | 2.6 | Bacteroides | finegoldii | 2.53 | A | 1 | 1 | 0 | Bacteroides finegoldii | Bacteroides finegoldii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides | fragilis | Bacteroides | fragilis | 2.35 | Bacteroides | fragilis | 2.33 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181001478 | 2; 0 | DT | 2018-05-02 | Bacteroides | pyogenes | Bacteroides | pyogenes | 2.48 | Bacteroides | pyogenes | 2.48 | A | 1 | 1 | 0 | Bacteroides pyogenes | Bacteroides pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella | heparinolytica | Prevotella | heparinolytica | 2.52 | Prevotella | heparinolytica | 2.49 | A | 1 | 1 | 0 | Prevotella heparinolytica | Prevotella heparinolytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium | gleum | Chryseobacterium | gleum | 1.94 | Chryseobacterium | gleum | 1.9 | B | 0 | 0 | 1 | Chryseobacterium gleum | Chryseobacterium gleum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 2.1 | Chryseobacterium | tructae | 2.07 | A | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium tructae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 1.85 | Chryseobacterium | indologenes | 1.8 | B | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium indologenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32947 | 3; confirmed proficency test | EFEx | 2021-02-18 | Bacteroides | fragilis | Bacteroides | fragilis | 2.22 | Bacteroides | fragilis | 2.2 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium | hibernum | Flavobacterium | hibernum | 2 | Flavobacterium | saccharophilum | 1.97 | A | 1 | 1 | 0 | Flavobacterium hibernum | Flavobacterium saccharophilum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium | sp | Flavobacterium | saccharophilum | 2.04 | Flavobacterium | hibernum | 1.97 | A | 0 | 0 | 1 | Flavobacterium saccharophilum | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium | chaponense | Chryseobacterium | chaponense | 1.81 | Chryseobacterium | chaponense | 1.65 | B | 0 | 0 | 1 | Chryseobacterium chaponense | Chryseobacterium chaponense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium | hominis | Chryseobacterium | hominis | 2.08 | Chryseobacterium | hominis | 1.82 | A | 1 | 1 | 0 | Chryseobacterium hominis | Chryseobacterium hominis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium | manosquense | Deinococcus | wulumuqiensis | 1.45 | Streptomyces | hygroscopicus | 1.23 | B | 0 | 0 | 1 | Deinococcus wulumuqiensis | Streptomyces hygroscopicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 185-U1-6 | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium | manosquense | Staphylococcus | simiae | 1.28 | Chryseobacterium | joostei | 1.25 | B | 0 | 0 | 1 | Staphylococcus simiae | Chryseobacterium joostei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11450.4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium | sp-I-7 | Chryseobacterium | geocarposphaerae | 1.42 | Chryseobacterium | indologenes | 1.3 | B | 0 | 0 | 1 | Chryseobacterium geocarposphaerae | Chryseobacterium indologenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | aquidurense | Flavobacterium | araucananum | 1.89 | Flavobacterium | piscis | 1.88 | B | 0 | 0 | 1 | Flavobacterium araucananum | Flavobacterium piscis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5322.2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | branchiarum | Flavobacterium | pectinovorum | 1.94 | Flavobacterium | hibernum | 1.93 | B | 0 | 0 | 1 | Flavobacterium pectinovorum | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | chungangense | Flavobacterium | plurextorum | 1.76 | Flavobacterium | hercynium | 1.72 | B | 0 | 0 | 1 | Flavobacterium plurextorum | Flavobacterium hercynium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium | columnare | Flavobacterium | columnare | 2.3 | Flavobacterium | columnare | 2.16 | A | 1 | 1 | 0 | Flavobacterium columnare | Flavobacterium columnare | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.04 | Flavobacterium | hibernum | 1.94 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4217,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 1.94 | Flavobacterium | hibernum | 1.88 | B | 0 | 0 | 1 | Flavobacterium frigidimaris | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium | piscis | Flavobacterium | piscis | 2.29 | Flavobacterium | piscis | 2.26 | A | 1 | 1 | 0 | Flavobacterium piscis | Flavobacterium piscis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium | plurextorum | Flavobacterium | plurextorum | 2.35 | Flavobacterium | oncorhynchi | 2.25 | B | 0 | 0 | 1 | Flavobacterium plurextorum | Flavobacterium oncorhynchi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5323.2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | psychoterrae | Flavobacterium | hydatis | 1.81 | Flavobacterium | hibernum | 1.68 | B | 0 | 0 | 1 | Flavobacterium hydatis | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11550,4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.52 | Flavobacterium | psychrophilum | 2.36 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.3 | Flavobacterium | psychrophilum | 2.11 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | reichenbachii | Flavobacterium | reichenbachii | 1.9 | Flavobacterium | cutihirudinis | 1.58 | B | 0 | 0 | 1 | Flavobacterium reichenbachii | Flavobacterium cutihirudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | sp-C2 | Paraburkholderia | phenazinium | 1.57 | Myroides | odoratimimus | 1.45 | B | 0 | 0 | 1 | Paraburkholderia phenazinium | Myroides odoratimimus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-10323 | Flavobacterium | chungbukense | 1.65 | Flavobacterium | anhuiense | 1.6 | B | 0 | 0 | 1 | Flavobacterium chungbukense | Flavobacterium anhuiense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium | sp-CVUAS-4146 | Flavobacterium | pectinovorum | 1.96 | Flavobacterium | hercynium | 1.95 | B | 0 | 0 | 1 | Flavobacterium pectinovorum | Flavobacterium hercynium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4216,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-4216,2 | Flavobacterium | degerlachei | 1.46 | Paraburkholderia | xenovorans | 1.4 | B | 0 | 0 | 1 | Flavobacterium degerlachei | Paraburkholderia xenovorans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-5506 | Flavobacterium | johnsoniae | 1.82 | Flavobacterium | anhuiense | 1.79 | B | 0 | 0 | 1 | Flavobacterium johnsoniae | Flavobacterium anhuiense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | sp-HMD2009 | Limosilactobacillus | oris | 1.4 | Lactobacillus | curvatus | 1.38 | B | 0 | 0 | 1 | Limosilactobacillus oris | Lactobacillus curvatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium | sp-JQMS | Flavobacterium | limicola | 1.45 | Malikia | spinosa | 1.33 | B | 0 | 0 | 1 | Flavobacterium limicola | Malikia spinosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33726 | 3; sequenced | EFEx | 2023-11-07 | Flavobacterium | sp-T157L | Flavobacterium | psychrolimnae | 1.57 | Flavobacterium | sinopsychrotolerans | 1.54 | B | 0 | 0 | 1 | Flavobacterium psychrolimnae | Flavobacterium sinopsychrotolerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium | succinicans | Flavobacterium | tiangeerense | 1.78 | Flavobacterium | sinopsychrotolerans | 1.6 | B | 0 | 0 | 1 | Flavobacterium tiangeerense | Flavobacterium sinopsychrotolerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | tiangeerense | Flavobacterium | tiangeerense | 2.02 | Flavobacterium | glaciei | 1.57 | A | 1 | 1 | 0 | Flavobacterium tiangeerense | Flavobacterium glaciei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 82 | 2; confirmed proficency test | EFEx | 2021-09-10 | Bergeyella | zoohelcum | Bergeyella | zoohelcum | 2.14 | Bergeyella | zoohelcum | 2.14 | A | 1 | 1 | 0 | Bergeyella zoohelcum | Bergeyella zoohelcum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2589,2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas | tenax | Chryseobacterium | aquaticum | 1.4 | Aromatoleum | aromaticum | 1.36 | B | 0 | 0 | 1 | Chryseobacterium aquaticum | Aromatoleum aromaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.27 | Ornithobacterium | rhinotracheale | 2.24 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.21 | Ornithobacterium | rhinotracheale | 2.04 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella | columbina | Lactobacillus | plantarum | 1.25 | Lactobacillus | plantarum | 1.21 | B | 0 | 0 | 1 | Lactobacillus plantarum | Lactobacillus plantarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34092 | 3; sequenced | EFex | 2022-08-05 | Weeksellaceae | sp-CVUAS-34092 | Empedobacter | falsenii | 1.39 | Empedobacter | tilapiae | 1.39 | B | 0 | 0 | 1 | Empedobacter falsenii | Empedobacter tilapiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9288.2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca | zopfii | Gordonia | sputi | 1.28 | Tissierella | praeacuta | 1.23 | B | 0 | 0 | 1 | Gordonia sputi | Tissierella praeacuta | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9875.2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca | zopfii | Herbaspirillum | huttiense | 1.43 | Bergeyella | zoohelcum | 1.37 | B | 0 | 0 | 1 | Herbaspirillum huttiense | Bergeyella zoohelcum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus | sp-CVUAS-11266 | Acidovorax | facilis | 1.23 | Lacticaseibacillus | paracasei | 1.17 | B | 0 | 0 | 1 | Acidovorax facilis | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium | necrophorum | Fusobacterium | necrophorum | 2.48 | Fusobacterium | necrophorum | 2.41 | A | 1 | 1 | 0 | Fusobacterium necrophorum | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium | varium | Fusobacterium | varium | 2.09 | Fusobacterium | varium | 2.04 | A | 1 | 1 | 0 | Fusobacterium varium | Fusobacterium varium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.47 | Streptobacillus | moniliformis | 2.43 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter | abscessus | Pseudoclavibacter | caeni | 1.3 | Anaerobiospirillum | succiniciproducens | 1.22 | B | 0 | 0 | 1 | Pseudoclavibacter caeni | Anaerobiospirillum succiniciproducens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga | salmonicida | Candida | parapsilosis | 1.66 | Candida | parapsilosis | 1.59 | B | 0 | 0 | 1 | Candida parapsilosis | Candida parapsilosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus | hongkongensis | Pseudomonas | trivialis | 1.46 | Pseudomonas | oleovorans | 1.44 | B | 0 | 0 | 1 | Pseudomonas trivialis | Pseudomonas oleovorans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella | termitidis | Pichia | kudriavzevii | 1.2 | Acinetobacter | radioresistens | 1.14 | B | 0 | 0 | 1 | Pichia kudriavzevii | Acinetobacter radioresistens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella | termitidis | Clostridium | chauvoei | 1.24 | Novosphingobium | resinovorum | 1.22 | B | 0 | 0 | 1 | Clostridium chauvoei | Novosphingobium resinovorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCUG 38322 | 3; public strain collection | EFEx | 2015-12-16 | Sneathia | sanguinegens | Nocardia | asteroides | 1.24 | Pseudomonas | umsongensis | 1.23 | B | 0 | 0 | 1 | Nocardia asteroides | Pseudomonas umsongensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCUG 41628 | 3; type strain | EFEx | 2020-10-08 | Sneathia | sanguinegens | Staphylococcus | aureus | 1.25 | Kandleria | vitulina | 1.22 | B | 0 | 0 | 1 | Staphylococcus aureus | Kandleria vitulina | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Arthrobacter | monumenti | 1.52 | Latilactobacillus | fuchuensis | 1.31 | B | 0 | 0 | 1 | Arthrobacter monumenti | Latilactobacillus fuchuensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Arthrobacter | monumenti | 1.42 | Paecilomyces | variotii | 1.32 | B | 0 | 0 | 1 | Arthrobacter monumenti | Paecilomyces variotii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus | felis | Streptobacillus | moniliformis | 1.33 | Paenibacillus | cellulosilyticus | 1.31 | B | 0 | 0 | 1 | Streptobacillus moniliformis | Paenibacillus cellulosilyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.58 | Streptobacillus | moniliformis | 2.52 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus | notomytis | Streptobacillus | moniliformis | 1.4 | Streptobacillus | moniliformis | 1.34 | B | 0 | 0 | 1 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus | ratti | Pseudomonas | abietaniphila | 1.33 | Malikia | spinosa | 1.32 | B | 0 | 0 | 1 | Pseudomonas abietaniphila | Malikia spinosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.55 | Ochrobactrum | intermedium | 2.32 | A | 1 | 1 | 0 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum | sp | Ochrobactrum | anthropi | 2.46 | Ochrobactrum | anthropi | 2.41 | A | 0 | 0 | 1 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum | sp | Ochrobactrum | intermedium | 1.78 | Ochrobactrum | anthropi | 1.68 | B | 0 | 0 | 1 | Ochrobactrum intermedium | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012040 | 2; 0 | DT | 2017-10-04 | Ochrobactrum | sp | Ochrobactrum | tritici | 1.78 | Ochrobactrum | intermedium | 1.65 | B | 0 | 0 | 1 | Ochrobactrum tritici | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia | coli | Escherichia | coli | 2.32 | Escherichia | coli | 2.29 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.4 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella | variicola | Klebsiella | variicola | 2.08 | Klebsiella | variicola | 2.05 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1816.2 | 3; sequenced | DT | 2013-07-24 | Basfia | succiniciproducens | Pasteurella | multocida | 1.6 | Pasteurella | stomatis | 1.54 | B | 0 | 0 | 1 | Pasteurella multocida | Pasteurella stomatis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas | alcaligenes | Pseudomonas | alcaligenes | 2.44 | Pseudomonas | alcaligenes | 2.35 | A | 1 | 1 | 0 | Pseudomonas alcaligenes | Pseudomonas alcaligenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas | anguilliseptica | Pseudomonas | anguilliseptica | 2.02 | Pseudomonas | borbori | 1.95 | A | 1 | 1 | 0 | Pseudomonas anguilliseptica | Pseudomonas borbori | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas | brenneri | Pseudomonas | brenneri | 2.14 | Pseudomonas | veronii | 1.84 | A | 1 | 1 | 0 | Pseudomonas brenneri | Pseudomonas veronii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas | fulva | Pseudomonas | fulva | 2.29 | Pseudomonas | fulva | 1.9 | A | 1 | 1 | 0 | Pseudomonas fulva | Pseudomonas fulva | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas | koreensis | Pseudomonas | koreensis | 2.26 | Pseudomonas | corrugata | 1.89 | A | 1 | 1 | 0 | Pseudomonas koreensis | Pseudomonas corrugata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas | monteilii | Pseudomonas | monteilii | 2.01 | Pseudomonas | putida | 1.92 | A | 1 | 1 | 0 | Pseudomonas monteilii | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1 A | 2; 0 | DT | 2017-08-31 | Pseudomonas | nitroreducens | Pseudomonas | nitroreducens | 2.07 | Pseudomonas | nitroreducens | 1.97 | A | 1 | 1 | 0 | Pseudomonas nitroreducens | Pseudomonas nitroreducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas | oryzihabitans | Pseudomonas | oryzihabitans | 2.12 | Pseudomonas | oryzihabitans | 2.09 | A | 1 | 1 | 0 | Pseudomonas oryzihabitans | Pseudomonas oryzihabitans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1 B | 2; 0 | DT | 2017-08-31 | Pseudomonas | putida | Pseudomonas | putida | 2.49 | Pseudomonas | putida | 2.31 | A | 1 | 1 | 0 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas | taetrolens | Pseudomonas | taetrolens | 2 | Pseudomonas | lundensis | 1.95 | A | 1 | 1 | 0 | Pseudomonas taetrolens | Pseudomonas lundensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2280,3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas | glaciistagni | Limosilactobacillus | gastricus | 1.28 | Agromyces | rhizospherae | 1.26 | B | 0 | 0 | 1 | Limosilactobacillus gastricus | Agromyces rhizospherae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 35685 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-23 | Bartonella | bacilliformis | Bartonella | elizabethae | 2.1 | Pseudomonas | plecoglossicida | 1.49 | A | 0 | 0 | 1 | Bartonella elizabethae | Pseudomonas plecoglossicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 700132 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | grahamii | Bartonella | queenslandensis | 1.9 | Bartonella | elizabethae | 1.47 | B | 0 | 0 | 1 | Bartonella queenslandensis | Bartonella elizabethae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 49882 DSM 28221 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | henselae | Bartonella | henselae | 2.24 | Lactobacillus | paralimentarius | 1.31 | A | 1 | 1 | 0 | Bartonella henselae | Lactobacillus paralimentarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BK8 | 2; strain collection of the RKI | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2022-05-10 | Bartonella | henselae | Schleiferilactobacillus | harbinensis | 1.34 | Schleiferilactobacillus | perolens | 1.18 | B | 0 | 0 | 1 | Schleiferilactobacillus harbinensis | Schleiferilactobacillus perolens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CIP 103739 | 3; public strain collection | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | quintana | Lentilactobacillus | diolivorans | 1.55 | Acinetobacter | baumannii | 1.3 | B | 0 | 0 | 1 | Lentilactobacillus diolivorans | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 700727 DSM 28218 | 3; type strain | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | vinsonii | Bartonella | vinsonii | 2.16 | Lactobacillus | jensenii | 1.22 | A | 1 | 1 | 0 | Bartonella vinsonii | Lactobacillus jensenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_08 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2012-04-02 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.42 | Ochrobactrum | anthropi | 2.39 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM373 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | abortus | Ligilactobacillus | ruminis | 1.43 | Ligilactobacillus | ruminis | 1.34 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Ligilactobacillus ruminis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM369 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | abortus | Staphylococcus | schleiferi | 1.39 | Ligilactobacillus | acidipiscis | 1.37 | B | 0 | 0 | 1 | Staphylococcus schleiferi | Ligilactobacillus acidipiscis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-11 | Brucella | canis | Agromyces | fucosus | 1.36 | Ochrobactrum | daejeonense | 1.35 | B | 0 | 0 | 1 | Agromyces fucosus | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-03-05 | Brucella | canis | Limosilactobacillus | reuteri | 1.46 | Ochrobactrum | daejeonense | 1.41 | B | 0 | 0 | 1 | Limosilactobacillus reuteri | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 007 1 002 34 | 2; 0 | EFEx | 2013-06-20 | Brucella | ceti | Agrobacterium | radiobacter | 1.39 | Ochrobactrum | daejeonense | 1.38 | B | 0 | 0 | 1 | Agrobacterium radiobacter | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM389 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | ceti | Ochrobactrum | daejeonense | 1.43 | Ochrobactrum | gallinifaecis | 1.42 | B | 0 | 0 | 1 | Ochrobactrum daejeonense | Ochrobactrum gallinifaecis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 10 1 001 44 | 2; 0 | EFEx | 2013-06-20 | Brucella | inopinata | Terrimonas | ferruginea | 1.46 | Flavobacterium | flevense | 1.41 | B | 0 | 0 | 1 | Terrimonas ferruginea | Flavobacterium flevense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A1309 | 2; 0 | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 28.07.2021¦ 1:1 HCCA¦Andy 3.0 | 2021-07-28 | Brucella | intermedia | Ochrobactrum | intermedium | 2.43 | Ochrobactrum | intermedium | 1.93 | A | 0 | 0 | 1 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM374 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | melitensis | Staphylococcus | sciuri | 1.34 | Agrobacterium | radiobacter | 1.32 | B | 0 | 0 | 1 | Staphylococcus sciuri | Agrobacterium radiobacter | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM407 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | melitensis | Ochrobactrum | anthropi | 1.42 | Bacteroides | thetaiotaomicron | 1.39 | B | 0 | 0 | 1 | Ochrobactrum anthropi | Bacteroides thetaiotaomicron | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM408 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | melitensis | Schleiferilactobacillus | harbinensis | 1.43 | Ochrobactrum | daejeonense | 1.42 | B | 0 | 0 | 1 | Schleiferilactobacillus harbinensis | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM410 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | melitensis | Bacteroides | fragilis | 1.37 | Ochrobactrum | tritici | 1.29 | B | 0 | 0 | 1 | Bacteroides fragilis | Ochrobactrum tritici | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 4915 CVUAS | 3; type strain | DT | 2013-11-26 | Brucella | microti | Lacticaseibacillus | pantheris | 1.5 | Ochrobactrum | daejeonense | 1.47 | B | 0 | 0 | 1 | Lacticaseibacillus pantheris | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 4915 RKI | 3; type strain | TFAextr | 2015-03-17 | Brucella | microti | Ligilactobacillus | aviarius | 1.35 | Ochrobactrum | daejeonense | 1.34 | B | 0 | 0 | 1 | Ligilactobacillus aviarius | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-20 | Brucella | neotomae | Klebsiella | pneumoniae | 1.52 | Ochrobactrum | daejeonense | 1.52 | B | 0 | 0 | 1 | Klebsiella pneumoniae | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM421 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | neotomae | Ligilactobacillus | ruminis | 1.49 | Stenotrophomonas | acidaminiphila | 1.48 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Stenotrophomonas acidaminiphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 06 1 007 32 | 2; 0 | EFEx | 2013-06-20 | Brucella | ovis | Ochrobactrum | daejeonense | 1.48 | Ochrobactrum | tritici | 1.38 | B | 0 | 0 | 1 | Ochrobactrum daejeonense | Ochrobactrum tritici | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-11 | Brucella | ovis | Ochrobactrum | tritici | 1.35 | Lactobacillus | pentosus | 1.31 | B | 0 | 0 | 1 | Ochrobactrum tritici | Lactobacillus pentosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-26 | Brucella | pinnipedialis | Pandoraea | norimbergensis | 1.4 | Pannonibacter | phragmitetus | 1.33 | B | 0 | 0 | 1 | Pandoraea norimbergensis | Pannonibacter phragmitetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM405 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | pinnipedialis | Staphylococcus | vitulinus | 1.5 | Corynebacterium | flavescens | 1.45 | B | 0 | 0 | 1 | Staphylococcus vitulinus | Corynebacterium flavescens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM431 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | suis | Ligilactobacillus | ruminis | 1.39 | Actinocorallia | libanotica | 1.38 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Actinocorallia libanotica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM427 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | suis | Proteus | vulgaris | 1.36 | Pseudomonas | putida | 1.34 | B | 0 | 0 | 1 | Proteus vulgaris | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella | suis | Ochrobactrum | anthropi | 1.23 | Ochrobactrum | anthropi | 1.19 | B | 0 | 0 | 1 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM428 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | suis | Lactobacillus | delbrueckii | 1.37 | Lactobacillus | delbrueckii | 1.34 | B | 0 | 0 | 1 | Lactobacillus delbrueckii | Lactobacillus delbrueckii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM430 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-15 | Brucella | suis | Ligilactobacillus | ruminis | 1.34 | Acinetobacter | calcoaceticus | 1.29 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Acinetobacter calcoaceticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2016-10-15 | Brucella | vulpis | Pannonibacter | phragmitetus | 1.36 | Paenibacillus | turicensis | 1.33 | B | 0 | 0 | 1 | Pannonibacter phragmitetus | Paenibacillus turicensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum | ovis | Falsochrobactrum | ovis | 2.3 | Shewanella | profunda | 1.21 | A | 1 | 1 | 0 | Falsochrobactrum ovis | Shewanella profunda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2016-06-03 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.54 | Ochrobactrum | anthropi | 2.42 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32748 | 3; sequenced | EFex | 2020-10-29 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.24 | Ochrobactrum | intermedium | 2.16 | A | 1 | 1 | 0 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter | aganoensis | Pseudomonas | mucidolens | 1.45 | Pseudomonas | orientalis | 1.44 | B | 0 | 0 | 1 | Pseudomonas mucidolens | Pseudomonas orientalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter | aminovorans | Lactobacillus | curvatus | 1.5 | Aromatoleum | toluolicum | 1.42 | B | 0 | 0 | 1 | Lactobacillus curvatus | Aromatoleum toluolicum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30768 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 16.09.2022¦ 1:1 HCCA¦Andy 3.1#EXT=Paracoccus yeei DSM 30768, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-10-07 | Paracoccus | yeei | Paracoccus | yeei | 2.37 | Paracoccus | yeei | 2.33 | A | 1 | 1 | 0 | Paracoccus yeei | Paracoccus yeei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-16 | Paracoccus | yeei | Paracoccus | yeei | 2.45 | Paracoccus | yeei | 2.24 | A | 1 | 1 | 0 | Paracoccus yeei | Paracoccus yeei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 103800 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 25.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Roseomonas gilardii DSM 103800, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-04-06 | Roseomonas | gilardii | Roseomonas | mucosa | 1.76 | Roseomonas | gilardii | 1.66 | B | 0 | 0 | 1 | Roseomonas mucosa | Roseomonas gilardii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas | mucosa | Roseomonas | mucosa | 2.53 | Roseomonas | mucosa | 2.49 | A | 1 | 1 | 0 | Roseomonas mucosa | Roseomonas mucosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2007-01-24 | Inquilinus | limosus | Inquilinus | limosus | 2.63 | Inquilinus | limosus | 2.62 | A | 1 | 1 | 0 | Inquilinus limosus | Inquilinus limosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1098 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Sphingomonas paucimobilis DSM 30198, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Sphingomonas | paucimobilis | Sphingomonas | paucimobilis | 2.5 | Sphingomonas | paucimobilis | 2.42 | A | 1 | 1 | 0 | Sphingomonas paucimobilis | Sphingomonas paucimobilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30198 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Sphingomonas paucimobilis DSM 30198, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Sphingomonas | paucimobilis | Sphingomonas | paucimobilis | 2.53 | Sphingomonas | paucimobilis | 2.45 | A | 1 | 1 | 0 | Sphingomonas paucimobilis | Sphingomonas paucimobilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.32 | Taylorella | equigenitalis | 1.32 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter | denitrificans | Achromobacter | denitrificans | 2.22 | Achromobacter | denitrificans | 2.04 | A | 1 | 1 | 0 | Achromobacter denitrificans | Achromobacter denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter | sp-DSM-30031 | Achromobacter | piechaudii | 1.95 | Achromobacter | spanius | 1.9 | B | 0 | 0 | 1 | Achromobacter piechaudii | Achromobacter spanius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-11-18 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.33 | Achromobacter | xylosoxidans | 2.16 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-24 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.26 | Achromobacter | xylosoxidans | 2.23 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30030 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Alcaligenes faecalis ssp. parafaecalis DSM 13975, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-29 | Alcaligenes | faecalis | Alcaligenes | faecalis | 2.5 | Alcaligenes | faecalis | 2.37 | A | 1 | 1 | 0 | Alcaligenes faecalis | Alcaligenes faecalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13975 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Alcaligenes faecalis subsp. Parafaecalis DSM 13975, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Alcaligenes | faecalis | Alcaligenes | faecalis | 2.5 | Alcaligenes | faecalis | 2.49 | A | 1 | 1 | 0 | Alcaligenes faecalis | Alcaligenes faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea | sp-CVUAS-31213 | Streptomyces | badius | 1.35 | Staphylococcus | aureus | 1.32 | B | 0 | 0 | 1 | Streptomyces badius | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella | avium | Bordetella | avium | 2.39 | Bordetella | avium | 2.29 | A | 1 | 1 | 0 | Bordetella avium | Bordetella avium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 21.2 | 2; confirmed proficency test | eDT | 2022-02-14 | Bordetella | trematum | Bordetella | trematum | 2.42 | Bordetella | trematum | 2.34 | A | 1 | 1 | 0 | Bordetella trematum | Bordetella trematum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32946 | 3; confirmed proficency test | EFEx | 2021-02-18 | Kerstersia | gyiorum | Kerstersia | gyiorum | 2.48 | Kerstersia | gyiorum | 2.48 | A | 1 | 1 | 0 | Kerstersia gyiorum | Kerstersia gyiorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 1.67 | Oligella | urethralis | 1.66 | B | 0 | 0 | 1 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 1.6 | Oligella | urethralis | 1.56 | B | 0 | 0 | 1 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga | kullae | Pigmentiphaga | daeguensis | 1.57 | Clostridium | spiroforme | 1.37 | B | 0 | 0 | 1 | Pigmentiphaga daeguensis | Clostridium spiroforme | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.09 | Taylorella | equigenitalis | 1.28 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.15 | Taylorella | equigenitalis | 2.1 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.35 | Taylorella | equigenitalis | 2.24 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-04 | Burkholderia | pseudomallei | Burkholderia | thailandensis | 2.03 | Staphylococcus | aureus | 1.49 | A | 0 | 0 | 1 | Burkholderia thailandensis | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-01 | Burkholderia | pseudomallei | Burkholderia | thailandensis | 2.03 | Paeniclostridium | sordellii | 1.49 | A | 0 | 0 | 1 | Burkholderia thailandensis | Paeniclostridium sordellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-02 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.38 | Burkholderia | ambifaria | 1.76 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.27 | Burkholderia | ambifaria | 2.24 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16087 LMG 19182 | 3; type starin | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.36 | Burkholderia | ambifaria | 2.15 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | anthina | Burkholderia | anthina | 2.35 | Burkholderia | anthina | 2.27 | A | 1 | 1 | 0 | Burkholderia anthina | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-08-06 | Burkholderia | anthina | Burkholderia | pyrrocinia | 2.17 | Burkholderia | anthina | 2.16 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-18 | Burkholderia | arboris | Burkholderia | pyrrocinia | 2.34 | Burkholderia | cepacia | 2.23 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | arboris | Burkholderia | cenocepacia | 2.29 | Burkholderia | vietnamiensis | 2.22 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.48 | Burkholderia | cenocepacia | 2.36 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-07-02 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.42 | Burkholderia | cenocepacia | 2.17 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-05-04 | Burkholderia | cepacia | Burkholderia | cenocepacia | 2.46 | Burkholderia | cenocepacia | 2.15 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-09 | Burkholderia | cepacia | Burkholderia | cepacia | 2.4 | Burkholderia | pyrrocinia | 2.34 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-03 | Burkholderia | contaminans | Burkholderia | cepacia | 2.28 | Burkholderia | cepacia | 2.16 | A | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 01 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-19 | Burkholderia | contaminans | Burkholderia | cepacia | 2.42 | Burkholderia | cenocepacia | 2.17 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-31 | Burkholderia | diffusa | Burkholderia | ambifaria | 2.35 | Burkholderia | vietnamiensis | 2.32 | B | 0 | 0 | 1 | Burkholderia ambifaria | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | diffusa | Burkholderia | cenocepacia | 2.31 | Burkholderia | ambifaria | 2.21 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | dolosa | Burkholderia | dolosa | 2.29 | Burkholderia | multivorans | 2.15 | B | 0 | 0 | 1 | Burkholderia dolosa | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-17 | Burkholderia | dolosa | Burkholderia | dolosa | 2.08 | Burkholderia | cenocepacia | 2.07 | B | 0 | 0 | 1 | Burkholderia dolosa | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-24 | Burkholderia | gladioli | Burkholderia | gladioli | 2.45 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia | gladioli | Burkholderia | gladioli | 2.42 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | glumae | Burkholderia | glumae | 2.33 | Burkholderia | plantarii | 2.06 | B | 0 | 0 | 1 | Burkholderia glumae | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-05-04 | Burkholderia | glumae | Burkholderia | glumae | 2.33 | Burkholderia | plantarii | 2.13 | B | 0 | 0 | 1 | Burkholderia glumae | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-31 | Burkholderia | lata | Burkholderia | cepacia | 2.49 | Burkholderia | lata | 2.41 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia lata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-08-23 | Burkholderia | lata | Burkholderia | cepacia | 2.47 | Burkholderia | lata | 2.38 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia lata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-20 | Burkholderia | latens | Burkholderia | latens | 2.32 | Burkholderia | vietnamiensis | 2.29 | B | 0 | 0 | 1 | Burkholderia latens | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-09-04 | Burkholderia | mallei | Burkholderia | thailandensis | 1.65 | Staphylococcus | aureus | 1.38 | B | 0 | 0 | 1 | Burkholderia thailandensis | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-08-24 | Burkholderia | mallei | Burkholderia | thailandensis | 1.73 | Lactobacillus | paralimentarius | 1.38 | B | 0 | 0 | 1 | Burkholderia thailandensis | Lactobacillus paralimentarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-01 | Burkholderia | metallica | Burkholderia | metallica | 2.4 | Burkholderia | cenocepacia | 2.27 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | metallica | Burkholderia | metallica | 2.31 | Burkholderia | cenocepacia | 2.23 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-13 | Burkholderia | multivorans | Burkholderia | multivorans | 2.22 | Burkholderia | multivorans | 2.21 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-08 | Burkholderia | multivorans | Burkholderia | multivorans | 2.13 | Burkholderia | multivorans | 2 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | oklahomensis | Burkholderia | multivorans | 1.77 | Burkholderia | ambifaria | 1.76 | B | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | oklahomensis | Burkholderia | thailandensis | 1.72 | Burkholderia | ambifaria | 1.65 | B | 0 | 0 | 1 | Burkholderia thailandensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | plantarii | Burkholderia | plantarii | 2.49 | Burkholderia | pyrrocinia | 1.89 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | plantarii | Burkholderia | plantarii | 2.5 | Burkholderia | pyrrocinia | 1.94 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-21 | Burkholderia | pseudomultivorans | Burkholderia | multivorans | 2.1 | Burkholderia | multivorans | 2.07 | A | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-21 | Burkholderia | pseudomultivorans | Burkholderia | multivorans | 2.03 | Burkholderia | cenocepacia | 1.93 | A | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CAMPA 1043 (HST 249) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-01 | Burkholderia | puraquae | Burkholderia | pyrrocinia | 2.27 | Burkholderia | cepacia_Group | 2.2 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | R-50214 (BP 78) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-18 | Burkholderia | puraquae | Burkholderia | pyrrocinia | 2.41 | Burkholderia | cepacia_Group | 2.32 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | pyrrocinia | Burkholderia | pyrrocinia | 2.57 | Burkholderia | cepacia_Group | 2.33 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | pyrrocinia | Burkholderia | cenocepacia | 2.13 | Burkholderia | pyrrocinia | 2.12 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia | seminalis | Burkholderia | seminalis | 2.41 | Burkholderia | cepacia | 2.32 | B | 0 | 0 | 1 | Burkholderia seminalis | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-11 | Burkholderia | seminalis | Burkholderia | cepacia | 2.29 | Burkholderia | seminalis | 2.11 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia seminalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.56 | Burkholderia | stabilis | 2.44 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-19 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.51 | Burkholderia | stabilis | 2.28 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-16 | Burkholderia | stagnalis | Burkholderia | anthina | 2.14 | Burkholderia | ambifaria | 2.11 | B | 0 | 0 | 1 | Burkholderia anthina | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia | stagnalis | Burkholderia | vietnamiensis | 2.16 | Burkholderia | ambifaria | 2.01 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Burkholderia | taxonK | Burkholderia | cepacia | 2.45 | Burkholderia | cenocepacia | 2.16 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-26 | Burkholderia | taxonK | Burkholderia | cepacia | 2.23 | Burkholderia | cenocepacia | 2.13 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia | territorii | Burkholderia | cenocepacia | 2.3 | Burkholderia | cenocepacia | 2.05 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-16 | Burkholderia | territorii | Burkholderia | cenocepacia | 2.28 | Burkholderia | vietnamiensis | 2.21 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13276 n | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2010-11-24 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.34 | Burkholderia | glumae | 1.59 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia | ubonensis | Burkholderia | cenocepacia | 2.19 | Burkholderia | cepacia | 2.07 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia | ubonensis | Burkholderia | cepacia_Group | 2.1 | Burkholderia | plantarii | 2.09 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9631 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2007-05-09 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.33 | Burkholderia | cenocepacia | 2.2 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10929 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-12-13 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.43 | Burkholderia | pyrrocinia | 2.14 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-13 | Caballeronia | glathei | Caballeronia | glathei | 2.55 | Caballeronia | turbans | 1.53 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-20 | Caballeronia | glathei | Caballeronia | glathei | 2.47 | Caballeronia | turbans | 1.56 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea | apista | Pandoraea | apista | 2.36 | Pandoraea | apista | 2.28 | A | 1 | 1 | 0 | Pandoraea apista | Pandoraea apista | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea | pnomenusa | Pandoraea | pnomenusa | 2.4 | Pandoraea | pnomenusa | 2.29 | A | 1 | 1 | 0 | Pandoraea pnomenusa | Pandoraea pnomenusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea | pulmonicola | Pandoraea | pulmonicola | 2.49 | Pandoraea | pulmonicola | 2.25 | A | 1 | 1 | 0 | Pandoraea pulmonicola | Pandoraea pulmonicola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.52 | Paraburkholderia | caledonica | 2.45 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.46 | Paraburkholderia | caledonica | 2.37 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-13 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.25 | Burkholderia | thailandensis | 1.6 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-13 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.23 | Burkholderia | thailandensis | 1.42 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-20 | Paraburkholderia | hospita | Lactobacillus | fructivorans | 1.71 | Herbaspirillum | frisingense | 1.44 | B | 0 | 0 | 1 | Lactobacillus fructivorans | Herbaspirillum frisingense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-20 | Paraburkholderia | hospita | Paraburkholderia | phenazinium | 1.6 | Paraburkholderia | caribensis | 1.58 | B | 0 | 0 | 1 | Paraburkholderia phenazinium | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-01 | Paraburkholderia | kururiensis | Burkholderia | pyrrocinia | 1.64 | Burkholderia | cenocepacia | 1.54 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-26 | Paraburkholderia | phymatum | Paraburkholderia | caribensis | 1.7 | Paraburkholderia | fungorum | 1.51 | B | 0 | 0 | 1 | Paraburkholderia caribensis | Paraburkholderia fungorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-19 | Paraburkholderia | phytofirmans | Paraburkholderia | xenovorans | 2.25 | Paraburkholderia | phymatum | 2 | B | 0 | 0 | 1 | Paraburkholderia xenovorans | Paraburkholderia phymatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-20 | Paraburkholderia | tuberum | Paraburkholderia | tuberum | 2.36 | Paraburkholderia | xenovorans | 1.59 | A | 1 | 1 | 0 | Paraburkholderia tuberum | Paraburkholderia xenovorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2010-11-26 | Paraburkholderia | xenovorans | Paraburkholderia | xenovorans | 2.49 | Paraburkholderia | phymatum | 2.06 | B | 0 | 0 | 1 | Paraburkholderia xenovorans | Paraburkholderia phymatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2016-05-13 | Ralstonia | insidiosa | Ralstonia | insidiosa | 2.38 | Ralstonia | insidiosa | 2.35 | A | 1 | 1 | 0 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2016-05-13 | Ralstonia | mannitolilytica | Ralstonia | mannitolilytica | 2.46 | Ralstonia | mannitolilytica | 2.22 | A | 1 | 1 | 0 | Ralstonia mannitolilytica | Ralstonia mannitolilytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-22 | Ralstonia | pickettii | Ralstonia | pickettii | 2.15 | Ralstonia | pickettii | 2.12 | A | 1 | 1 | 0 | Ralstonia pickettii | Ralstonia pickettii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2009-05-07 | Ralstonia | pickettii | Ralstonia | insidiosa | 2.07 | Ralstonia | insidiosa | 2.07 | A | 0 | 0 | 1 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia | acidovorans | Delftia | acidovorans | 2.54 | Delftia | tsuruhatensis | 2.44 | B | 0 | 0 | 1 | Delftia acidovorans | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia | tsuruhatensis | Delftia | tsuruhatensis | 2.31 | Delftia | tsuruhatensis | 2.23 | A | 1 | 1 | 0 | Delftia tsuruhatensis | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga | intermedia | Hydrogenophaga | intermedia | 2.4 | Novosphingobium | resinovorum | 1.34 | A | 1 | 1 | 0 | Hydrogenophaga intermedia | Novosphingobium resinovorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-07-25 | Xenophilus | sp | Aromatoleum | evansii | 1.43 | Ligilactobacillus | salivarius | 1.32 | B | 0 | 0 | 1 | Aromatoleum evansii | Ligilactobacillus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 24713 EFex | 3; type strain | Extr | 2016-02-02 | Deefgea | chitinilytica | Mannheimia | haemolytica | 1.42 | Clostridium | cochlearium | 1.32 | B | 0 | 0 | 1 | Mannheimia haemolytica | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 24713 eDT | 3; type strain | eDT | 2015-09-22 | Deefgea | chitinilytica | Pichia | cactophila | 1.24 | Prevotella | oralis | 1.14 | B | 0 | 0 | 1 | Pichia cactophila | Prevotella oralis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea | rivuli | Arthrobacter | woluwensis | 1.45 | Methylobacterium | mesophilicum | 1.41 | B | 0 | 0 | 1 | Arthrobacter woluwensis | Methylobacterium mesophilicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 08-F17,2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Lactobacillus | crispatus | 1.31 | Agromyces | salentinus | 1.27 | B | 0 | 0 | 1 | Lactobacillus crispatus | Agromyces salentinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Lactobacillus | crispatus | 1.38 | Pseudomonas | putida | 1.33 | B | 0 | 0 | 1 | Lactobacillus crispatus | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6995,3 | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea | sp-CVUAS-6995,3 | Pseudomonas | chlororaphis | 1.5 | Pseudomonas | orientalis | 1.45 | B | 0 | 0 | 1 | Pseudomonas chlororaphis | Pseudomonas orientalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-6995,3 | Pseudomonas | orientalis | 1.48 | Corynebacterium | ulcerans | 1.36 | B | 0 | 0 | 1 | Pseudomonas orientalis | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-9920 | Flavobacterium | flevense | 1.27 | Bacteroides | pyogenes | 1.27 | B | 0 | 0 | 1 | Flavobacterium flevense | Bacteroides pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea | sp-CVUAS-9920 | Raoultella | ornithinolytica | 1.22 | Pediococcus | acidilactici | 1.21 | B | 0 | 0 | 1 | Raoultella ornithinolytica | Pediococcus acidilactici | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter | limnosediminis | Iodobacter | fluviatilis | 1.45 | Aromatoleum | anaerobicus | 1.36 | B | 0 | 0 | 1 | Iodobacter fluviatilis | Aromatoleum anaerobicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella | kingae | Kingella | kingae | 2.49 | Kingella | kingae | 2.41 | A | 1 | 1 | 0 | Kingella kingae | Kingella kingae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10067,2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria | animaloris | Neisseria | animaloris | 2.17 | Neisseria | animaloris | 2.12 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria | animaloris | Neisseria | animaloris | 2.42 | Neisseria | animaloris | 2.4 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-4931 | Neisseria | animaloris | 1.42 | Neisseria | animaloris | 1.4 | B | 0 | 0 | 1 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-8997 | Neisseria_flavescens_subflava | group | 1.48 | Providencia | rettgeri | 1.45 | B | 0 | 0 | 1 | Neisseria_flavescens_subflava group | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-LHL-1296 | Neisseria | animaloris | 1.61 | Neisseria | cinerea | 1.37 | B | 0 | 0 | 1 | Neisseria animaloris | Neisseria cinerea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5699,3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria | sp-N95-16 | Bergeriella | denitrificans | 1.99 | Neisseria_flavescens_subflava | group | 1.73 | C | 0 | 0 | 1 | Bergeriella denitrificans | Neisseria_flavescens_subflava group | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae | sp-CVUAS-8636 | Pseudomonas | taetrolens | 1.43 | Pseudomonas | fluorescens | 1.37 | B | 0 | 0 | 1 | Pseudomonas taetrolens | Pseudomonas fluorescens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter | fetus | Campylobacter | fetus | 2.34 | Campylobacter | fetus | 2.28 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.37 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.19 | Campylobacter | hyointestinalis | 2.12 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | lari | Campylobacter | lari | 2.19 | Campylobacter | lari | 2.04 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.04 | Campylobacter | lari | 1.82 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.2 | Campylobacter | lari | 2.11 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Campylobacter | lari | 1.43 | A | 1 | 1 | 0 | Campylobacter peloridis | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-7048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Burkholderia | plantarii | 1.42 | A | 1 | 1 | 0 | Campylobacter peloridis | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter | cryaerophilus | Arcobacter | cryaerophilus | 2.03 | Arcobacter | cryaerophilus | 2.01 | A | 0 | 0 | 1 | Arcobacter cryaerophilus | Arcobacter cryaerophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.06 | Aliarcobacter | butzleri | 1.94 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.24 | Aliarcobacter | butzleri | 2.21 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter | coli | Campylobacter | coli | 2.18 | Campylobacter | coli | 1.97 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4689 | 3; type strain | eDT | 2015-12-30 | Campylobacter | coli | Campylobacter | coli | 2.24 | Campylobacter | coli | 2.21 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5361 | 3; type strain | DT | 2013-09-24 | Campylobacter | fetus | Campylobacter | fetus | 2.51 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6673,2 | 3; sequenced | DT | 2013-08-28 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.15 | Campylobacter | hyointestinalis | 2.06 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter | jejuni | Campylobacter | jejuni | 2.19 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter | jejuni | Campylobacter | jejuni | 2.39 | Campylobacter | jejuni | 2.32 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter | jejuni | Campylobacter | jejuni | 2.16 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 11375 | 3; type strain | eDT | 2015-12-30 | Campylobacter | lari | Campylobacter | lari | 2.42 | Campylobacter | lari | 2.19 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter | sputorum | Campylobacter | sputorum | 2.01 | Campylobacter | sputorum | 1.9 | A | 1 | 1 | 0 | Campylobacter sputorum | Campylobacter sputorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter | sp-W1 | Campylobacter | jejuni | 1.49 | Campylobacter | jejuni | 1.48 | B | 0 | 0 | 1 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter | upsaliensis | Campylobacter | upsaliensis | 2.05 | Campylobacter | upsaliensis | 1.96 | A | 1 | 1 | 0 | Campylobacter upsaliensis | Campylobacter upsaliensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11185.3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter | canadensis | Clostridium | cochlearium | 1.42 | Helicobacter | pullorum | 1.4 | B | 0 | 0 | 1 | Clostridium cochlearium | Helicobacter pullorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas | bestiarum | Aeromonas | bestiarum | 2.34 | Aeromonas | popoffii | 2.15 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas popoffii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas | bivalvium | Aeromonas | veronii | 1.93 | Aeromonas | ichthiosmia | 1.9 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas ichthiosmia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas | caviae | Aeromonas | caviae | 2.29 | Aeromonas | caviae | 2.25 | A | 1 | 1 | 0 | Aeromonas caviae | Aeromonas caviae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | hydrophila | 2.03 | Aeromonas | jandaei | 1.96 | A | 0 | 0 | 1 | Aeromonas hydrophila | Aeromonas jandaei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | hydrophila | 2.13 | Aeromonas | hydrophila | 2.11 | A | 0 | 0 | 1 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30187 | 4; type strain, WGS sequenced (BfR) | EFEx | 2015-08-07 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.57 | Aeromonas | hydrophila | 2.45 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.25 | Aeromonas | hydrophila | 2.09 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas | jandaei | Aeromonas | jandaei | 2.05 | Aeromonas | hydrophila | 1.97 | A | 1 | 1 | 0 | Aeromonas jandaei | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | media | Aeromonas | media | 2.2 | Aeromonas | veronii | 2.13 | B | 0 | 0 | 1 | Aeromonas media | Aeromonas veronii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas | molluscorum | Aeromonas | molluscorum | 2.15 | Aeromonas | molluscorum | 1.89 | A | 1 | 1 | 0 | Aeromonas molluscorum | Aeromonas molluscorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | popoffii | Aeromonas | popoffii | 2.11 | Aeromonas | bestiarum | 2.05 | B | 0 | 0 | 1 | Aeromonas popoffii | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 1.89 | Aeromonas | salmonicida | 1.86 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.1 | Aeromonas | bestiarum | 2.02 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | bestiarum | 2.1 | Aeromonas | popoffii | 2.1 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas popoffii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas | salmonicida | Aeromonas | bestiarum | 2.04 | Aeromonas | eucrenophila | 2.04 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8400,2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.19 | Aeromonas | bestiarum | 2.05 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.14 | Aeromonas | bestiarum | 2.12 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas | salmonicida | Aeromonas | bestiarum | 2.11 | Aeromonas | bestiarum | 2.02 | A | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.01 | Aeromonas | bestiarum | 1.96 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas | simiae | Aeromonas | simiae | 2.07 | Rhizorhabdus | wittichii | 1.68 | A | 1 | 1 | 0 | Aeromonas simiae | Rhizorhabdus wittichii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas | veronii | Aeromonas | ichthiosmia | 2.18 | Aeromonas | veronii | 2.17 | B | 0 | 0 | 1 | Aeromonas ichthiosmia | Aeromonas veronii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7386 | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas | veronii | Aeromonas | veronii | 1.94 | Aeromonas | enteropelogenes | 1.91 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas enteropelogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas | pectinilytica | Aeromonas | eucrenophila | 1.4 | Pseudaeromonas | sharmana | 1.38 | B | 0 | 0 | 1 | Aeromonas eucrenophila | Pseudaeromonas sharmana | 0 | 0 | 0 | 0
|  |  | 1 | 0 | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella | viscosa | Moritella | japonica | 1.97 | Moritella | dasanensis | 1.51 | B | 0 | 0 | 1 | Moritella japonica | Moritella dasanensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella | viscosa | Moritella | japonica | 1.62 | Pseudomonas | tolaasii | 1.27 | B | 0 | 0 | 1 | Moritella japonica | Pseudomonas tolaasii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | profunda | 1.46 | Lactobacillus | paraplantarum | 1.45 | B | 0 | 0 | 1 | Shewanella profunda | Lactobacillus paraplantarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | baltica | 1.42 | Shewanella | profunda | 1.41 | B | 0 | 0 | 1 | Shewanella baltica | Shewanella profunda | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 81,2 | 3; sequenced | EFEx | 2013-11-28 | Shewanella | sp-RKVR22 | Shewanella | corallii | 1.59 | Shewanella | corallii | 1.55 | B | 0 | 0 | 1 | Shewanella corallii | Shewanella corallii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter | condimenti | Cronobacter | sp | 2.14 | Cronobacter | sp | 2.13 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter | dublinensis | Cronobacter | sp | 2.2 | Cronobacter | sp | 2.19 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18702 | 3; type strain | DT | 2019-03-05 | Cronobacter | malonaticus | Cronobacter | sp | 2.46 | Cronobacter | sp | 2.42 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter | turicensis | Cronobacter | sp | 2.36 | Cronobacter | sp | 2.29 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2011-11-10 | Enterobacter | cloacae | Enterobacter | hormaechei | 2.06 | Enterobacter | hormaechei | 2.04 | A | 0 | 0 | 1 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Citrobacter | amalonaticus | Citrobacter | amalonaticus | 2.41 | Citrobacter | amalonaticus | 2.3 | A | 1 | 1 | 0 | Citrobacter amalonaticus | Citrobacter amalonaticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17342 | 3; public strain collection | see [0185] | 2022-08-29 | Citrobacter | amalonaticus | Citrobacter | amalonaticus | 2.44 | Citrobacter | amalonaticus | 2.25 | A | 1 | 1 | 0 | Citrobacter amalonaticus | Citrobacter amalonaticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | freundii | 2.39 | Citrobacter | freundii | 2.37 | A | 1 | 1 | 0 | Citrobacter freundii | Citrobacter freundii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | braakii | 2.29 | Citrobacter | braakii | 2.23 | A | 0 | 0 | 1 | Citrobacter braakii | Citrobacter braakii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25408 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2008-04-11 | Citrobacter | koseri | Citrobacter | koseri | 2.52 | Citrobacter | koseri | 2.47 | A | 1 | 1 | 0 | Citrobacter koseri | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4570 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Citrobacter koseri DSM 4570, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-29 | Citrobacter | koseri | Citrobacter | koseri | 2.58 | Citrobacter | koseri | 2.49 | A | 1 | 1 | 0 | Citrobacter koseri | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-14 | Cronobacter | malonaticus | Cronobacter | sp | 2.29 | Cronobacter | sp | 2.19 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32945 | 3; confirmed proficency test | EFEx | 2021-02-18 | Cronobacter | sakazakii | Cronobacter | sp | 2.14 | Cronobacter | sp | 1.99 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter | sakazakii | Cronobacter | sp | 2.44 | Cronobacter | sp | 2.33 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30054 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Enterobacter | cloacae | Enterobacter | cloacae | 2.34 | Enterobacter | cloacae | 2.08 | A | 1 | 1 | 0 | Enterobacter cloacae | Enterobacter cloacae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14563 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter | hormaechei | Enterobacter | hormaechei | 2.1 | Enterobacter | hormaechei | 1.9 | A | 1 | 1 | 0 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12409 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter | hormaechei | Enterobacter | hormaechei | 1.95 | Enterobacter | hormaechei | 1.74 | B | 0 | 0 | 1 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16691 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter | hormaechei | Enterobacter | hormaechei | 1.82 | Enterobacter | hormaechei | 1.61 | B | 0 | 0 | 1 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2023000186 | 2; corresponds to expected result of PT | EFEx | 2023-09-08 | Enterobacter | sp | Enterobacter | bugandensis | 2.27 | Enterobacter | asburiae | 2.25 | B | 0 | 0 | 1 | Enterobacter bugandensis | Enterobacter asburiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Enterobacter | sp-ATCC-27988 | Kosakonia | radicincitans | 1.94 | Raoultella | ornithinolytica | 1.68 | B | 0 | 0 | 1 | Kosakonia radicincitans | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae | sp-CVUAS-30072 | Kosakonia | cowanii | 1.45 | Escherichia | coli | 1.42 | B | 0 | 0 | 1 | Kosakonia cowanii | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0500,2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia | coli | Escherichia | coli | 2.41 | Escherichia | coli | 2.34 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30053 | 3; type strain | EFex | 2019-02-12 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.53 | Klebsiella | aerogenes | 2.46 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NM20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.48 | Klebsiella | aerogenes | 2.29 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella | grimontii | Klebsiella | oxytoca | 2.42 | Klebsiella | oxytoca | 2.24 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2486.3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella | grimontii | Klebsiella | oxytoca | 2.3 | Klebsiella | oxytoca | 2.27 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2404.3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella | michiganensis | Klebsiella | oxytoca | 2.35 | Klebsiella | oxytoca | 2.27 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella | michiganensis | Klebsiella | oxytoca | 2.28 | Klebsiella | oxytoca | 2.22 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-11-01 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.42 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 2008-04-11 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.33 | Klebsiella | pneumoniae | 2.33 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30104 | 3; type strain | EFEx | 2013-09-17 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.52 | Klebsiella | pneumoniae | 2.39 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | pneumoniae | 2.22 | Klebsiella | pneumoniae | 2.14 | A | 0 | 0 | 1 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | pneumoniae | 1.86 | Klebsiella | pneumoniae | 1.85 | B | 0 | 0 | 1 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | variicola | Klebsiella | variicola | 2.42 | Klebsiella | variicola | 2.32 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia | sacchari | Salmonella | sp | 1.72 | Salmonella | sp | 1.7 | B | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.36 | Plesiomonas | shigelloides | 2.34 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.42 | Plesiomonas | shigelloides | 2.41 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Pluralibacter | gergoviae | Pluralibacter | gergoviae | 2.32 | Pluralibacter | gergoviae | 2.24 | A | 1 | 1 | 0 | Pluralibacter gergoviae | Pluralibacter gergoviae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7763 | 3; NRL Salmonella BfR | DT | 2012-10-16 | Salmonella | enterica | Salmonella | sp | 2.32 | Salmonella | sp | 2.23 | A | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Salmonella | enterica | Salmonella | sp | 2.36 | Salmonella | sp | 2.36 | A | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | II 06 06820 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Shigella | boydii | Escherichia | coli | 2.43 | Escherichia | coli | 2.39 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 3A 07 02503 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Shigella | flexneri | Escherichia | coli | 2.46 | Escherichia | coli | 2.37 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea | agglomerans | Pantoea | agglomerans | 2.52 | Pantoea | agglomerans | 2.47 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 273 | 2; strain collection of the RKI | TFAextr | 2021-11-19 | Erwinia | amylovora | Erwinia | amylovora | 2.35 | Erwinia | amylovora | 2.07 | A | 1 | 1 | 0 | Erwinia amylovora | Erwinia amylovora | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3493 CVUAS | 3; type strain | eDT | 2015-07-09 | Pantoea | agglomerans | Pantoea | agglomerans | 2.35 | Pantoea | agglomerans | 2.31 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea | sp | Pantoea | sp | 1.99 | Pantoea | agglomerans | 1.93 | B | 0 | 0 | 1 | Pantoea sp | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4263,2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea | sp-FF5 | Pantoea | septica | 1.84 | Pantoea | septica | 1.45 | B | 0 | 0 | 1 | Pantoea septica | Pantoea septica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella | sp | Edwardsiella | tarda | 2.47 | Edwardsiella | tarda | 2.47 | A | 0 | 0 | 1 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Edwardsiella | tarda | Edwardsiella | tarda | 2.49 | Edwardsiella | tarda | 2.37 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-18 | Hafnia | alvei | Hafnia | alvei | 2.42 | Hafnia | alvei | 2.32 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia | alvei | Hafnia | alvei | 2.36 | Hafnia | alvei | 2.29 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea | myxofaciens | Cosenzaea | myxofaciens | 2.54 | Cosenzaea | myxofaciens | 2.37 | A | 1 | 1 | 0 | Cosenzaea myxofaciens | Cosenzaea myxofaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella | morganii | Morganella | morganii | 2.56 | Morganella | morganii | 2.55 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella | morganii | Morganella | morganii | 2.68 | Morganella | morganii | 2.61 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | hauseri | Proteus | hauseri | 2.55 | Proteus | hauseri | 2.42 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus | hauseri | Proteus | hauseri | 2.45 | Proteus | hauseri | 2.36 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus | mirabilis | Proteus | mirabilis | 2.56 | Proteus | mirabilis | 2.36 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4479 | 3; typestrain, sequenced (rpoB) | TFAextr | 2022-10-07 | Proteus | mirabilis | Proteus | mirabilis | 2.46 | Proteus | mirabilis | 2.41 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | penneri | Proteus | penneri | 2.43 | Proteus | penneri | 2.4 | A | 1 | 1 | 0 | Proteus penneri | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | vulgaris | 2.24 | Proteus | penneri | 2.13 | B | 0 | 0 | 1 | Proteus vulgaris | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | vulgaris | 2.45 | Proteus | hauseri | 2.33 | B | 0 | 0 | 1 | Proteus vulgaris | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | hauseri | 2.39 | Proteus | hauseri | 2.36 | A | 0 | 0 | 1 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus | terrae | Proteus | hauseri | 2.26 | Proteus | hauseri | 2.23 | A | 0 | 0 | 1 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | terrae | Proteus | hauseri | 2.31 | Proteus | vulgaris | 2.28 | B | 0 | 0 | 1 | Proteus hauseri | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Proteus | vulgaris | Proteus | vulgaris | 2.28 | Proteus | vulgaris | 2.24 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2140 | 3; strain from public collection; sequenced (rpoB) | eDT | 2017-11-30 | Proteus | vulgaris | Proteus | vulgaris | 2.38 | Proteus | vulgaris | 2.33 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia | rettgeri | Providencia | rettgeri | 2.57 | Providencia | rettgeri | 2.52 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia | rettgeri | Providencia | rettgeri | 2.54 | Providencia | rettgeri | 2.31 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25827 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Providencia | stuartii | Providencia | stuartii | 2.32 | Providencia | stuartii | 2.31 | A | 1 | 1 | 0 | Providencia stuartii | Providencia stuartii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4539 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 25.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Providencia stuartii DSM 4539 (T), Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-04-06 | Providencia | stuartii | Providencia | stuartii | 2.4 | Providencia | stuartii | 2.24 | A | 1 | 1 | 0 | Providencia stuartii | Providencia stuartii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.24 | Yersinia | enterocolitica | 2.08 | B | 0 | 0 | 1 | Yersinia frederiksenii | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia | mollaretii | Yersinia | mollaretii | 2.32 | Yersinia | mollaretii | 2.25 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 2028 | 3; public strain collection | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.48 | Yersinia | pseudotuberculosis | 2.48 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.57 | Yersinia | pseudotuberculosis | 2.43 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia | wautersii | Yersinia | pseudotuberculosis | 2.23 | Yersinia | pseudotuberculosis | 2.22 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7849 | 2; sequenced | eDT | 2022-08-23 | Rahnella | aquatilis | Rahnella | aquatilis | 2.51 | Rahnella | aquatilis | 2.46 | A | 1 | 1 | 0 | Rahnella aquatilis | Rahnella aquatilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia | fonticola | Serratia | fonticola | 2.56 | Serratia | fonticola | 2.55 | A | 1 | 1 | 0 | Serratia fonticola | Serratia fonticola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-11-04 | Serratia | grimesii | Serratia | grimesii | 2.42 | Serratia | liquefaciens | 2.03 | B | 0 | 0 | 1 | Serratia grimesii | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.19 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.13 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia | quinivorans | Serratia | proteamaculans | 2.16 | Serratia | liquefaciens | 2.12 | B | 0 | 0 | 1 | Serratia proteamaculans | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6221 | 4; sequenced | DT | 2013-01-08 | Yersinia | aldovae | Yersinia | aldovae | 2.33 | Yersinia | aldovae | 2.3 | A | 1 | 1 | 0 | Yersinia aldovae | Yersinia aldovae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6319 eDT | 2; sequenced (16S) | eDT | 2024-12-04 | Yersinia | aldovae | Yersinia | aldovae | 2.16 | Yersinia | aldovae | 2.04 | A | 1 | 1 | 0 | Yersinia aldovae | Yersinia aldovae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2007-03-09 | Yersinia | aleksiciae | Yersinia | aleksiciae | 2.28 | Yersinia | aleksiciae | 2.28 | A | 1 | 1 | 0 | Yersinia aleksiciae | Yersinia aleksiciae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMB 4355 | 2; 0 | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦ 1:1 HCCA-TA2 | 2007-07-13 | Yersinia | aleksiciae | Yersinia | aleksiciae | 2.23 | Yersinia | aleksiciae | 2.22 | A | 1 | 1 | 0 | Yersinia aleksiciae | Yersinia aleksiciae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia | bercovieri | Yersinia | bercovieri | 2.33 | Yersinia | bercovieri | 2.25 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.56 | Yersinia | bercovieri | 2.45 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4780 | 3; type strain | DT | 2013-05-08 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.49 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.48 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.38 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.34 | Yersinia | enterocolitica | 2.31 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.45 | Yersinia | enterocolitica | 2.42 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.57 | Yersinia | enterocolitica | 2.46 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | entomophaga | Yersinia | ruckeri | 2.06 | Yersinia | ruckeri | 2.06 | A | 0 | 0 | 1 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 21 | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-22 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.41 | Yersinia | frederiksenii | 2.36 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BP 76 | 2; 0 | eDT | 2024-12-04 | Yersinia | intermedia | Yersinia | intermedia | 2.47 | Yersinia | intermedia | 2.37 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia | intermedia | Yersinia | intermedia | 2.51 | Yersinia | intermedia | 2.45 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.57 | Yersinia | kristensenii | 2.53 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 302 | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.46 | Yersinia | kristensenii | 2.39 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-11-25 | Yersinia | massiliensis | Yersinia | massiliensis | 2.33 | Yersinia | massiliensis | 2.06 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB 925699 | 3; sequenced | eDT | 2024-12-04 | Yersinia | massiliensis | Yersinia | massiliensis | 2.39 | Yersinia | massiliensis | 2.15 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18520 | 3; type strain | DT | 16.01.2013 | Yersinia | mollaretii | Yersinia | mollaretii | 2.13 | Yersinia | mollaretii | 2.11 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | nurmii | Yersinia | frederiksenii | 1.89 | Yersinia | ruckeri | 1.87 | B | 0 | 0 | 1 | Yersinia frederiksenii | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | pekkanenii | Yersinia | intermedia | 1.66 | Yersinia | enterocolitica | 1.62 | B | 0 | 0 | 1 | Yersinia intermedia | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 304.2 | 4; WGS sequence (BfR) | EFEx | 2023-06-21 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.2 | Yersinia | pseudotuberculosis | 2.17 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8992 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2009-04-03 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.69 | Yersinia | pseudotuberculosis | 2.67 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia | rohdei | Yersinia | rohdei | 2.46 | Yersinia | rohdei | 2.41 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2007-03-02 | Yersinia | rohdei | Yersinia | rohdei | 2.83 | Yersinia | rohdei | 2.31 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2772 | 3; sequenced | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.52 | Yersinia | ruckeri | 2.35 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.45 | Yersinia | ruckeri | 2.31 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia | similis | Yersinia | pseudotuberculosis | 2.36 | Yersinia | pseudotuberculosis | 2.36 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMB 4354 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2007-07-13 | Yersinia | similis | Yersinia | pseudotuberculosis | 2.52 | Yersinia | pseudotuberculosis | 2.49 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | wautersii | Yersinia | pseudotuberculosis | 2.39 | Yersinia | pseudotuberculosis | 2.39 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella | londiniensis | Lactobacillus | pentosus | 1.33 | Streptococcus | anginosus | 1.23 | B | 0 | 0 | 1 | Lactobacillus pentosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella | pneumophila | Legionella | pneumophila | 1.73 | Legionella | pneumophila | 1.61 | B | 0 | 0 | 1 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4122 | 3; 0 | eDT | 2016-01-20 | Legionella | pneumophila | Legionella | pneumophila | 2.33 | Legionella | pneumophila | 2.29 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella | taurinensis | Legionella | sp | 1.86 | Legionella | erythra | 1.43 | B | 0 | 0 | 1 | Legionella sp | Legionella erythra | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10873,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | aquamarina | 1.54 | Halomonas | aquamarina | 1.41 | B | 0 | 0 | 1 | Halomonas aquamarina | Halomonas aquamarina | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10874,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Dermacoccus | sp[3] | 1.35 | Halomonas | aquamarina | 1.24 | B | 0 | 0 | 1 | Dermacoccus sp[3] | Halomonas aquamarina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LLBB 23UGF306 | 3; sequenced | OS Extr | 2023-10-11 | Volucribacter | psittacicida | Pasteurella | multocida | 1.4 | Bisgaardia | miroungae | 1.33 | B | 0 | 0 | 1 | Pasteurella multocida | Bisgaardia miroungae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) | caballi | Actinobacillus | rossii | 1.63 | Actinobacillus | rossii | 1.59 | B | 0 | 0 | 1 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) | caballi | Actinobacillus | rossii | 2.13 | Actinobacillus | rossii | 2.06 | A | 0 | 0 | 1 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32909 | 3; sequenced | eDT | 2021-02-05 | (Pasteurella) | mairii | Actinobacillus | seminis | 1.99 | Actinobacillus | seminis | 1.9 | B | 0 | 0 | 1 | Actinobacillus seminis | Actinobacillus seminis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) | mairii | Pasteurella | mairii | 2.11 | Actinobacillus | seminis | 2.06 | C | 0 | 0 | 1 | Pasteurella mairii | Actinobacillus seminis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) | testudinis | Chelonobacter | oris | 1.52 | Acinetobacter | junii | 1.49 | B | 0 | 0 | 1 | Chelonobacter oris | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 957 | 3; sequenced | eDT | 2015-12-02 | (Pasteurella) | testudinis | Lentilactobacillus | parabuchneri | 1.33 | Escherichia | coli | 1.33 | B | 0 | 0 | 1 | Lentilactobacillus parabuchneri | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | anseriformium | Actinobacillus | suis | 2.01 | Actinobacillus | suis | 1.99 | A | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UGF 198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus | anseriformium | Actinobacillus | suis | 1.96 | Actinobacillus | suis | 1.93 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus | arthritidis | Actinobacillus | arthritidis | 1.93 | Actinobacillus | suis | 1.69 | B | 0 | 0 | 1 | Actinobacillus arthritidis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.14 | Actinobacillus | equuli | 1.88 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.41 | Actinobacillus | suis | 2.21 | B | 0 | 0 | 1 | Actinobacillus capsulatus | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.34 | Actinobacillus | suis | 2.23 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.28 | Actinobacillus | suis | 2.22 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.36 | Actinobacillus | equuli | 2.3 | A | 1 | 1 | 0 | Actinobacillus equuli | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus | equuli | Actinobacillus | suis | 2.52 | Actinobacillus | suis | 2.46 | A | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus | lignieresii | Actinobacillus | lignieresii | 2.33 | Actinobacillus | pleuropneumoniae | 2.27 | B | 0 | 0 | 1 | Actinobacillus lignieresii | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9945.2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | lignieresii | Actinobacillus | lignieresii | 2.37 | Actinobacillus | pleuropneumoniae | 2.36 | B | 0 | 0 | 1 | Actinobacillus lignieresii | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus | rossii | Actinobacillus | rossii | 2 | Actinobacillus | rossii | 1.94 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus | rossii | Actinobacillus | rossii | 2.64 | Actinobacillus | rossii | 2.56 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32505,2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus | sp-CVUAS-32505,2 | Actinobacillus | lignieresii | 2.08 | Actinobacillus | pleuropneumoniae | 2.07 | B | 0 | 0 | 1 | Actinobacillus lignieresii | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | pleuropneumoniae | 2.1 | Actinobacillus | pleuropneumoniae | 2.05 | A | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | pleuropneumoniae | 2.08 | Actinobacillus | lignieresii | 2.01 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus | sp-MK-2012 | Actinobacillus | suis | 2.22 | Actinobacillus | equuli | 2.15 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-MK-2012 | Actinobacillus | suis | 2.27 | Actinobacillus | equuli | 2.18 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus | suis | Actinobacillus | suis | 2.6 | Actinobacillus | suis | 2.55 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus | suis | Actinobacillus | suis | 2.5 | Actinobacillus | suis | 2.47 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium | gallinarum | Avibacterium | volantium | 1.74 | Avibacterium | endocarditidis | 1.68 | B | 0 | 0 | 1 | Avibacterium volantium | Avibacterium endocarditidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium | gallinarum | Avibacterium | paragallinarum | 1.74 | Avibacterium | paragallinarum | 1.61 | B | 0 | 0 | 1 | Avibacterium paragallinarum | Avibacterium paragallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.52 | Bibersteinia | trehalosi | 2.48 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.47 | Bibersteinia | trehalosi | 2.39 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter | oris | Chelonobacter | oris | 2.41 | Chelonobacter | oris | 2.4 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter | oris | Chelonobacter | oris | 1.96 | Chelonobacter | oris | 1.89 | B | 0 | 0 | 1 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3802,2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter | osteomyelitidis | Haemophilus | haemolyticus | 1.43 | Aggregatibacter | segnis | 1.41 | B | 0 | 0 | 1 | Haemophilus haemolyticus | Aggregatibacter segnis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31212 | 3; sequenced | EFEx | 2018-10-16 | Exercitatus | varius | Actinobacillus | rossii | 1.59 | Actinobacillus | rossii | 1.57 | B | 0 | 0 | 1 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Exercitatus | varius | Actinobacillus | succinogenes | 1.59 | Actinobacillus | rossii | 1.56 | B | 0 | 0 | 1 | Actinobacillus succinogenes | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia | canicola | Frederiksenia | canicola | 1.98 | Frederiksenia | canicola | 1.92 | B | 0 | 0 | 1 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia | canicola | Frederiksenia | canicola | 2.32 | Frederiksenia | canicola | 2.18 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium | anatis | Gallibacterium | anatis | 2.07 | Gallibacterium | anatis | 2.04 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium | anatis | Gallibacterium | anatis | 2.45 | Gallibacterium | anatis | 2.36 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium | sp-F151 | Gallibacterium | anatis | 1.44 | Ligilactobacillus | murinus | 1.33 | B | 0 | 0 | 1 | Gallibacterium anatis | Ligilactobacillus murinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium | trehalosifermentans | Gallibacterium | anatis | 1.58 | Gallibacterium | anatis | 1.58 | B | 0 | 0 | 1 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus | haemoglobinophilus | Haemophilus | haemoglobinophilus | 1.74 | Haemophilus | haemoglobinophilus | 1.55 | B | 0 | 0 | 1 | Haemophilus haemoglobinophilus | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4690 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 03.06.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Haemophilus influenzae DSM 4690, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-04 | Haemophilus | influenzae | Haemophilus | influenzae | 2.43 | Haemophilus | influenzae | 2.31 | A | 1 | 1 | 0 | Haemophilus influenzae | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | haemoglobinophilus | 2.28 | Haemophilus | haemoglobinophilus | 1.99 | A | 0 | 0 | 1 | Haemophilus haemoglobinophilus | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | haemoglobinophilus | 2.38 | Haemophilus | haemoglobinophilus | 2.34 | A | 0 | 0 | 1 | Haemophilus haemoglobinophilus | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | P 1066 2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus | somni | Histophilus | somni | 1.82 | Histophilus | somni | 1.77 | B | 0 | 0 | 1 | Histophilus somni | Histophilus somni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia | caviae | Mannheimia | caviae | 2.22 | Mannheimia | haemolytica | 1.76 | A | 1 | 1 | 0 | Mannheimia caviae | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1064,4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia | glucosida | Mannheimia | haemolytica | 2.09 | Mannheimia | haemolytica | 2.08 | A | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia | glucosida | Mannheimia | haemolytica | 2.36 | Mannheimia | haemolytica | 2.36 | A | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia | granulomatis | Mannheimia | granulomatis | 2.05 | Mannheimia | haemolytica | 1.77 | A | 1 | 1 | 0 | Mannheimia granulomatis | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia | granulomatis | Mannheimia | granulomatis | 2.03 | Actinobacillus | lignieresii | 1.87 | C | 0 | 0 | 1 | Mannheimia granulomatis | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1738 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.45 | Mannheimia | haemolytica | 2.27 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3952 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.49 | Mannheimia | haemolytica | 2.29 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.47 | Mannheimia | haemolytica | 2.44 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6197.2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.52 | Mannheimia | haemolytica | 2.51 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia | pernigra | Mannheimia | haemolytica | 1.5 | Mannheimia | haemolytica | 1.39 | B | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia | pernigra | Mannheimia | caviae | 1.33 | Glaesserella | parasuis | 1.3 | B | 0 | 0 | 1 | Mannheimia caviae | Glaesserella parasuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.3 | Mannheimia | ruminalis | 2.22 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.19 | Mannheimia | ruminalis | 2.13 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10396.2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia | sp-10-12559-1 | Mannheimia | caviae | 2.03 | Mannheimia | haemolytica | 1.73 | A | 0 | 0 | 1 | Mannheimia caviae | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10183.3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia | sp-14UCF195 | Mannheimia | granulomatis | 2.24 | Mannheimia | haemolytica | 1.73 | A | 0 | 0 | 1 | Mannheimia granulomatis | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia | sp-14UCF195 | Mannheimia | granulomatis | 2.02 | Mannheimia | haemolytica | 1.85 | A | 0 | 0 | 1 | Mannheimia granulomatis | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia | sp-15USF15 | Mannheimia | caviae | 2.07 | Mannheimia | varigena | 2 | B | 0 | 0 | 1 | Mannheimia caviae | Mannheimia varigena | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia | sp-15USF15 | Mannheimia | caviae | 2.13 | Mannheimia | haemolytica | 1.98 | A | 0 | 0 | 1 | Mannheimia caviae | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia | sp-CVUAS-11240 | Mannheimia | varigena | 1.87 | Mannheimia | haemolytica | 1.74 | B | 0 | 0 | 1 | Mannheimia varigena | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia | sp-CVUAS-1496 | Mannheimia | haemolytica | 2.31 | Mannheimia | haemolytica | 2.24 | A | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.43 | Haemophilus | haemoglobinophilus | 1.74 | C | 0 | 0 | 1 | Necropsobacter rosorum | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.31 | Haemophilus | haemoglobinophilus | 1.75 | C | 0 | 0 | 1 | Necropsobacter rosorum | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella | semolina | Actinobacillus | equuli | 1.57 | Bisgaardia | hudsonensis | 1.44 | B | 0 | 0 | 1 | Actinobacillus equuli | Bisgaardia hudsonensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella | semolina | Mannheimia | caviae | 1.49 | Glaesserella | parasuis | 1.4 | B | 0 | 0 | 1 | Mannheimia caviae | Glaesserella parasuis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4060.3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella | aerogenes | Pasteurella | aerogenes | 2.01 | Actinobacillus | seminis | 2 | C | 0 | 0 | 1 | Pasteurella aerogenes | Actinobacillus seminis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | canis | Pasteurella | canis | 2.07 | Pasteurella | canis | 2.04 | A | 1 | 1 | 0 | Pasteurella canis | Pasteurella canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella | multocida | Pasteurella | multocida | 2.32 | Pasteurella | multocida | 2.3 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32055 | 3; specific PCR | EFEx | 2024-09-25 | Pasteurella | multocida | Pasteurella | multocida | 2.35 | Pasteurella | multocida | 2.35 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella | multocida | Pasteurella | multocida | 2.45 | Pasteurella | multocida | 2.4 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.32 | Pasteurella | multocida | 2.28 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.42 | Pasteurella | multocida | 2.32 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella | multocida | Pasteurella | multocida | 2.35 | Pasteurella | multocida | 2.31 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | Bisgaard-Taxon-40 | Haemophilus | parainfluenzae | 1.36 | Pseudomonas | marginalis | 1.31 | B | 0 | 0 | 1 | Haemophilus parainfluenzae | Pseudomonas marginalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9987,2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae | canine-oral-taxon-080 | Frederiksenia | canicola | 1.83 | Frederiksenia | canicola | 1.8 | B | 0 | 0 | 1 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 261,4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae | sp-CVUAS-261,4 | Klebsiella | oxytoca | 1.47 | Pseudomonas | savastanoi | 1.42 | B | 0 | 0 | 1 | Klebsiella oxytoca | Pseudomonas savastanoi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae | sp-CVUAS-31988 | Haemophilus | haemolyticus | 1.65 | Haemophilus | haemolyticus | 1.61 | B | 0 | 0 | 1 | Haemophilus haemolyticus | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34093 | 2; sequenced | EFex | 2022-08-05 | Pasteurellaceae | sp-CVUAS-34093 | Actinobacillus | lignieresii | 1.55 | Mannheimia | caviae | 1.49 | B | 0 | 0 | 1 | Actinobacillus lignieresii | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | sp-CVUAS-4868 | Enterococcus | faecalis | 1.34 | Paucilactobacillus | suebicus | 1.3 | B | 0 | 0 | 1 | Enterococcus faecalis | Paucilactobacillus suebicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 724,3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae | sp-CVUAS-724,3 | Rodentibacter | heylii | 1.47 | Rodentibacter | heylii | 1.44 | B | 0 | 0 | 1 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae | sp-CVUAS-9305 | Actinobacillus | suis | 1.33 | Actinobacillus | suis | 1.3 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9967,2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae | sp-CVUAS-9967,2 | Haemophilus | haemolyticus | 1.74 | Haemophilus | parainfluenzae | 1.62 | B | 0 | 0 | 1 | Haemophilus haemolyticus | Haemophilus parainfluenzae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae | sp-LHL-U-309 | Lactobacillus | pentosus | 1.45 | Chelonobacter | oris | 1.39 | B | 0 | 0 | 1 | Lactobacillus pentosus | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL U-309 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-309 | Chelonobacter | oris | 1.66 | Chelonobacter | oris | 1.58 | B | 0 | 0 | 1 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL U-3444 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-3444 | Chelonobacter | oris | 1.54 | Chelonobacter | oris | 1.51 | B | 0 | 0 | 1 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae | sp-MCCM-2539 | Haemophilus | pittmaniae | 1.53 | Citrobacter | koseri | 1.51 | B | 0 | 0 | 1 | Haemophilus pittmaniae | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.4 | Rodentibacter | heylii | 2.25 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.5 | Rodentibacter | heylii | 2.48 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium | mucosae | Staphylococcus | sciuri | 1.42 | Photorhabdus | akhurstii | 1.36 | B | 0 | 0 | 1 | Staphylococcus sciuri | Photorhabdus akhurstii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium | mucosae | Sphingomonas | aurantiaca | 1.28 | Xanthomonas | hortorum | 1.26 | B | 0 | 0 | 1 | Sphingomonas aurantiaca | Xanthomonas hortorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8966 | 3; sequenced | DT | 2013-07-23 | Volucribacter | amazonae | Haemophilus | influenzae | 1.53 | Haemophilus | influenzae | 1.49 | B | 0 | 0 | 1 | Haemophilus influenzae | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 308,2 | 3; sequenced | DT | 2013-11-06 | Psittacicella | hinzii | Agromyces | cerinus | 1.43 | Clostridium | tetani | 1.29 | B | 0 | 0 | 1 | Agromyces cerinus | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-14 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.51 | Acinetobacter | baumannii | 2.48 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-09-03 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.46 | Acinetobacter | baumannii | 2.41 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter | defluvii | Acinetobacter | defluvii | 2.16 | Acinetobacter | defluvii | 2.05 | A | 1 | 1 | 0 | Acinetobacter defluvii | Acinetobacter defluvii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | guillouiae | Acinetobacter | guillouiae | 2.34 | Acinetobacter | guillouiae | 2.24 | A | 1 | 1 | 0 | Acinetobacter guillouiae | Acinetobacter guillouiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | johnsonii | Acinetobacter | johnsonii | 2.3 | Acinetobacter | johnsonii | 2.26 | A | 1 | 1 | 0 | Acinetobacter johnsonii | Acinetobacter johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | junii | Acinetobacter | junii | 2.33 | Acinetobacter | junii | 2.31 | A | 1 | 1 | 0 | Acinetobacter junii | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7714 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | junii | Acinetobacter | junii | 2.49 | Acinetobacter | junii | 2.45 | A | 1 | 1 | 0 | Acinetobacter junii | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | lwoffii | Acinetobacter | lwoffii | 2.35 | Acinetobacter | lwoffii | 2.27 | A | 1 | 1 | 0 | Acinetobacter lwoffii | Acinetobacter lwoffii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30502 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter lwoffii DSM 30502, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Acinetobacter | lwoffii | Acinetobacter | lwoffii | 2.13 | Acinetobacter | lwoffii | 1.9 | A | 1 | 1 | 0 | Acinetobacter lwoffii | Acinetobacter lwoffii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9312 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 28.04.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter pittii DSM 9312, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-05-06 | Acinetobacter | pittii | Acinetobacter | pittii | 2.4 | Acinetobacter | pittii | 2.36 | A | 1 | 1 | 0 | Acinetobacter pittii | Acinetobacter pittii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9318 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter pittii DSM 9318, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Acinetobacter | pittii | Acinetobacter | lactucae | 2.28 | Acinetobacter | lactucae | 2.14 | A | 0 | 0 | 1 | Acinetobacter lactucae | Acinetobacter lactucae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7379,3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.29 | Acinetobacter | pseudolwoffii | 2.24 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.08 | Acinetobacter | pseudolwoffii | 2.07 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32366 | 3; sequenced | eDT | 2020-06-19 | Acinetobacter | sp-CVUAS-32366 | Clostridium | baratii | 1.47 | Acinetobacter | cumulans | 1.47 | B | 0 | 0 | 1 | Clostridium baratii | Acinetobacter cumulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | variabilis | Acinetobacter | variabilis | 2.17 | Acinetobacter | variabilis | 2.15 | A | 1 | 1 | 0 | Acinetobacter variabilis | Acinetobacter variabilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | bovoculi | 2.16 | Moraxella | ovis | 2.1 | B | 0 | 0 | 1 | Moraxella bovoculi | Moraxella ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | bovoculi | 2.06 | Moraxella | ovis | 1.96 | A | 1 | 1 | 0 | Moraxella bovoculi | Moraxella ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella | catarrhalis | Moraxella | catarrhalis | 2.53 | Moraxella | catarrhalis | 2.51 | A | 1 | 1 | 0 | Moraxella catarrhalis | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11674,3 | 3; sequenced | eDT | 2022-03-22 | Moraxella | cuniculi | Moraxella | catarrhalis | 1.63 | Moraxella | catarrhalis | 1.51 | B | 0 | 0 | 1 | Moraxella catarrhalis | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella | osloensis | Moraxella | osloensis | 1.94 | Moraxella | osloensis | 1.87 | B | 0 | 0 | 1 | Moraxella osloensis | Moraxella osloensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella | ovis | Moraxella | bovoculi | 2.32 | Moraxella | ovis | 1.97 | A | 0 | 0 | 1 | Moraxella bovoculi | Moraxella ovis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | frigidicola | Pseudomonas | rhodesiae | 1.33 | Acinetobacter | parvus | 1.24 | B | 0 | 0 | 1 | Pseudomonas rhodesiae | Acinetobacter parvus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter | pasteurii | Psychrobacter | pasteurii | 2.04 | Psychrobacter | phenylpyruvicus | 1.92 | A | 1 | 1 | 0 | Psychrobacter pasteurii | Psychrobacter phenylpyruvicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.35 | Psychrobacter | faecalis | 2.19 | B | 0 | 0 | 1 | Psychrobacter pulmonis | Psychrobacter faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | maritimus | 1.78 | Psychrobacter | pulmonis | 1.71 | B | 0 | 0 | 1 | Psychrobacter maritimus | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas | (alkylphenolica) | Pseudomonas | japonica | 1.59 | Pseudomonas | graminis | 1.57 | B | 0 | 0 | 1 | Pseudomonas japonica | Pseudomonas graminis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas | [baetica] | Pseudomonas | koreensis | 2.06 | Pseudomonas | rhodesiae | 1.94 | A | 0 | 0 | 1 | Pseudomonas koreensis | Pseudomonas rhodesiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.33 | Pseudomonas | aeruginosa | 2.28 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.46 | Pseudomonas | aeruginosa | 2.38 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-02-20 | Pseudomonas | chlororaphis | Pseudomonas | chlororaphis | 2.44 | Pseudomonas | chlororaphis | 2.3 | A | 1 | 1 | 0 | Pseudomonas chlororaphis | Pseudomonas chlororaphis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 23974 | 3; public strain collection | TFAextr | 2021-09-08 | Pseudomonas | putida | Pseudomonas | putida | 2.33 | Pseudomonas | putida | 2.22 | A | 1 | 1 | 0 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5190 | 3; type strain | TFAextr | 2022-06-04 | Pseudomonas | stutzeri | Pseudomonas | stutzeri | 2.61 | Pseudomonas | stutzeri | 2.27 | A | 1 | 1 | 0 | Pseudomonas stutzeri | Pseudomonas stutzeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2012-04-02 | Francisella | tularensis | Arthrobacter | pyridinolis | 1.43 | Lactococcus | lactis | 1.33 | B | 0 | 0 | 1 | Arthrobacter pyridinolis | Lactococcus lactis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-12-12 | Allofrancisella | guangzhouensis | Paracoccus | denitrificans | 1.33 | Enterococcus | faecalis | 1.29 | B | 0 | 0 | 1 | Paracoccus denitrificans | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-11-18 | Francisella | hispaniensis | Lactobacillus | plantarum | 1.42 | Francisella | philomiragia | 1.38 | B | 0 | 0 | 1 | Lactobacillus plantarum | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18777 | 3; strain collection of the RKI | EFEx | 2013-06-20 | Francisella | noatunensis | Francisella | philomiragia | 1.78 | Francisella | philomiragia | 1.69 | B | 0 | 0 | 1 | Francisella philomiragia | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-11-25 | Francisella | orientalis | Francisella | philomiragia | 1.57 | Francisella | philomiragia | 1.53 | B | 0 | 0 | 1 | Francisella philomiragia | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7535 | 3; type strain | EFEx | 2013-06-20 | Francisella | philomiragia | Francisella | philomiragia | 1.9 | Francisella | philomiragia | 1.86 | B | 0 | 0 | 1 | Francisella philomiragia | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4680,2 | 3; FLI: specific PCR | DT | 2013-04-02 | Francisella | tularensis | Morganella | morganii | 1.27 | Morganella | morganii | 1.26 | B | 0 | 0 | 1 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 2013-04-23 | Francisella | tularensis | Clostridium | beijerinckii | 1.49 | Loigolactobacillus | coryniformis | 1.46 | B | 0 | 0 | 1 | Clostridium beijerinckii | Loigolactobacillus coryniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 2013-04-26 | Francisella | tularensis | Agrobacterium | radiobacter | 1.46 | Latilactobacillus | fuchuensis | 1.43 | B | 0 | 0 | 1 | Agrobacterium radiobacter | Latilactobacillus fuchuensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | F064 | 3; strain collection of the FLI | EFEx | 2013-06-20 | Francisella | tularensis | Lentilactobacillus | parabuchneri | 1.6 | Haemophilus | influenzae | 1.41 | B | 0 | 0 | 1 | Lentilactobacillus parabuchneri | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 2017-06-01 | Francisella | tularensis | Lactobacillus | plantarum | 1.29 | Pandoraea | nosoerga | 1.27 | B | 0 | 0 | 1 | Lactobacillus plantarum | Pandoraea nosoerga | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio | aestuarianus | Vibrio | aestuarianus | 2.39 | Vibrio | aestuarianus | 2.39 | A | 1 | 1 | 0 | Vibrio aestuarianus | Vibrio aestuarianus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio | anguillarum | Vibrio | anguillarum | 2.34 | Vibrio | anguillarum | 2.3 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio | campbelli | Vibrio | harveyi | 2.03 | Vibrio | harveyi | 1.97 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio | campbelli | Vibrio | harveyi | 2.17 | Vibrio | harveyi | 2.13 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16-VB00067-0_BfR 1_eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio | cholerae | Vibrio | mimicus | 1.93 | Vibrio | albensis | 1.87 | B | 0 | 0 | 1 | Vibrio mimicus | Vibrio albensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 45V066-10 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 1.52 | Vibrio | alginolyticus | 1.4 | B | 0 | 0 | 1 | Vibrio cincinnatiensis | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 2.14 | Vibrio | europaeus | 1.44 | A | 1 | 1 | 0 | Vibrio cincinnatiensis | Vibrio europaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.08 | Vibrio | diazotrophicus | 1.98 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio | fluvialis | Vibrio | fluvialis | 2.42 | Vibrio | fluvialis | 2.16 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio | fluvialis | Vibrio | fluvialis | 2.22 | Vibrio | fluvialis | 2.18 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 2012-08-16 | Vibrio | harveyi | Vibrio | harveyi | 2.28 | Vibrio | harveyi | 2.18 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio | harveyi | Vibrio | harveyi | 2.32 | Vibrio | harveyi | 2.04 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | P340-7 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio | jasicida | Vibrio | harveyi | 2.49 | Vibrio | harveyi | 2.43 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | P340-9 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio | jasicida | Vibrio | harveyi | 2.29 | Vibrio | harveyi | 2.25 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 12-VBH0498 | 4; WGS sequenced (BfR) | DT | 2021-02-02 | Vibrio | metoecus | Vibrio | mimicus | 2.34 | Vibrio | diazotrophicus | 1.58 | A | 0 | 0 | 1 | Vibrio mimicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00016 | 4; WGS sequenced (BfR) | DT | 2021-02-24 | Vibrio | metoecus | Vibrio | mimicus | 2.06 | Pseudomonas | cichorii | 1.46 | A | 0 | 0 | 1 | Vibrio mimicus | Pseudomonas cichorii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metschnikovii | Vibrio | metschnikovii | 1.67 | Vibrio | metschnikovii | 1.58 | B | 0 | 0 | 1 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00299 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | navarrensis | Vibrio | navarrensis | 2.33 | Vibrio | navarrensis | 2.21 | A | 1 | 1 | 0 | Vibrio navarrensis | Vibrio navarrensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | navarrensis | Vibrio | navarrensis | 2.16 | Vibrio | navarrensis | 1.99 | A | 1 | 1 | 0 | Vibrio navarrensis | Vibrio navarrensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio | owensii | Vibrio | harveyi | 2.19 | Vibrio | harveyi | 2.17 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00254-0_BfR 18_DT | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.28 | Vibrio | parahaemolyticus | 2.26 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio | vulnificus | Vibrio | vulnificus | 2.32 | Vibrio | vulnificus | 2.3 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33180 | 3; sequenced | DT | 2021-08-04 | Photobacterium | damselae | Photobacterium | damselae | 2.22 | Photobacterium | damselae | 2.02 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium | damselae | Photobacterium | damselae | 2.14 | Photobacterium | damselae | 2.07 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1266,2 | 3; sequenced | EFEx | 2019-09-25 | Photobacterium | piscicola | Photobacterium | iliopiscarium | 1.57 | Photobacterium | leiognathi | 1.56 | B | 0 | 0 | 1 | Photobacterium iliopiscarium | Photobacterium leiognathi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 842.2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium | swingsii | Providencia | alcalifaciens | 1.37 | Arthrobacter | luteolus | 1.37 | B | 0 | 0 | 1 | Providencia alcalifaciens | Arthrobacter luteolus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7079 | 3; spezific PCR (CVFR) | eDT | 2022-06-23 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.06 | Vibrio | alginolyticus | 1.91 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2171 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 2012-07-12 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.3 | Vibrio | alginolyticus | 2.29 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | A8-12 | 3; gyrB sequence | eDT | 2022-06-23 | Vibrio | anguillarum | Vibrio | anguillarum | 2.22 | Vibrio | anguillarum | 2.2 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10227 | 3; sequenced | eDT | 2022-06-23 | Vibrio | cholerae | Vibrio | mimicus | 2.04 | Vibrio | albensis | 1.95 | A | 0 | 0 | 1 | Vibrio mimicus | Vibrio albensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8995 | 3; sequenced | eDT | 2013-04-10 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.14 | Vibrio | diazotrophicus | 1.9 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1754 | 3; sequenced | eDT | 2013-04-16 | Vibrio | furnissii | Vibrio | furnissii | 2.2 | Vibrio | furnissii | 2.05 | A | 1 | 1 | 0 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14383 | 3; public strain collection, sequenced | eDT | 2022-06-23 | Vibrio | furnissii | Vibrio | furnissii | 2.21 | Vibrio | furnissii | 2.06 | A | 1 | 1 | 0 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 619 | 2; 0 | eDT | 2022-06-23 | Vibrio | metschnikovii | Vibrio | metschnikovii | 1.82 | Vibrio | metschnikovii | 1.64 | B | 0 | 0 | 1 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7872 | 3; sequenced | eDT | 2022-06-23 | Vibrio | mimicus | Vibrio | mimicus | 2.22 | Vibrio | mimicus | 1.34 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19130 | 4; type strain; WGS sequenced | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 2011-01-24 | Vibrio | mimicus | Vibrio | mimicus | 2.08 | Vibrio | mimicus | 1.98 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10027 | 3; type strain | eDT | 2022-06-23 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.21 | Vibrio | parahaemolyticus | 2.16 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 871,2 | 3; sequenced | eDT | 2014-08-21 | Vibrio | sp-CVUAS-871,2 | Vibrio | mimicus | 1.45 | Weissella | viridescens | 1.33 | B | 0 | 0 | 1 | Vibrio mimicus | Weissella viridescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778,2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio | tubiashii | Vibrio | europaeus | 1.65 | Vibrio | brasiliensis | 1.56 | B | 0 | 0 | 1 | Vibrio europaeus | Vibrio brasiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778,2 | 3; sequenced | eDT | 2019-10-24 | Vibrio | tubiashii | Vibrio | europaeus | 1.77 | Vibrio | brasiliensis | 1.52 | B | 0 | 0 | 1 | Vibrio europaeus | Vibrio brasiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LVM05 LVU05 | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-14 | Vibrio | vulnificus | Vibrio | vulnificus | 2.46 | Vibrio | vulnificus | 2.31 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas | sp-G9077 | Wohlfahrtiimonas | chitiniclastica | 2.35 | Wohlfahrtiimonas | chitiniclastica | 2.31 | A | 0 | 0 | 1 | Wohlfahrtiimonas chitiniclastica | Wohlfahrtiimonas chitiniclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria | indica | Ignatzschineria | indica | 2.18 | Ignatzschineria | indica | 1.92 | A | 1 | 1 | 0 | Ignatzschineria indica | Ignatzschineria indica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria | ureiclastica | Ignatzschineria | ureiclastica | 2.42 | Ignatzschineria | larvae | 2.15 | B | 0 | 0 | 1 | Ignatzschineria ureiclastica | Ignatzschineria larvae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas | chelatiphaga | Stenotrophomonas | maltophilia | 1.61 | Stenotrophomonas | acidaminiphila | 1.54 | B | 0 | 0 | 1 | Stenotrophomonas maltophilia | Stenotrophomonas acidaminiphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-09-03 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.26 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-05-14 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.22 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 1.79 | Stenotrophomonas | maltophilia | 1.63 | B | 0 | 0 | 1 | Stenotrophomonas rhizophila | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-18 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.4 | Stenotrophomonas | acidaminiphila | 1.51 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas acidaminiphila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4239,2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas | tumulicola | Stenotrophomonas | acidaminiphila | 1.49 | Stenotrophomonas | maltophilia | 1.43 | B | 0 | 0 | 1 | Stenotrophomonas acidaminiphila | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0

## Report (Species)
Targeted Identification
following the Guidelines for Validating Species Identifications using MALDI-TOF-MS pursuant to §64 of the German Food and Feed code (LFGB), coordinated by the Federal Office of Consumer Protection and Food Safety, Berlin, Germany (BVL) 28 Oct 2022
Parameter: Corynebacterium pseudotuberculosis
Version: | 4
generated by | Dyk | checked by | Rau
generated on | 2025-02-19 | checked on | 2025-03-10
Application | isolates from food & veterinary diagnostics
Environment | Food microbiology, veterinary diagnostics
Applied database / software version |  | MBT(K) / BRU 11897 | BT 4.0
Validation set
Number of target parameter isolates: 183 |  | number of non-target parameter isolates: 1408
cultivation conditions: SBA, 30°/37° C, 24h, aerob/anaerob/mikroaerophil
preparation acc. to manufacturer instructions: DT, eDT, EFex; RKI-specific preparation TFA-method
Only identified samples are used for determination of the inclusivity and exclusivity.
True-positive: 1. Hit = targeted parameter shows classifcation score ≥ 2 and score 2. Hit < 1,7 or score 2. Hit ≥ 1,7 & genus is the same as 1. Hit or score 2. Hit ≥ 2 & species is the same as 1. Hit |  | True-negative: 1. Hit without assignment to the target parameter with score ≥ 2
False-negative: true-positive criteria are not fulfilled |  | False-positive: true-negative criteria are not fullfilled
Validation result | Target parameter: | Corynebacterium pseudotuberculosis |  | target | fullfilled
Total number of spectra | 80
of which with ID result | 80 | Rate of identified spectra | 100% | 20 | ja | 1
Total number of own spectra | 74 |  |  | 16 | ja | 1
With respect to identified spectra: |  | Inclusivity
positive accordance (PA) | 80 | True positive rate | 100% | 95% | ja | 1
negative deviation (ND) | 0 | False negative rate | 0% | 1% | ja | 1
Non-Target-Parameter
Total number of spectra | 1529
of which with ID result | 723 | Rate of identified spectra | 47.29% | 30 | ja | 1
Total number of own spectra | 425 |  |  | 144.6 | ja | 1
With respect to identified spectra: |  | Exclusivity
negative accordance (NA) | 723 | True negative rate | 100% | 99% | ja | 1
positive deviation (PD) | 0 | False-Positive Rate | 0% | 1% | ja | 1
Using the complete database, out of 80 identified target parameter spectra 80 (=100%) were identified correctly (inclusivity). 0 (=0%) of identified spectra of the parameter were assigned to a wrong species.
Using the complete database, out of 723 identified non-target parameter spectra 100% were counted correctly as negative (exclusivity). Out of these 723 spectra 0 (=0%) were allocated incorrectly to the parameter.
Criteria for the | targeted identification | are thus | fulfilled
comment:
This validation is described in Rau et al., 2025
Released for site | For department | On (date) | By (signature)
CVUA Stuttgart / Germany | T, MT, D | 2025-03-10 | Rau
The procedure is suitable for the intended purpose and is thereby released for usage at site.
Analytical parameter added to LIMS:
A signed version of this validation report is kept within the generating lab.
Other microbes that may be present in this matrix.
Other germs growing on the used agar
Non parameter isolates/materials that can provide similar spectra.
yes
no
Other microbes that may be present in this matrix.
Other germs growing on the used agar
Non parameter isolates/materials that can provide similar spectra.
yes
no
Other microbes that may be present in this matrix.
Other germs growing on the used agar
Non parameter isolates/materials that can provide similar spectra.
yes
no
Other microbes that may be present in this matrix.
Other germs growing on the used agar
Non parameter isolates/materials that can provide similar spectra.
yes
no