## Settings
Settings für die Validierung
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Institution / Labor | CVUA S
Parameter | Gattung (Genus) |  | Art (Spezies)
|  | Staphylococcus | haemolyticus-borealis
Unterscheidungsgrenzen
| Score zur Unterscheidung der Gattung | 1.7 | Score zur Unterscheidung der Spezies | 2

## Parameter (Spezies)
Parameter: | Staphylococcus | haemolyticus-borealis | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | nicht zugeordnet | richtig-positiv | falsch-negativ | falsch zugeordnet
|  |  | 0 | 0 | 09-00512 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-05 | Staphylococcus | haemolyticus | 2.23 | Staphylococcus | haemolyticus | 2.06 | A | 1 | 0 | 1 | 0 | 0
Zahl aller Spektren | 23 |  | 1 | 1 | 191e-2 | 3; 16S rDNA-sequencing | eDT | 2016-08-31 | Staphylococcus | haemolyticus | 2.17 | Staphylococcus | haemolyticus | 2.13 | A | 1 | 0 | 1 | 0 | 0
davon Identifiziert | 22 |  | 0 | 0 | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2011-08-04 | Staphylococcus | haemolyticus | 2.46 | Staphylococcus | haemolyticus | 2.43 | A | 1 | 0 | 1 | 0 | 0
nicht identifiziert: | 1 |  | 1 | 1 | coa008a | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | haemolyticus | 2.16 | Staphylococcus | haemolyticus | 1.73 | A | 1 | 0 | 1 | 0 | 0
von allen identifizierten Isolaten |  |  | 1 | 1 | coa015 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | haemolyticus | 2.11 | Staphylococcus | haemolyticus | 2.11 | A | 1 | 0 | 1 | 0 | 0
richtig-positiv: | 22 |  | 1 | 1 | coa022 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus | haemolyticus | 2.32 | Staphylococcus | haemolyticus | 2.06 | A | 1 | 0 | 1 | 0 | 0
falsch zugeordnet | 0 |  | 1 | 1 | coa095 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | haemolyticus | 2.28 | Staphylococcus | haemolyticus | 1.67 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | coa098 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | haemolyticus | 2.51 | Staphylococcus | haemolyticus | 1.96 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 0 | coa127 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | haemolyticus | 1.91 | Staphylococcus | borealis | 1.89 | B | 0 | 1 | 0 | 0 | 0
Scorewertgrenzen: | 2 | Species | 1 | 1 | coa142 | 3; 16S rDNA-sequencing | DT | 2013-03-20 | Staphylococcus | haemolyticus | 2.01 | Staphylococcus | haemolyticus | 1.93 | A | 1 | 0 | 1 | 0 | 0
| 1.7 | Genus | 1 | 1 | CVUAS 2927 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | haemolyticus | 2.36 | Staphylococcus | haemolyticus | 2.25 | A | 1 | 0 | 1 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 1 | 1 | CVUAS 3076 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | borealis | 2.33 | Staphylococcus | borealis | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3443 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | borealis | 2.21 | Staphylococcus | haemolyticus | 2.2 | B | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3444 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | haemolyticus | 2.24 | Staphylococcus | haemolyticus | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3564 | 3; 16S rDNA-sequencing | DT | 2013-04-10 | Staphylococcus | haemolyticus | 2.25 | Staphylococcus | haemolyticus | 1.98 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3693 | 3; 16S rDNA-sequencing | DT | 2013-02-05 | Staphylococcus | haemolyticus | 2.28 | Staphylococcus | haemolyticus | 2.05 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3694 | 3; 16S rDNA-sequencing | DT | 2013-02-05 | Staphylococcus | haemolyticus | 2.17 | Staphylococcus | borealis | 1.67 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 3777 | 3; 16S rDNA-sequencing | DT | 2015-11-05 | Staphylococcus | haemolyticus | 2.22 | Staphylococcus | haemolyticus | 2.03 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4637 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | haemolyticus | 2.12 | Staphylococcus | haemolyticus | 2.1 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | haemolyticus | 2.29 | Staphylococcus | haemolyticus | 2.22 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TS 062 | 2; biochemical tests (API Staph ID32) | DT | 2013-04-30 | Staphylococcus | haemolyticus | 2.23 | Staphylococcus | haemolyticus | 1.64 | A | 1 | 0 | 1 | 0 | 0
davon mit ID-Ergebnis | 22 |  | 1 | 1 | TS 171 | 2; biochemical tests (API Staph ID32) | DT | 2013-04-30 | Staphylococcus | haemolyticus | 2.32 | Staphylococcus | haemolyticus | 1.97 | A | 1 | 0 | 1 | 0 | 0
Positive Übereinstimmungen (PA) | 22 |  | 1 | 1 | TS 301 | 2; biochemical tests (API Staph ID32) | DT | 2013-04-30 | Staphylococcus | haemolyticus | 2.43 | Staphylococcus | haemolyticus | 2.02 | A | 1 | 0 | 1 | 0 | 0
Negative Abweichungen (ND) | 0
manuel erfasst Rau
nicht identifiziert in % | 4.348%
falsch zugeordnert in % | 0%

## #Parameter (Spezies) DB
Parameter: | Staphylococcus haemolyticus-borealis |  | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | verwendete Gattung | verwendete Art | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | richtige ID | keine richtige ID | 1. Hit | 2. Hit | 1. Hit Treffer Parameter | 2. Hit Treffer Parameter | falsch als Parameter identifiziert | PM unter den ersten 2 Hits
|  |  | 0 | 0 | A281_1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-05-31 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 2.37 | Shouchella | clausii | 1.98 | C | 0 | 0 | 1 | Alkalihalobacillus clausii | Shouchella clausii | 0 | 0 | 0 | 0
Anzahl Isolate: | 2391 |  | 1 | 1 | CVUAS 31071.2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 2.23 | Shouchella | clausii | 1.69 | A | 1 | 1 | 0 | Alkalihalobacillus clausii | Shouchella clausii | 0 | 0 | 0 | 0
davon Identifiziert | 1599 |  | 0 | 0 | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 2.37 | Anoxybacillus | contaminans | 1.21 | A | 1 | 1 | 0 | Alkalihalobacillus clausii | Anoxybacillus contaminans | 0 | 0 | 0 | 0
von allen identifizierten Isolaten |  |  | 0 | 0 | Oman 24 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-06-08 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 2.12 | Shouchella | clausii | 1.6 | A | 1 | 1 | 0 | Alkalihalobacillus clausii | Shouchella clausii | 0 | 0 | 0 | 0
richtige ID | 1599 |  | 0 | 0 | DSM 13666 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Alkalihalobacillus okuhidensis DSM 13666: preparation and measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-19 | Alkalihalobacillus | okuhidensis | Alkalihalobacillus | okuhidensis | 2.79 | Alkalihalobacillus | halodurans | 2.41 | B | 0 | 0 | 1 | Alkalihalobacillus okuhidensis | Alkalihalobacillus halodurans | 0 | 0 | 0 | 0
falsch als Parameter identifiziert | 0 |  | 0 | 0 | BT2.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-13 | Bacillus | altitudinis | Bacillus | altitudinis | 2.1 | Bacillus | pumilus | 2.07 | B | 0 | 0 | 1 | Bacillus altitudinis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21631 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Strain obtained from DSMZ strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-05-09 | Bacillus | altitudinis | Bacillus | altitudinis | 2.58 | Bacillus | altitudinis | 2.39 | A | 1 | 1 | 0 | Bacillus altitudinis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26896 | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-05 | Bacillus | altitudinis | Bacillus | altitudinis | 2.81 | Bacillus | altitudinis | 2.33 | A | 1 | 1 | 0 | Bacillus altitudinis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 2015-02-13 | Bacillus | alveayuensis | Bacillus | alveayuensis | 2.52 | Bacillus | alveayuensis | 2.07 | A | 1 | 1 | 0 | Bacillus alveayuensis | Bacillus alveayuensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7 B159 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2016-03-01 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.23 | Bacillus | velezensis | 2.11 | B | 0 | 0 | 1 | Bacillus amyloliquefaciens | Bacillus velezensis | 0 | 0 | 0 | 0
Parameter wird getroffen mit: |  |  | 0 | 0 | DSM 7 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2016-03-01 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.24 | Bacillus | velezensis | 2.17 | B | 0 | 0 | 1 | Bacillus amyloliquefaciens | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7 CVUAS | 3; type strain | DT | 2014-01-29 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.15 | Bacillus | amyloliquefaciens | 2.15 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
1. Hit Treffer Parameter | 0 |  | 0 | 0 | IGV 015 IGV | 3; sequenced | DT | 2016-11-30 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.33 | Bacillus | siamensis | 1.73 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus siamensis | 0 | 0 | 0 | 0
2. Hit Treffer Parameter | 0 |  | 0 | 0 | IGV 016 IGV | 3; sequenced | DT | 2016-11-30 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 1.93 | Bacillus | amyloliquefaciens | 1.58 | B | 0 | 0 | 1 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
PM unter den ersten 2 Hits | 0 |  | 0 | 0 | ABI02A | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-07-14 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.54 | Bacillus | atrophaeus | 2.16 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
in % | 0% |  | 0 | 0 | ABI05A | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 21.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-11 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.4 | Bacillus | atrophaeus | 2.24 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ALEQ 1 | 2; strain collection of the RKI | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2022-05-24 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.06 | Bacillus | atrophaeus | 2.02 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 9372 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus | atrophaeus | Bacillus | atrophaeus | 1.93 | Bacillus | atrophaeus | 1.9 | B | 0 | 0 | 1 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
Scorewertgrenzen: | 2 | Species | 1 | 1 | B 4 CVUAS | 2; 0 | DT | 2014-01-22 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.14 | Bacillus | atrophaeus | 2.14 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
| 1.7 | Genus | 0 | 0 | B 4 RKI | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.23 | Bacillus | atrophaeus | 1.96 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 0 | 0 | CCM 878 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.13 | Bacillus | atrophaeus | 2.11 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CKL1 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-03-02 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.48 | Bacillus | atrophaeus | 2.23 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
davon mit ID-Ergebnis | 1599 |  | 0 | 0 | CKL1 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 19.04.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-08 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.11 | Bacillus | atrophaeus | 1.96 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
Negative Übereinstimmungen (NA) | 1599 |  | 0 | 0 | DSM 2277 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-03-05 | Bacillus | atrophaeus | Bacillus | atrophaeus | 1.88 | Bacillus | atrophaeus | 1.86 | B | 0 | 0 | 1 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
Positive Abweichungen (PD) | 0 |  | 0 | 0 | DSM 675 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-04-13 | Bacillus | atrophaeus | Bacillus | atrophaeus | 1.51 | Bacillus | atrophaeus | 1.51 | B | 0 | 0 | 1 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-15 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.43 | Bacillus | atrophaeus | 2.36 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
nicht identifiziert in % | 33.12% |  | 0 | 0 | EZ05-02 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2018-08-27 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.43 | Bacillus | atrophaeus | 2.06 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
falsch zugeordnert in % | 0% |  | 0 | 0 | EZ05-02 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 26.09.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2022-09-27 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.46 | Bacillus | atrophaeus | 2.44 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSC56 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 22.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-09 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.46 | Bacillus | atrophaeus | 2.38 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LSSC22 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 15.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-09 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.67 | Bacillus | atrophaeus | 2.28 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LSSC3 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 02.08.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2019-01-09 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.4 | Bacillus | atrophaeus | 2.36 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMTD54 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 15.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-09 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.23 | Bacillus | atrophaeus | 2.16 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | TS1 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 26.09.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-09-27 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.39 | Bacillus | atrophaeus | 2.28 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | TS1 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 31.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-19 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.37 | Bacillus | atrophaeus | 2.28 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | TS1 (03) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 31.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-16 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.43 | Bacillus | atrophaeus | 2.37 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DGL6 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 07.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-19 | Bacillus | axarquiensis | Bacillus | axarquiensis | 2.68 | Bacillus | mojavensis | 2.62 | B | 0 | 0 | 1 | Bacillus axarquiensis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNXM37 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-16 | Bacillus | axarquiensis | Bacillus | axarquiensis | 2.56 | Bacillus | halotolerans | 2.23 | B | 0 | 0 | 1 | Bacillus axarquiensis | Bacillus halotolerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNXM78 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 16.06.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2016-06-16 | Bacillus | axarquiensis | Bacillus | axarquiensis | 2.45 | Bacillus | mojavensis | 2.29 | B | 0 | 0 | 1 | Bacillus axarquiensis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCEB 623 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-10 | Bacillus | badius | Bacillus | badius | 2.06 | Bacillus | badius | 1.99 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Bacillus | badius | Bacillus | badius | 2.45 | Bacillus | badius | 2.26 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111456 | 3; type strain | eDT | 2025-03-18 | Bacillus | albus | Bacillus | cereus | 2.36 | Bacillus | cereus_thuringiensis_PG_III | 2.32 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28875 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-01-25 | Bacillus | albus | Bacillus | albus | 2.82 | Bacillus | tropicus | 2.41 | B | 0 | 0 | 1 | Bacillus albus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-11 | Bacillus | anthracis | Bacillus | luti | 2.16 | Bacillus | pacificus | 2.14 | B | 0 | 0 | 1 | Bacillus luti | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A10 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.18 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A100 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus | anthracis | Bacillus | toyonensis | 2.01 | Bacillus | cereus | 2 | B | 0 | 0 | 1 | Bacillus toyonensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A101 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus | anthracis | Bacillus | tropicus | 2.1 | Bacillus | cereus | 2.01 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A102 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus | anthracis | Bacillus | cereus | 2.14 | Bacillus | tropicus | 2.02 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A103 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus | anthracis | Bacillus | tropicus | 2.11 | Bacillus | tropicus | 2.11 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A106 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-03 | Bacillus | anthracis | Bacillus | pacificus | 2.17 | Bacillus | luti | 2.12 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A107 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-03 | Bacillus | anthracis | Bacillus | cereus | 2.26 | Bacillus | cereus_thuringiensis_PG_IV | 2.2 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A108 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus | anthracis | Bacillus | pacificus | 2.06 | Bacillus | cereus | 2.01 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A11 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-29 | Bacillus | anthracis | Bacillus | tropicus | 1.81 | Bacillus | albus | 1.8 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A110 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-06 | Bacillus | anthracis | Bacillus | pacificus | 1.94 | Bacillus | cereus | 1.92 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A111 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus | anthracis | Bacillus | pacificus | 2.02 | Bacillus | cereus | 2.01 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A112 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-04 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | cereus | 1.98 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A112-1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-02-08 | Bacillus | anthracis | Bacillus | albus | 1.91 | Bacillus | tropicus | 1.9 | B | 0 | 0 | 1 | Bacillus albus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A113 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-09 | Bacillus | anthracis | Bacillus | luti | 2.19 | Bacillus | proteolyticus | 2.1 | B | 0 | 0 | 1 | Bacillus luti | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A118 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus | anthracis | Bacillus | luti | 2.2 | Bacillus | cereus | 2.16 | B | 0 | 0 | 1 | Bacillus luti | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A12 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus | anthracis | Bacillus | cereus | 2.18 | Bacillus | pacificus | 2.06 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A120 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.17 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A122 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus | anthracis | Bacillus | pacificus | 2.17 | Bacillus | cereus | 2.15 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A123 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus | anthracis | Bacillus | cereus | 2.07 | Bacillus | cereus_thuringiensis_PG_IV | 1.94 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A124 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | luti | 1.98 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A125 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus | anthracis | Bacillus | cereus | 2.06 | Bacillus | cereus_thuringiensis_PG_III | 1.99 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A126 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus | anthracis | Bacillus | cereus | 2.26 | Bacillus | cereus_thuringiensis_PG_IV | 2.11 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A127 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-10 | Bacillus | anthracis | Bacillus | cereus | 1.92 | Bacillus | toyonensis | 1.88 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A128 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-04 | Bacillus | anthracis | Bacillus | cereus | 2.16 | Bacillus | luti | 2.05 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A130 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-11 | Bacillus | anthracis | Bacillus | tropicus | 2.15 | Bacillus | tropicus | 2.07 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A131 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | paranthracis | 2.21 | Bacillus | tropicus | 2.11 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A132 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-28 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | cereus | 2.07 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A133 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-05 | Bacillus | anthracis | Bacillus | pacificus | 2.28 | Bacillus | luti | 2.21 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A134 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-28 | Bacillus | anthracis | Bacillus | paranthracis | 2.07 | Bacillus | pacificus | 2.02 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A135 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | tropicus | 2.25 | Bacillus | paranthracis | 2.16 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A136 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-28 | Bacillus | anthracis | Bacillus | albus | 2.13 | Bacillus | tropicus | 2.08 | B | 0 | 0 | 1 | Bacillus albus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A137 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-13 | Bacillus | anthracis | Bacillus | paranthracis | 2.21 | Bacillus | tropicus | 2.18 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A138-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-13 | Bacillus | anthracis | Bacillus | paranthracis | 2.11 | Bacillus | tropicus | 2.07 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A15 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-01 | Bacillus | anthracis | Bacillus | cereus | 2.08 | Bacillus | luti | 2.04 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A16 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-01 | Bacillus | anthracis | Bacillus | wiedmannii | 2.12 | Bacillus | luti | 2.1 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A173-2 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-05-30 | Bacillus | anthracis | Bacillus | tropicus | 2.32 | Bacillus | paranthracis | 2.24 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A18 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-05 | Bacillus | anthracis | Bacillus | luti | 2.24 | Bacillus | proteolyticus | 2.16 | B | 0 | 0 | 1 | Bacillus luti | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus | anthracis | Bacillus | pseudomycoides_PG_I | 2 | Bacillus | cereus | 2 | B | 0 | 0 | 1 | Bacillus pseudomycoides_PG_I | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A22 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus | anthracis | Bacillus | paranthracis | 1.93 | Bacillus | pacificus | 1.93 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A23 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus | anthracis | Bacillus | paranthracis | 2.01 | Bacillus | pacificus | 2.01 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A24 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus | anthracis | Bacillus | tropicus | 2.07 | Bacillus | paranthracis | 1.99 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A25 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-05 | Bacillus | anthracis | Bacillus | cereus | 1.68 | Bacillus | paranthracis | 1.62 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A269-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-02 | Bacillus | anthracis | Bacillus | tropicus | 2.23 | Bacillus | cereus | 2.19 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A27 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus | anthracis | Bacillus | cereus | 2.23 | Bacillus | cereus | 2.21 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A28 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus | anthracis | Bacillus | wiedmannii | 2.18 | Bacillus | proteolyticus | 2.14 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus | anthracis | Bacillus | tropicus | 2.19 | Bacillus | luti | 2.09 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A294 1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus | anthracis | Bacillus | tropicus | 1.99 | Bacillus | tropicus | 1.94 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A3 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-19 | Bacillus | anthracis | Bacillus | cereus | 2.06 | Bacillus | nitratireducens | 2.06 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A30 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-05 | Bacillus | anthracis | Bacillus | cereus | 1.9 | Bacillus | cereus_thuringiensis_PG_III | 1.81 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A300-1a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-07 | Bacillus | anthracis | Bacillus | paranthracis | 1.92 | Bacillus | tropicus | 1.84 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A300-1b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-14 | Bacillus | anthracis | Bacillus | nitratireducens | 1.78 | Bacillus | cereus | 1.77 | B | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A300-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-10-23 | Bacillus | anthracis | Bacillus | paranthracis | 1.82 | Bacillus | pseudomycoides | 1.69 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A300-2b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-13 | Bacillus | anthracis | Bacillus | cereus | 2.02 | Bacillus | cereus | 1.79 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A302-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-11 | Bacillus | anthracis | Bacillus | paranthracis | 2.01 | Bacillus | tropicus | 1.97 | A | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A306-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-12 | Bacillus | anthracis | Bacillus | tropicus | 1.87 | Bacillus | tropicus | 1.86 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A315-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-12 | Bacillus | anthracis | Bacillus | paranthracis | 1.93 | Bacillus | albus | 1.83 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A315-2 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-12 | Bacillus | anthracis | Bacillus | tropicus | 1.77 | Bacillus | tropicus | 1.73 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A32 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus | anthracis | Bacillus | luti | 2.08 | Bacillus | wiedmannii | 2.05 | B | 0 | 0 | 1 | Bacillus luti | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A33 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus | anthracis | Bacillus | pacificus | 2.03 | Bacillus | tropicus | 1.96 | A | 0 | 0 | 1 | Bacillus pacificus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A34 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus | anthracis | Bacillus | tropicus | 2.16 | Bacillus | paranthracis | 2.12 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A35 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-06 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.09 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A36 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-06 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | tropicus | 2.05 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A37 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus | anthracis | Bacillus | tropicus | 1.97 | Bacillus | cereus | 1.95 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A38 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus | anthracis | Bacillus | proteolyticus | 2.12 | Bacillus | luti | 2.09 | B | 0 | 0 | 1 | Bacillus proteolyticus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A39 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus | anthracis | Bacillus | tropicus | 2.18 | Bacillus | pacificus | 2.17 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A4 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus | anthracis | Bacillus | paranthracis | 2.1 | Bacillus | tropicus | 2.04 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A40 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-06 | Bacillus | anthracis | Bacillus | nitratireducens | 1.95 | Bacillus | cereus | 1.92 | B | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A41 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus | anthracis | Bacillus | tropicus | 2.17 | Bacillus | pacificus | 2.14 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A42 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus | anthracis | Bacillus | tropicus | 2.11 | Bacillus | nitratireducens | 1.96 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A43 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus | anthracis | Bacillus | tropicus | 2.16 | Bacillus | pacificus | 2.03 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A44 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus | anthracis | Bacillus | tropicus | 1.96 | Bacillus | cereus | 1.95 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A45 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus | anthracis | Bacillus | tropicus | 2.13 | Bacillus | pacificus | 2.11 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A46 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus | anthracis | Bacillus | tropicus | 2.1 | Bacillus | pacificus | 1.94 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A49 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus | anthracis | Bacillus | luti | 2 | Bacillus | wiedmannii | 1.97 | A | 0 | 0 | 1 | Bacillus luti | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A5 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus | anthracis | Bacillus | luti | 2.09 | Bacillus | proteolyticus | 2.09 | B | 0 | 0 | 1 | Bacillus luti | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A58 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | proteolyticus | 2.08 | Bacillus | nitratireducens | 2.07 | B | 0 | 0 | 1 | Bacillus proteolyticus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A59 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | tropicus | 2.14 | Bacillus | pacificus | 2.07 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A6 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-19 | Bacillus | anthracis | Bacillus | tropicus | 2.16 | Bacillus | pacificus | 2.15 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A60 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | tropicus | 2.08 | Bacillus | pacificus | 2.03 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A61 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | pacificus | 2.2 | Bacillus | tropicus | 2.13 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A62 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | cereus | 2.13 | Bacillus | nitratireducens | 2.08 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A63 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-10 | Bacillus | anthracis | Bacillus | luti | 2.08 | Bacillus | cereus | 2.05 | B | 0 | 0 | 1 | Bacillus luti | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A64 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | cereus | 2.09 | Bacillus | cereus | 1.96 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A65 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | albus | 1.9 | Bacillus | paranthracis | 1.9 | B | 0 | 0 | 1 | Bacillus albus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A66 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | nitratireducens | 2.15 | Bacillus | luti | 2.1 | B | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A67 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus | anthracis | Bacillus | pacificus | 2.11 | Bacillus | proteolyticus | 2.1 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A68 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-12 | Bacillus | anthracis | Bacillus | cereus | 2.21 | Bacillus | cereus_thuringiensis_PG_III | 2.11 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A69 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus | anthracis | Bacillus | pacificus | 2.23 | Bacillus | tropicus | 2.22 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A7 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus | anthracis | Bacillus | cereus | 2.17 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A72 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-12 | Bacillus | anthracis | Bacillus | cereus | 2.08 | Bacillus | cereus_thuringiensis_PG_V | 1.95 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A73 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-12 | Bacillus | anthracis | Bacillus | cereus | 2.3 | Bacillus | cereus_thuringiensis_PG_III | 2.17 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A74 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus | anthracis | Bacillus | cereus | 2.06 | Bacillus | cereus | 2.05 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A8 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-19 | Bacillus | anthracis | Bacillus | pacificus | 2.27 | Bacillus | luti | 2.17 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A80 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus | anthracis | Bacillus | paranthracis | 2.08 | Bacillus | tropicus | 1.97 | A | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A81 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus | anthracis | Bacillus | tropicus | 2.15 | Bacillus | pacificus | 2.06 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A82 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus | anthracis | Bacillus | pacificus | 2.23 | Bacillus | luti | 2.18 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A83 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus | anthracis | Bacillus | cereus | 2.11 | Bacillus | cereus | 2.09 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A84 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus | anthracis | Bacillus | cereus | 2.05 | Bacillus | luti | 2.02 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A85 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus | anthracis | Bacillus | tropicus | 1.82 | Bacillus | tropicus | 1.77 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A86 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-03 | Bacillus | anthracis | Bacillus | pacificus | 2.19 | Bacillus | proteolyticus | 2.17 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A87 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus | anthracis | Bacillus | paranthracis | 2.03 | Bacillus | albus | 1.99 | A | 0 | 0 | 1 | Bacillus paranthracis | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A88 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus | anthracis | Bacillus | cereus | 2.22 | Bacillus | cereus_thuringiensis_PG_IV | 2.1 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A88 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus | anthracis | Bacillus | cereus | 2.18 | Bacillus | cereus | 2.11 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A89 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus | anthracis | Bacillus | cereus | 2.21 | Bacillus | paranthracis | 2.18 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A89 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus | anthracis | Bacillus | luti | 2.13 | Bacillus | pacificus | 2.1 | B | 0 | 0 | 1 | Bacillus luti | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A9 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus | anthracis | Bacillus | cereus | 2.2 | Bacillus | cereus_thuringiensis_PG_III | 2.09 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A90 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus | anthracis | Bacillus | cereus | 1.95 | Bacillus | tropicus | 1.87 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A90 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus | anthracis | Bacillus | proteolyticus | 2.02 | Bacillus | cereus | 1.98 | A | 0 | 0 | 1 | Bacillus proteolyticus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A91 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus | anthracis | Bacillus | paranthracis | 1.99 | Bacillus | tropicus | 1.91 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A92 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus | anthracis | Bacillus | tropicus | 2.06 | Bacillus | pacificus | 1.98 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A93 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus | anthracis | Bacillus | cereus | 2.04 | Bacillus | cereus_thuringiensis_PG_IV | 2.03 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AF 28 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | tropicus | 2.22 | Bacillus | paranthracis | 2.19 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AF 29 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | albus | 2.21 | Bacillus | paranthracis | 2.2 | B | 0 | 0 | 1 | Bacillus albus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AF 31 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | tropicus | 2.24 | Bacillus | paranthracis | 2.22 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AF 71 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | tropicus | 2.33 | Bacillus | cereus | 2.32 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AF 96 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | tropicus | 2.25 | Bacillus | paranthracis | 2.17 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-09 | Bacillus | anthracis | Bacillus | luti | 2.1 | Bacillus | cereus | 2.05 | B | 0 | 0 | 1 | Bacillus luti | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14185 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus | anthracis | Bacillus | cereus | 2.02 | Bacillus | nitratireducens | 1.99 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14578 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus | anthracis | Bacillus | tropicus | 2.11 | Bacillus | pacificus | 2.07 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-26 | Bacillus | anthracis | Bacillus | cereus | 2.36 | Bacillus | cereus | 2.25 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SA 20 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus | anthracis | Bacillus | paranthracis | 2.23 | Bacillus | tropicus | 2.19 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Sterne | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus | anthracis | Bacillus | tropicus | 2.2 | Bacillus | pacificus | 2.19 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | vollum | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-03-06 | Bacillus | anthracis | Bacillus | albus | 1.95 | Bacillus | cereus | 1.89 | B | 0 | 0 | 1 | Bacillus albus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Wirt | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus | anthracis | Bacillus | tropicus | 2.27 | Bacillus | paranthracis | 2.2 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111020 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | bingmayongensis | Bacillus | pseudomycoides | 2.1 | Bacillus | bingmayongensis | 2.03 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus bingmayongensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25427 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 05.10.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus bingmayongensis DSM 25427, preparation and measurements: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-11-05 | Bacillus | bingmayongensis | Bacillus | bingmayongensis | 2.8 | Bacillus | wiedmannii | 1.95 | A | 1 | 1 | 0 | Bacillus bingmayongensis | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | AaCol | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-03-06 | Bacillus | cereus | Bacillus | paranthracis | 2 | Bacillus | pacificus | 1.96 | A | 0 | 0 | 1 | Bacillus paranthracis | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SK1b | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-03-06 | Bacillus | cereus | Bacillus | pacificus | 2.21 | Bacillus | paranthracis | 2.2 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B194 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus | cereus | Bacillus | cereus | 2.62 | Bacillus | tropicus | 2.39 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B195 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus | cereus | Bacillus | cereus | 2.6 | Bacillus | tropicus | 2.37 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B196 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | tropicus | 2.38 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34551 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus | cereus | Bacillus | cereus | 2.49 | Bacillus | cereus_thuringiensis_PG_IV | 2.47 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34552 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus | cereus | Bacillus | thuringiensis | 2.01 | Bacillus | cereus | 2 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B21 | 2; strain collection of the RKI | see [0185] | 2015-12-02 | Bacillus | cereus | Bacillus | pseudomycoides | 1.86 | Bacillus | pseudomycoides_PG_I | 1.85 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pseudomycoides_PG_I | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B22 | 2; strain collection of the RKI | see [0185] | 2015-12-02 | Bacillus | cereus | Bacillus | wiedmannii | 2.03 | Bacillus | wiedmannii | 1.9 | A | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 340 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus | cereus | Bacillus | mycoides | 2.3 | Bacillus | toyonensis | 2.27 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 341 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus | cereus | Bacillus | cereus | 2.37 | Bacillus | paranthracis | 2.24 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 342 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus | cereus | Bacillus | thuringiensis | 2.33 | Bacillus | thuringiensis | 2.25 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 343 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus | cereus | Bacillus | wiedmannii | 2.45 | Bacillus | wiedmannii | 2.06 | A | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-071-2-A-2 Bac 25 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus sp. A-071; 2-A-2 Bac 25: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-26 | Bacillus | cereus | Bacillus | wiedmannii | 2.45 | Bacillus | toyonensis | 2.41 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-082-2 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 29.02.2022 | Bacillus | cereus | Bacillus | tropicus | 2.43 | Bacillus | cereus | 2.38 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-082-4 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.04.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 29.02.2022 | Bacillus | cereus | Bacillus | tropicus | 2.45 | Bacillus | tropicus | 2.45 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-175-2 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2021-03-31 | Bacillus | cereus | Bacillus | cereus | 2.46 | Bacillus | tropicus | 2.31 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-550-5 (SC6) Bac 50 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus sp. A-550; 5 (SC6) Bac 50: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-12 | Bacillus | cereus | Bacillus | pacificus | 2.3 | Bacillus | tropicus | 2.21 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BT250c Bac 33 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus sp. A-550; 5 (SC6) Bac 50: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-12 | Bacillus | cereus | Bacillus | cereus | 2.34 | Bacillus | proteolyticus | 2.29 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CD 3-1a | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-05-11 | Bacillus | cereus | Bacillus | toyonensis | 2.29 | Bacillus | pacificus | 2.25 | B | 0 | 0 | 1 | Bacillus toyonensis | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CD 3-1c | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-05-11 | Bacillus | cereus | Bacillus | cereus | 2.28 | Bacillus | pacificus | 2.26 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CD3.5 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-11-16 | Bacillus | cereus | Bacillus | cereus | 2.41 | Bacillus | albus | 2.37 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD1.3 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-08-01 | Bacillus | cereus | Bacillus | tropicus | 2.32 | Bacillus | tropicus | 2.31 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KN17-s3053 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=Source: AKRA Austria, Dr. Karner#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=Dr. Karger AKRA (Austria)# | 2022-05-11 | Bacillus | cereus | Bacillus | toyonensis | 2.38 | Bacillus | proteolyticus | 2.36 | B | 0 | 0 | 1 | Bacillus toyonensis | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4222 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | cereus | Bacillus | cereus_thuringiensis_PG_III | 2.26 | Bacillus | pseudomycoides | 2.23 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004439-1 | 2; 0 | DT | 2016-04-12 | Bacillus | cereus | Bacillus | thuringiensis | 2.51 | Bacillus | thuringiensis | 2.51 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004439-2 | 2; 0 | DT | 2016-04-12 | Bacillus | cereus | Bacillus | cereus | 2.34 | Bacillus | cereus_thuringiensis_PG_IV | 2.23 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161005557-1 | 2; 0 | DT | 2016-08-11 | Bacillus | cereus | Bacillus | cereus_thuringiensis_PG_III | 2.29 | Bacillus | cereus_thuringiensis_PG_III | 2.25 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012589-5 | 2; 0 | DT | 2017-09-19 | Bacillus | cereus | Bacillus | cereus_thuringiensis_PG_III | 2.27 | Bacillus | cereus_thuringiensis_PG_III | 2.27 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181015909-7 | 2; 0 | DT | 2018-11-02 | Bacillus | cereus | Bacillus | cereus | 2.07 | Bacillus | cereus | 2.06 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2617 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-03-31 | Bacillus | cereus | Bacillus | tropicus | 2.51 | Bacillus | albus | 2.38 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 301101RA0432 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-12 | Bacillus | cereus | Bacillus | cereus | 2.41 | Bacillus | tropicus | 2.29 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-049 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2022-03-31 | Bacillus | cereus | Bacillus | tropicus | 2.5 | Bacillus | paranthracis | 2.35 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-1607 12B | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus cereus A-1607; 12B: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Bacillus | cereus | Bacillus | pacificus | 2.38 | Bacillus | cereus | 2.32 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A22 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam I)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus | cereus | Bacillus | albus | 2.32 | Bacillus | cereus | 2.26 | B | 0 | 0 | 1 | Bacillus albus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A24 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus | cereus | Bacillus | cereus | 2.38 | Bacillus | thuringiensis | 2.3 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A31 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus | cereus | Bacillus | cereus | 2.45 | Bacillus | pacificus | 2.37 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A42 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam I)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus | cereus | Bacillus | cereus | 2.32 | Bacillus | albus | 2.29 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A8 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus | cereus | Bacillus | cereus | 2.44 | Bacillus | albus | 2.41 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10876 | 3; public strain collection | TFA Extr | 2015-04-29 | Bacillus | cereus | Bacillus | cereus | 2.25 | Bacillus | albus | 2.23 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11778 IGV | 3; public strain collection | DT | 2016-02-03 | Bacillus | cereus | Bacillus | cereus | 2.25 | Bacillus | cereus | 2.21 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11778 LHL | 3; public strain collection | DT | 2015-06-23 | Bacillus | cereus | Bacillus | cereus_thuringiensis_PG_IV | 2.3 | Bacillus | cereus_thuringiensis_PG_IV | 2.29 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11778 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 2007-03-05 | Bacillus | cereus | Bacillus | tropicus | 2.33 | Bacillus | pacificus | 2.32 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14579 IGV | 3; type strain | DT | 2016-04-06 | Bacillus | cereus | Bacillus | cereus | 2.79 | Bacillus | thuringiensis | 2.43 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14579 RKI | 3; type strain | in 80% TFA/ 1:1 HCCATA2 (A) | 2007-03-05 | Bacillus | cereus | Bacillus | cereus | 2.13 | Bacillus | tropicus | 2.03 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | cereus_thuringiensis_PG_V | 2.38 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33018 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-04 | Bacillus | cereus | Bacillus | wiedmannii | 2.25 | Bacillus | toyonensis | 2.17 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33019 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-22 | Bacillus | cereus | Bacillus | cereus | 2.45 | Bacillus | tropicus | 2.25 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 6464 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus | cereus | Bacillus | thuringiensis | 2.16 | Bacillus | pacificus | 2.15 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 7004 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus | cereus | Bacillus | cereus | 2.5 | Bacillus | tropicus | 2.37 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 9139 | 3; reference strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus | cereus | Bacillus | cereus | 2.42 | Bacillus | thuringiensis | 2.19 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B10 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-01 | Bacillus | cereus | Bacillus | wiedmannii | 2.14 | Bacillus | thuringiensis | 2.08 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B108 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-15 | Bacillus | cereus | Bacillus | cereus | 2.42 | Bacillus | cereus | 2.38 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B109 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-13 | Bacillus | cereus | Bacillus | tropicus | 2.54 | Bacillus | tropicus | 2.38 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B110 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-15 | Bacillus | cereus | Bacillus | tropicus | 2.37 | Bacillus | tropicus | 2.3 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B146 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-27 | Bacillus | cereus | Bacillus | pacificus | 2.33 | Bacillus | paranthracis | 2.27 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B16 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-02 | Bacillus | cereus | Bacillus | thuringiensis | 2.12 | Bacillus | wiedmannii | 2.08 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B197 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus | cereus | Bacillus | cereus | 2.49 | Bacillus | tropicus | 2.45 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B198 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-14 | Bacillus | cereus | Bacillus | pseudomycoides | 2.38 | Bacillus | pacificus | 2.35 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B199 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus | cereus | Bacillus | paranthracis | 2.37 | Bacillus | cereus | 2.32 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B200 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus | cereus | Bacillus | cereus | 2.54 | Bacillus | tropicus | 2.38 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B232 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-18 | Bacillus | cereus | Bacillus | cereus | 2.35 | Bacillus | thuringiensis | 2.25 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 239 | 3; public strain collection | see [0185] | 2015-04-22 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | tropicus | 2.28 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B252 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus | cereus | Bacillus | wiedmannii | 2.31 | Bacillus | thuringiensis | 2.2 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B253 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus | cereus | Bacillus | wiedmannii | 2.28 | Bacillus | wiedmannii | 2.22 | A | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B254 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus | cereus | Bacillus | toyonensis | 2.25 | Bacillus | nitratireducens | 1.99 | A | 0 | 0 | 1 | Bacillus toyonensis | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B255 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus | cereus | Bacillus | paranthracis | 2.46 | Bacillus | cereus | 2.45 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B257 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | albus | 2.25 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B258 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus | cereus | Bacillus | cereus | 2.39 | Bacillus | thuringiensis | 2.21 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B259 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus | cereus | Bacillus | cereus | 2.43 | Bacillus | thuringiensis | 2.27 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B261 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus | cereus | Bacillus | paranthracis | 2.42 | Bacillus | cereus | 2.32 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B262 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-14 | Bacillus | cereus | Bacillus | cereus | 2.25 | Bacillus | tropicus | 2.22 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B272 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus | cereus | Bacillus | cereus | 2.27 | Bacillus | paranthracis | 2.11 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B273 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus | cereus | Bacillus | paranthracis | 2.28 | Bacillus | cereus | 2.26 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B274 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus | cereus | Bacillus | wiedmannii | 2.52 | Bacillus | wiedmannii | 2.42 | A | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B275 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus | cereus | Bacillus | cereus | 2.19 | Bacillus | paranthracis | 2.13 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B277 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | tropicus | 2.41 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B278 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | tropicus | 2.19 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B282 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus | cereus | Bacillus | paranthracis | 2.41 | Bacillus | cereus | 2.38 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B293 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-19 | Bacillus | cereus | Bacillus | cereus | 2.44 | Bacillus | thuringiensis | 2.27 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B294 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus | cereus | Bacillus | tropicus | 2.28 | Bacillus | pacificus | 2.26 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B295 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus | cereus | Bacillus | cereus | 2.56 | Bacillus | tropicus | 2.47 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B296 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus | cereus | Bacillus | cereus | 2.62 | Bacillus | tropicus | 2.4 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B297 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | paranthracis | 2.37 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B298 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | tropicus | 2.42 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B299 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-19 | Bacillus | cereus | Bacillus | cereus | 2.49 | Bacillus | tropicus | 2.35 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B3 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-11 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | paranthracis | 2.4 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B300 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.6 | Bacillus | tropicus | 2.48 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B302 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.54 | Bacillus | tropicus | 2.39 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B303 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.59 | Bacillus | paranthracis | 2.38 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B304 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | tropicus | 2.42 | Bacillus | cereus | 2.38 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B305 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.53 | Bacillus | tropicus | 2.41 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B306 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.56 | Bacillus | tropicus | 2.45 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B308 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | tropicus | 2.41 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B309 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-22 | Bacillus | cereus | Bacillus | cereus | 2.3 | Bacillus | albus | 2.24 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B310 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus | cereus | Bacillus | tropicus | 2.39 | Bacillus | cereus | 2.37 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B311 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus | cereus | Bacillus | cereus | 2.37 | Bacillus | tropicus | 2.33 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B316 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus | cereus | Bacillus | cereus | 2.38 | Bacillus | albus | 2.26 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B329 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 1:10 diluted , spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Cultivation & sample preparation by W. Beyer (University Hohenheim); Arrival: 09.08.2006; Preparation for MALDI-ToF MS#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Dr. Beyer (Uni Hohenheim)# | 2015-05-07 | Bacillus | cereus | Bacillus | cereus | 2.53 | Bacillus | tropicus | 2.4 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B33 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-10 | Bacillus | cereus | Bacillus | cereus | 2.54 | Bacillus | tropicus | 2.41 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B332 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus | cereus | Bacillus | thuringiensis | 2.34 | Bacillus | wiedmannii | 2.3 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 88 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-14 | Bacillus | cereus | Bacillus | thuringiensis | 2.32 | Bacillus | cereus | 2.25 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BW-A | 2; strain collection of the RKI | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2021-06-21 | Bacillus | cereus | Bacillus | cereus | 2.2 | Bacillus | proteolyticus | 2.17 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BW-B | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus | cereus | Bacillus | tropicus | 2.31 | Bacillus | tropicus | 2.16 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 99 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus | cereus | Bacillus | cereus | 2.2 | Bacillus | cereus | 2.04 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1107 | 2; 0 | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.45 | Bacillus | cereus | 2.32 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3901 | 3; 0 | eDT | 2016-04-26 | Bacillus | cereus | Bacillus | pseudomycoides | 2.46 | Bacillus | cereus | 2.36 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1274 | 3; public strain collection ? | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | pacificus | 2.21 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2299 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-16 | Bacillus | cereus | Bacillus | cereus | 2.31 | Bacillus | tropicus | 2.2 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 2301 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | cereus | Bacillus | cereus_thuringiensis_PG_III | 2.39 | Bacillus | cereus_thuringiensis_PG_III | 2.35 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2301 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus cereus DSM 2301: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Bacillus | cereus | Bacillus | cereus | 2.24 | Bacillus | cereus | 2.22 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2302 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus | cereus | Bacillus | cereus | 2.39 | Bacillus | cereus | 2.35 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 31 B181 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-30 | Bacillus | cereus | Bacillus | cereus | 2.38 | Bacillus | cereus | 2.27 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 31 B69 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-02-06 | Bacillus | cereus | Bacillus | cereus | 2.31 | Bacillus | cereus | 2.28 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 31 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | cereus | 2.55 | Bacillus | cereus_thuringiensis_PG_III | 2.34 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 31 RKI | 3; type strain | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-09-21 | Bacillus | cereus | Bacillus | cereus | 2.21 | Bacillus | cereus | 2.16 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 3101 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-20 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | albus | 2.32 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 318 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-20 | Bacillus | cereus | Bacillus | tropicus | 2.5 | Bacillus | tropicus | 2.43 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 336 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-18 | Bacillus | cereus | Bacillus | tropicus | 2.38 | Bacillus | tropicus | 2.37 | A | 0 | 0 | 1 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 345 B170 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-30 | Bacillus | cereus | Bacillus | cereus | 2.36 | Bacillus | nitratireducens | 2.25 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 345 B23 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-11 | Bacillus | cereus | Bacillus | cereus | 2.5 | Bacillus | paranthracis | 2.43 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 345 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | cereus | 2.43 | Bacillus | cereus_thuringiensis_PG_IV | 2.41 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 345 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 02.12.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus cereus DSM 345, preparation and measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-12-02 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | thuringiensis | 2.38 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-18 | Bacillus | cereus | Bacillus | thuringiensis | 2.44 | Bacillus | thuringiensis | 2.29 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 3648 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | cereus | Bacillus | cereus_thuringiensis_PG_III | 2.26 | Bacillus | cereus_thuringiensis_PG_IV | 2.25 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4218 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-20 | Bacillus | cereus | Bacillus | wiedmannii | 2.13 | Bacillus | wiedmannii | 2.08 | A | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4282 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | cereus | Bacillus | cereus | 2.5 | Bacillus | cereus_thuringiensis_PG_IV | 2.44 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4312 B269 RKI | 3; public strain collection | TFA Extr | 2006-09-15 | Bacillus | cereus | Bacillus | cereus | 2.52 | Bacillus | tropicus | 2.4 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4312 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.47 | Bacillus | cereus_thuringiensis_PG_III | 2.37 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4312 IGV | 3; public strain collection | DT | 2016-04-13 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | tropicus | 2.36 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4312 RKI | 3; public strain collection | TFA Extr | 2016-06-30 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | paranthracis | 2.45 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4313 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.62 | Bacillus | cereus_thuringiensis_PG_III | 2.26 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4313 IGV | 3; public strain collection | DT | 2016-04-06 | Bacillus | cereus | Bacillus | cereus | 2.32 | Bacillus | tropicus | 2.29 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4384 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | cereus | Bacillus | cereus | 2.49 | Bacillus | cereus_thuringiensis_PG_III | 2.44 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4490 | 3; public strain collection | in 80% TFA, Zentrifugation und Sterilfiltration, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-09-21 | Bacillus | cereus | Bacillus | cereus | 2.34 | Bacillus | nitratireducens | 2.3 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 487 RKI | 3; public strain collection | see [0185] | 2021-12-03 | Bacillus | cereus | Bacillus | paranthracis | 2.47 | Bacillus | cereus | 2.38 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 487 B169 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | paranthracis | 2.43 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 487 B185 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus | cereus | Bacillus | cereus | 2.43 | Bacillus | tropicus | 2.33 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 508 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-04-07 | Bacillus | cereus | Bacillus | cereus | 2.42 | Bacillus | nitratireducens | 2.3 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 609 | 3; strain from public collection | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-09-21 | Bacillus | cereus | Bacillus | thuringiensis | 2.33 | Bacillus | wiedmannii | 2.29 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B241 | 3; public strain collection | TFA Extr | 2006-09-14 | Bacillus | cereus | Bacillus | cereus | 2.55 | Bacillus | tropicus | 2.43 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6127 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | cereus | Bacillus | cereus_thuringiensis_PG_III | 2.42 | Bacillus | pseudomycoides | 2.41 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6127 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 02.12.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus cereus DSM 6127, preparation and measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-12-02 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | tropicus | 2.44 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 626 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus | cereus | Bacillus | cereus | 2.33 | Bacillus | albus | 2.22 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 679 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-14 | Bacillus | cereus | Bacillus | thuringiensis | 2.23 | Bacillus | cereus | 2.19 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 689 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-14 | Bacillus | cereus | Bacillus | tropicus | 2.37 | Bacillus | cereus | 2.37 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8438 | 3; public strain collection | TFAextr | 2015-04-23 | Bacillus | cereus | Bacillus | pacificus | 2.19 | Bacillus | luti | 2.13 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 9378 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | cereus | Bacillus | cereus | 2.24 | Bacillus | cereus_thuringiensis_PG_III | 2.15 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HB3.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus | cereus | Bacillus | cereus | 2.33 | Bacillus | pacificus | 2.19 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD1.4B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-16 | Bacillus | cereus | Bacillus | cereus | 2.19 | Bacillus | thuringiensis | 2.09 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Hohenheim | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-16 | Bacillus | cereus | Bacillus | paranthracis | 2.13 | Bacillus | cereus | 2.06 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MS2.1B | 2; strain collection of the RKI | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2020-08-12 | Bacillus | cereus | Bacillus | cereus | 2.02 | Bacillus | paranthracis | 1.96 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | NCIMB 11925 | 3; reference strain collection | DT | 2015-07-22 | Bacillus | cereus | Bacillus | cereus | 2.22 | Bacillus | cereus_thuringiensis_PG_III | 2.18 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 10320 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus | cereus | Bacillus | cereus | 2.55 | Bacillus | paranthracis | 2.37 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 10989 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus | cereus | Bacillus | cereus | 2.29 | Bacillus | cereus | 2.23 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 6474 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus | cereus | Bacillus | thuringiensis | 2.44 | Bacillus | thuringiensis | 2.33 | A | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 9939 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus | cereus | Bacillus | cereus | 2.47 | Bacillus | paranthracis | 2.4 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 9945. ATCC 10876a (B264. Beyer) | 3; public strain collection | TFA Extr | 2015-04-28 | Bacillus | cereus | Bacillus | albus | 2.25 | Bacillus | cereus | 2.25 | B | 0 | 0 | 1 | Bacillus albus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | NiMe_1 | 3; 16S, ces-pos | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.59 | Bacillus | cereus_thuringiensis_PG_III | 2.32 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 1 | 0 | SL2009001164 | 3; 0 | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.39 | Bacillus | cereus | 2.26 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | SL2009005808 | 2; 0 | eDT | 2019-09-23 | Bacillus | cereus | Bacillus | cereus | 2.59 | Bacillus | cereus_thuringiensis_PG_IV | 2.35 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 1 | 0 | SL2009007190 | 2; 0 | eDT | 2019-09-25 | Bacillus | cereus | Bacillus | cereus | 2.53 | Bacillus | cereus_thuringiensis_PG_III | 2.43 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 1 | 0 | SL2014002835 | 2; CytK2 pos, Krist neg | eDT | 2019-09-25 | Bacillus | cereus | Bacillus | cereus | 2.46 | Bacillus | cereus_thuringiensis_PG_III | 2.37 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SN4-3 | 2; strain collection of the RKI | see [0185] | 2020-10-11 | Bacillus | cereus | Bacillus | albus | 2.4 | Bacillus | cereus | 2.4 | B | 0 | 0 | 1 | Bacillus albus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Wasserprobe Bac RKI | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus | cereus | Bacillus | cereus | 2.4 | Bacillus | pacificus | 2.3 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WIS-St. Nr. 29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus | cereus | Bacillus | wiedmannii | 1.85 | Bacillus | wiedmannii | 1.85 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WIS-St. Nr. 311 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus | cereus | Bacillus | cereus | 2.13 | Bacillus | tropicus | 2.05 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10028 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | cereus | Bacillus | cereus | 2.46 | Bacillus | pacificus | 2.29 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus pacificus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10030 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | cereus | Bacillus | cereus | 2.32 | Bacillus | paranthracis | 2.31 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10286 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus | cereus | Bacillus | cereus | 2.48 | Bacillus | paranthracis | 2.38 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10288 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | cereus | Bacillus | cereus | 2.51 | Bacillus | tropicus | 2.23 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10441 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | cereus | Bacillus | tropicus | 2.52 | Bacillus | albus | 2.18 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | WSBC 10925 | 3; well defined isolat, ces-pos | eDT | 2016-01-27 | Bacillus | cereus | Bacillus | pseudomycoides | 2.53 | Bacillus | cereus | 2.31 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 46776 CVUAS | 3; specific PCR, growth at 52°C | eDT | 2017-05-31 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.23 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 46776 MVZ | 3; specific PCR, growth at 52°C | DT | 2017-05-23 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.43 | Bacillus | cytotoxicus_PG_VII | 2.43 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 46778 CVUAS | 3; specific PCR, growth at 52°C | eDT | 2017-05-31 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.35 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 46778 MVZ | 3; specific PCR, growth at 52°C | DT | 2017-05-23 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus_PG_VII | 2.38 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Ad 2163 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.43 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Ad 2164 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus | 2.37 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Ad 2466 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.45 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | Ad 2469 | 3; specific PCR | eDT | 2017-12-15 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.25 | Bacillus | cytotoxicus_PG_VII | 2.22 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-BA-01360 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.11 | Bacillus | cytotoxicus | 2.1 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-BA-01430 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 1.69 | Bacillus | cytotoxicus | 1.65 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-BA-01431 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.1 | Bacillus | cytotoxicus | 1.99 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-BA-01492 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.11 | Bacillus | cytotoxicus | 2.1 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-BA-01522 | 3; sequenced | 0 | 2020-10-06 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 1.95 | Bacillus | cytotoxicus | 1.92 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH1 | 4; WGS | 0 | 0 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.51 | Bacillus | cytotoxicus | 2.41 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH13 | 4; WGS | 0 | 0 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus_PG_VII | 2.42 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH134 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.34 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH15 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.38 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH16 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.59 | Bacillus | cytotoxicus | 2.47 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH2 | 4; WGS | 0 | 0 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.46 | Bacillus | cytotoxicus | 2.42 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH212 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.38 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH213 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.38 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH214 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.39 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH216 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.37 | Bacillus | cytotoxicus_PG_VII | 2.37 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH217 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.57 | Bacillus | cytotoxicus | 2.39 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH218 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.34 | Bacillus | cytotoxicus | 2.33 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH23 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.41 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH25 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.44 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH3 | 4; WGS | 0 | 0 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.51 | Bacillus | cytotoxicus | 2.47 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH348 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.39 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH38 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.34 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH39 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.37 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH4 | 4; WGS | 0 | 0 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.36 | Bacillus | cytotoxicus_PG_VII | 2.35 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH647 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.33 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH755 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH834 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.45 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH835 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.27 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH836 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.27 | Bacillus | cytotoxicus_PG_VII | 2.22 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH837 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.33 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH838 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.22 | Bacillus | cytotoxicus | 2.19 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH839 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.35 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH840 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.32 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH841 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.31 | Bacillus | cytotoxicus | 2.29 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH842 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.39 | Bacillus | cytotoxicus | 2.38 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH843 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.3 | Bacillus | cytotoxicus | 2.3 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH844 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.38 | Bacillus | cytotoxicus_PG_VII | 2.36 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH845 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.35 | Bacillus | cytotoxicus | 2.33 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH846 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.39 | Bacillus | cytotoxicus | 2.31 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH847 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH848 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.33 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH849 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.35 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH850 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH851 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.06 | Bacillus | cytotoxicus | 1.96 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH852 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.29 | Bacillus | cytotoxicus_PG_VII | 2.27 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH853 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.34 | Bacillus | cytotoxicus_PG_VII | 2.34 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH854 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.38 | Bacillus | cytotoxicus | 2.34 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH855 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus_PG_VII | 2.34 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH856 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.27 | Bacillus | cytotoxicus | 2.27 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH857 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.35 | Bacillus | cytotoxicus_PG_VII | 2.27 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH858 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus_PG_VII | 2.39 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH859 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.32 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH860 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.4 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH861 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus_PG_VII | 2.4 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH862 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.49 | Bacillus | cytotoxicus_PG_VII | 2.45 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH863 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.35 | Bacillus | cytotoxicus_PG_VII | 2.33 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH864 | 3; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.44 | Bacillus | cytotoxicus_PG_VII | 2.34 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH865 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.35 | Bacillus | cytotoxicus | 2.34 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH866 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.46 | Bacillus | cytotoxicus | 2.39 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH867 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.47 | Bacillus | cytotoxicus_PG_VII | 2.43 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH868 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus_PG_VII | 2.42 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH869 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.51 | Bacillus | cytotoxicus_PG_VII | 2.43 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH870 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus_PG_VII | 2.31 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH871 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.4 | Bacillus | cytotoxicus_PG_VII | 2.36 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH872 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.43 | Bacillus | cytotoxicus_PG_VII | 2.38 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH873 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.37 | Bacillus | cytotoxicus | 2.34 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH874 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.4 | Bacillus | cytotoxicus_PG_VII | 2.37 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH875 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.38 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH876 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.45 | Bacillus | cytotoxicus | 2.43 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH877 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.5 | Bacillus | cytotoxicus_PG_VII | 2.42 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH878 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.4 | Bacillus | cytotoxicus_PG_VII | 2.37 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH879 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.08 | Bacillus | cytotoxicus_PG_VII | 2.04 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH880 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.47 | Bacillus | cytotoxicus_PG_VII | 2.42 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH881 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.3 | Bacillus | cytotoxicus_PG_VII | 2.29 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH882 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.18 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH883 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.27 | Bacillus | cytotoxicus | 2.26 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH884 | 4; WGS | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.33 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CH886 | 3; PCR | 0 | 2023-01-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.34 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0644 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.11 | Bacillus | cytotoxicus | 2.08 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 0644 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.79 | Bacillus | cytotoxicus | 2.29 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10162 | 3; sequenced | eDT | 2015-11-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.35 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10577 | 4; WGS | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.38 | Bacillus | cytotoxicus | 2.3 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11025 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.44 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11326.2 | 4; WGS | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.44 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2831 | 3; sequenced; Wt 55°C, cytK1 | DT | 2013-01-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.03 | Bacillus | cytotoxicus | 2 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2833 CVUAS | 4; 16S, draft genome [0009] | eDT | 2016-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.53 | Bacillus | cytotoxicus | 2.42 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 2833 RKI | 4; 16S, draft genome [0009] | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.55 | Bacillus | cytotoxicus | 2.28 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2894 | 3; sequenced; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2905 | 3; sequenced; Wt 55°C, cytK1 | DT | 2013-09-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.3 | Bacillus | cytotoxicus | 2.29 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2979 CVUAS | 4; WGS, 16S | eDT | 2016-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.45 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 2979 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.77 | Bacillus | cytotoxicus | 2.31 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2982 | 3; sequenced; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.3 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3017.3 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.32 | Bacillus | cytotoxicus_PG_VII | 2.28 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30572 | 1; 0 | EFex [BRU03] | 2017-11-22 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.51 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30605 | 1; 0 | eDT | 2017-12-01 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.55 | Bacillus | cytotoxicus | 2.43 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30626 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus_PG_VII | 2.41 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31073 | 4; WGS, specific PCR | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.38 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31094 | 4; WGS, specific PCR | eDT | 2018-08-14 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.56 | Bacillus | cytotoxicus | 2.5 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 35239 | 4; WGS; specific PCR | eDT | 2024-09-24 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.39 | Bacillus | cytotoxicus_PG_VII | 2.14 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 35361 | 4; WGS; specific PCR | eDT | 2024-09-24 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.29 | Bacillus | cytotoxicus | 2.27 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 355.3 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.26 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 356.3 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.31 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 363.4 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.38 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3827 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.26 | Bacillus | cytotoxicus | 2.23 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 3827 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.56 | Bacillus | cytotoxicus | 2.16 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3828 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus | 2.42 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 3828 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.68 | Bacillus | cytotoxicus | 2.22 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3896 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-14 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.32 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4117 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-14 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.54 | Bacillus | cytotoxicus_PG_VII | 2.39 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4545 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.29 | Bacillus | cytotoxicus | 2.1 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4546.4 | 3; cyt K1 | eDT | 2015-05-20 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.31 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5334 | 3; Wt 55°C, cytK1 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.44 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5335 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.3 | Bacillus | cytotoxicus | 2.3 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5403 | 2; Wt 55°C | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.45 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5404 | 2; Wt 55°C | eDT | 2016-11-22 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.16 | Bacillus | cytotoxicus | 2.09 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5405 | 2; Wt 55°C | eDT | 2016-11-22 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.15 | Bacillus | cytotoxicus | 2.11 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6081 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.2 | Bacillus | cytotoxicus | 2.07 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6082 CVUAS | 4; WGS; Wt 50°C, cytK1 | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.3 | Bacillus | cytotoxicus | 2.27 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 6082 RKI | 4; WGS; Wt 50°C, cytK1 | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.7 | Bacillus | cytotoxicus | 2.25 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6083 | 3; Wt 50°C, cytK1 | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.28 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6092 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus_PG_VII | 2.35 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6128 | 3; Wt 50°C, cytK1 | EFEx | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.09 | Bacillus | cytotoxicus | 2 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6129 | 3; Wt 50°C, cytK1 | EFEx | 2015-01-15 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.5 | Bacillus | cytotoxicus | 2.4 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6130 | 3; Wt 50°C, cytK1 | EFEx | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.23 | Bacillus | cytotoxicus | 2.15 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6183 CVUAS | 3; Wt 55°C, cytK1 | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 6183 RKI | 3; Wt 55°C, cytK1 | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.66 | Bacillus | cytotoxicus | 2.27 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6204 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.23 | Bacillus | cytotoxicus | 2.19 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6205 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-14 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus_PG_VII | 2.39 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6206 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.43 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6207 | 3; Wt 50°C, cytK1 | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.34 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6208 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.25 | Bacillus | cytotoxicus | 2.21 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6209 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-03-17 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.2 | Bacillus | cytotoxicus | 2.2 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6210 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.32 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6211 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.06 | Bacillus | cytotoxicus | 2.03 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6238 | 3; sequenced; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.21 | Bacillus | cytotoxicus | 2.17 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6239 | 3; sequenced; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.2 | Bacillus | cytotoxicus | 2.17 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6240 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.3 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6241 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.46 | Bacillus | cytotoxicus | 2.35 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6242 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.54 | Bacillus | cytotoxicus | 2.51 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6243 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.37 | Bacillus | cytotoxicus | 2.25 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6244 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.33 | Bacillus | cytotoxicus | 2.25 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 6244 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.63 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6245 | 3; Wt 50°C, cytK1 | eDT | 2022-11-23 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.4 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6292 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.44 | Bacillus | cytotoxicus | 2.26 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6336 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.25 | Bacillus | cytotoxicus_PG_VII | 2.23 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6433.2 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.47 | Bacillus | cytotoxicus | 2.47 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6435 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.42 | Bacillus | cytotoxicus | 2.26 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6475 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.19 | Bacillus | cytotoxicus | 2.12 | B | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6608 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.25 | Bacillus | cytotoxicus_PG_VII | 2.2 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6655 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.25 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6656 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-03-11 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.32 | Bacillus | cytotoxicus | 2.16 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6657 | 3; Wt 50°C, cytK1 | eDT | 2015-03-25 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.18 | Bacillus | cytotoxicus | 2.09 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6658 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.36 | Bacillus | cytotoxicus | 2.24 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6659 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2016-04-26 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus_PG_VII | 2.27 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6660 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.13 | Bacillus | cytotoxicus | 2.07 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6661 | 4; Wt 55°C, cytK1 | eDT | 2015-12-09 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.54 | Bacillus | cytotoxicus | 2.39 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6688 CVUAS | 4; WGS; Wt 55°C, cytK1 | eDT | 2019-09-25 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.47 | Bacillus | cytotoxicus | 2.36 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVUAS 6688 RKI | 4; WGS; Wt 55°C, cytK1 | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.72 | Bacillus | cytotoxicus | 2.19 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6689 | 3; cytK1 | EFex [BRU03] | 2015-03-11 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.48 | Bacillus | cytotoxicus | 2.39 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6690 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.14 | Bacillus | cytotoxicus | 2.05 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6691 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.18 | Bacillus | cytotoxicus | 2.12 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6692 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.27 | Bacillus | cytotoxicus | 2.19 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6693 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.34 | Bacillus | cytotoxicus | 2.3 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6694 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.24 | Bacillus | cytotoxicus_PG_VII | 2.18 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 701.3 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.37 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7208 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.53 | Bacillus | cytotoxicus | 2.35 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7209 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.28 | Bacillus | cytotoxicus | 2.26 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7500 | 3; specific PCR | EFex [BRU03] | 2015-01-13 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.37 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8754 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.41 | Bacillus | cytotoxicus | 2.35 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8755 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.39 | Bacillus | cytotoxicus | 2.38 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9297 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.4 | Bacillus | cytotoxicus | 2.32 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9298 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.52 | Bacillus | cytotoxicus | 2.42 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9299 | 1; 0 | eDT | 2022-11-18 | Bacillus | cytotoxicus | Bacillus | cytotoxicus_PG_VII | 2.43 | Bacillus | cytotoxicus_PG_VII | 2.42 | A | 0 | 0 | 1 | Bacillus cytotoxicus_PG_VII | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22905 RKI | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.68 | Bacillus | cytotoxicus | 2.22 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 22905 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.56 | Bacillus | cytotoxicus | 2.4 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | IGV 003 | 3; sequenced | eDT | 2016-08-12 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.43 | Bacillus | cytotoxicus_PG_VII | 2.4 | B | 0 | 0 | 1 | Bacillus cytotoxicus | Bacillus cytotoxicus_PG_VII | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NVH 391-98 | 3; type strain | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.49 | Bacillus | cytotoxicus | 2.46 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111458 | 3; type strain | eDT | 2025-03-18 | Bacillus | luti | Bacillus | cereus | 2.46 | Bacillus | cereus | 2.41 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28872 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-04 | Bacillus | luti | Bacillus | luti | 2.76 | Bacillus | nitratireducens | 2.43 | B | 0 | 0 | 1 | Bacillus luti | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 111019 | 3; type strain | eDT | 2025-03-18 | Bacillus | manliponensis | Bacillus | manliponensis | 2.37 | Bacillus | manliponensis | 1.57 | A | 1 | 1 | 0 | Bacillus manliponensis | Bacillus manliponensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 26473 CVUAS | 3; type strain | eDT | 2025-03-18 | Bacillus | manliponensis | Bacillus | manliponensis | 2.48 | Bacillus | manliponensis | 1.97 | A | 1 | 1 | 0 | Bacillus manliponensis | Bacillus manliponensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26473 RKI | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-07-25 | Bacillus | manliponensis | Bacillus | manliponensis | 2.42 | Bacillus | manliponensis | 2.14 | A | 1 | 1 | 0 | Bacillus manliponensis | Bacillus manliponensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111457 | 3; type strain | eDT | 2025-03-18 | Bacillus | mobilis | Bacillus | cereus_thuringiensis_PG_II | 2.1 | Bacillus | cereus_thuringiensis_PG_II | 2.08 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_II | Bacillus cereus_thuringiensis_PG_II | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28877 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-04 | Bacillus | mobilis | Bacillus | mobilis | 2.8 | Bacillus | thuringiensis | 2.12 | B | 0 | 0 | 1 | Bacillus mobilis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 12826 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus | mycoides | Bacillus | tropicus | 1.98 | Bacillus | cereus | 1.9 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 103 | 2; 0 | DT | 2015-07-22 | Bacillus | mycoides | Bacillus | cereus | 2.31 | Bacillus | cereus_thuringiensis_PG_III | 2.27 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104109 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.27 | Bacillus | mycoides_PG_VI | 2.22 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 104111 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.14 | Bacillus | mycoides | 2.04 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104112 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.28 | Bacillus | mycoides_PG_VI | 2.24 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104113 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.29 | Bacillus | cereus_thuringiensis_PG_V | 2.2 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104118 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.24 | Bacillus | cereus | 2.22 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104119 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.32 | Bacillus | cereus | 2.25 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 104119 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-05-09 | Bacillus | mycoides | Bacillus | mycoides | 2.76 | Bacillus | nitratireducens | 2.27 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104121 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | cereus_thuringiensis_PG_V | 2.14 | Bacillus | cereus_thuringiensis_PG_V | 2.08 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_V | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104122 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | cereus_thuringiensis_PG_V | 2.22 | Bacillus | cereus | 2.2 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_V | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 104135 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.29 | Bacillus | mycoides | 2.15 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 104137 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.26 | Bacillus | cereus_thuringiensis_PG_V | 2.19 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 11821 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus | mycoides | Bacillus | mycoides_PG_VI | 2.19 | Bacillus | mycoides | 2.18 | B | 0 | 0 | 1 | Bacillus mycoides_PG_VI | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 11821 RKI | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-05-03 | Bacillus | mycoides | Bacillus | mycoides_PG_VI | 2 | Bacillus | mycoides | 1.97 | A | 0 | 0 | 1 | Bacillus mycoides_PG_VI | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 12001 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides_PG_VI | 2.21 | Bacillus | mycoides_PG_VI | 2.18 | A | 0 | 0 | 1 | Bacillus mycoides_PG_VI | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12001 RKI | 3; public strain collection | see [0185] | 2021-12-02 | Bacillus | mycoides | Bacillus | mycoides | 2.36 | Bacillus | mycoides_PG_VI | 2.26 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 2048 CVUAS | 3; type strain | eDT | 09.12.2015 | Bacillus | mycoides | Bacillus | mycoides | 2.45 | Bacillus | mycoides_PG_VI | 2.39 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2048 RKI | 3; type strain | EFex [BRU03] | 2022-01-13 | Bacillus | mycoides | Bacillus | mycoides | 2.15 | Bacillus | mycoides_PG_VI | 2.02 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 299 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | mycoides | 2.29 | Bacillus | mycoides_PG_VI | 2.23 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 299 IGV | 3; public strain collection | DT | 2016-04-08 | Bacillus | mycoides | Bacillus | mycoides | 2.41 | Bacillus | mycoides_PG_VI | 2.27 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 303 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | cereus | 1.99 | Bacillus | cereus_thuringiensis_PG_V | 1.98 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | mycoides | Bacillus | mycoides_PG_VI | 2 | Bacillus | mycoides | 2 | B | 0 | 0 | 1 | Bacillus mycoides_PG_VI | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 30963 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | nitratireducens | 2.18 | Bacillus | cereus_thuringiensis_PG_V | 2.17 | B | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 30966 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | cereus_thuringiensis_PG_V | 2.14 | Bacillus | nitratireducens | 2.11 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_V | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 30969 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | mycoides | Bacillus | nitratireducens | 2.04 | Bacillus | cereus_thuringiensis_PG_V | 1.93 | A | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WS 2641 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA-TA2#EXT=Cultivation & preparation: 02.11.2006, preparation for MALDI MS 06.11.2006: M. Staemmler#LAS=LL26 pct. 600, "LL70 pct." 10 sh. ¦ Laser attenuator offset: 65 pct.¦Range: 20 pct.#CAL=ProteinStandard 1 (Bruker)#CUS=RKI ZBS6# | 2006-07-11 | Bacillus | mycoides | Bacillus | mycoides | 2.46 | Bacillus | mycoides_PG_VI | 2.33 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10204 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus | mycoides | Bacillus | mycoides | 2.37 | Bacillus | mycoides_PG_VI | 2.31 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus mycoides_PG_VI | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10207 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus | mycoides | Bacillus | mycoides | 2.31 | Bacillus | mycoides | 2.3 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10277 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | mycoides | Bacillus | mycoides | 2.16 | Bacillus | mycoides | 2.02 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10279 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | mycoides | Bacillus | mycoides | 2.17 | Bacillus | mycoides | 2.17 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10291 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | mycoides | Bacillus | pseudomycoides_PG_I | 2.16 | Bacillus | pseudomycoides | 2.15 | B | 0 | 0 | 1 | Bacillus pseudomycoides_PG_I | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10295 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus | mycoides | Bacillus | mycoides | 2.47 | Bacillus | proteolyticus | 2.09 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10360 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus | mycoides | Bacillus | pseudomycoides | 2.08 | Bacillus | pseudomycoides_PG_I | 1.96 | A | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pseudomycoides_PG_I | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10363 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus | mycoides | Bacillus | mycoides | 2.36 | Bacillus | nitratireducens | 2.04 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 10387 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus | mycoides | Bacillus | mycoides | 2.62 | Bacillus | proteolyticus | 2.14 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28871 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-01-25 | Bacillus | nitratireducens | Bacillus | nitratireducens | 2.77 | Bacillus | proteolyticus | 2.51 | B | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111460 | 3; type strain | eDT | 2025-03-18 | Bacillus | nitratireduscens | Bacillus | nitratireducens | 2.27 | Bacillus | cereus_thuringiensis_PG_V | 2.12 | B | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10987 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-10-23 | Bacillus | pacificus | Bacillus | cereus | 2.26 | Bacillus | cereus | 2.06 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SN4-1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-04-11 | Bacillus | pacificus | Bacillus | pacificus | 2.78 | Bacillus | cereus | 2.37 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111542 | 3; type strain | eDT | 2025-03-18 | Bacillus | pacificus | Bacillus | cereus_thuringiensis_PG_III | 2.48 | Bacillus | cereus_thuringiensis_PG_III | 2.4 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111461 | 3; type strain | eDT | 2025-03-18 | Bacillus | paramycoides | Bacillus | cereus_thuringiensis_PG_V | 2.31 | Bacillus | cereus_thuringiensis_PG_V | 2.3 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_V | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28876 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 3.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Bacillus paramycoides LMG 28876, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-02-04 | Bacillus | paramycoides | Bacillus | nitratireducens | 2.36 | Bacillus | paramycoides | 2.3 | B | 0 | 0 | 1 | Bacillus nitratireducens | Bacillus paramycoides | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111454 | 3; type strain | eDT | 2025-03-18 | Bacillus | paranthracis | Bacillus | cereus_thuringiensis_PG_III | 2.39 | Bacillus | cereus | 2.38 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP A30 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | paranthracis | Bacillus | cereus_thuringiensis_PG_III | 2.33 | Bacillus | cereus_thuringiensis_PG_III | 2.24 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28873 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-11 | Bacillus | paranthracis | Bacillus | paranthracis | 2.75 | Bacillus | cereus | 2.33 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111459 | 3; type strain | eDT | 2025-03-18 | Bacillus | proteolyticus | Bacillus | cereus_thuringiensis_PG_V | 2.22 | Bacillus | cereus_thuringiensis_PG_V | 2.22 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_V | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28870 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-01 | Bacillus | proteolyticus | Bacillus | proteolyticus | 2.77 | Bacillus | nitratireducens | 2.53 | B | 0 | 0 | 1 | Bacillus proteolyticus | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.53 | Bacillus | cereus | 2.48 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 12442 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.3 | Bacillus | pseudomycoides | 2.2 | A | 1 | 1 | 0 | Bacillus pseudomycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12442 RKI | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2007-05-03 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.23 | Bacillus | pseudomycoides_PG_I | 2.21 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pseudomycoides_PG_I | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 12443 CVUAS | 3; public strain collection | eDT | 09.12.2015 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.64 | Bacillus | pseudomycoides_PG_I | 2.54 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pseudomycoides_PG_I | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12443 RKI | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-05 | Bacillus | pseudomycoides | Bacillus | pseudomycoides | 2.34 | Bacillus | pseudomycoides_PG_I | 2.28 | B | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pseudomycoides_PG_I | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.64 | Bacillus | thuringiensis | 2.52 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.67 | Bacillus | thuringiensis | 2.64 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.55 | Bacillus | thuringiensis | 2.51 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 5724 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_IV | 2.39 | Bacillus | cereus_thuringiensis_PG_IV | 2.33 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 5815 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_IV | 2.31 | Bacillus | thuringiensis | 2.25 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6022 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.5 | Bacillus | thuringiensis | 2.48 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5815 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 02.12.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus thuringiensis DSM 5815, preparation and measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-12-02 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.45 | Bacillus | thuringiensis | 2.35 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.62 | Bacillus | thuringiensis | 2.54 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 5725 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus | 2.37 | Bacillus | cereus_thuringiensis_PG_III | 2.29 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.57 | Bacillus | thuringiensis | 2.57 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 8 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-01 | Bacillus | thuringiensis | Bacillus | cereus | 2.46 | Bacillus | thuringiensis | 2.31 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10203 | 4; Cry-pos, Cyt-K2, Krist. Pos | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.53 | Bacillus | thuringiensis | 2.51 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5760 | 4; Cry-pos, Krist. Pos | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.6 | Bacillus | thuringiensis | 2.58 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5761 | 4; Cry-pos, Krist. Pos | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.54 | Bacillus | thuringiensis | 2.48 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2046 (B188. Beyer) | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus | thuringiensis | Bacillus | wiedmannii | 2.17 | Bacillus | cereus_thuringiensis_PG_II | 2.17 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus cereus_thuringiensis_PG_II | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 2046 CVUAS | 3; type strain | DT | 2014-01-22 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_IV | 2.25 | Bacillus | cereus_thuringiensis_PG_IV | 2.22 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2046 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus | thuringiensis | Bacillus | toyonensis | 2.07 | Bacillus | thuringiensis | 2.06 | B | 0 | 0 | 1 | Bacillus toyonensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 350 WIS-St. Nr. 315 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus | thuringiensis | Bacillus | toyonensis | 1.96 | Bacillus | thuringiensis | 1.87 | B | 0 | 0 | 1 | Bacillus toyonensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-03 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.2 | Bacillus | nitratireducens | 2.1 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus nitratireducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6022 IGV | 3; public strain collection | DT | 2016-04-06 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.43 | Bacillus | thuringiensis | 2.28 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6024 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus | 2.24 | Bacillus | cereus_thuringiensis_PG_IV | 2.15 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6030 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_IV | 2.25 | Bacillus | cereus_thuringiensis_PG_IV | 2.25 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6033 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.32 | Bacillus | cereus_thuringiensis_PG_IV | 2.29 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6073 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-12-04 | Bacillus | thuringiensis | Bacillus | cereus | 2.13 | Bacillus | cereus | 2.1 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6087 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_IV | 2.36 | Bacillus | thuringiensis | 2.32 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6099 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_III | 2.23 | Bacillus | cereus_thuringiensis_PG_III | 2.21 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6100 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus | 2.3 | Bacillus | cereus_thuringiensis_PG_III | 2.29 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6101 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.11 | Bacillus | thuringiensis | 2.05 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6102 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus | 2.09 | Bacillus | thuringiensis | 2.08 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6106 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_IV | 2.25 | Bacillus | thuringiensis | 2.24 | B | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 6890 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus | thuringiensis | Bacillus | cereus_thuringiensis_PG_IV | 2.27 | Bacillus | cereus_thuringiensis_PG_IV | 2.23 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_IV | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6890 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-12-02 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.42 | Bacillus | proteolyticus | 2.25 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus proteolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ05-05 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 05.10.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-10-09 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.4 | Bacillus | thuringiensis | 2.26 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSC29 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 21.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-01 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.34 | Bacillus | tropicus | 2.08 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Konkukian | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 1:10 diluted , spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Cells from solid cultures#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2015-02-03 | Bacillus | thuringiensis | Bacillus | tropicus | 2.54 | Bacillus | albus | 2.45 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMTD81 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=inactivation date 15.03.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2021-08-21 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.37 | Bacillus | thuringiensis | 2.31 | A | 1 | 1 | 0 | Bacillus thuringiensis | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WS 2614 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus | thuringiensis | Bacillus | paranthracis | 2.3 | Bacillus | cereus | 2.28 | B | 0 | 0 | 1 | Bacillus paranthracis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WS 2621 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.09 | Bacillus | cereus_thuringiensis_PG_IV | 2.08 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus_thuringiensis_PG_IV | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WS 2623 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus | thuringiensis | Bacillus | cereus | 2.23 | Bacillus | paranthracis | 2.19 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-11-07 | Bacillus | thuringiensis | Bacillus | cereus | 2.33 | Bacillus | paranthracis | 2.25 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus paranthracis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 43 Paciflor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus | toyonensis | Bacillus | toyonensis | 1.91 | Bacillus | toyonensis | 1.86 | B | 0 | 0 | 1 | Bacillus toyonensis | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 8 Toyocerin | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2022-06-28 | Bacillus | toyonensis | Bacillus | toyonensis | 2.25 | Bacillus | toyonensis | 2.2 | A | 1 | 1 | 0 | Bacillus toyonensis | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B260 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus | toyonensis | Bacillus | toyonensis | 2.44 | Bacillus | toyonensis | 2.34 | A | 1 | 1 | 0 | Bacillus toyonensis | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP A31 | 3; public strain collection | eDT | 2025-03-18 | Bacillus | toyonensis | Bacillus | cereus | 2.52 | Bacillus | cereus_thuringiensis_PG_V | 2.5 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus_thuringiensis_PG_V | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMTD92 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 03.11.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-01 | Bacillus | toyonensis | Bacillus | toyonensis | 2.7 | Bacillus | luti | 2.28 | B | 0 | 0 | 1 | Bacillus toyonensis | Bacillus luti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Paciflor | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2022-06-21 | Bacillus | toyonensis | Bacillus | toyonensis | 2.74 | Bacillus | toyonensis | 2.55 | A | 1 | 1 | 0 | Bacillus toyonensis | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Toyocerin | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus | toyonensis | Bacillus | toyonensis | 2.76 | Bacillus | toyonensis | 2.62 | A | 1 | 1 | 0 | Bacillus toyonensis | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-08 | Bacillus | tropicus | Bacillus | pacificus | 2.42 | Bacillus | tropicus | 2.4 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CD 3-2 | 2; strain collection of the RKI | see [0185] | 2022-05-11 | Bacillus | tropicus | Bacillus | pacificus | 2.23 | Bacillus | cereus | 2.16 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111455 | 3; type-strain | eDT | 2025-03-18 | Bacillus | tropicus | Bacillus | cereus_thuringiensis_PG_III | 2.32 | Bacillus | cereus_thuringiensis_PG_III | 2.24 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_III | Bacillus cereus_thuringiensis_PG_III | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28874 | 3; type-strain | see [0185] | 2022-02-04 | Bacillus | tropicus | Bacillus | tropicus | 2.78 | Bacillus | albus | 2.45 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus albus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SN1 | 2; strain collection of the RKI | see [0185] | 2020-12-15 | Bacillus | tropicus | Bacillus | tropicus | 2.8 | Bacillus | tropicus | 2.58 | A | 1 | 1 | 0 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | VM11 (01) | 2; strain collection of the RKI | see [0185] | 2018-11-14 | Bacillus | tropicus | Bacillus | tropicus | 2.47 | Bacillus | tropicus | 2.41 | A | 1 | 1 | 0 | Bacillus tropicus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | VM21 | 2; strain collection of the RKI | see [0185] | 2018-08-17 | Bacillus | tropicus | Bacillus | tropicus | 2.57 | Bacillus | pseudomycoides | 2.27 | B | 0 | 0 | 1 | Bacillus tropicus | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B256 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.38 | Bacillus | toyonensis | 2.19 | B | 0 | 0 | 1 | Bacillus wiedmannii | Bacillus toyonensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 102050 CVUAS | 3; type strain | eDT | 2025-03-18 | Bacillus | wiedmannii | Bacillus | cereus_thuringiensis_PG_II | 2.38 | Bacillus | cereus_thuringiensis_PG_II | 2.37 | A | 0 | 0 | 1 | Bacillus cereus_thuringiensis_PG_II | Bacillus cereus_thuringiensis_PG_II | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102050 RKI | 3; type strain | see [0185] | 2018-05-23 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.75 | Bacillus | wiedmannii | 2.45 | A | 1 | 1 | 0 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSC29 (01) | 2; strain collection of the RKI | see [0185] | 2018-05-17 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.65 | Bacillus | wiedmannii | 2.39 | A | 1 | 1 | 0 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSC45 | 2; strain collection of the RKI | see [0185] | 2018-08-23 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.43 | Bacillus | wiedmannii | 2.4 | A | 1 | 1 | 0 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LLCG23 | 2; strain collection of the RKI | see [0185] | 2018-05-17 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.44 | Bacillus | wiedmannii | 2.38 | A | 1 | 1 | 0 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMSL88 | 2; strain collection of the RKI | see [0185] | 2018-08-21 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.66 | Bacillus | wiedmannii | 2.33 | A | 1 | 1 | 0 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMSW16 (02) | 2; strain collection of the RKI | see [0185] | 2018-01-10 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.48 | Bacillus | wiedmannii | 2.35 | A | 1 | 1 | 0 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMSW23 | 2; strain collection of the RKI | see [0185] | 2018-08-21 | Bacillus | wiedmannii | Bacillus | wiedmannii | 2.52 | Bacillus | wiedmannii | 2.38 | A | 1 | 1 | 0 | Bacillus wiedmannii | Bacillus wiedmannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 338 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus | cereus | Bacillus | pacificus | 2.46 | Bacillus | tropicus | 2.3 | B | 0 | 0 | 1 | Bacillus pacificus | Bacillus tropicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-31 | Bacillus | ectoiniformans | Bacillus | ectoiniformans | 2.78 | Streptococcus | dysgalactiae | 1.49 | A | 1 | 1 | 0 | Bacillus ectoiniformans | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 111018 | 3; type strain | eDT | 2025-03-18 | Bacillus | gaemokensis | Bacillus | pseudomycoides | 2.2 | Bacillus | pseudomycoides | 2.19 | A | 0 | 0 | 1 | Bacillus pseudomycoides | Bacillus pseudomycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29500 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Strain obtained from DSMZ strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-09-05 | Bacillus | gobiensis | Bacillus | gobiensis | 2.69 | Stenotrophomonas | pictorum | 1.52 | A | 1 | 1 | 0 | Bacillus gobiensis | Stenotrophomonas pictorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-30 | Bacillus | halotolerans | Bacillus | mojavensis | 2.8 | Bacillus | axarquiensis | 2.55 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus axarquiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-12 | Bacillus | halotolerans | Bacillus | halotolerans | 2.73 | Bacillus | mojavensis | 2.23 | B | 0 | 0 | 1 | Bacillus halotolerans | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21200 | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-28 | Bacillus | inaquosorum | Bacillus | inaquosorum | 2.82 | Bacillus | inaquosorum | 2.14 | A | 1 | 1 | 0 | Bacillus inaquosorum | Bacillus inaquosorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-12-02 | Bacillus | inaquosorum | Bacillus | inaquosorum | 2.78 | Bacillus | subtilis | 2.27 | B | 0 | 0 | 1 | Bacillus inaquosorum | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-12 | Bacillus | infantis | Bacillus | infantis | 2.46 | Bacillus | infantis | 2.39 | A | 1 | 1 | 0 | Bacillus infantis | Bacillus infantis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141014442-3 | 3; VETQAS | DT | 2014-11-07 | Bacillus | licheniformis | Bacillus | licheniformis | 2.02 | Bacillus | licheniformis | 1.99 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141015909-3 | 2; 0 | DT | 2014-12-12 | Bacillus | licheniformis | Bacillus | licheniformis | 2.21 | Bacillus | licheniformis | 2.15 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004445 | 2; 0 | DT | 2016-04-12 | Bacillus | licheniformis | Bacillus | licheniformis | 2.09 | Bacillus | licheniformis | 2.05 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161007249-3 | 2; 0 | DT | 2016-06-14 | Bacillus | licheniformis | Bacillus | licheniformis | 2.15 | Bacillus | licheniformis | 2.11 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14580 | 3; type strain | DT | 2015-06-23 | Bacillus | licheniformis | Bacillus | licheniformis | 2.11 | Bacillus | licheniformis | 2.11 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus | licheniformis | Bacillus | licheniformis | 2.17 | Bacillus | licheniformis | 2.15 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-26 | Bacillus | licheniformis | Bacillus | licheniformis | 2.19 | Bacillus | paralicheniformis | 1.55 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus paralicheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 116 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-26 | Bacillus | licheniformis | Bacillus | licheniformis | 1.87 | Bacillus | paralicheniformis | 1.87 | B | 0 | 0 | 1 | Bacillus licheniformis | Bacillus paralicheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 118 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-26 | Bacillus | licheniformis | Bacillus | licheniformis | 2.3 | Bacillus | licheniformis | 1.47 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 62 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus | licheniformis | Bacillus | licheniformis | 2.18 | Bacillus | licheniformis | 1.71 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 72 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-14 | Bacillus | licheniformis | Bacillus | licheniformis | 2.19 | Bacillus | licheniformis | 1.76 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BRLK08_04 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-04-17 | Bacillus | licheniformis | Bacillus | licheniformis | 2.3 | Bacillus | licheniformis | 1.89 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13 B178 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus | licheniformis | Bacillus | licheniformis | 2.26 | Bacillus | licheniformis | 1.73 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus | licheniformis | Bacillus | licheniformis | 2.27 | Bacillus | licheniformis | 1.81 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-12 | Bacillus | licheniformis | Bacillus | licheniformis | 1.74 | Bacillus | licheniformis | 1.61 | B | 0 | 0 | 1 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-10 | Bacillus | mojavensis | Bacillus | mojavensis | 2.78 | Bacillus | halotolerans | 2.36 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus halotolerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Bacillus | mojavensis | Bacillus | halotolerans | 2.42 | Bacillus | mojavensis | 2.33 | B | 0 | 0 | 1 | Bacillus halotolerans | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 12759 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus | paralicheniformis | Staphylococcus | epidermidis | 1.45 | Staphylococcus | carnosus | 1.45 | B | 0 | 0 | 1 | Staphylococcus epidermidis | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ01-02 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-02-24 | Bacillus | paralicheniformis | Bacillus | paralicheniformis | 2.37 | Bacillus | licheniformis | 1.7 | A | 1 | 1 | 0 | Bacillus paralicheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ01-05 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-02-27 | Bacillus | paralicheniformis | Bacillus | paralicheniformis | 2.65 | Bacillus | licheniformis | 1.78 | A | 1 | 1 | 0 | Bacillus paralicheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ10-02 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2018-03-05 | Bacillus | paralicheniformis | Bacillus | paralicheniformis | 2.5 | Bacillus | licheniformis | 1.66 | A | 1 | 1 | 0 | Bacillus paralicheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ15-08 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.07.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-07-23 | Bacillus | paralicheniformis | Bacillus | paralicheniformis | 2.78 | Bacillus | licheniformis | 1.8 | A | 1 | 1 | 0 | Bacillus paralicheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMSW12 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2027-04-10 | Bacillus | paralicheniformis | Bacillus | paralicheniformis | 2.2 | Bacillus | licheniformis | 1.69 | A | 1 | 1 | 0 | Bacillus paralicheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34924 | 3; sequenced | EFEx | 2023-02-17 | Bacillus | pumilus | Bacillus | pumilus | 2.3 | Bacillus | pumilus | 2 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LGA MB 2022000019 | 2; confirmed proficency test | eDT | 2022-02-14 | Bacillus | pumilus | Bacillus | altitudinis | 2.1 | Bacillus | altitudinis | 2.08 | A | 0 | 0 | 1 | Bacillus altitudinis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 17 CVUAS | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus | pumilus | Bacillus | pumilus | 2.15 | Bacillus | pumilus | 1.96 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 17 RKI | 3; sequenced (CVUAS) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Bacillus | pumilus | Bacillus | pumilus | 2.13 | Bacillus | pumilus | 2.12 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 31 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-11 | Bacillus | pumilus | Bacillus | safensis | 2.2 | Bacillus | pumilus | 2.04 | B | 0 | 0 | 1 | Bacillus safensis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 91 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-14 | Bacillus | pumilus | Bacillus | pumilus | 2.71 | Bacillus | pumilus | 2.2 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 340 RKI | 3; reference strain collection | see [0185] | 2015-02-09 | Bacillus | pumilus | Bacillus | pumilus | 2.83 | Bacillus | pumilus | 2.23 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 340 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Bacillus | pumilus | Bacillus | pumilus | 2.04 | Bacillus | pumilus | 1.95 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CJCL2 | 2; strain collection of the RKI | see [0185] | 2018-01-16 | Bacillus | pumilus | Bacillus | pumilus | 2.51 | Bacillus | pumilus | 2.27 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13835 LMTG62N07 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-23 | Bacillus | pumilus | Bacillus | pumilus | 1.83 | Bacillus | pumilus | 1.77 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13835 RKI | 3; public strain collection | see [0185] | 2022-08-05 | Bacillus | pumilus | Bacillus | pumilus | 1.81 | Bacillus | pumilus | 1.81 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1794 | 3; public strain collection | see [0185] | 2022-08-05 | Bacillus | pumilus | Bacillus | pumilus | 2.84 | Bacillus | altitudinis | 2.06 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Bacillus | pumilus | Bacillus | pumilus | 2.18 | Bacillus | pumilus | 2.13 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27 Neukauf RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-15 | Bacillus | pumilus | Bacillus | pumilus | 2.27 | Bacillus | pumilus | 2.09 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 27 CVUAS | 3; type strain | DT | 2014-01-22 | Bacillus | pumilus | Bacillus | pumilus | 2.22 | Bacillus | pumilus | 2.02 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 355 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus | pumilus | Bacillus | pumilus | 2.27 | Bacillus | safensis | 2.09 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus safensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Bacillus | pumilus | Bacillus | pumilus | 2.64 | Bacillus | pumilus | 2.31 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 492 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-16 | Bacillus | pumilus | Bacillus | pumilus | 2.52 | Bacillus | pumilus | 2.3 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8786 | 3; public strain collection | see [0185] | 2022-08-05 | Bacillus | pumilus | Bacillus | pumilus | 2.84 | Bacillus | safensis | 2.55 | B | 0 | 0 | 1 | Bacillus pumilus | Bacillus safensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ01-05 (01) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-02-07 | Bacillus | pumilus | Bacillus | pumilus | 2.22 | Bacillus | pumilus | 2.21 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBAC46 (01) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-01-10 | Bacillus | pumilus | Bacillus | pumilus | 2.36 | Bacillus | pumilus | 2.33 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSC66 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-08-14 | Bacillus | pumilus | Bacillus | pumilus | 2.27 | Bacillus | pumilus | 2.24 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSW19 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 | Bacillus | pumilus | Bacillus | pumilus | 2.39 | Bacillus | pumilus | 2.39 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSW2 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-08 | Bacillus | pumilus | Bacillus | pumilus | 2.66 | Bacillus | pumilus | 2.26 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LDZX38 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-08-15 | Bacillus | pumilus | Bacillus | zhangzhouensis | 2.19 | Bacillus | pumilus | 2.16 | B | 0 | 0 | 1 | Bacillus zhangzhouensis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LLTC96 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-16 | Bacillus | pumilus | Bacillus | zhangzhouensis | 2.2 | Bacillus | pumilus | 2.18 | B | 0 | 0 | 1 | Bacillus zhangzhouensis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNTW65 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 | Bacillus | pumilus | Bacillus | safensis | 2.23 | Bacillus | pumilus | 2.16 | B | 0 | 0 | 1 | Bacillus safensis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNXM10 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 | Bacillus | pumilus | Bacillus | pumilus | 2.49 | Bacillus | pumilus | 2.3 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNXM12 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-07-12 | Bacillus | pumilus | Bacillus | pumilus | 2.37 | Bacillus | pumilus | 2.3 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNXM65 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-05-08 | Bacillus | pumilus | Bacillus | pumilus | 2.4 | Bacillus | pumilus | 2.12 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LNXM70 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-03-09 | Bacillus | pumilus | Bacillus | safensis | 2.08 | Bacillus | pumilus | 2.04 | B | 0 | 0 | 1 | Bacillus safensis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MM 171 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus | pumilus | Bacillus | pumilus | 2.42 | Bacillus | pumilus | 2.19 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MZGC1 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-10 | Bacillus | pumilus | Bacillus | pumilus | 2.54 | Bacillus | pumilus | 2.32 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMCC4 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 26.07.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 | Bacillus | pumilus | Bacillus | pumilus | 2.32 | Bacillus | pumilus | 2.25 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMCC46 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Inactivation date 08.03.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-08-17 | Bacillus | pumilus | Bacillus | pumilus | 2.37 | Bacillus | pumilus | 2.34 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMCN1 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-08-15 | Bacillus | pumilus | Bacillus | pumilus | 2.04 | Bacillus | pumilus | 1.99 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMCN6 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-08-15 | Bacillus | pumilus | Bacillus | pumilus | 2.09 | Bacillus | pumilus | 2.07 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMSW10 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 01.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-09 | Bacillus | pumilus | Bacillus | pumilus | 2.03 | Bacillus | pumilus | 2.03 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMSX4 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-02 | Bacillus | pumilus | Bacillus | pumilus | 2.41 | Bacillus | pumilus | 2.29 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMTD17 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.07.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-07-18 | Bacillus | pumilus | Bacillus | pumilus | 2.33 | Bacillus | pumilus | 2.16 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NVH-819-05 | 3; well described isolate [0017] | DT | 2015-04-21 | Bacillus | pumilus | Bacillus | pumilus | 2.11 | Bacillus | pumilus | 1.95 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | YBWC18 | 2; strain collection of the RKI | see [0185] | 2018-08-16 | Bacillus | pumilus | Bacillus | pumilus | 2.51 | Bacillus | pumilus | 2.28 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 103081 | 3; public strain collection | see [0185] | 2023-03-02 | Bacillus | safensis | Bacillus | safensis | 2.85 | Bacillus | pumilus | 2.41 | B | 0 | 0 | 1 | Bacillus safensis | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19292 | 3; type strain | see [0185] | 2023-03-02 | Bacillus | safensis | Bacillus | safensis | 2.74 | Bacillus | altitudinis | 2.29 | B | 0 | 0 | 1 | Bacillus safensis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSW22 | 2; strain collection of the RKI | see [0185] | 2023-03-02 | Bacillus | safensis | Bacillus | safensis | 2.33 | Bacillus | altitudinis | 2.11 | B | 0 | 0 | 1 | Bacillus safensis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-05-24 | Bacillus | siamensis | Bacillus | siamensis | 2.68 | Bacillus | velezensis | 2.2 | B | 0 | 0 | 1 | Bacillus siamensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-15 | Bacillus | sonorensis | Bacillus | sonorensis | 2.6 | Bacillus | sp. | 2.05 | B | 0 | 0 | 1 | Bacillus sonorensis | Bacillus sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 324 | 2; strain collection of the RKI | TFAextr | 2015-05-07 | Bacillus | sp | Bacillus | subtilis | 2.36 | Bacillus | subtilis | 2.28 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12369 RKI | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Bacillus | sp-DSM-12369 | Bacillus | sp. | 2.83 | Bacillus | sonorensis | 2.03 | B | 0 | 0 | 1 | Bacillus sp. | Bacillus sonorensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 6633 RKI B | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-01 | Bacillus | spizizenii | Bacillus | subtilis | 2.12 | Bacillus | subtilis | 1.88 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 6633 RKI A | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus | spizizenii | Bacillus | mojavensis | 1.84 | Bacillus | subtilis | 1.63 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 6633 CVUAS | 3; public strain collection | eDT | 2015-07-22 | Bacillus | spizizenii | Bacillus | subtilis | 2.08 | Bacillus | subtilis | 1.96 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-02 | Bacillus | spizizenii | Bacillus | subtilis | 2.21 | Bacillus | subtilis | 2.08 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 061 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-01-07 | Bacillus | spizizenii | Bacillus | subtilis | 1.99 | Bacillus | subtilis | 1.82 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1087 (B156) | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2017-12-18 | Bacillus | spizizenii | Bacillus | subtilis | 2.37 | Bacillus | subtilis | 2.08 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 1087 CVUAS | 3; public strain collection | DT | 2014-01-22 | Bacillus | spizizenii | Bacillus | subtilis | 2.02 | Bacillus | subtilis | 1.96 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1087 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus | spizizenii | Bacillus | subtilis | 2.17 | Bacillus | subtilis | 2.05 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 347 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2007-04-23 | Bacillus | spizizenii | Bacillus | subtilis | 2.04 | Bacillus | inaquosorum | 1.8 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus inaquosorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 618 | 4; public strain collection | eDT | 2016-01-27 | Bacillus | spizizenii | Bacillus | subtilis | 2.09 | Bacillus | subtilis | 2.04 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6395 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 17.11.2021¦ 1:1 HCCA¦Andy 3.0#EXT | 2021-11-17 | Bacillus | spizizenii | Bacillus | subtilis | 2.24 | Bacillus | subtilis | 2.07 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WIS-St. Nr. 491 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus | spizizenii | Bacillus | tequilensis | 1.81 | Bacillus | subtilis | 1.81 | B | 0 | 0 | 1 | Bacillus tequilensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 130 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus | subtilis | Bacillus | subtilis | 1.98 | Bacillus | subtilis | 1.96 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 402 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-28 | Bacillus | subtilis | Bacillus | subtilis | 2.45 | Bacillus | subtilis | 2.37 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 6051 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus | subtilis | Bacillus | subtilis | 2.03 | Bacillus | subtilis | 1.9 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 014 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-11 | Bacillus | subtilis | Bacillus | subtilis | 1.62 | Bacillus | mojavensis | 1.51 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 020 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-02 | Bacillus | subtilis | Bacillus | subtilis | 2.26 | Bacillus | subtilis | 1.88 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 129 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus | subtilis | Bacillus | subtilis | 2.03 | Bacillus | subtilis | 1.87 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-23 | Bacillus | subtilis | Bacillus | subtilis | 2.21 | Bacillus | subtilis | 1.98 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BRLK08_02 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Bacillus | subtilis | Bacillus | subtilis | 2.59 | Bacillus | subtilis | 2.1 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 110 B068 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.33 | Bacillus | subtilis | 2.15 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 10 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-23 | Bacillus | subtilis | Bacillus | subtilis | 2.41 | Bacillus | subtilis | 2.28 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1092 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus | subtilis | Bacillus | subtilis | 2.38 | Bacillus | subtilis | 2.31 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1092 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.23 | Bacillus | subtilis | 2.2 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | E rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-23 | Bacillus | subtilis | Bacillus | subtilis | 2.3 | Bacillus | subtilis | 1.99 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GR2.1 | 2; strain collection of the RKI | see [0185] | 2020-01-16 | Bacillus | subtilis | Bacillus | subtilis | 2.45 | Bacillus | subtilis | 2.32 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HY 1152 | 2; strain collection of the RKI | see [0185] | 2022-05-24 | Bacillus | subtilis | Bacillus | subtilis | 2.48 | Bacillus | inaquosorum | 2.05 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus inaquosorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | I rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-11 | Bacillus | subtilis | Bacillus | subtilis | 2.2 | Bacillus | subtilis | 1.89 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 013 IGV | 3; sequenced | DT | 2016-11-09 | Bacillus | subtilis | Bacillus | subtilis | 2.27 | Bacillus | subtilis | 1.94 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LVM03 | 3; sequenced (LGL) | eDT | 2015-07-09 | Bacillus | subtilis | Bacillus | subtilis | 2.1 | Bacillus | subtilis | 1.94 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MM 508 subt RKI | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus | subtilis | Bacillus | velezensis | 2.08 | Bacillus | velezensis | 1.94 | A | 0 | 0 | 1 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NMSX4 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-06-16 | Bacillus | subtilis | Bacillus | subtilis | 2.78 | Bacillus | subtilis | 2.16 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RKI-w2gr | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-04-04 | Bacillus | subtilis | Bacillus | subtilis | 2.45 | Bacillus | subtilis | 2.04 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SYST2 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-06-16 | Bacillus | subtilis | Bacillus | subtilis | 2.79 | Bacillus | inaquosorum | 2.12 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus inaquosorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 12 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=in 80 pct. TFA, diluted by 1:1 HCCA:TA2 | 2014-12-01 | Bacillus | subtilis | Bacillus | subtilis | 2.29 | Bacillus | subtilis | 1.61 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16466 | 3; public strain collection | see [0185] | 2022-08-19 | Bacillus | taeanensis | Bacillus | taeanensis | 2.81 | Microbacterium | schleiferi | 1.46 | A | 1 | 1 | 0 | Bacillus taeanensis | Microbacterium schleiferi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WIS-St. Nr. 494 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦ 1:1 HCCA-TA2 (A)#EXT=Cultivation & sample preparation by WIS; Shipment: 20.12.2005; Preparation for MALDI-ToF Maren Staemmler#LAS=LL25 pct. 600, "LL70 pct." 10 sh. offset: 63 pct.¦Range: 20 pct. Laser neu Defok.#CAL=ProteinStandard 1 (Bruker)#CUS=RKI ZBS6 ¦WIS (Dr. Niederwoehrmeier, Munster)# | 2007-01-15 | Bacillus | thuringiensis | Bacillus | cereus | 2.33 | Bacillus | thuringiensis | 2.26 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2786.2 | 3; sequenced | DT | 2013-07-31 | Bacillus | timorensis | Bacillus | timorensis | 2.59 | Bacillus | timonensis | 1.49 | A | 1 | 1 | 0 | Bacillus timorensis | Bacillus timonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-07-20 | Bacillus | vallismortis | Bacillus | vallismortis | 2.37 | Bacillus | inaquosorum | 1.97 | A | 1 | 1 | 0 | Bacillus vallismortis | Bacillus inaquosorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JX-1 (01) | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2017-07-14 | Bacillus | velezensis | Bacillus | velezensis | 2.68 | Bacillus | velezensis | 2.55 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JX-1 (02) | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2017-07-14 | Bacillus | velezensis | Bacillus | velezensis | 2.41 | Bacillus | velezensis | 2.32 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A25 vel | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus | velezensis | Bacillus | velezensis | 2.61 | Bacillus | velezensis | 2.41 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A35 vel | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-10-19 | Bacillus | velezensis | Bacillus | velezensis | 2.66 | Bacillus | velezensis | 2.29 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 032 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-10 | Bacillus | velezensis | Bacillus | velezensis | 2.37 | Bacillus | velezensis | 2.13 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP1.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus | velezensis | Bacillus | velezensis | 2.41 | Bacillus | velezensis | 2.39 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP1.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus | velezensis | Bacillus | velezensis | 2.59 | Bacillus | velezensis | 2.27 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP1.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus | velezensis | Bacillus | velezensis | 2.61 | Bacillus | velezensis | 2.32 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP1.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus | velezensis | Bacillus | velezensis | 2.68 | Bacillus | velezensis | 2.39 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP2.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.59 | Bacillus | velezensis | 2.41 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP2.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.61 | Bacillus | velezensis | 2.51 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP2.3 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.68 | Bacillus | velezensis | 2.51 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BP3.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.62 | Bacillus | velezensis | 2.42 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BT2.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-13 | Bacillus | velezensis | Bacillus | velezensis | 2.67 | Bacillus | velezensis | 2.28 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BT2.4 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-13 | Bacillus | velezensis | Bacillus | velezensis | 2.61 | Bacillus | velezensis | 2.43 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CB1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-17 | Bacillus | velezensis | Bacillus | velezensis | 2.61 | Bacillus | velezensis | 2.3 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CB1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus | velezensis | Bacillus | velezensis | 2.4 | Bacillus | velezensis | 2.25 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CF1.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.79 | Bacillus | velezensis | 2.49 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CF2.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.66 | Bacillus | velezensis | 2.34 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CF2.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-23 | Bacillus | velezensis | Bacillus | velezensis | 2.72 | Bacillus | velezensis | 2.32 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CF2.3 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-23 | Bacillus | velezensis | Bacillus | velezensis | 2.67 | Bacillus | velezensis | 2.21 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP1.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Bacillus | velezensis | Bacillus | velezensis | 2.69 | Bacillus | velezensis | 2.45 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP5.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-29 | Bacillus | velezensis | Bacillus | velezensis | 2.64 | Bacillus | velezensis | 2.28 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP5.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-29 | Bacillus | velezensis | Bacillus | velezensis | 2.46 | Bacillus | velezensis | 2.36 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP6 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-17 | Bacillus | velezensis | Bacillus | velezensis | 2.46 | Bacillus | velezensis | 2.38 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP7.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-02 | Bacillus | velezensis | Bacillus | velezensis | 2.36 | Bacillus | velezensis | 2.23 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP7.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-02 | Bacillus | velezensis | Bacillus | velezensis | 2.43 | Bacillus | velezensis | 2.35 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP7.1C | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-03 | Bacillus | velezensis | Bacillus | velezensis | 2.7 | Bacillus | velezensis | 2.37 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP7.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-03 | Bacillus | velezensis | Bacillus | velezensis | 2.37 | Bacillus | velezensis | 2.18 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CP7.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-18 | Bacillus | velezensis | Bacillus | velezensis | 1.86 | Bacillus | velezensis | 1.46 | B | 0 | 0 | 1 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 35818 | 3; sequenced | eDT | 2023-11-02 | Bacillus | velezensis | Bacillus | velezensis | 2.26 | Bacillus | velezensis | 1.96 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DJFZ40 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-03-07 | Bacillus | velezensis | Bacillus | velezensis | 2.35 | Bacillus | velezensis | 2.34 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DL2.3A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.54 | Bacillus | velezensis | 2.36 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DL2.3B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus | velezensis | Bacillus | velezensis | 2.56 | Bacillus | velezensis | 2.36 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DP1.3B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-18 | Bacillus | velezensis | Bacillus | velezensis | 2.68 | Bacillus | velezensis | 2.07 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DP2.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-08 | Bacillus | velezensis | Bacillus | velezensis | 2.82 | Bacillus | velezensis | 2.52 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23117 | 3; public strain collection | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2015-03-10 | Bacillus | velezensis | Bacillus | velezensis | 2.83 | Bacillus | velezensis | 2.22 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28326 | 3; public strain collection | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-05-24 | Bacillus | velezensis | Bacillus | velezensis | 2.75 | Bacillus | velezensis | 2.17 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EG4.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus | velezensis | Bacillus | velezensis | 2.7 | Bacillus | velezensis | 2.4 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EG4.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-16 | Bacillus | velezensis | Bacillus | velezensis | 2.78 | Bacillus | velezensis | 2.38 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EG5.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus | velezensis | Bacillus | velezensis | 2.29 | Bacillus | velezensis | 2.26 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EG5.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus | velezensis | Bacillus | velezensis | 2.12 | Bacillus | velezensis | 2.05 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ01-02 (01) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-08-31 | Bacillus | velezensis | Bacillus | velezensis | 2.65 | Bacillus | velezensis | 2.55 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ05-04 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-08-31 | Bacillus | velezensis | Bacillus | velezensis | 2.7 | Bacillus | velezensis | 2.54 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ10-02 (02) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-08-31 | Bacillus | velezensis | Bacillus | velezensis | 2.65 | Bacillus | velezensis | 2.46 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ15-07 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-09-21 | Bacillus | velezensis | Bacillus | velezensis | 2.67 | Bacillus | velezensis | 2.5 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | EZ15-09 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-09-21 | Bacillus | velezensis | Bacillus | velezensis | 2.53 | Bacillus | velezensis | 2.52 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | FZB24 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2014-10-15 | Bacillus | velezensis | Bacillus | velezensis | 2.57 | Bacillus | velezensis | 2.46 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | FZB45 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2017-07-14 | Bacillus | velezensis | Bacillus | velezensis | 2.48 | Bacillus | velezensis | 2.3 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GBSW11 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2018-03-07 | Bacillus | velezensis | Bacillus | velezensis | 2.62 | Bacillus | velezensis | 2.47 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | GS3.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-16 | Bacillus | velezensis | Bacillus | velezensis | 2.41 | Bacillus | velezensis | 2.33 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HB1.4A | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2019-11-27 | Bacillus | velezensis | Bacillus | velezensis | 2.72 | Bacillus | velezensis | 2.27 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD2.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2019-12-18 | Bacillus | velezensis | Bacillus | velezensis | 2.48 | Bacillus | velezensis | 2.22 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD2.2 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2019-12-18 | Bacillus | velezensis | Bacillus | velezensis | 2.14 | Bacillus | velezensis | 1.97 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD2.3 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-08 | Bacillus | velezensis | Bacillus | velezensis | 2.75 | Bacillus | velezensis | 2.28 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus | velezensis | Bacillus | velezensis | 2.81 | Bacillus | velezensis | 2.26 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus | velezensis | Bacillus | velezensis | 2.76 | Bacillus | velezensis | 2.49 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus | velezensis | Bacillus | velezensis | 2.84 | Bacillus | velezensis | 2.27 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus | velezensis | Bacillus | velezensis | 2.64 | Bacillus | velezensis | 2.25 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.3B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus | velezensis | Bacillus | velezensis | 2.81 | Bacillus | velezensis | 2.48 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.3C | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus | velezensis | Bacillus | velezensis | 2.76 | Bacillus | velezensis | 2.44 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD5.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus | velezensis | Bacillus | velezensis | 2.78 | Bacillus | velezensis | 2.43 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD5.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus | velezensis | Bacillus | velezensis | 2.81 | Bacillus | velezensis | 2.27 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD5.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus | velezensis | Bacillus | velezensis | 2.77 | Bacillus | velezensis | 2.39 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD5 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-11-21 | Bacillus | velezensis | Bacillus | velezensis | 2.8 | Bacillus | velezensis | 2.38 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JS1.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus | velezensis | Bacillus | velezensis | 2.75 | Bacillus | velezensis | 2.32 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KKD2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-01-23 | Bacillus | velezensis | Bacillus | velezensis | 2.61 | Bacillus | velezensis | 2.56 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KT1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2021-10-14 | Bacillus | velezensis | Bacillus | velezensis | 2.58 | Bacillus | velezensis | 2.42 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LLCG43 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2018-08-17 | Bacillus | velezensis | Bacillus | velezensis | 2.68 | Bacillus | velezensis | 2.57 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MM 508 vel RKI | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2007-03-05 | Bacillus | velezensis | Bacillus | velezensis | 2.1 | Bacillus | velezensis | 2.04 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MR2.1B | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-17 | Bacillus | velezensis | Bacillus | velezensis | 2.34 | Bacillus | velezensis | 2.3 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MS2.3A | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-17 | Bacillus | velezensis | Bacillus | velezensis | 2.48 | Bacillus | velezensis | 2.23 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MS2.3B | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-17 | Bacillus | velezensis | Bacillus | velezensis | 2.71 | Bacillus | velezensis | 2.42 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | OL1.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-13 | Bacillus | velezensis | Bacillus | velezensis | 2.56 | Bacillus | velezensis | 2.46 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | OL2.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-14 | Bacillus | velezensis | Bacillus | velezensis | 2.78 | Bacillus | velezensis | 2.42 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | OR2.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-13 | Bacillus | velezensis | Bacillus | velezensis | 2.64 | Bacillus | velezensis | 2.26 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | OR2.2 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-14 | Bacillus | velezensis | Bacillus | velezensis | 2.64 | Bacillus | velezensis | 2.35 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RJGP41 (02) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2017-08-01 | Bacillus | velezensis | Bacillus | velezensis | 2.45 | Bacillus | velezensis | 2.3 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RJGP41 (03) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-01-11 | Bacillus | velezensis | Bacillus | velezensis | 2.71 | Bacillus | velezensis | 2.52 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RJGP41 (04) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2017-08-01 | Bacillus | velezensis | Bacillus | velezensis | 2.73 | Bacillus | velezensis | 2.58 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RS1.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus | velezensis | Bacillus | velezensis | 2.62 | Bacillus | velezensis | 2.29 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RS1.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus | velezensis | Bacillus | velezensis | 2.37 | Bacillus | velezensis | 2.13 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SC3.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus | velezensis | Bacillus | velezensis | 2.56 | Bacillus | velezensis | 2.26 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | SC3.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus | velezensis | Bacillus | velezensis | 2.67 | Bacillus | velezensis | 2.26 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | TK2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2021-10-14 | Bacillus | velezensis | Bacillus | velezensis | 2.49 | Bacillus | velezensis | 2.44 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | VM10 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-11-14 | Bacillus | velezensis | Bacillus | velezensis | 2.61 | Bacillus | velezensis | 2.45 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | VM49 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-11-21 | Bacillus | velezensis | Bacillus | velezensis | 2.58 | Bacillus | velezensis | 2.43 | A | 1 | 1 | 0 | Bacillus velezensis | Bacillus velezensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29883 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2018-05-24 | Bacillus | xiamenensis | Bacillus | xiamenensis | 2.79 | Bacillus | safensis | 2.1 | B | 0 | 0 | 1 | Bacillus xiamenensis | Bacillus safensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LLTC93 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-07-18 | Bacillus | zhangzhouensis | Bacillus | zhangzhouensis | 2.76 | Bacillus | altitudinis | 2.01 | B | 0 | 0 | 1 | Bacillus zhangzhouensis | Bacillus altitudinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Cytobacillus | firmus | Cytobacillus | firmus | 2.45 | Cytobacillus | firmus | 2.24 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 48 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-01-07 | Cytobacillus | firmus | Cytobacillus | firmus | 2.33 | Cytobacillus | firmus | 2.2 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCEB 627 DSM 12 | 3; type strain | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.19 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 37 B 88 | 3; reference strain collection | 0 | 2015-02-09 | Cytobacillus | firmus | Cytobacillus | firmus | 2.06 | Cytobacillus | firmus | 1.98 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 37 | 3; reference strain collection | DT | 2014-01-22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.33 | Cytobacillus | firmus | 2.13 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 359 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-09 | Cytobacillus | firmus | Cytobacillus | firmus | 2.02 | Cytobacillus | firmus | 1.78 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-15 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.65 | Cytobacillus | horneckiae | 2.11 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-15 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.46 | Cytobacillus | horneckiae | 2.32 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus | kochii | Cytobacillus | kochii | 2.33 | Cytobacillus | kochii | 1.86 | A | 1 | 1 | 0 | Cytobacillus kochii | Cytobacillus kochii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-02-19 | Domibacillus | aminovorans | Domibacillus | aminovorans | 2.81 | Pseudomonas | aeruginosa | 1.37 | A | 1 | 1 | 0 | Domibacillus aminovorans | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-02-15 | Domibacillus | aminovorans | Domibacillus | aminovorans | 2.63 | Staphylococcus | lugdunensis | 1.46 | A | 1 | 1 | 0 | Domibacillus aminovorans | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-19 | Evansella | clarkii | Evansella | clarkii | 2.81 | Evansella | clarkii | 2.35 | A | 1 | 1 | 0 | Evansella clarkii | Evansella clarkii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Evansella | clarkii | Evansella | clarkii | 2.77 | Evansella | clarkii | 2.29 | A | 1 | 1 | 0 | Evansella clarkii | Evansella clarkii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus | stearothermophilus | Geobacillus | stearothermophilus | 2.39 | Geobacillus | stearothermophilus | 2.29 | A | 1 | 1 | 0 | Geobacillus stearothermophilus | Geobacillus stearothermophilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-25 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.39 | Halalkalibacterium | halodurans | 2.19 | C | 0 | 0 | 1 | Alkalihalobacillus halodurans | Halalkalibacterium halodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6939 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.67 | Alkalihalobacillus | okuhidensis | 2.45 | B | 0 | 0 | 1 | Alkalihalobacillus halodurans | Alkalihalobacillus okuhidensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.74 | Alkalihalobacillus | okuhidensis | 2.37 | B | 0 | 0 | 1 | Alkalihalobacillus halodurans | Alkalihalobacillus okuhidensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Heyndrickxia | coagulans | Weizmannia | coagulans | 2.45 | Weizmannia | coagulans | 2.12 | A | 0 | 0 | 1 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2308 | 2; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Weizmannia coagulans DSM 2308: Preparation & measurement: N. Fischer#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-22 | Heyndrickxia | coagulans | Weizmannia | coagulans | 2.32 | Weizmannia | coagulans | 2.04 | A | 0 | 0 | 1 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2311 | 2; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Weizmannia coagulans DSM 2308: Preparation & measurement: N. Fischer#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Heyndrickxia | coagulans | Weizmannia | coagulans | 2.29 | Weizmannia | coagulans | 1.76 | A | 0 | 0 | 1 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30760 | 2; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Weizmannia coagulans DSM 2308: Preparation & measurement: N. Fischer#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-02 | Heyndrickxia | coagulans | Weizmannia | coagulans | 2.15 | Weizmannia | coagulans | 2.02 | A | 0 | 0 | 1 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia | sporothermodurans | Bacillus | sporothermodurans | 2.17 | Bacillus | sporothermodurans | 2.14 | A | 0 | 0 | 1 | Bacillus sporothermodurans | Bacillus sporothermodurans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 106 | 3; sequenced | DT | 2015-07-22 | Heyndrickxia | sporothermodurans | Bacillus | sporothermodurans | 2.22 | Bacillus | sporothermodurans | 2.2 | A | 0 | 0 | 1 | Bacillus sporothermodurans | Bacillus sporothermodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 10599 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Heyndrickxia | sporothermodurans | Heyndrickxia | sporothermodurans | 2.79 | Heyndrickxia | sporothermodurans | 2.35 | A | 1 | 1 | 0 | Heyndrickxia sporothermodurans | Heyndrickxia sporothermodurans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-20 | Lederbergia | lenta | Lederbergia | lenta | 2.43 | Lederbergia | lenta | 2.33 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9 CVUAS | 3; type strain | DT | 2014-01-22 | Lederbergia | lenta | Lederbergia | lenta | 2.22 | Lederbergia | lenta | 2.08 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-13 | Lederbergia | lenta | Pichia | fermentans | 1.51 | Lactobacillus | crispatus | 1.39 | B | 0 | 0 | 1 | Pichia fermentans | Lactobacillus crispatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus | bataviensis | Neobacillus | bataviensis | 2.76 | Neobacillus | bataviensis | 2.29 | A | 1 | 1 | 0 | Neobacillus bataviensis | Neobacillus bataviensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-20 | Neobacillus | cucumis | Neobacillus | cucumis | 2.74 | Neobacillus | bataviensis | 2.02 | B | 0 | 0 | 1 | Neobacillus cucumis | Neobacillus bataviensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 216 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Niallia | circulans | Niallia | circulans | 2.13 | Niallia | circulans | 1.94 | A | 1 | 1 | 0 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-10 | Niallia | circulans | Niallia | circulans | 1.98 | Niallia | circulans | 1.88 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 11 | 3; type strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-04-13 | Niallia | circulans | Niallia | circulans | 2.05 | Niallia | circulans | 1.61 | A | 1 | 1 | 0 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-03-05 | Niallia | circulans | Niallia | circulans | 1.9 | Niallia | circulans | 1.65 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2582.2 | 3; sequenced | eDT | 2015-09-23 | Niallia | nealsonii | Niallia | nealsonii | 2.34 | Niallia | taxi | 2.04 | B | 0 | 0 | 1 | Niallia nealsonii | Niallia taxi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33253 | 2; sequenced | EFex | 2023-01-12 | Niallia | sp | Niallia | nealsonii | 2.62 | Niallia | taxi | 2.1 | B | 0 | 0 | 1 | Niallia nealsonii | Niallia taxi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-10 | Peribacillus | muralis | Peribacillus | muralis | 2.38 | Peribacillus | simplex | 2.09 | B | 0 | 0 | 1 | Peribacillus muralis | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-18 | Peribacillus | psychrosaccharolyticus | Peribacillus | psychrosaccharolyticus | 2.21 | Staphylococcus | simiae | 1.44 | A | 1 | 1 | 0 | Peribacillus psychrosaccharolyticus | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus | simplex | Peribacillus | simplex | 2.08 | Peribacillus | simplex | 1.74 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-09-03 | Peribacillus | simplex | Peribacillus | simplex | 2.57 | Peribacillus | simplex | 1.89 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-12-15 | Peribacillus | simplex | Peribacillus | simplex | 2.68 | Peribacillus | simplex | 2.19 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 110302 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2021-11-17 | Priestia | endophytica | Priestia | endophytica | 2.3 | Priestia | endophytica | 2.08 | A | 1 | 1 | 0 | Priestia endophytica | Priestia endophytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13796 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2021-11-17 | Priestia | endophytica | Priestia | endophytica | 2.26 | Priestia | endophytica | 2.01 | A | 1 | 1 | 0 | Priestia endophytica | Priestia endophytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1316 | 3; public strain collection | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from DSMZ strain collection#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Priestia | flexa | Priestia | flexa | 2.32 | Priestia | flexa | 2.09 | A | 1 | 1 | 0 | Priestia flexa | Priestia flexa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-18 | Priestia | flexa | Priestia | flexa | 2.54 | Priestia | flexa | 2.17 | A | 1 | 1 | 0 | Priestia flexa | Priestia flexa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Priestia | flexa | Priestia | megaterium | 2.01 | Priestia | megaterium | 1.95 | A | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30607 | 3; public strain collection | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from DSMZ strain collection#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-22 | Priestia | flexa | Priestia | flexa | 2.24 | Priestia | flexa | 1.96 | A | 1 | 1 | 0 | Priestia flexa | Priestia flexa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 14581 | 3; type strain | DT | 2014-01-29 | Priestia | megaterium | Priestia | megaterium | 2.32 | Bacillus | arybhattai | 2.28 | C | 0 | 0 | 1 | Priestia megaterium | Bacillus arybhattai | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-26 | Priestia | megaterium | Priestia | megaterium | 1.94 | Priestia | megaterium | 1.92 | B | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 2 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-01 | Priestia | megaterium | Priestia | megaterium | 2.33 | Priestia | megaterium | 2.03 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 333 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-11 | Priestia | megaterium | Priestia | megaterium | 2.42 | Priestia | megaterium | 1.95 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia | megaterium | Priestia | megaterium | 2.06 | Priestia | megaterium | 2.06 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 32 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Priestia | megaterium | Priestia | megaterium | 2.82 | Priestia | megaterium | 2.08 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 333 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Priestia | megaterium | Priestia | megaterium | 2.31 | Priestia | megaterium | 2.05 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 509 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Priestia | megaterium | Priestia | megaterium | 2.08 | Bacillus | arybhattai | 1.9 | C | 0 | 0 | 1 | Priestia megaterium | Bacillus arybhattai | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 90 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-13 | Priestia | megaterium | Priestia | megaterium | 2.04 | Priestia | megaterium | 1.89 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_01 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-11 | Priestia | megaterium | Priestia | megaterium | 2.44 | Priestia | megaterium | 2.08 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Salibacterium | salarium | Salibacterium | salarium | 2.05 | Aeromonas | salmonicida | 1.35 | A | 1 | 1 | 0 | Salibacterium salarium | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6307 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.56 | Sutcliffiella | cohnii | 2.24 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.21 | Sutcliffiella | cohnii | 2.13 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-25 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.8 | Sutcliffiella | cohnii | 2.03 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus | saccharophilus | Terribacillus | saccharophilus | 2.4 | Lodderomyces | elongisporus | 1.45 | A | 1 | 1 | 0 | Terribacillus saccharophilus | Lodderomyces elongisporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Virgibacillus | pantothenticus | Virgibacillus | panthotenticus | 2.82 | Virgibacillus | pantothenticus | 2.01 | B | 0 | 0 | 1 | Virgibacillus panthotenticus | Virgibacillus pantothenticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus | proomii | Virgibacillus | proomii | 2.02 | Virgibacillus | proomii | 1.99 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3616.2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus | proomii | Virgibacillus | proomii | 2.2 | Virgibacillus | proomii | 1.83 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 71.2 | 3; sequenced | DT | 2013-12-13 | Weizmannia | ginsengihumi | Weizmannia | ginsengihumi | 2.4 | Weizmannia | ginsengihumi | 2.19 | A | 1 | 1 | 0 | Weizmannia ginsengihumi | Weizmannia ginsengihumi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | halophilus | Jeotgalicoccus | halophilus | 2.45 | Listeria | innocua | 1.68 | A | 1 | 1 | 0 | Jeotgalicoccus halophilus | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | huakuii | Jeotgalicoccus | huakuii | 2.6 | Macrococcus | canis | 1.37 | A | 1 | 1 | 0 | Jeotgalicoccus huakuii | Macrococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | psychrophilus | Jeotgalicoccus | psychrophilus | 2.37 | Jeotgalicoccus | psychrophilus | 1.84 | A | 1 | 1 | 0 | Jeotgalicoccus psychrophilus | Jeotgalicoccus psychrophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 2013-03-14 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.08 | Macrococcus | caseolyticus | 1.75 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.19 | Macrococcus | caseolyticus | 1.87 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4448 | 3; 16S rDNA-sequencing | DT | 2013-03-14 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.41 | Macrococcus | caseolyticus | 1.87 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4454 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.19 | Macrococcus | caseolyticus | 1.76 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4456 | 3; 16S rDNA-sequencing | DT | 2013-03-14 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.35 | Macrococcus | caseolyticus | 1.67 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4503 | 3; sequenced | DT | 2013-03-14 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.42 | Macrococcus | caseolyticus | 1.81 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4506 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.25 | Macrococcus | caseolyticus | 1.89 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4510 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.03 | Macrococcus | caseolyticus | 1.58 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4513 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.25 | Macrococcus | caseolyticus | 1.59 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4636 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.33 | Macrococcus | caseolyticus | 1.93 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus | carouselicus | Macrococcus | carouselicus | 2.59 | Staphylococcus | simulans | 1.4 | A | 1 | 1 | 0 | Macrococcus carouselicus | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus | lentus | Mammaliicoccus | lentus | 2.4 | Mammaliicoccus | lentus | 1.8 | A | 1 | 1 | 0 | Mammaliicoccus lentus | Mammaliicoccus lentus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 4363 | 2; API Staph ID32 | DT | 2013-04-30 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 1.89 | Mammaliicoccus | sciuri | 1.72 | B | 0 | 0 | 1 | Mammaliicoccus sciuri | Mammaliicoccus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | A 4778 | 2; API Staph ID32 | DT | 2013-04-30 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 1.97 | Mammaliicoccus | sciuri | 1.84 | B | 0 | 0 | 1 | Mammaliicoccus sciuri | Mammaliicoccus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | coa145 | 2; API Staph ID32 | DT | 2013-04-30 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 1.89 | Mammaliicoccus | sciuri | 1.72 | B | 0 | 0 | 1 | Mammaliicoccus sciuri | Mammaliicoccus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3442 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 2.23 | Mammaliicoccus | sciuri | 2.15 | A | 1 | 1 | 0 | Mammaliicoccus sciuri | Mammaliicoccus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 2.76 | Mammaliicoccus | sciuri | 2.37 | A | 1 | 1 | 0 | Mammaliicoccus sciuri | Mammaliicoccus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus | sciuri | Mammaliicoccus | sciuri | 2.13 | Mammaliicoccus | sciuri | 2.12 | A | 1 | 1 | 0 | Mammaliicoccus sciuri | Mammaliicoccus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9937.2 DT | 3; 16S rDNA-sequencing | DT | 2015-12-01 | Mammaliicoccus | vitulinus | Mammaliicoccus | vitulinus | 2.26 | Mammaliicoccus | vitulinus | 1.77 | A | 1 | 1 | 0 | Mammaliicoccus vitulinus | Mammaliicoccus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9937.2 extr | 3; 16S rDNA-sequencing | EFex | 2016-04-06 | Mammaliicoccus | vitulinus | Mammaliicoccus | vitulinus | 2.31 | Mammaliicoccus | vitulinus | 1.87 | A | 1 | 1 | 0 | Mammaliicoccus vitulinus | Mammaliicoccus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 343.2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus | vitulinus | Mammaliicoccus | vitulinus | 2.23 | Mammaliicoccus | vitulinus | 2.23 | A | 1 | 1 | 0 | Mammaliicoccus vitulinus | Mammaliicoccus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5687 sir | 3; sequenced | EFex | 2025-07-08 | Staphylococcus | agnetis | Staphylococcus | agnetis | 2.62 | Staphylococcus | agnetis | 2.42 | A | 1 | 1 | 0 | Staphylococcus agnetis | Staphylococcus agnetis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5687 LT | 3; sequenced | DT | 2013-04-16 | Staphylococcus | agnetis | Staphylococcus | agnetis | 2.35 | Staphylococcus | hyicus | 2.25 | B | 0 | 0 | 1 | Staphylococcus agnetis | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 622.3 | 2; biochemical | EFex | 2025-07-08 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.3 | Staphylococcus | hyicus | 2.3 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 23656 LT | 3; type strain [0003] | DT | 2013-04-30 | Staphylococcus | agnetis | Staphylococcus | hyicus | 1.98 | Staphylococcus | hyicus | 1.62 | B | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 23656 sir | 3; type strain [0003] | EFex | 2025-07-08 | Staphylococcus | agnetis | Staphylococcus | agnetis | 2.78 | Staphylococcus | agnetis | 2.59 | A | 1 | 1 | 0 | Staphylococcus agnetis | Staphylococcus agnetis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 7385 sir | 3; sequenced | EFex | 2025-07-08 | Staphylococcus | agnetis | Staphylococcus | agnetis | 2.41 | Staphylococcus | agnetis | 2.32 | A | 1 | 1 | 0 | Staphylococcus agnetis | Staphylococcus agnetis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TUM 7385 LT | 3; sequenced | DT | 02.10.2013 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.24 | Staphylococcus | hyicus | 2.22 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.37 | Staphylococcus | argenteus | 2.28 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.39 | Staphylococcus | argenteus | 2.38 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.01 | Staphylococcus | arlettae | 1.72 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 265U3-3 | 2; sequenced | DT | 2016-08-31 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.02 | Staphylococcus | arlettae | 1.71 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa031 | 2; API Staph ID32; 66342360 | DT | 2013-03-20 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.43 | Nocardia | exalbida-gamkensis | 1.49 | A | 1 | 1 | 0 | Staphylococcus arlettae | Nocardia exalbida-gamkensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.13 | Staphylococcus | arlettae | 1.63 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20672 | 3; type strain [0003] | DT | 04.10.2013 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.21 | Staphylococcus | arlettae | 1.97 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KNS 35 | 2; API Staph ID32: 66340261 | DT | 2013-02-13 | Staphylococcus | arlettae | Staphylococcus | arlettae | 2.16 | Staphylococcus | arlettae | 1.58 | A | 1 | 1 | 0 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TUM MG88 | 3; sequenced | DT | 02.10.2013 | Staphylococcus | arlettae | Staphylococcus | arlettae | 1.95 | Staphylococcus | arlettae | 1.7 | B | 0 | 0 | 1 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AGES 001 | 3; 0 | EFEx | 2019-03-20 | Staphylococcus | aureus | Staphylococcus | aureus | 2.68 | Staphylococcus | aureus | 2.21 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20231 CASO | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-02-25 | Staphylococcus | aureus | Staphylococcus | aureus | 2.46 | Staphylococcus | aureus | 2.45 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 96-01678 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-09-28 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.35 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00981 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.33 | Staphylococcus | aureus | 2.28 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-01243 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.36 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 03-00397 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.34 | Staphylococcus | aureus | 2.32 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-00602 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-01-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.31 | Staphylococcus | aureus | 2.29 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04-03100 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-10-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.35 | Staphylococcus | aureus | 2.27 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-03338 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.33 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 03-00220 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.37 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 06-00896 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.49 | Staphylococcus | aureus | 2.44 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-03347 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.37 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00186 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.43 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00416 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-17 | Staphylococcus | aureus | Staphylococcus | aureus | 2.38 | Staphylococcus | aureus | 2.37 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-01140 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02135 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-05-17 | Staphylococcus | aureus | Staphylococcus | aureus | 2.46 | Staphylococcus | aureus | 2.42 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-01104 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.35 | Staphylococcus | aureus | 2.27 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-01238 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.42 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-01949 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-28 | Staphylococcus | aureus | Staphylococcus | aureus | 2.53 | Staphylococcus | aureus | 2.48 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-01531 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.55 | Staphylococcus | aureus | 2.49 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02446 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2022-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.49 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02476 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-28 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.46 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02431 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.51 | Staphylococcus | aureus | 2.51 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02434 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.53 | Staphylococcus | aureus | 2.51 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02444 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.41 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-00401 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-00004 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.36 | Staphylococcus | aureus | 2.29 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-02172 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.34 | Staphylococcus | aureus | 2.19 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-02337 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.31 | Staphylococcus | aureus | 2.3 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-02874 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.27 | Staphylococcus | aureus | 2.24 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-02989 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.36 | Staphylococcus | aureus | 2.32 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-02727 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.27 | Staphylococcus | aureus | 2.25 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00558 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.42 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | N315 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.32 | Staphylococcus | aureus | 2.29 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Mu50 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.56 | Staphylococcus | aureus | 2.51 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-00170 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-02020 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.51 | Staphylococcus | aureus | 2.45 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 06-00603 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.33 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 02-02424 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-17 | Staphylococcus | aureus | Staphylococcus | aureus | 2.38 | Staphylococcus | aureus | 2.32 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 06-00953-1 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.41 | Staphylococcus | aureus | 2.37 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04-02232 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.44 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 06-01243 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 02-03179 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.33 | Staphylococcus | aureus | 2.31 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04-02981 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.49 | Staphylococcus | aureus | 2.42 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-00392 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-05-21 | Staphylococcus | aureus | Staphylococcus | aureus | 2.37 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00825 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-21 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.45 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-00463 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-21 | Staphylococcus | aureus | Staphylococcus | aureus | 2.44 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-00265 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-00010 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.31 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-00250 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.23 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-00993 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-10-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.36 | Staphylococcus | aureus | 2.28 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08-01718 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.29 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00001 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-10-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.36 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-01189 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.42 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02378 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.28 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00615 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-02069 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.35 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-01345 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.38 | Staphylococcus | aureus | 2.37 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00212 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.3 | Staphylococcus | aureus | 2.28 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-01791-1 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.34 | Staphylococcus | aureus | 2.21 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 07-02652 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.48 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20231 TSA | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-02-25 | Staphylococcus | aureus | Staphylococcus | aureus | 2.34 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-00783 | 2; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-04 | Staphylococcus | aureus | Staphylococcus | aureus | 2.45 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-00902 | 2; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.48 | Staphylococcus | aureus | 2.47 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-00961 | 2; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-04 | Staphylococcus | aureus | Staphylococcus | aureus | 2.56 | Staphylococcus | aureus | 2.21 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-01100 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.56 | Staphylococcus | aureus | 2.49 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2017-4101-008269 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.48 | Staphylococcus | aureus | 2.47 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2018-4101-001305-003 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.58 | Staphylococcus | aureus | 2.5 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2018-4101-001376 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.35 | Staphylococcus | aureus | 2.32 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2018-4101-002922-005 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.35 | Staphylococcus | aureus | 2.32 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2018-4101-003040 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.51 | Staphylococcus | aureus | 2.49 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2018-4104-000006 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.47 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 3321-2017-000527 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.43 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 4201-2017-002411 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.54 | Staphylococcus | aureus | 2.52 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 4201-2017-003006 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.36 | Staphylococcus | aureus | 2.36 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 4201-2017-004220 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.55 | Staphylococcus | aureus | 2.49 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 4201-2017-004348 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.41 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 4201-2018-001040 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.54 | Staphylococcus | aureus | 2.5 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 4201-2018-001081 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.47 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 4201-2018-001446 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.44 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 29213 CVUAS | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 29213 RKI | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-04 | Staphylococcus | aureus | Staphylococcus | aureus | 2.48 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C3658-06 | 3; sequenced (LGL) | EFex | 2015-04-15 | Staphylococcus | aureus | Staphylococcus | aureus | 2.23 | Staphylococcus | aureus | 2.22 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | COL | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.35 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10197.2 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.37 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10507 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus | aureus | Staphylococcus | aureus | 2.48 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1083 | 4; sequenztype | DT | 05.02.2013 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1304.2 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.42 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1305.2 | 4; sequenztype | DT | 2014-08-20 | Staphylococcus | aureus | Staphylococcus | aureus | 2.05 | Staphylococcus | aureus | 2.04 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1894 | 4; sequenztype | DT | 2013-02-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.38 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2645 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus | aureus | Staphylococcus | aureus | 2.3 | Staphylococcus | aureus | 2.28 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3072 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus | aureus | Staphylococcus | aureus | 2.37 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3194 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus | aureus | Staphylococcus | aureus | 2.46 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3195 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus | aureus | Staphylococcus | aureus | 2.59 | Staphylococcus | aureus | 2.54 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3196 | 4; sequenztype | DT | 13.01.2015 | Staphylococcus | aureus | Staphylococcus | aureus | 2.48 | Staphylococcus | aureus | 2.45 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3199 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus | aureus | Staphylococcus | aureus | 2.44 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3204 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus | aureus | Staphylococcus | aureus | 2.51 | Staphylococcus | aureus | 2.51 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3206 | 4; sequenztype | DT | 31.10.2012 | Staphylococcus | aureus | Staphylococcus | aureus | 2.58 | Staphylococcus | aureus | 2.49 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3260 | 4; sequenztype | DT | 31.10.2012 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.47 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3320 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus | aureus | Staphylococcus | aureus | 2.41 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3330 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus | aureus | Staphylococcus | aureus | 2.49 | Staphylococcus | aureus | 2.46 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3433 | 4; sequenztype | DT | 31.10.2012 | Staphylococcus | aureus | Staphylococcus | aureus | 2.45 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3452.2 | 4; sequenztype | DT | 10.09.2013 | Staphylococcus | aureus | Staphylococcus | aureus | 2.31 | Staphylococcus | aureus | 2.3 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3740.2 | 4; sequenztype | DT | 2014-08-20 | Staphylococcus | aureus | Staphylococcus | aureus | 2.28 | Staphylococcus | aureus | 2.26 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3771 | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.43 | Staphylococcus | aureus | 2.42 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3776 | 3; 16S rDNA-sequencing | DT | 23.10.2012 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.45 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4645 | 3; 16S rDNA-sequencing | DT | 2013-02-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.46 | Staphylococcus | aureus | 2.44 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5756 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus | aureus | Staphylococcus | aureus | 2.41 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6080 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus | aureus | Staphylococcus | aureus | 2.48 | Staphylococcus | aureus | 2.45 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6459 | 3; 16S rDNA-sequencing | DT | 2014-05-14 | Staphylococcus | aureus | Staphylococcus | aureus | 2.35 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6963 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.3 | Staphylococcus | aureus | 2.27 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7112 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus | aureus | Staphylococcus | aureus | 2.34 | Staphylococcus | aureus | 2.27 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7665 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus | aureus | Staphylococcus | aureus | 2.46 | Staphylococcus | aureus | 2.42 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7779 | 3; 16S rDNA-sequencing | DT | 2012-09-12 | Staphylococcus | aureus | Staphylococcus | aureus | 2.42 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7958.2 | 3; sequenztype | DT | 2013-04-30 | Staphylococcus | aureus | Staphylococcus | aureus | 2.18 | Staphylococcus | aureus | 2.08 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7963.2 | 3; sequenztype | DT | 2013-04-30 | Staphylococcus | aureus | Staphylococcus | aureus | 2.33 | Staphylococcus | aureus | 2.25 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8134 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.36 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8135 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus | aureus | Staphylococcus | aureus | 2.55 | Staphylococcus | aureus | 2.54 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8312 | 4; sequenztype | DT | 2012-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.54 | Staphylococcus | aureus | 2.49 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8900 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.43 | Staphylococcus | aureus | 2.41 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9019 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus | aureus | Staphylococcus | aureus | 2.42 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9197 | 4; sequenztype | DT | 2013-02-13 | Staphylococcus | aureus | Staphylococcus | aureus | 2.49 | Staphylococcus | aureus | 2.46 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 935 | 3; 16S rDNA-sequencing | DT | 2014-05-19 | Staphylococcus | aureus | Staphylococcus | aureus | 2.38 | Staphylococcus | aureus | 2.33 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9373 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus | aureus | Staphylococcus | aureus | 2.48 | Staphylococcus | aureus | 2.44 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9388 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.44 | Staphylococcus | aureus | 2.42 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9642 | 2; biochemical tube tests | DT | 2012-12-04 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9681 | 4; Sequencetyp (NRL-Staph, BfR) | DT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.15 | Staphylococcus | aureus | 2.13 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9682 | 4; Sequencetyp (NRL-Staph, BfR) | DT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.3 | Staphylococcus | aureus | 2.28 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9704 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus | aureus | Staphylococcus | aureus | 2.5 | Staphylococcus | aureus | 2.49 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9790 | 4; Sequencetyp (NRL-Staph, BfR) | eDT | 2015-11-05 | Staphylococcus | aureus | Staphylococcus | aureus | 2.43 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9914.2 sir | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9914.2 LT | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.51 | Staphylococcus | aureus | 2.44 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1104 | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.23 | Staphylococcus | aureus | 2.22 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 11729 | 3; public strain collection | DT | 2013-05-14 | Staphylococcus | aureus | Staphylococcus | aureus | 2.24 | Staphylococcus | aureus | 2.24 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12463 | 3; public strain collection | DT | 2019-11-21 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.44 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20491 | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.32 | Staphylococcus | aureus | 2.29 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 346 | 3; public strain collection | eDT | 2013-03-12 | Staphylococcus | aureus | Staphylococcus | aureus | 2.42 | Staphylococcus | aureus | 2.39 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4910 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-05-24 | Staphylococcus | aureus | Staphylococcus | aureus | 2.55 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MSSA 129 | 4; 16S rDNA-sequencing | DT | 2013-03-12 | Staphylococcus | aureus | Staphylococcus | aureus | 2.27 | Staphylococcus | aureus | 2.17 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RKI WG PS 29 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-04-09 | Staphylococcus | aureus | Staphylococcus | aureus | 2.43 | Staphylococcus | aureus | 2.34 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RKI WG PS 42D | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-02-20 | Staphylococcus | aureus | Staphylococcus | aureus | 2.59 | Staphylococcus | aureus | 2.48 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RKI WG PS 94 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-02-20 | Staphylococcus | aureus | Staphylococcus | aureus | 2.4 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RKI WG PS 96 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-04-02 | Staphylococcus | aureus | Staphylococcus | aureus | 2.41 | Staphylococcus | aureus | 2.4 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI4 | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.46 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S6C | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus | aureus | Staphylococcus | aureus | 2.45 | Staphylococcus | aureus | 2.33 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WDCM 00131 | 3; public strain collection | DT | 2018-06-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.44 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus | auricularis | Staphylococcus | auricularis | 2.14 | Staphylococcus | auricularis | 1.85 | A | 1 | 1 | 0 | Staphylococcus auricularis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2023000185 | 2; corresponds to expected result of PT | EFEx | 2023-09-07 | Staphylococcus | auricularis | Staphylococcus | auricularis | 2.5 | Staphylococcus | auricularis | 2.46 | A | 1 | 1 | 0 | Staphylococcus auricularis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6717 | 3; type strain [0003] | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0 | 2022-09-01 | Staphylococcus | capitis | Staphylococcus | capitis | 2.29 | Staphylococcus | capitis | 2.22 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20326 | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.3 | Staphylococcus | capitis | 2.25 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.25 | Staphylococcus | capitis | 2.2 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 21-1 | 2; confirmed proficency test | eDT | 2022-02-14 | Staphylococcus | caprae | Staphylococcus | caprae | 2.41 | Staphylococcus | caprae | 2.4 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20608 RWTH | 3; type strain [0003] | EFEx | 2023-06-20 | Staphylococcus | caprae | Staphylococcus | caprae | 2.33 | Staphylococcus | caprae | 2.14 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20608 CVUAS | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.33 | Staphylococcus | caprae | 2.17 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.38 | Staphylococcus | caprae | 2.36 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.18 | Staphylococcus | carnosus | 2.08 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.32 | Staphylococcus | carnosus | 2.24 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.42 | Staphylococcus | chromogenes | 2.3 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa009a sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.23 | Staphylococcus | chromogenes | 2.08 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa009a LT | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.37 | Staphylococcus | chromogenes | 2.28 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa051 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.26 | Staphylococcus | chromogenes | 2.2 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa051 LT | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.54 | Staphylococcus | chromogenes | 2.47 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa096 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.47 | Staphylococcus | chromogenes | 2.33 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa121 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.33 | Staphylococcus | chromogenes | 2.23 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa143 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.33 | Staphylococcus | chromogenes | 2.29 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3348 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.34 | Staphylococcus | chromogenes | 2.25 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3348 LT | 3; 16S rDNA-sequencing | DT | 2013-02-05 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.42 | Staphylococcus | chromogenes | 2.32 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3349 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.47 | Staphylococcus | chromogenes | 2.45 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3706 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.29 | Staphylococcus | chromogenes | 2.26 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3706 LT | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.34 | Staphylococcus | chromogenes | 2.29 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3784 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.26 | Staphylococcus | chromogenes | 2.23 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3784 LT | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.34 | Staphylococcus | chromogenes | 2.32 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KNS 44 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.53 | Staphylococcus | chromogenes | 2.42 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KNS 60 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.55 | Staphylococcus | chromogenes | 2.45 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KNS 73 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.55 | Staphylococcus | chromogenes | 2.43 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MKNS 02 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.44 | Staphylococcus | chromogenes | 2.37 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MKNS 26 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.18 | Staphylococcus | chromogenes | 2.16 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 084 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.26 | Staphylococcus | chromogenes | 2.22 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 084 LT | 3; 16S rDNA-sequencing | DT | 2013-03-20 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.49 | Staphylococcus | chromogenes | 2.44 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.33 | Staphylococcus | chromogenes | 2.28 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.36 | Staphylococcus | cohnii | 1.99 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.72 | Staphylococcus | cohnii | 2.25 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 217b-3 | 3; 16S rDNA-sequencing | DT | 2016-08-31 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.33 | Staphylococcus | cohnii | 2.32 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MKNS 18 | 2; API Staph ID32 | DT | 2013-03-20 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.55 | Staphylococcus | ureilyticus | 2.4 | B | 0 | 0 | 1 | Staphylococcus cohnii | Staphylococcus ureilyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MKNS 24 | 3; sequenced, biochemical profile | DT | 20.03.2013 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.62 | Staphylococcus | ureilyticus | 2.49 | B | 0 | 0 | 1 | Staphylococcus cohnii | Staphylococcus ureilyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM MG98 | 3; rpoB-sequencing | DT | 2013-10-24 | Staphylococcus | cohnii | Staphylococcus | cohnii | 2.37 | Staphylococcus | cohnii | 2.03 | A | 1 | 1 | 0 | Staphylococcus cohnii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus | delphini | Staphylococcus | delphini | 2.53 | Staphylococcus | delphini | 1.9 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 0 | 0 | x Sta RKI | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-04 | Staphylococcus | delphini | Staphylococcus | delphini | 2.32 | Staphylococcus | delphini | 2.01 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-00903 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2011-08-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.45 | Staphylococcus | epidermidis | 2.34 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 2124 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-05-19 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.42 | Staphylococcus | epidermidis | 2.39 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa076 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.45 | Staphylococcus | epidermidis | 2.44 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3134 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.37 | Staphylococcus | epidermidis | 2.35 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3439 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.39 | Staphylococcus | epidermidis | 2.39 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3441 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.27 | Staphylococcus | epidermidis | 2.25 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3560 | 3; 16S rDNA-sequencing | DT | 2012-12-12 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.31 | Staphylococcus | epidermidis | 2.29 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3701 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.28 | Staphylococcus | epidermidis | 2.24 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.37 | Staphylococcus | epidermidis | 2.31 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3787 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.44 | Staphylococcus | epidermidis | 2.44 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5742 | 2; 0 | eDT | 2015-11-05 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.13 | Staphylococcus | epidermidis | 2.1 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6867 | 2; biochemical tube tests | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.43 | Staphylococcus | epidermidis | 2.39 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1798 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2008-02-20 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.24 | Staphylococcus | epidermidis | 2.14 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.31 | Staphylococcus | epidermidis | 2.19 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK07_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.28 | Staphylococcus | epidermidis | 2.16 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | equorum | Staphylococcus | equorum | 2.16 | Staphylococcus | equorum | 1.8 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa034 | 2; API Staph ID32 | DT | 2013-03-20 | Staphylococcus | equorum | Staphylococcus | equorum | 2.21 | Staphylococcus | equorum | 2.04 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20674 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | equorum | Staphylococcus | equorum | 2.11 | Staphylococcus | equorum | 1.88 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM MG847 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus | equorum | Staphylococcus | equorum | 2.08 | Staphylococcus | equorum | 1.93 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | felis | Staphylococcus | felis | 2.13 | Staphylococcus | felis | 1.96 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | felis | Staphylococcus | felis | 2.14 | Staphylococcus | felis | 2 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A269-3b | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-02 | Staphylococcus | hominis | Staphylococcus | hominis | 2.42 | Staphylococcus | hominis | 2.27 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 789.2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus | hominis | Staphylococcus | hominis | 2.13 | Staphylococcus | hominis | 2.01 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus | hominis | Staphylococcus | hominis | 2.29 | Staphylococcus | hominis | 2.27 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20330 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus | hominis | Staphylococcus | hominis | 2.35 | Staphylococcus | hominis | 2.26 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK01_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-15 | Staphylococcus | hominis | Staphylococcus | hominis | 2.46 | Staphylococcus | hominis | 2.46 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_04 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Staphylococcus | hominis | Staphylococcus | hominis | 2.53 | Staphylococcus | hominis | 2.47 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | coa179 sir | 2; biochemical tests (API Staph ID32) | EFex | 2025-07-08 | Staphylococcus | hyicus | Staphylococcus | agnetis | 2.26 | Staphylococcus | agnetis | 2.22 | A | 0 | 0 | 1 | Staphylococcus agnetis | Staphylococcus agnetis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa179 LT | 2; biochemical tests (API Staph ID32) | DT | 2013-07-02 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.32 | Staphylococcus | hyicus | 2.28 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1346 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.44 | Staphylococcus | hyicus | 2.27 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1347 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.42 | Staphylococcus | hyicus | 2.2 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1525 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.33 | Staphylococcus | hyicus | 2.26 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1526 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.41 | Staphylococcus | hyicus | 2.18 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1552 | 2; biochemical tube tests | DT | 2015-08-26 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.44 | Staphylococcus | hyicus | 2.24 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1553 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.14 | Staphylococcus | hyicus | 2 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1702 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.48 | Staphylococcus | hyicus | 2.25 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1742 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.45 | Staphylococcus | hyicus | 2.26 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1917 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.3 | Staphylococcus | hyicus | 2.07 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1984 sir | 2; biochemical tube tests | EFex | 2025-07-08 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.32 | Staphylococcus | hyicus | 2.29 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1984 LT | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.42 | Staphylococcus | hyicus | 2.18 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 212 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.45 | Staphylococcus | hyicus | 2.33 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2148 | 2; biochemical tube tests | DT | 2013-04-16 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.12 | Staphylococcus | hyicus | 1.96 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2152 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.45 | Staphylococcus | hyicus | 2.35 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.55 | Staphylococcus | hyicus | 2.36 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32680 | 2; biochemical tube tests | EFex | 2025-07-08 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.15 | Staphylococcus | hyicus | 2.03 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 330 | 2; biochemical tube tests | DT | 2013-04-16 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.38 | Staphylococcus | hyicus | 2.19 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3767 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus | hyicus | Staphylococcus | agnetis | 2.72 | Staphylococcus | agnetis | 2.54 | A | 0 | 0 | 1 | Staphylococcus agnetis | Staphylococcus agnetis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3767 LT | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.42 | Staphylococcus | hyicus | 2.4 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 593 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.38 | Staphylococcus | hyicus | 2.23 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.15 | Staphylococcus | intermedius | 2.02 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.07 | Staphylococcus | intermedius | 1.97 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.5 | Staphylococcus | lugdunensis | 2.43 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00272 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.34 | Staphylococcus | lugdunensis | 2.28 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00274 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.28 | Staphylococcus | lugdunensis | 2.26 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00276 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.37 | Staphylococcus | lugdunensis | 2.35 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00283 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-04-04 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.29 | Staphylococcus | lugdunensis | 2.26 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00285 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.46 | Staphylococcus | lugdunensis | 2.42 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00513 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.34 | Staphylococcus | lugdunensis | 2.25 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00515 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.44 | Staphylococcus | lugdunensis | 2.4 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.4 | Staphylococcus | lugdunensis | 2.4 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | microti | Staphylococcus | microti | 2.48 | Staphylococcus | microti | 1.85 | A | 1 | 1 | 0 | Staphylococcus microti | Staphylococcus microti | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | pasteuri | Staphylococcus | warneri | 2.2 | Staphylococcus | warneri | 2.18 | A | 0 | 0 | 1 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3086 | 3; 16S rDNA-sequencing | DT | 2013-04-10 | Staphylococcus | pasteuri | Staphylococcus | warneri | 2.19 | Staphylococcus | warneri | 2.16 | A | 0 | 0 | 1 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus | pasteuri | Staphylococcus | pasteuri | 2.21 | Staphylococcus | pasteuri | 2.02 | A | 1 | 1 | 0 | Staphylococcus pasteuri | Staphylococcus pasteuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MKNS 30 | 3; sequenced | DT | 2013-09-04 | Staphylococcus | pasteuri | Staphylococcus | warneri | 1.97 | Staphylococcus | warneri | 1.92 | B | 0 | 0 | 1 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TUM 5823 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | pasteuri | Staphylococcus | pasteuri | 1.98 | Staphylococcus | pasteuri | 1.94 | B | 0 | 0 | 1 | Staphylococcus pasteuri | Staphylococcus pasteuri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB969149 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.19 | Staphylococcus | pseudintermedius | 2.16 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB969390 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.2 | Staphylococcus | pseudintermedius | 2.11 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB971580 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.34 | Staphylococcus | pseudintermedius | 2.33 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB971904 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.19 | Staphylococcus | pseudintermedius | 2.14 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-0127-2 | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-03-02 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.14 | Staphylococcus | delphini | 2.09 | B | 0 | 0 | 1 | Staphylococcus pseudintermedius | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2375 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.17 | Staphylococcus | pseudintermedius | 2.15 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3688 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.26 | Staphylococcus | pseudintermedius | 2.19 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3831 | 2; biochemical tube tests | DT | 2013-07-05 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.32 | Staphylococcus | pseudintermedius | 2.22 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3856 | 3; 16S rDNA-sequencing (CVUAS) | eDT | 2015-04-21 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.42 | Staphylococcus | pseudintermedius | 2.37 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5457 | 2; biochemical tube tests | DT | 2013-07-05 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.21 | Staphylococcus | pseudintermedius | 2.19 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5594 | 2; biochemical tube tests | DT | 2013-07-05 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.11 | Staphylococcus | pseudintermedius | 2.08 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | pseudintermedius | Staphylococcus | delphini | 2.1 | Staphylococcus | pseudintermedius | 2.06 | B | 0 | 0 | 1 | Staphylococcus delphini | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-03-02 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.76 | Staphylococcus | pseudintermedius | 2.15 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6260.2 eDT | 3; sequenced | eDT | 2023-10-19 | Staphylococcus | ratti | Staphylococcus | ratti | 2.18 | Staphylococcus | hyicus | 1.65 | A | 1 | 1 | 0 | Staphylococcus ratti | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6260.2 extr | 3; sequenced | EFex | 2023-10-19 | Staphylococcus | ratti | Staphylococcus | ratti | 2.59 | Staphylococcus | chromogenes | 1.94 | A | 1 | 1 | 0 | Staphylococcus ratti | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.03 | Staphylococcus | saprophyticus | 1.84 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | A 436 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.4 | Staphylococcus | saprophyticus | 2.4 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | A 446 | 2; biochemical tube tests | DT | 2013-09-11 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.13 | Staphylococcus | saprophyticus | 2.01 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa045 | 2; biochemical tube tests | DT | 2013-03-20 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.33 | Staphylococcus | saprophyticus | 2.31 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 503.2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.28 | Staphylococcus | saprophyticus | 2.23 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus | schleiferi | Staphylococcus | coagulans | 2.31 | Staphylococcus | schleiferi | 2.3 | B | 0 | 0 | 1 | Staphylococcus coagulans | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus | schweitzeri | Staphylococcus | schweitzeri | 2.48 | Staphylococcus | schweitzeri | 2.17 | A | 1 | 1 | 0 | Staphylococcus schweitzeri | Staphylococcus schweitzeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | simulans | Staphylococcus | simulans | 2.38 | Staphylococcus | simulans | 2.33 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4919 | 2; API Staph ID32 | DT | 2013-07-02 | Staphylococcus | simulans | Staphylococcus | simulans | 2.42 | Staphylococcus | simulans | 2.39 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa037 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | simulans | Staphylococcus | simulans | 2.46 | Staphylococcus | simulans | 2.41 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa117 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus | simulans | Staphylococcus | simulans | 2.47 | Staphylococcus | simulans | 2.45 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa126 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | simulans | Staphylococcus | simulans | 2.44 | Staphylococcus | simulans | 2.4 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa150 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | simulans | Staphylococcus | simulans | 2.42 | Staphylococcus | simulans | 2.39 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20322 | 3; type strain | DT | 2013-10-02 | Staphylococcus | simulans | Staphylococcus | simulans | 2.02 | Staphylococcus | simulans | 2.01 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KNS 05 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | simulans | Staphylococcus | simulans | 2.26 | Staphylococcus | simulans | 2.21 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KNS 27 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | simulans | Staphylococcus | simulans | 2.27 | Staphylococcus | simulans | 2.27 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MKNS 04 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | simulans | Staphylococcus | simulans | 2.41 | Staphylococcus | simulans | 2.35 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MKNS 08 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | simulans | Staphylococcus | simulans | 2.26 | Staphylococcus | simulans | 2.16 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | simulans | Staphylococcus | simulans | 2.31 | Staphylococcus | simulans | 2.23 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2439.2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.26 | Staphylococcus | succinus | 2.23 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 761.2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.28 | Staphylococcus | succinus | 2.21 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 4379 | 2; API Staph ID32 | DT | 2013-03-19 | Staphylococcus | warneri | Staphylococcus | pasteuri | 2.08 | Staphylococcus | pasteuri | 1.98 | A | 0 | 0 | 1 | Staphylococcus pasteuri | Staphylococcus pasteuri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4380 | 2; API Staph ID32 | DT | 2013-04-10 | Staphylococcus | warneri | Staphylococcus | warneri | 2.15 | Staphylococcus | warneri | 2.09 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa032 | 2; API Staph ID32 | DT | 2013-04-10 | Staphylococcus | warneri | Staphylococcus | warneri | 2.05 | Staphylococcus | warneri | 2.01 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | coa083 | 3; rpoB-sequencing | DT | 2013-04-10 | Staphylococcus | warneri | Staphylococcus | warneri | 2.11 | Staphylococcus | warneri | 2.08 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus | warneri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 2.13 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 687 | 3; 16S rDNA-sequencing | DT | 2013-04-10 | Staphylococcus | warneri | Staphylococcus | warneri | 2.1 | Staphylococcus | warneri | 2.1 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20316 CVUAS | 3; type strain | DT | 2013-10-02 | Staphylococcus | warneri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 2.03 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20316 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus | warneri | Staphylococcus | warneri | 2.31 | Staphylococcus | warneri | 1.89 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 6412 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | warneri | Staphylococcus | warneri | 2.12 | Staphylococcus | warneri | 2.04 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.04 | Staphylococcus | xylosus | 2.02 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 6124 | 2; biochemical tube tests | EFEx | 2016-09-22 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.35 | Staphylococcus | xylosus | 1.75 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | coa157 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.88 | Staphylococcus | xylosus | 1.85 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3440 | 3; 16S, Xyl pos | DT | 2013-03-26 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.08 | Staphylococcus | xylosus | 2.03 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3562 | 3; 16S, Xyl pos | DT | 2013-03-26 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.91 | Staphylococcus | xylosus | 1.85 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3563 | 3; sequenced | DT | 2012-11-14 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.07 | Staphylococcus | xylosus | 2.04 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3566 | 3; 16S, Xyl pos | DT | 2013-03-26 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.22 | Staphylococcus | xylosus | 2.21 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3692 | 3; 16S, Xyl pos | DT | 2012-10-31 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.19 | Staphylococcus | xylosus | 2.14 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3699 | 3; 16S, Xyl pos | DT | 2013-07-24 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.14 | Staphylococcus | xylosus | 2.06 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 974.2 | 3; 16S, Xyl pos | eDT | 2015-08-25 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.1 | Staphylococcus | xylosus | 2.02 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | FLI-ING 4922 | 3; 16S, Xyl pos | EFEx | 2016-06-01 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.46 | Staphylococcus | xylosus | 1.89 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | KNS 61 | 2; biochemical tube tests | DT | 2013-03-20 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.98 | Staphylococcus | xylosus | 1.92 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MKNS 06 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.18 | Staphylococcus | xylosus | 2.15 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 227 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.1 | Staphylococcus | xylosus | 2.05 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.88 | Staphylococcus | xylosus | 1.76 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus | devriesei | Staphylococcus | devriesei | 2.33 | Staphylococcus | capitis | 1.46 | A | 1 | 1 | 0 | Staphylococcus devriesei | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | bovis | Actinomyces | bovis | 2.05 | Actinomyces | bovis | 1.92 | A | 1 | 1 | 0 | Actinomyces bovis | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces | denticolens | Actinomyces | denticolens | 2.4 | Actinomyces | denticolens | 2.07 | A | 1 | 1 | 0 | Actinomyces denticolens | Actinomyces denticolens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces | sp-CVUAS-31303 | Actinomyces | sp-CVUAS-31303 | 2.6 | Campylobacter | jejuni | 1.26 | A | 1 | 1 | 0 | Actinomyces sp-CVUAS-31303 | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces | viscosus | Actinomyces | viscosus | 2.04 | Actinomyces | viscosus | 1.87 | A | 1 | 1 | 0 | Actinomyces viscosus | Actinomyces viscosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | weissii | Actinomyces | weissii | 2.48 | Actinomyces | weissii | 2.35 | A | 1 | 1 | 0 | Actinomyces weissii | Actinomyces weissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium | bovis | Arcanobacterium | bovis | 2.73 | Arcanobacterium | pluranimalium | 1.77 | A | 1 | 1 | 0 | Arcanobacterium bovis | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium | canis | Arcanobacterium | canis | 2.34 | Arcanobacterium | canis | 2.19 | A | 1 | 1 | 0 | Arcanobacterium canis | Arcanobacterium canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium | haemolyticum | Arcanobacterium | haemolyticum | 2.46 | Arcanobacterium | haemolyticum | 2.3 | A | 1 | 1 | 0 | Arcanobacterium haemolyticum | Arcanobacterium haemolyticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-2056 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.64 | Arcanobacterium | hippocoleae | 2.54 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.32 | Arcanobacterium | hippocoleae | 2.3 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10002 | 3; type strain | EFex | 2017-10-04 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.66 | Arcanobacterium | phocae | 2.5 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.48 | Arcanobacterium | phocae | 2.4 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2698 | 3; type strain | DT | 2013-08-07 | Arcanobacterium | phocisimile | Arcanobacterium | phocisimile | 2.15 | Arcanobacterium | phocisimile | 2.13 | A | 1 | 1 | 0 | Arcanobacterium phocisimile | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.63 | Arcanobacterium | pinnipediorum | 2.47 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium pinnipediorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.75 | Arcanobacterium | pinnipediorum | 2.56 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium pinnipediorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13483 | 3; type strain | EFex | 2017-10-04 | Arcanobacterium | pluranimalium | Arcanobacterium | pluranimalium | 2.47 | Arcanobacterium | pluranimalium | 2.1 | A | 1 | 1 | 0 | Arcanobacterium pluranimalium | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | sp-17-2065.02 | 2.72 | Arcanobacterium | sp-17-2065.02 | 2.62 | A | 1 | 1 | 0 | Arcanobacterium sp-17-2065.02 | Arcanobacterium sp-17-2065.02 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | sp-17-2065.02 | 2.63 | Arcanobacterium | sp-17-2065.02 | 2.61 | A | 1 | 1 | 0 | Arcanobacterium sp-17-2065.02 | Arcanobacterium sp-17-2065.02 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium | sp-2701 | Arcanobacterium | sp-2701 | 2.7 | Arcanobacterium | sp-2701 | 2.57 | A | 1 | 1 | 0 | Arcanobacterium sp-2701 | Arcanobacterium sp-2701 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium | sp-2701 | Arcanobacterium | sp-2701 | 2.78 | Arcanobacterium | sp-2701 | 2.64 | A | 1 | 1 | 0 | Arcanobacterium sp-2701 | Arcanobacterium sp-2701 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1060.2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium | wilhelmae | Arcanobacterium | wilhelmae | 2.32 | Arcanobacterium | wilhelmae | 2.18 | A | 1 | 1 | 0 | Arcanobacterium wilhelmae | Arcanobacterium wilhelmae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 647 | 4; type strain | DT | 2014-09-04 | Arcanobacterium | wilhelmae | Arcanobacterium | wilhelmae | 2.43 | Arcanobacterium | wilhelmae | 2.25 | A | 1 | 1 | 0 | Arcanobacterium wilhelmae | Arcanobacterium wilhelmae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia | sp-CVUAS-30561 | Boudabousia | sp-CVUAS-30561 | 2.65 | Streptomyces | polychromogenes | 1.39 | A | 1 | 1 | 0 | Boudabousia sp-CVUAS-30561 | Streptomyces polychromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia | canis | Schaalia | canis | 2.41 | Schaalia | canis | 2.26 | A | 1 | 1 | 0 | Schaalia canis | Schaalia canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia | cardiffensis | Schaalia | cardiffensis | 2.49 | Schaalia | cardiffensis | 1.4 | A | 1 | 1 | 0 | Schaalia cardiffensis | Schaalia cardiffensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.45 | Schaalia | hyovaginalis | 2.43 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.48 | Schaalia | hyovaginalis | 2.4 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia | sp-CVUAS-32183 | Schaalia | sp-CVUAS-32183 | 2.55 | Schaalia | hyovaginalis | 1.47 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-32183 | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia | sp-CVUAS-8688 | Schaalia | sp-CVUAS-8688 | 2.57 | Schaalia | sp-CVUAS-8688 | 2.35 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-8688 | Schaalia sp-CVUAS-8688 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31845.2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia | sp-CVUAS-8688 | Schaalia | sp-CVUAS-8688 | 2.73 | Schaalia | sp-CVUAS-8688 | 2.3 | A | 1 | 1 | 0 | Schaalia sp-CVUAS-8688 | Schaalia sp-CVUAS-8688 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella | abortisuis | Trueperella | abortisuis | 2.27 | Trueperella | abortisuis | 2.2 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | abortisuis | Trueperella | abortisuis | 2.36 | Trueperella | abortisuis | 2.16 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella | bernardiae | Trueperella | bernardiae | 2.61 | Trueperella | bernardiae | 2.11 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella | bernardiae | Trueperella | bernardiae | 2.23 | Trueperella | bernardiae | 2.15 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella | bialowiezensis | Trueperella | bialowiezensis | 2.02 | Oceanobacillus | picturae | 1.3 | A | 1 | 1 | 0 | Trueperella bialowiezensis | Oceanobacillus picturae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella | bonasi | Trueperella | bonasi | 2.47 | Trueperella | bonasi | 2.32 | A | 1 | 1 | 0 | Trueperella bonasi | Trueperella bonasi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pecoris | 2.31 | Trueperella | pecoris | 1.98 | A | 1 | 1 | 0 | Trueperella pecoris | Trueperella pecoris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pecoris | 2.37 | Trueperella | pecoris | 1.65 | A | 1 | 1 | 0 | Trueperella pecoris | Trueperella pecoris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.42 | Trueperella | pyogenes | 2.39 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.33 | Trueperella | pyogenes | 2.32 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.19 | Corynebacterium | amycolatum | 2.15 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-17 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.3 | Corynebacterium | amycolatum | 2.29 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium | aurimucosum | Corynebacterium | aurimucosum | 2.07 | Corynebacterium | aurimucosum | 2.05 | A | 1 | 1 | 0 | Corynebacterium aurimucosum | Corynebacterium aurimucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | belfantii | Corynebacterium | belfantii | 2.75 | Corynebacterium | belfantii | 2.59 | A | 1 | 1 | 0 | Corynebacterium belfantii | Corynebacterium belfantii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | belfantii | Corynebacterium | belfantii | 2.66 | Corynebacterium | belfantii | 2.62 | A | 1 | 1 | 0 | Corynebacterium belfantii | Corynebacterium belfantii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium | bovis | Corynebacterium | bovis | 2.35 | Corynebacterium | bovis | 2.33 | A | 1 | 1 | 0 | Corynebacterium bovis | Corynebacterium bovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium | camporealensis | Corynebacterium | camporealensis | 2.58 | Corynebacterium | camporealensis | 1.96 | A | 1 | 1 | 0 | Corynebacterium camporealensis | Corynebacterium camporealensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium | casei | Corynebacterium | casei | 1.98 | Corynebacterium | stationis | 1.64 | B | 0 | 0 | 1 | Corynebacterium casei | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium | confusum | Corynebacterium | confusum | 2.49 | Corynebacterium | confusum | 1.93 | A | 1 | 1 | 0 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | confusum | Corynebacterium | confusum | 2.14 | Corynebacterium | confusum | 2.01 | A | 1 | 1 | 0 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium | cystitidis | Corynebacterium | cystitidis | 2.5 | Corynebacterium | cystitidis | 2.36 | A | 1 | 1 | 0 | Corynebacterium cystitidis | Corynebacterium cystitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.66 | Corynebacterium | diphtheriae | 2.59 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.58 | Corynebacterium | diphtheriae | 2.58 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.71 | Corynebacterium | diphtheriae | 2.65 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0956 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.04 | Corynebacterium | diphtheriae | 2 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.32 | Corynebacterium | frankenforstense | 2.2 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Corynebacterium frankenforstense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.74 | Corynebacterium | frankenforstense | 2.48 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Corynebacterium frankenforstense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium | freiburgense | Corynebacterium | freiburgense | 2.46 | Pseudomonas | gessardii | 1.27 | A | 1 | 1 | 0 | Corynebacterium freiburgense | Pseudomonas gessardii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | hansenii | Corynebacterium | hansenii | 2.35 | Corynebacterium | hansenii | 2.3 | A | 1 | 1 | 0 | Corynebacterium hansenii | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | hansenii | Corynebacterium | hansenii | 2.41 | Corynebacterium | hansenii | 2.2 | A | 1 | 1 | 0 | Corynebacterium hansenii | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | humireducens | Corynebacterium | humireducens | 2.51 | Corynebacterium | sp-CVUAS-31533 | 1.98 | A | 1 | 1 | 0 | Corynebacterium humireducens | Corynebacterium sp-CVUAS-31533 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | kutscheri | Corynebacterium | kutscheri | 2.57 | Corynebacterium | kutscheri | 2.56 | A | 1 | 1 | 0 | Corynebacterium kutscheri | Corynebacterium kutscheri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | lactis | Corynebacterium | lactis | 2.38 | Corynebacterium | lactis | 2.29 | A | 1 | 1 | 0 | Corynebacterium lactis | Corynebacterium lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | lactis | Corynebacterium | lactis | 2.7 | Corynebacterium | lactis | 2.54 | A | 1 | 1 | 0 | Corynebacterium lactis | Corynebacterium lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium | phoceense | Corynebacterium | phoceense | 2.61 | Corynebacterium | phoceense | 2.19 | A | 1 | 1 | 0 | Corynebacterium phoceense | Corynebacterium phoceense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | pilosum | Corynebacterium | pilosum | 2.3 | Corynebacterium | pilosum | 1.78 | A | 1 | 1 | 0 | Corynebacterium pilosum | Corynebacterium pilosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.48 | Corynebacterium | pseudotuberculosis | 2.46 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20689 CVUAS | 3; type strain | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.6 | Corynebacterium | pseudotuberculosis | 2.56 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C104 | 4; sequenced (rpoB) UGent; NGS (CVUAS) | DT | 2020-06-18 | Corynebacterium | ramonii | Corynebacterium | ramonii | 2.5 | Corynebacterium | ulcerans | 2.47 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 112226 EFex | 3; type strain | EFEx | 2024-08-13 | Corynebacterium | ramonii | Corynebacterium | ramonii | 2.74 | Corynebacterium | ramonii | 2.69 | A | 1 | 1 | 0 | Corynebacterium ramonii | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.73 | Corynebacterium | rouxii | 2.71 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3559.2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.7 | Corynebacterium | rouxii | 2.66 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | rouxii | Corynebacterium | rouxii | 2.74 | Corynebacterium | rouxii | 2.65 | A | 1 | 1 | 0 | Corynebacterium rouxii | Corynebacterium rouxii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.71 | Corynebacterium | silvaticum | 2.66 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium | silvaticum | Corynebacterium | silvaticum | 2.64 | Corynebacterium | silvaticum | 2.55 | A | 1 | 1 | 0 | Corynebacterium silvaticum | Corynebacterium silvaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium | sp-208-U2-6 | Corynebacterium | sp-208-U2-6 | 2.63 | Sphingomonas | yabuuchiae | 1.39 | A | 1 | 1 | 0 | Corynebacterium sp-208-U2-6 | Sphingomonas yabuuchiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32165.2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium | sp-2300500 | Corynebacterium | sp-2300500 | 2.59 | Corynebacterium | sp-2300500 | 2.56 | A | 1 | 1 | 0 | Corynebacterium sp-2300500 | Corynebacterium sp-2300500 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5775.2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium | sp-2300500 | Corynebacterium | sp-2300500 | 2.24 | Corynebacterium | sp-2300500 | 2.22 | A | 1 | 1 | 0 | Corynebacterium sp-2300500 | Corynebacterium sp-2300500 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2554.2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium | sp-2396 | Corynebacterium | sp-2396 | 2.44 | Corynebacterium | sp | 1.92 | A | 1 | 1 | 0 | Corynebacterium sp-2396 | Corynebacterium sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium | sp-243-z-3 | Corynebacterium | sp-243-z-3 | 2.57 | Corynebacterium | resistens | 1.78 | A | 1 | 1 | 0 | Corynebacterium sp-243-z-3 | Corynebacterium resistens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | sp-257-y-4 | Corynebacterium | sp-257-y-4 | 2.57 | Proteus | mirabilis | 1.36 | A | 1 | 1 | 0 | Corynebacterium sp-257-y-4 | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | sp-284z-2 | Corynebacterium | sp-284z-2 | 2.5 | Proteus | mirabilis | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-284z-2 | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium | sp-CVUAS-11664 | Corynebacterium | sp-CVUAS-11664 | 2.74 | Pseudomonas | aeruginosa | 1.32 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-11664 | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | sp-CVUAS-1888 | Corynebacterium | sp-CVUAS-1888 | 2.62 | Corynebacterium | striatum | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-1888 | Corynebacterium striatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1721.3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium | sp-CVUAS-2484.4 | Corynebacterium | sp-CVUAS-2484.4 | 2.61 | Corynebacterium | sp-CVUAS-2484.4 | 2.11 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-2484.4 | Corynebacterium sp-CVUAS-2484.4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2484.4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388; 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium | sp-CVUAS-2484.4 | Corynebacterium | sp-CVUAS-2484.4 | 2.37 | Corynebacterium | sp-CVUAS-2484.4 | 1.99 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-2484.4 | Corynebacterium sp-CVUAS-2484.4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp-CVUAS-30404 | 2.49 | Corynebacterium | sp-CVUAS-30404 | 2.27 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-30404 | Corynebacterium sp-CVUAS-30404 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp-CVUAS-30404 | 2.12 | Corynebacterium | sp-CVUAS-30404 | 2.04 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-30404 | Corynebacterium sp-CVUAS-30404 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium | sp-CVUAS-31533 | Corynebacterium | sp-CVUAS-31533 | 2.4 | Corynebacterium | humireducens | 2.22 | B | 0 | 0 | 1 | Corynebacterium sp-CVUAS-31533 | Corynebacterium humireducens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32278.2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium | sp-CVUAS-32278.2 | Corynebacterium | sp-CVUAS-32278.2 | 2.72 | Shewanella | sp-RKVR22 | 1.3 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-32278.2 | Shewanella sp-RKVR22 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium | sp-CVUAS-874,4 | Corynebacterium | sp-CVUAS-874,4 | 2.14 | Corynebacterium | sp-CVUAS-874,4 | 1.98 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-874,4 | Corynebacterium sp-CVUAS-874,4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 874.4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium | sp-CVUAS-874,4 | Corynebacterium | sp-CVUAS-874,4 | 2.48 | Corynebacterium | sp-CVUAS-874,4 | 2.32 | A | 1 | 1 | 0 | Corynebacterium sp-CVUAS-874,4 | Corynebacterium sp-CVUAS-874,4 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 21TRD0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-LALLF-21TRD0553 | Corynebacterium | sp-LALLF-20TRB0553 | 2.39 | Cryptococcus | neoformans | 1.38 | A | 0 | 0 | 1 | Corynebacterium sp-LALLF-20TRB0553 | Cryptococcus neoformans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium | sp-LHL-0497 | Corynebacterium | sp-LHL-0497 | 2.22 | Haemophilus | influenzae | 1.26 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-0497 | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | sp-LHL-0558 | Corynebacterium | sp-LHL-0558 | 2.37 | Nocardia | cyriacigeorgica | 1.32 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-0558 | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium | sp-LHL-1322 | Corynebacterium | sp-LHL-1322 | 2.58 | Agromyces | humatus | 1.41 | A | 1 | 1 | 0 | Corynebacterium sp-LHL-1322 | Agromyces humatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | stationis | Corynebacterium | stationis_JRA_G+R-52 | 2.37 | Corynebacterium | stationis | 2.27 | B | 0 | 0 | 1 | Corynebacterium stationis_JRA_G+R-52 | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium | testudinoris | Corynebacterium | testudinoris | 2.56 | Bacteroides | fragilis | 1.29 | A | 1 | 1 | 0 | Corynebacterium testudinoris | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 2823 | 4; reference strain collection, WGS (ANI) CVUAS | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.52 | Corynebacterium | ramonii | 2.44 | A | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ramonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.6 | Corynebacterium | ulcerans | 2.59 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.53 | Corynebacterium | ulcerans | 2.53 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.34 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ramonii | 2.49 | Corynebacterium | ulcerans | 2.44 | B | 0 | 0 | 1 | Corynebacterium ramonii | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6522.4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.43 | Corynebacterium | ulcerans | 2.43 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5437.3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium | variabile | Corynebacterium | variabile | 2.17 | Corynebacterium | variabile | 1.62 | A | 1 | 1 | 0 | Corynebacterium variabile | Corynebacterium variabile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium | xerosis | Corynebacterium | xerosis | 2.39 | Corynebacterium | xerosis | 1.57 | A | 1 | 1 | 0 | Corynebacterium xerosis | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium | xerosis-group | Corynebacterium | xerosis-group | 2.33 | Corynebacterium | xerosis | 1.58 | A | 1 | 1 | 0 | Corynebacterium xerosis-group | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2379.2 | 3; sequenced | EFex | 2021-04-08 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.29 | Gordonia | paraffinivorans | 2.27 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.27 | Gordonia | paraffinivorans | 2.14 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia | rubripertincta | Gordonia | rubripertincta | 2.61 | Gordonia | rubripertincta | 2.25 | A | 1 | 1 | 0 | Gordonia rubripertincta | Gordonia rubripertincta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10185.3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium | smegmatis | Mycolicibacterium | smegmatis | 2.32 | Tsukamurella | sp | 1.49 | A | 1 | 1 | 0 | Mycolicibacterium smegmatis | Tsukamurella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31296.3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia | cyriacigeorgica | Nocardia | cyriacigeorgica | 2.45 | Nocardia | cyriacigeorgica | 2.28 | A | 1 | 1 | 0 | Nocardia cyriacigeorgica | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia | farcinica | Nocardia | farcinica | 2.36 | Nocardia | kroppenstedtii | 2.26 | B | 0 | 0 | 1 | Nocardia farcinica | Nocardia kroppenstedtii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.46 | Rhodococcus | equi | 2.42 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | equi | 2.41 | Rhodococcus | equi | 2.35 | A | 1 | 1 | 0 | Rhodococcus equi | Rhodococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1447.3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella | sp | Tsukamurella | sp | 2.44 | Tsukamurella | tyrosinosolvens | 1.58 | A | 1 | 1 | 0 | Tsukamurella sp | Tsukamurella tyrosinosolvens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2895.3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium | luteolum | Brevibacterium | luteolum | 2.41 | Brevibacterium | luteolum | 1.92 | A | 1 | 1 | 0 | Brevibacterium luteolum | Brevibacterium luteolum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas | flavigena | Cellulomonas | flavigena | 2.76 | Bacteroides | fragilis | 1.43 | A | 1 | 1 | 0 | Cellulomonas flavigena | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium | sp-234z-2 | Brachybacterium | sp-234z-2 | 2.32 | Glutamicibacter | protophormiae | 1.31 | A | 1 | 1 | 0 | Brachybacterium sp-234z-2 | Glutamicibacter protophormiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium | sp-CVUAS-30560 | Brachybacterium | sp-CVUAS-30560 | 2.66 | Raoultella | ornithinolytica | 1.39 | A | 1 | 1 | 0 | Brachybacterium sp-CVUAS-30560 | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus | sedentarius | Kytococcus | sedentarius | 2.1 | Kytococcus | sedentarius | 1.84 | A | 1 | 1 | 0 | Kytococcus sedentarius | Kytococcus sedentarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus | congolensis | Dermatophilus | congolensis | 2.41 | Dermatophilus | congolensis | 2.19 | A | 1 | 1 | 0 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus | congolensis | Dermatophilus | congolensis | 2.65 | Dermatophilus | congolensis | 2.48 | A | 1 | 1 | 0 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium | flaccumfaciens | Curtobacterium | flaccumfaciens | 2.42 | Curtobacterium | flaccumfaciens | 2.32 | A | 1 | 1 | 0 | Curtobacterium flaccumfaciens | Curtobacterium flaccumfaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia | sp-MUVA-A-508948 | Leifsonia | sp-MUVA-A-508948 | 2.62 | Staphylococcus | aureus | 1.46 | A | 1 | 1 | 0 | Leifsonia sp-MUVA-A-508948 | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium | sp | Microbacterium | phyllosphaerae | 1.84 | Microbacterium | foliorum | 1.84 | B | 0 | 0 | 1 | Microbacterium phyllosphaerae | Microbacterium foliorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34837 | 3; sequenced | EFEx | 2023-01-19 | Microbacterium | sp-CVUAS-34837 | Microbacterium | sp-CVUAS-34837 | 2.6 | Microbacterium | lacticum | 1.39 | A | 1 | 1 | 0 | Microbacterium sp-CVUAS-34837 | Microbacterium lacticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34838 | 3; sequenced | EFEx | 2023-01-19 | Microbacterium | sp-G5-3 | Microbacterium | sp-G5-3 | 2.5 | Microbacterium | sorbitolivorans | 1.48 | A | 1 | 1 | 0 | Microbacterium sp-G5-3 | Microbacterium sorbitolivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2011-09-01 | Glutamicibacter | nicotianae | Glutamicibacter | nicotianae | 2.75 | Lactobacillus | pentosus | 1.33 | A | 1 | 1 | 0 | Glutamicibacter nicotianae | Lactobacillus pentosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter | protophormiae | Glutamicibacter | protophormiae | 2.55 | Lactobacillus | sakei | 1.37 | A | 1 | 1 | 0 | Glutamicibacter protophormiae | Lactobacillus sakei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria | atrinae | Kocuria | atrinae | 2.27 | Kocuria | atrinae | 2 | A | 1 | 1 | 0 | Kocuria atrinae | Kocuria atrinae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 56a-5 | 3; sequenced | EFEx | 2020-05-20 | Kocuria | gwangalensis | Kocuria | gwangalliensis | 2.73 | Kocuria | carniphila | 1.64 | A | 0 | 0 | 1 | Kocuria gwangalliensis | Kocuria carniphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-14 | Kocuria | palustris | Kocuria | palustris | 2.38 | Kocuria | palustris | 2.33 | A | 1 | 1 | 0 | Kocuria palustris | Kocuria palustris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20030 | 3; type strain | #CON=3-5 loops in 100 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 2:2 HCCA¦TA2 (A) | 2011-08-25 | Micrococcus | luteus | Micrococcus | luteus | 2.86 | Micrococcus | luteus | 1.98 | A | 1 | 1 | 0 | Micrococcus luteus | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Micrococcus | luteus | Micrococcus | luteus | 2.37 | Micrococcus | luteus | 2.24 | A | 1 | 1 | 0 | Micrococcus luteus | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20032 CVUAS | 3; type strain | eDT | 2015-07-09 | Rothia | kristinae | Rothia | kristinae | 2.29 | Rothia | kristinae | 2.26 | A | 1 | 1 | 0 | Rothia kristinae | Rothia kristinae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10432 DSM 20127 | 3; type strain | unknown | 2011-09-08 | Sinomonas | atrocyanea | Sinomonas | atrocyanea | 2.52 | Sinomonas | atrocyanea | 2.24 | A | 1 | 1 | 0 | Sinomonas atrocyanea | Sinomonas atrocyanea | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium | cellulans | Cellulosimicrobium | cellulans | 2.05 | Cellulosimicrobium | cellulans | 1.97 | A | 1 | 1 | 0 | Cellulosimicrobium cellulans | Cellulosimicrobium cellulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 749.2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter | inulinus | Sanguibacter | inulinus | 2.51 | Sanguibacter | inulinus | 2.46 | A | 1 | 1 | 0 | Sanguibacter inulinus | Sanguibacter inulinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108415 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Cutibacterium acnes DSM 108415: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Cutibacterium | acnes | Cutibacterium | acnes | 2.24 | Cutibacterium | acnes | 2.17 | A | 1 | 1 | 0 | Cutibacterium acnes | Cutibacterium acnes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1079 | 3; public strain collection | EFEx | 2017-01-20 | Penicillium | roqueforti | Penicillium | roqueforti | 2.7 | Penicillium | camemberti | 1.55 | A | 1 | 1 | 0 | Penicillium roqueforti | Penicillium camemberti | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10354.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | verrucosum | 2.35 | Trichophyton | verrucosum | 2.22 | A | 1 | 1 | 0 | Trichophyton verrucosum | Trichophyton verrucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10357.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | verrucosum | 2.65 | Trichophyton | verrucosum | 2.06 | A | 1 | 1 | 0 | Trichophyton verrucosum | Trichophyton verrucosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1386 | 3; public strain collection | Fextr | 2016-03-10 | Candida | albicans | Candida | albicans | 2.2 | Candida | albicans | 2.14 | A | 1 | 1 | 0 | Candida albicans | Candida albicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31652 | 3; standard identification | EFEx | 2019-03-25 | Saccharomyces | cerevisiae | Saccharomyces | cerevisiae | 2.18 | Saccharomyces | cerevisiae | 1.92 | A | 1 | 1 | 0 | Saccharomyces cerevisiae | Saccharomyces cerevisiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 011 | 3; sequenced | EFEx | 2016-12-08 | Sarocladium | strictum | Sarocladium | strictum | 2.67 | Sarocladium | strictum | 1.54 | A | 1 | 1 | 0 | Sarocladium strictum | Sarocladium strictum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2520.4 | 3; sequenced (ITS) | F-Extr | 2018-07-26 | Fusarium | sambucinum | Fusarium | sambucinum | 2.74 | Kodamaea | ohmeri | 1.23 | A | 1 | 1 | 0 | Fusarium sambucinum | Kodamaea ohmeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.48 | Kurthia | gibsonii | 2.47 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.34 | Kurthia | gibsonii | 2.29 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia | zopfii | Kurthia | zopfii | 2.31 | Kurthia | zopfii | 1.66 | A | 1 | 1 | 0 | Kurthia zopfii | Kurthia zopfii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 100506 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 03.11.2021¦ 1:1 HCCA | 2021-11-03 | Lysinibacillus | endophyticus | Lysinibacillus | endophyticus | 2.45 | Enterococcus | faecium | 1.64 | A | 1 | 1 | 0 | Lysinibacillus endophyticus | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.34 | Lysinibacillus | fusiformis | 2.23 | A | 1 | 1 | 0 | Lysinibacillus fusiformis | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.44 | Lysinibacillus | sphaericus | 2.36 | B | 0 | 0 | 1 | Lysinibacillus fusiformis | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7192.2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus | macroides | Lysinibacillus | macroides | 2.64 | Lysinibacillus | xylanilyticus | 2.23 | B | 0 | 0 | 1 | Lysinibacillus macroides | Lysinibacillus xylanilyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 54 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-03 | Lysinibacillus | macroides | Lysinibacillus | macroides | 2.67 | Lysinibacillus | macroides | 1.8 | A | 1 | 1 | 0 | Lysinibacillus macroides | Lysinibacillus macroides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24784 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-15 | Lysinibacillus | pakistanensis | Lysinibacillus | pakistanensis | 2.3 | Lysinibacillus | fusiformis | 1.34 | A | 1 | 1 | 0 | Lysinibacillus pakistanensis | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.34 | Lysinibacillus | sphaericus | 2.31 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.11 | Lysinibacillus | sphaericus | 1.96 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23493 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-15 | Lysinibacillus | xylanilyticus | Lysinibacillus | xylanilyticus | 2.74 | Lysinibacillus | xylanilyticus | 1.69 | A | 1 | 1 | 0 | Lysinibacillus xylanilyticus | Lysinibacillus xylanilyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus | saudimassiliensis | Metalysinibacillus | saudimassiliensis | 2.53 | Lysinibacillus | saudimassiliensis | 2.01 | C | 0 | 0 | 1 | Metalysinibacillus saudimassiliensis | Lysinibacillus saudimassiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus | odysseyi | Solibacillus | odysseyi | 2.2 | Ureibacillus | sinduriensis | 1.47 | A | 1 | 1 | 0 | Solibacillus odysseyi | Ureibacillus sinduriensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium | acetylicum | Exiguobacterium | acetylicum | 2.63 | Exiguobacterium | enclense | 2.1 | B | 0 | 0 | 1 | Exiguobacterium acetylicum | Exiguobacterium enclense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium | enclense | Exiguobacterium | enclense | 2.19 | Exiguobacterium | acetylicum | 2.03 | B | 0 | 0 | 1 | Exiguobacterium enclense | Exiguobacterium acetylicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium | profundum | Exiguobacterium | profundum | 2.44 | Exiguobacterium | marinum | 2.18 | B | 0 | 0 | 1 | Exiguobacterium profundum | Exiguobacterium marinum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.36 | Brochothrix | thermosphacta | 2.32 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.39 | Brochothrix | thermosphacta | 2.37 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria | aquatica | Listeria | aquatica | 2.33 | Listeria | aquatica | 2.28 | A | 1 | 1 | 0 | Listeria aquatica | Listeria aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria | aquatica | Listeria | aquatica | 2.5 | Listeria | aquatica | 2.43 | A | 1 | 1 | 0 | Listeria aquatica | Listeria aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.25 | Listeria | fleischmannii | 2.23 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LU2006-3 | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.2 | Listeria | fleischmannii | 2.19 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105447 | 3; type strain | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.32 | Listeria | grayi | 2.32 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20596 | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.25 | Listeria | grayi | 2.12 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10132.2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | innocua | Listeria | innocua | 2.45 | Listeria | innocua | 2.44 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8569 | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | innocua | Listeria | innocua | 2.48 | Listeria | innocua | 2.46 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria | ivanovii | Listeria | ivanovii | 2.33 | Listeria | ivanovii | 2.32 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32100 | 3; sequenced | EFex | 2019-12-13 | Listeria | ivanovii | Listeria | ivanovii | 2.41 | Listeria | ivanovii | 2.35 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 23813 | 3; type strain | EFEx | 2017-02-15 | Listeria | marthii | Listeria | marthii | 2.38 | Listeria | monocytogenes | 2.28 | B | 0 | 0 | 1 | Listeria marthii | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.42 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 873.2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.43 | Listeria | monocytogenes | 2.41 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 240.2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.29 | Listeria | monocytogenes | 2.24 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CLIP 70676 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SLCC 2479 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.19 | Listeria | monocytogenes | 2.19 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.22 | Listeria | monocytogenes | 2.21 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.37 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7366.2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.54 | Listeria | monocytogenes | 2.51 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria | newyorkensis | Listeria | newyorkensis | 2.47 | Listeria | sp | 2.17 | B | 0 | 0 | 1 | Listeria newyorkensis | Listeria sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 36192 | 3; sequenced | EFex | 2024-03-06 | Listeria | newyorkensis | Listeria | newyorkensis | 2.21 | Listeria | sp | 2.17 | B | 0 | 0 | 1 | Listeria newyorkensis | Listeria sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 22097 | 3; type strain | EFEx | 2017-02-15 | Listeria | rocourtiae | Listeria | rocourtiae | 2.45 | Listeria | rocourtiae | 2.1 | A | 1 | 1 | 0 | Listeria rocourtiae | Listeria rocourtiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2763 | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | seeligeri | Listeria | seeligeri | 2.45 | Listeria | seeligeri | 2.45 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | seeligeri | Listeria | seeligeri | 2.36 | Listeria | seeligeri | 2.3 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria | sp | Listeria | sp | 2.31 | Listeria | sp | 2.29 | A | 1 | 1 | 0 | Listeria sp | Listeria sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2412.3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | welshimeri | Listeria | welshimeri | 2.38 | Listeria | welshimeri | 2.28 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria | welshimeri | Listeria | welshimeri | 2.44 | Listeria | welshimeri | 2.41 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.31 | Brevibacillus | laterosporus | 1.95 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.42 | Brevibacillus | laterosporus | 2.25 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD1.4A | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-08 | Brevibacillus | parabrevis | Brevibacillus | parabrevis | 2.4 | Brevibacillus | parabrevis | 1.98 | A | 1 | 1 | 0 | Brevibacillus parabrevis | Brevibacillus parabrevis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HD3.3A | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Brevibacillus | parabrevis | Brevibacillus | parabrevis | 2.78 | Brevibacillus | parabrevis | 2.1 | A | 1 | 1 | 0 | Brevibacillus parabrevis | Brevibacillus parabrevis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HB1.1 | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Brevibacillus | porteri | Brevibacillus | porteri | 2.38 | Brevibacillus | choshinensis | 1.69 | A | 1 | 1 | 0 | Brevibacillus porteri | Brevibacillus choshinensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HB1.2 | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Brevibacillus | porteri | Brevibacillus | porteri | 2.74 | Brevibacillus | choshinensis | 1.81 | A | 1 | 1 | 0 | Brevibacillus porteri | Brevibacillus choshinensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 210 | 3; eudo | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Paenibacillus | alvei | Paenibacillus | alvei | 1.82 | Paenibacillus | curdlanolyticus | 1.45 | B | 0 | 0 | 1 | Paenibacillus alvei | Paenibacillus curdlanolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus | amylolyticus | Paenibacillus | amylolyticus | 2.26 | Paenibacillus | amylolyticus | 2.25 | A | 1 | 1 | 0 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus | amylolyticus | Paenibacillus | xylanexedens | 2.22 | Paenibacillus | amylolyticus | 2.17 | B | 0 | 0 | 1 | Paenibacillus xylanexedens | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.53 | Paenibacillus | barengoltzii | 2.5 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.45 | Paenibacillus | barengoltzii | 2.43 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus | chibensis | Paenibacillus | chibensis | 2.28 | Paenibacillus | chibensis | 2.15 | A | 1 | 1 | 0 | Paenibacillus chibensis | Paenibacillus chibensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.18 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.25 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus | larvae | Paenibacillus | larvae | 2.46 | Paenibacillus | larvae | 2.38 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.42 | Paenibacillus | larvae | 2.39 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.38 | Paenibacillus | larvae | 2.34 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.18 | Paenibacillus | larvae | 2.18 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus | macerans | Paenibacillus | macerans | 2.28 | Paenibacillus | macerans | 2.07 | A | 1 | 1 | 0 | Paenibacillus macerans | Paenibacillus macerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus | nuruki | Paenibacillus | nuruki | 2.54 | Ochrobactrum | anthropi | 1.29 | A | 1 | 1 | 0 | Paenibacillus nuruki | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.12 | Paenibacillus | pabuli | 2.02 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.22 | Paenibacillus | pabuli | 2.15 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus | pasadenensis | Paenibacillus | pasadenensis | 2.52 | Paenibacillus | pasadenensis | 2.05 | A | 1 | 1 | 0 | Paenibacillus pasadenensis | Paenibacillus pasadenensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-09-12 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 1.9 | Paenibacillus | polymyxa | 1.78 | B | 0 | 0 | 1 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-24 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 2.81 | Paenibacillus | polymyxa | 2.1 | A | 1 | 1 | 0 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus | sp-CVUAS-6369 | Paenibacillus | sp-CVUAS-6369 | 2.08 | Lentilactobacillus | parabuchneri | 1.35 | A | 1 | 1 | 0 | Paenibacillus sp-CVUAS-6369 | Lentilactobacillus parabuchneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus | sp-E8a | Paenibacillus | sp-E8a | 2.28 | Paenibacillus | amylolyticus | 2.24 | B | 0 | 0 | 1 | Paenibacillus sp-E8a | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus | taichungensis | Paenibacillus | taichungensis | 2.29 | Paenibacillus | taichungensis | 1.9 | A | 1 | 1 | 0 | Paenibacillus taichungensis | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | FLI 003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus | taichungensis | Paenibacillus | taichungensis | 2.81 | Paenibacillus | taichungensis | 1.87 | A | 1 | 1 | 0 | Paenibacillus taichungensis | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus | tundrae | Paenibacillus | tundrae | 2.09 | Paenibacillus | taichungensis | 1.87 | A | 1 | 1 | 0 | Paenibacillus tundrae | Paenibacillus taichungensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus | validus | Paenibacillus | validus | 2.26 | Paenibacillus | validus | 2.16 | A | 1 | 1 | 0 | Paenibacillus validus | Paenibacillus validus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus | sp-185-U1-2 | Aerococcus | sp-185-U1-2 | 2.41 | Xanthomonas | hortorum | 1.38 | A | 1 | 1 | 0 | Aerococcus sp-185-U1-2 | Xanthomonas hortorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus | suis | Aerococcus | suis | 2.36 | Aerococcus | suis | 2.3 | A | 1 | 1 | 0 | Aerococcus suis | Aerococcus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus | urinaeequi | Aerococcus | urinaeequi | 2.49 | Aerococcus | viridans | 1.72 | A | 1 | 1 | 0 | Aerococcus urinaeequi | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 2016-09-27 | Aerococcus | viridans | Aerococcus | viridans | 2.33 | Aerococcus | viridans | 2.04 | A | 1 | 1 | 0 | Aerococcus viridans | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter | phocae | Atopobacter | phocae | 2.33 | Atopobacter | phocae | 2.31 | A | 1 | 1 | 0 | Atopobacter phocae | Atopobacter phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2442.2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium | divergens | Carnobacterium | divergens | 2.44 | Carnobacterium | divergens | 2.1 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium | divergens | Carnobacterium | divergens | 2.62 | Carnobacterium | divergens | 2.37 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6273 | 3; sequenced | eDT | 2016-01-13 | Carnobacterium | maltaromaticum | Carnobacterium | maltaromaticum | 2.55 | Carnobacterium | maltaromaticum | 2.54 | A | 1 | 1 | 0 | Carnobacterium maltaromaticum | Carnobacterium maltaromaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1519.3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella | adiacens | Granulicatella | adiacens | 2.3 | Granulicatella | adiacens | 2.24 | A | 1 | 1 | 0 | Granulicatella adiacens | Granulicatella adiacens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2616 | 3; sequenced | DT | 2013-09-25 | Marinilactibacillus | sp-CVUAS-2616 | Marinilactibacillus | sp-CVUAS-2616 | 2.02 | Marinilactibacillus | sp-CVUAS-2616 | 1.83 | A | 1 | 1 | 0 | Marinilactibacillus sp-CVUAS-2616 | Marinilactibacillus sp-CVUAS-2616 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8980.3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus | sp-CVUAS-2616 | Marinilactibacillus | sp-CVUAS-2616 | 2.54 | Marinilactibacillus | sp-CVUAS-2616 | 2.44 | A | 1 | 1 | 0 | Marinilactibacillus sp-CVUAS-2616 | Marinilactibacillus sp-CVUAS-2616 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16 Ec | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.3 | Enterococcus | casseliflavus | 2.2 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.33 | Enterococcus | casseliflavus | 2.32 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5030 | 3; sequenced | eDT | 2014-07-23 | Enterococcus | cecorum | Enterococcus | cecorum | 2.34 | Enterococcus | cecorum | 2.29 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8897 | 3; sequenced | eDT | 2014-07-23 | Enterococcus | cecorum | Enterococcus | cecorum | 2.82 | Enterococcus | cecorum | 1.41 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus | durans | Enterococcus | hirae | 2.09 | Enterococcus | hirae | 1.8 | A | 0 | 0 | 1 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus | durans | Enterococcus | durans | 2 | Enterococcus | durans | 1.94 | A | 1 | 1 | 0 | Enterococcus durans | Enterococcus durans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus | faecalis | Enterococcus | faecalis | 2.51 | Enterococcus | faecalis | 2.41 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus | faecalis | Enterococcus | faecalis | 2.39 | Enterococcus | faecalis | 2.38 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2011-07-11 | Enterococcus | faecium | Enterococcus | faecium | 2.09 | Enterococcus | faecium | 2.05 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MR11-207 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus | faecium | Enterococcus | faecium | 2.42 | Enterococcus | faecium | 2.1 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-23 | Enterococcus | gallinarum | Enterococcus | gallinarum | 2.04 | Enterococcus | gallinarum | 1.98 | A | 1 | 1 | 0 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus | gallinarum | Enterococcus | gallinarum | 2.42 | Enterococcus | gallinarum | 2.38 | A | 1 | 1 | 0 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-23 | Enterococcus | hirae | Enterococcus | hirae | 2.33 | Enterococcus | hirae | 2.16 | A | 1 | 1 | 0 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34954 | 3; sequenced | eDT | 2023-11-30 | Enterococcus | hirae | Enterococcus | hirae | 2.39 | Enterococcus | hirae | 2.37 | A | 1 | 1 | 0 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34219 | 3; sequenced | eDT | 2023-11-30 | Enterococcus | mundtii | Enterococcus | mundtii | 2.39 | Enterococcus | mundtii | 2.37 | A | 1 | 1 | 0 | Enterococcus mundtii | Enterococcus mundtii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11490 | 3; sequenced | eDT | 2023-11-30 | Enterococcus | phoeniculicola | Enterococcus | phoeniculicola | 2.13 | Brevundimonas | vesicularis | 1.45 | A | 1 | 1 | 0 | Enterococcus phoeniculicola | Brevundimonas vesicularis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus | solitarius | Tetragenococcus | solitarius | 2.14 | Halomonas | elongata | 1.38 | A | 1 | 1 | 0 | Tetragenococcus solitarius | Halomonas elongata | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1549.2 | 3; sequenced | DT | 2016-09-27 | Vagococcus | salmoninarum | Vagococcus | salmoninarum | 2.3 | Vagococcus | salmoninarum | 2.09 | A | 1 | 1 | 0 | Vagococcus salmoninarum | Vagococcus salmoninarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2023000188 | 2; corresponds to expected result of PT | EFEx | 2023-09-07 | Lacticaseibacillus | rhamnosus | Lacticaseibacillus | rhamnosus | 2.6 | Lacticaseibacillus | rhamnosus | 2.52 | A | 1 | 1 | 0 | Lacticaseibacillus rhamnosus | Lacticaseibacillus rhamnosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus | acidophilus | Lactobacillus | acidophilus | 2.32 | Lactobacillus | acidophilus | 2.24 | A | 1 | 1 | 0 | Lactobacillus acidophilus | Lactobacillus acidophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus | johnsonii | Lactobacillus | johnsonii | 2.42 | Lactobacillus | johnsonii | 2.35 | A | 1 | 1 | 0 | Lactobacillus johnsonii | Lactobacillus johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32277.2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus | sp-CVUAS-32277,2 | Lactobacillus | sp-CVUAS-32277,2 | 2.64 | Limosilactobacillus | gorillae | 1.25 | A | 1 | 1 | 0 | Lactobacillus sp-CVUAS-32277,2 | Limosilactobacillus gorillae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus | claussenii | Pediococcus | clausenii | 2.41 | Pediococcus | claussenii | 2.25 | B | 0 | 0 | 1 | Pediococcus clausenii | Pediococcus claussenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1213 | 2; 0 | DT | 2013-07-09 | Leuconostoc | mesenteroides | Leuconostoc | mesenteroides | 2.28 | Leuconostoc | mesenteroides | 2.09 | A | 1 | 1 | 0 | Leuconostoc mesenteroides | Leuconostoc mesenteroides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1146 | 2; 0 | DT | 2013-07-09 | Lactococcus | lactis | Lactococcus | lactis | 2.42 | Lactococcus | lactis | 2.4 | A | 1 | 1 | 0 | Lactococcus lactis | Lactococcus lactis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20622 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | acidominimus | Streptococcus | acidominimus | 2.24 | Streptococcus | acidominimus | 1.89 | A | 1 | 1 | 0 | Streptococcus acidominimus | Streptococcus acidominimus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.45 | Streptococcus | agalactiae | 2.44 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6784 CVUAS | 3; public strain collection | eDT | 2019-07-09 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.6 | Streptococcus | agalactiae | 2.56 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20728 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | alactolyticus | Streptococcus | alactolyticus | 2.04 | Streptococcus | alactolyticus | 2.04 | A | 1 | 1 | 0 | Streptococcus alactolyticus | Streptococcus alactolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20563 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | anginosus | Streptococcus | anginosus | 2.47 | Streptococcus | anginosus | 2.21 | A | 1 | 1 | 0 | Streptococcus anginosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25818 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | anginosus | Streptococcus | anginosus | 2.13 | Streptococcus | anginosus | 2.08 | A | 1 | 1 | 0 | Streptococcus anginosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15627 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | australis | Streptococcus | australis | 2.09 | Streptococcus | parasanguinis | 1.88 | A | 1 | 1 | 0 | Streptococcus australis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 103678 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | azizii | Streptococcus | acidominimus | 1.76 | Streptococcus | pneumoniae | 1.64 | B | 0 | 0 | 1 | Streptococcus acidominimus | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 69277 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | bovimastitidis | Streptococcus | penaeicida | 1.89 | Streptococcus | uberis | 1.69 | B | 0 | 0 | 1 | Streptococcus penaeicida | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19004 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | caballi | Streptococcus | caballi | 2.44 | Streptococcus | mitis_oralis | 1.53 | A | 1 | 1 | 0 | Streptococcus caballi | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 73951 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | caledonicus | Streptococcus | porci | 1.58 | Streptococcus | ratti | 1.53 | B | 0 | 0 | 1 | Streptococcus porci | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105751 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | cameli | Streptococcus | pneumoniae | 1.53 | Streptococcus | pneumoniae | 1.52 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus | canis | Streptococcus | canis | 2.53 | Streptococcus | canis | 2.41 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL Mi-845-17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus | canis | Streptococcus | canis | 2.5 | Streptococcus | canis | 2.49 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 67170 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | caprae | Streptococcus | pyogenes | 1.51 | Streptococcus | alactolyticus | 1.5 | B | 0 | 0 | 1 | Streptococcus pyogenes | Streptococcus alactolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.72 | Streptococcus | castoreus | 2.64 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17536 CVUAS | 3; type strain | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.68 | Streptococcus | castoreus | 2.61 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus | catagoni | Streptococcus | catagoni | 2.83 | Streptococcus | catagoni | 2.81 | A | 1 | 1 | 0 | Streptococcus catagoni | Streptococcus catagoni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 99-1-2017 | 3; type strain | EFEx | 2018-05-03 | Streptococcus | catagoni | Streptococcus | catagoni | 2.81 | Streptococcus | catagoni | 2.8 | A | 1 | 1 | 0 | Streptococcus catagoni | Streptococcus catagoni | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102819 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | caviae | Streptococcus | devriesei | 2.11 | Streptococcus | orisasini | 2.1 | B | 0 | 0 | 1 | Streptococcus devriesei | Streptococcus orisasini | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 106182 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | chenjunshii | Streptococcus | canis | 1.59 | Streptococcus | ferus | 1.55 | B | 0 | 0 | 1 | Streptococcus canis | Streptococcus ferus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14507 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | constellatus | Streptococcus | constellatus | 2.31 | Streptococcus | constellatus | 2.24 | A | 1 | 1 | 0 | Streptococcus constellatus | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17475 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | constellatus | Streptococcus | constellatus | 2.52 | Streptococcus | constellatus | 1.99 | A | 1 | 1 | 0 | Streptococcus constellatus | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25819 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | constellatus | Streptococcus | constellatus | 2.32 | Streptococcus | constellatus | 2.1 | A | 1 | 1 | 0 | Streptococcus constellatus | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20562 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | criceti | Streptococcus | criceti | 2.51 | Streptococcus | criceti | 2.35 | A | 1 | 1 | 0 | Streptococcus criceti | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8249 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | cristatus | Streptococcus | cristatus | 2.36 | Streptococcus | cristatus | 2.35 | A | 1 | 1 | 0 | Streptococcus cristatus | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 65085 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | cuniculi | Streptococcus | suis | 1.7 | Streptococcus | acidominimus | 1.67 | B | 0 | 0 | 1 | Streptococcus suis | Streptococcus acidominimus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 66496 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | cuniculipharyngis | Pseudomonas | mucidolens | 1.54 | Pseudomonas | migulae | 1.51 | B | 0 | 0 | 1 | Pseudomonas mucidolens | Pseudomonas migulae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22233 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | danieliae | Streptococcus | pneumoniae | 1.53 | Streptococcus | pneumoniae | 1.49 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21999 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | dentapri | Streptococcus | equi | 1.7 | Streptococcus | canis | 1.61 | B | 0 | 0 | 1 | Streptococcus equi | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27381 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | dentiloxodontae | Streptococcus | equi_ssp_zooepidemicus | 1.6 | Streptococcus | salivarius | 1.59 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18963 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | dentirousetti | Streptococcus | dentirousetti | 2.15 | Streptococcus | downei | 2.14 | B | 0 | 0 | 1 | Streptococcus dentirousetti | Streptococcus downei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25137 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | dentisani | Streptococcus | downei | 1.76 | Streptococcus | criceti | 1.63 | B | 0 | 0 | 1 | Streptococcus downei | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 47155 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | devriesei | Streptococcus | devriesei | 2.59 | Streptococcus | devriesei | 2.41 | A | 1 | 1 | 0 | Streptococcus devriesei | Streptococcus devriesei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15616 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | didelphis | Streptococcus | didelphis | 2.28 | Streptococcus | pyogenes | 1.59 | A | 1 | 1 | 0 | Streptococcus didelphis | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5635 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | downei | Streptococcus | downei | 2.4 | Streptococcus | dentirousetti | 1.9 | A | 1 | 1 | 0 | Streptococcus downei | Streptococcus dentirousetti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 73139 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | downii | Streptococcus | mitis_oralis | 2.46 | Streptococcus | mitis_oralis | 2.38 | A | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.15 | Streptococcus | dysgalactiae | 2.14 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.26 | Streptococcus | dysgalactiae | 2.12 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.43 | Streptococcus | dysgalactiae | 2.35 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.46 | Streptococcus | dysgalactiae | 2.33 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14446 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | entericus | Streptococcus | entericus | 2.36 | Streptococcus | pyogenes | 1.42 | A | 1 | 1 | 0 | Streptococcus entericus | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32462.2 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi | 2.26 | Streptococcus | equi | 2.18 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32770 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi | 2.42 | Streptococcus | equi | 2.41 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus | equi | Streptococcus | equi | 2.42 | Streptococcus | equi | 2.35 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus | equi | Streptococcus | equi | 2.25 | Streptococcus | equi | 2.23 | A | 1 | 1 | 0 | Streptococcus equi | Streptococcus equi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20558 RWTH | 3; type strain | EFEx | 2023-09-05 | Streptococcus | equinus | Streptococcus | equinus | 2.37 | Streptococcus | equinus | 2.31 | A | 1 | 1 | 0 | Streptococcus equinus | Streptococcus equinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20646 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ferus | Streptococcus | ferus | 2.45 | Streptococcus | ferus | 2.24 | A | 1 | 1 | 0 | Streptococcus ferus | Streptococcus ferus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15349 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | gallinaceus | Streptococcus | gallinaceus | 2.39 | Streptococcus | mitis_oralis | 1.51 | A | 1 | 1 | 0 | Streptococcus gallinaceus | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16831 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | gallolyticus | Streptococcus | gallolyticus | 2.39 | Streptococcus | gallolyticus | 2.29 | A | 1 | 1 | 0 | Streptococcus gallolyticus | Streptococcus gallolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15879 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | gallolyticus | Streptococcus | gallolyticus | 2.27 | Streptococcus | gallolyticus | 2.1 | A | 1 | 1 | 0 | Streptococcus gallolyticus | Streptococcus gallolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15351 RWTH | 3; type strain | EFEx | 2023-09-06 | Streptococcus | gallolyticus | Streptococcus | gallolyticus | 2.42 | Streptococcus | gallolyticus | 2.12 | A | 1 | 1 | 0 | Streptococcus gallolyticus | Streptococcus gallolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6777 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | gordonii | Streptococcus | gordonii | 2.22 | Streptococcus | gordonii | 2.22 | A | 1 | 1 | 0 | Streptococcus gordonii | Streptococcus gordonii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17028 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | halichoeri | Streptococcus | halichoeri | 2.4 | Streptococcus | equi_ssp_zooepidemicus | 1.75 | A | 1 | 1 | 0 | Streptococcus halichoeri | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101996 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | halotolerans | Streptococcus | pluranimalium | 1.69 | Streptococcus | pluranimalium | 1.67 | B | 0 | 0 | 1 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19005 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | henryi | Streptococcus | henryi | 2.43 | Streptococcus | sp-CVUAS-35986 | 1.92 | A | 1 | 1 | 0 | Streptococcus henryi | Streptococcus sp-CVUAS-35986 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 107951 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | hillyeri | Streptococcus | agalactiae | 1.53 | Streptococcus | dysgalactiae | 1.49 | B | 0 | 0 | 1 | Streptococcus agalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101997 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | himalayensis | Streptococcus | sanguinis | 1.62 | Pseudomonas | luteola | 1.6 | B | 0 | 0 | 1 | Streptococcus sanguinis | Pseudomonas luteola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26014 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | hongkongensis | Streptococcus | equi_ssp_zooepidemicus | 1.84 | Streptococcus | agalactiae | 1.84 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20770 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | hyointestinalis | Streptococcus | hyointestinalis | 2.56 | Streptococcus | hyointestinalis | 2.3 | A | 1 | 1 | 0 | Streptococcus hyointestinalis | Streptococcus hyointestinalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12219 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | hyovaginalis | Streptococcus | hyovaginalis | 2.43 | Streptococcus | hyovaginalis | 2.3 | A | 1 | 1 | 0 | Streptococcus hyovaginalis | Streptococcus hyovaginalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 52536 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ictaluri | Streptococcus | canis | 1.71 | Streptococcus | castoreus | 1.7 | B | 0 | 0 | 1 | Streptococcus canis | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22960 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | infantarius | Streptococcus | infantarius | 2.03 | Streptococcus | infantarius | 1.91 | A | 1 | 1 | 0 | Streptococcus infantarius | Streptococcus infantarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 39817 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | infantis | Streptococcus | infantis | 2.3 | Streptococcus | infantis | 1.92 | A | 1 | 1 | 0 | Streptococcus infantis | Streptococcus infantis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus | iniae | Streptococcus | iniae | 2.02 | Streptococcus | iniae | 2.02 | A | 1 | 1 | 0 | Streptococcus iniae | Streptococcus iniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20576 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | iniae | Streptococcus | iniae | 2.27 | Streptococcus | iniae | 2.19 | A | 1 | 1 | 0 | Streptococcus iniae | Streptococcus iniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20573 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | intermedius | Streptococcus | intermedius | 2.14 | Streptococcus | intermedius | 2.05 | A | 1 | 1 | 0 | Streptococcus intermedius | Streptococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23027 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | lactarius | Streptococcus | parasanguinis | 2.07 | Streptococcus | parasanguinis | 1.94 | A | 0 | 0 | 1 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27382 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | loxodontisalivarius | Streptococcus | hyointestinalis | 1.65 | Streptococcus | infantarius | 1.64 | B | 0 | 0 | 1 | Streptococcus hyointestinalis | Streptococcus infantarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15350 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | lutetiensis | Streptococcus | lutetiensis | 2.39 | Streptococcus | lutetiensis | 2.33 | A | 1 | 1 | 0 | Streptococcus lutetiensis | Streptococcus lutetiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20724 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | macacae | Streptococcus | macacae | 2.25 | Streptococcus | ratti | 1.51 | A | 1 | 1 | 0 | Streptococcus macacae | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18627 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | marimammalium | Streptococcus | marimammalium | 2.03 | Streptococcus | zalophi | 1.75 | A | 1 | 1 | 0 | Streptococcus marimammalium | Streptococcus zalophi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101995 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | marmotae | Streptococcus | danieliae | 1.84 | Streptococcus | suis | 1.76 | B | 0 | 0 | 1 | Streptococcus danieliae | Streptococcus suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 49690 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | massiliensis | Streptococcus | massiliensis | 2.34 | Streptococcus | massiliensis | 2.28 | A | 1 | 1 | 0 | Streptococcus massiliensis | Streptococcus massiliensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19192 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | merionis | Streptococcus | merionis | 2.43 | Limosilactobacillus | antri | 1.46 | A | 1 | 1 | 0 | Streptococcus merionis | Limosilactobacillus antri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17118 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | minor | Streptococcus | minor | 2.35 | Streptococcus | minor | 2.27 | A | 1 | 1 | 0 | Streptococcus minor | Streptococcus minor | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12643 RWTH | 3; type strain | EFEx | 2023-09-05 | Streptococcus | mitis | Streptococcus | mitis_oralis | 2.25 | Streptococcus | pneumoniae | 2.23 | B | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12643 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 11.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus agalactiae DSM 2134, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-02-18 | Streptococcus | mitis | Streptococcus | pneumoniae | 2.14 | Streptococcus | pneumoniae | 2.06 | A | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105143 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | moroccensis | Streptococcus | moroccensis | 2.13 | Streptococcus | sanguinis | 1.43 | A | 1 | 1 | 0 | Streptococcus moroccensis | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20523 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | mutans | Streptococcus | mutans | 2.27 | Streptococcus | mutans | 2.14 | A | 1 | 1 | 0 | Streptococcus mutans | Streptococcus mutans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27088 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oralis | Streptococcus | mitis_oralis | 2.38 | Streptococcus | mitis_oralis | 2.35 | A | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20627 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | oralis | Streptococcus | mitis_oralis | 2.24 | Streptococcus | mitis_oralis | 2.23 | A | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24864 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oralis | Streptococcus | pneumoniae | 2.34 | Streptococcus | mitis_oralis | 2.32 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 100101 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oricebi | Streptococcus | troglodytidis | 1.9 | Streptococcus | sanguinis | 1.64 | B | 0 | 0 | 1 | Streptococcus troglodytidis | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27377 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | oriloxodontae | Streptococcus | criceti | 2.36 | Streptococcus | criceti | 2.28 | A | 0 | 0 | 1 | Streptococcus criceti | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus | orisasini | Streptococcus | orisasini | 2.86 | Streptococcus | orisasini | 2.44 | A | 1 | 1 | 0 | Streptococcus orisasini | Streptococcus orisasini | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25193 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | orisasini | Streptococcus | orisasini | 2.34 | Streptococcus | orisasini | 2.21 | A | 1 | 1 | 0 | Streptococcus orisasini | Streptococcus orisasini | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 43577 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | orisratti | Streptococcus | orisratti | 2.32 | Streptococcus | sp-CVUAS-35986 | 2.07 | B | 0 | 0 | 1 | Streptococcus orisratti | Streptococcus sp-CVUAS-35986 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18307 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | orisuis | Streptococcus | criceti | 2.45 | Streptococcus | criceti | 2.43 | A | 0 | 0 | 1 | Streptococcus criceti | Streptococcus criceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 39485 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | ovis | Streptococcus | ovis | 2.06 | Streptococcus | ovis | 1.86 | A | 1 | 1 | 0 | Streptococcus ovis | Streptococcus ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 69612 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | ovuberis | Streptococcus | gallolyticus | 1.55 | Streptococcus | cristatus | 1.48 | B | 0 | 0 | 1 | Streptococcus gallolyticus | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 111148 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | pacificus | Streptococcus | pacificus | 2.32 | Streptococcus | zalophi | 1.91 | A | 1 | 1 | 0 | Streptococcus pacificus | Streptococcus zalophi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29921 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | panodentis | Streptococcus | pneumoniae | 1.62 | Streptococcus | sanguinis | 1.58 | B | 0 | 0 | 1 | Streptococcus pneumoniae | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102135 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | pantholopis | Streptococcus | constellatus | 1.49 | Streptococcus | cristatus | 1.48 | B | 0 | 0 | 1 | Streptococcus constellatus | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 9056 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 11.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus agalactiae DSM 2134, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-09-01 | Streptococcus | parasanguinis | Streptococcus | parasanguinis | 2.3 | Streptococcus | parasanguinis | 2.26 | A | 1 | 1 | 0 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6778 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | parasanguinis | Streptococcus | parasanguinis | 2.27 | Streptococcus | parasanguinis | 2.25 | A | 1 | 1 | 0 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29126 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | parasuis | Streptococcus | parasuis | 2.14 | Streptococcus | parasuis | 2.05 | A | 1 | 1 | 0 | Streptococcus parasuis | Streptococcus parasuis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | parauberis | Streptococcus | parauberis_JRA_G+R-39 | 2.18 | Streptococcus | parauberis | 2.1 | B | 0 | 0 | 1 | Streptococcus parauberis_JRA_G+R-39 | Streptococcus parauberis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6631 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | parauberis | Streptococcus | parauberis | 2.33 | Streptococcus | parauberis_JRA_G+R-39 | 2.16 | B | 0 | 0 | 1 | Streptococcus parauberis | Streptococcus parauberis_JRA_G+R-39 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26545 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | penaeicida | Streptococcus | penaeicida | 2.41 | Streptococcus | dysgalactiae | 1.47 | A | 1 | 1 | 0 | Streptococcus penaeicida | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 39814 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | peroris | Streptococcus | mitis_oralis | 1.78 | Streptococcus | mitis_oralis | 1.74 | B | 0 | 0 | 1 | Streptococcus mitis_oralis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 16735 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | phocae | Streptococcus | phocae | 2.23 | Streptococcus | phocae | 2.15 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24768 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | phocae | Streptococcus | phocae | 2.34 | Streptococcus | phocae | 2.29 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | phocae | Streptococcus | phocae | 2.33 | Streptococcus | phocae | 2.3 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 12-13862-001-01 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus | pluranimalium | Streptococcus | pluranimalium | 2.2 | Streptococcus | pluranimalium | 2.18 | A | 1 | 1 | 0 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14177 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | pluranimalium | Streptococcus | pluranimalium | 2.15 | Streptococcus | pluranimalium | 2.12 | A | 1 | 1 | 0 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22810 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | plurextorum | Streptococcus | plurextorum | 2.4 | Streptococcus | merionis | 1.54 | A | 1 | 1 | 0 | Streptococcus plurextorum | Streptococcus merionis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.48 | Streptococcus | pneumoniae | 2.44 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.16 | Streptococcus | pneumoniae | 2.16 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.29 | Streptococcus | pneumoniae | 2.23 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.49 | Streptococcus | pneumoniae | 2.46 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23759 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | porci | Streptococcus | porci | 2.33 | Streptococcus | porci | 1.92 | A | 1 | 1 | 0 | Streptococcus porci | Streptococcus porci | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus | porcinus | Streptococcus | porcinus | 2.27 | Streptococcus | porcinus | 2.1 | A | 1 | 1 | 0 | Streptococcus porcinus | Streptococcus porcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20725 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | porcinus | Streptococcus | porcinus | 2.51 | Streptococcus | porcinus | 2.47 | A | 1 | 1 | 0 | Streptococcus porcinus | Streptococcus porcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28302 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | porcorum | Streptococcus | porcorum | 2.29 | Streptococcus | pyogenes | 1.66 | A | 1 | 1 | 0 | Streptococcus porcorum | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 49455 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | pseudopneumoniae | Streptococcus | pseudopneumoniae | 2.26 | Streptococcus | pseudopneumoniae | 2.25 | A | 1 | 1 | 0 | Streptococcus pseudopneumoniae | Streptococcus pseudopneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18513 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | pseudoporcinus | Streptococcus | pseudoporcinus | 2.15 | Streptococcus | pseudoporcinus | 2.08 | A | 1 | 1 | 0 | Streptococcus pseudoporcinus | Streptococcus pseudoporcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 1567 | 2; strain collection of the RKI | #CON=approx. 3 loops per Eppend. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, Inactivation date: 19.08.2021¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus pneumoniae B2046, preparation: C. Hartl, measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-09-02 | Streptococcus | pyogenes | Streptococcus | pyogenes | 2.5 | Streptococcus | pyogenes | 2.5 | A | 1 | 1 | 0 | Streptococcus pyogenes | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-04 | Streptococcus | pyogenes | Streptococcus | pyogenes | 2.43 | Streptococcus | pyogenes | 2.4 | A | 1 | 1 | 0 | Streptococcus pyogenes | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20564 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | ratti | Streptococcus | ratti | 2.54 | Streptococcus | ratti | 2.36 | A | 1 | 1 | 0 | Streptococcus ratti | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101998 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | respiraculi | Streptococcus | danieliae | 1.9 | Streptococcus | acidominimus | 1.89 | B | 0 | 0 | 1 | Streptococcus danieliae | Streptococcus acidominimus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105174 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | rifensis | Streptococcus | moroccensis | 1.83 | Streptococcus | downei | 1.43 | B | 0 | 0 | 1 | Streptococcus moroccensis | Streptococcus downei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26920 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | rubneri | Streptococcus | parasanguinis | 2.13 | Streptococcus | parasanguinis | 1.95 | A | 0 | 0 | 1 | Streptococcus parasanguinis | Streptococcus parasanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 104980 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ruminantium | Streptococcus | suis | 1.64 | Streptococcus | mitis_oralis | 1.63 | B | 0 | 0 | 1 | Streptococcus suis | Streptococcus mitis_oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28303 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | rupicaprae | Streptococcus | gallolyticus | 1.4 | Streptococcus | parasuis | 1.35 | B | 0 | 0 | 1 | Streptococcus gallolyticus | Streptococcus parasuis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20560 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | salivarius | Streptococcus | salivarius | 2.4 | Streptococcus | salivarius | 2.19 | A | 1 | 1 | 0 | Streptococcus salivarius | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20617 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | salivarius | Streptococcus | thermophilus | 2.43 | Streptococcus | thermophilus | 2.35 | A | 0 | 0 | 1 | Streptococcus thermophilus | Streptococcus thermophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10145.3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus | salivarius | Streptococcus | salivarius | 2.13 | Streptococcus | salivarius | 2.13 | A | 1 | 1 | 0 | Streptococcus salivarius | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27513 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | saliviloxodontae | Streptococcus | equi_ssp_zooepidemicus | 1.88 | Streptococcus | castoreus | 1.69 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus castoreus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20068 | 3; public strain collection | TFAextr | 2022-02-18 | Streptococcus | sanguinis | Streptococcus | sanguinis | 1.96 | Streptococcus | sanguinis | 1.75 | B | 0 | 0 | 1 | Streptococcus sanguinis | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20567 RWTH | 3; type strain | EFEx | 2023-08-30 | Streptococcus | sanguinis | Streptococcus | sanguinis | 2.31 | Streptococcus | sanguinis | 2.22 | A | 1 | 1 | 0 | Streptococcus sanguinis | Streptococcus sanguinis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14990 | 3; type strain | EFEx | 2023-08-30 | Streptococcus | sinensis | Streptococcus | sinensis | 2.14 | Streptococcus | cristatus | 2.11 | B | 0 | 0 | 1 | Streptococcus sinensis | Streptococcus cristatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33478 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | sobrinus | Streptococcus | sobrinus | 2.45 | Streptococcus | sobrinus | 2.41 | A | 1 | 1 | 0 | Streptococcus sobrinus | Streptococcus sobrinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus | sp-CVUAS-32349 | Streptococcus | sp-CVUAS-32349 | 2.58 | Streptococcus | dysgalactiae | 1.52 | A | 1 | 1 | 0 | Streptococcus sp-CVUAS-32349 | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9682 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | suis | Streptococcus | suis | 2.15 | Streptococcus | suis | 2.13 | A | 1 | 1 | 0 | Streptococcus suis | Streptococcus suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105173 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | tangierensis | Streptococcus | equi_ssp_zooepidemicus | 1.46 | Loigolactobacillus | coryniformis | 1.45 | B | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Loigolactobacillus coryniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 32906 | 3; type strain | EFEx | 2023-09-05 | Streptococcus | thoraltensis | Streptococcus | thoraltensis | 2.35 | Streptococcus | thoraltensis | 2.29 | A | 1 | 1 | 0 | Streptococcus thoraltensis | Streptococcus thoraltensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25324 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | troglodytae | Streptococcus | mutans | 2.18 | Streptococcus | mutans | 2.15 | A | 0 | 0 | 1 | Streptococcus mutans | Streptococcus mutans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2260.4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus | troglodytidis | Streptococcus | troglodytidis | 2.52 | Streptococcus | troglodytidis | 2.21 | A | 1 | 1 | 0 | Streptococcus troglodytidis | Streptococcus troglodytidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KCTC 33006 | 3; type strain | EFEx | 2023-06-20 | Streptococcus | troglodytidis | Streptococcus | troglodytidis | 2.34 | Streptococcus | troglodytidis | 2.17 | A | 1 | 1 | 0 | Streptococcus troglodytidis | Streptococcus troglodytidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.03 | Streptococcus | uberis | 1.81 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.23 | Streptococcus | uberis | 2.18 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 41590 | 3; type strain | EFEx | 2023-09-06 | Streptococcus | urinalis | Streptococcus | urinalis | 2.43 | Streptococcus | urinalis | 2.23 | A | 1 | 1 | 0 | Streptococcus urinalis | Streptococcus urinalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22768 | 3; type strain | EFEx | 2023-09-07 | Streptococcus | ursoris | Streptococcus | ratti | 2.17 | Streptococcus | ratti | 2.1 | A | 0 | 0 | 1 | Streptococcus ratti | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5636 | 3; type strain | EFEx | 2023-08-29 | Streptococcus | vestibularis | Streptococcus | vestibularis | 2.39 | Streptococcus | vestibularis | 2.08 | A | 1 | 1 | 0 | Streptococcus vestibularis | Streptococcus vestibularis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 110741 | 3; type strain; WGS (PMID: 33999792) | EFEx | 2023-06-20 | Streptococcus | vicugnae | Streptococcus | vicugnae | 2.17 | Streptococcus | lamae | 2.17 | B | 0 | 0 | 1 | Streptococcus vicugnae | Streptococcus lamae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 110742 | 3; type strain, WGS (PMID: 33999792) | EFEx | 2023-06-20 | Streptococcus | zalophi | Streptococcus | zalophi | 2.33 | Streptococcus | pacificus | 1.84 | A | 1 | 1 | 0 | Streptococcus zalophi | Streptococcus pacificus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 1.94 | Clostridioides | difficile | 1.83 | B | 0 | 0 | 1 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.2 | Clostridioides | difficile | 2.17 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.1 | Clostridium | sporogenes | 2.05 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.92 | Clostridium | sporogenes | 1.88 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 2.01 | Clostridium | botulinum | 1.99 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium botulinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 2.16 | Clostridium | sporogenes | 1.59 | A | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.99 | Clostridium | sporogenes | 1.67 | B | 0 | 0 | 1 | Clostridium botulinum | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | botulinum | 1.84 | Arthrobacter | ramosus | 1.51 | B | 0 | 0 | 1 | Clostridium botulinum | Arthrobacter ramosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium | barati | Clostridium | budayi | 2.31 | Clostridium | baratii | 2.28 | B | 0 | 0 | 1 | Clostridium budayi | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium | botulinum | Clostridium | botulinum | 2.36 | Bacillus | subtilis | 1.49 | A | 1 | 1 | 0 | Clostridium botulinum | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium | butyricum | Clostridium | butyricum | 2.61 | Clostridium | butyricum | 2.48 | A | 1 | 1 | 0 | Clostridium butyricum | Clostridium butyricum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium | chauvoei | Clostridium | chauvoei | 2.13 | Clostridium | chauvoei | 2.01 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium | chauvoei | Clostridium | chauvoei | 2.38 | Clostridium | chauvoei | 2.17 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002729 | 2; 0 | DT | 2017-02-28 | Clostridium | cochlearium | Clostridium | cochlearium | 2.51 | Clostridium | cochlearium | 2.5 | A | 1 | 1 | 0 | Clostridium cochlearium | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium | colicanis | Clostridium | colicanis | 2.29 | Bifidobacterium | boum | 1.42 | A | 1 | 1 | 0 | Clostridium colicanis | Bifidobacterium boum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 849.2 | 3; sequenced | DT | 2014-12-17 | Clostridium | isatidis | Clostridium | isatidis | 2.13 | Clostridium | isatidis | 1.6 | A | 1 | 1 | 0 | Clostridium isatidis | Clostridium isatidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108256 | 3; Public strain collection | EFEx | 0 | Clostridium | leptum | Clostridium | leptum | 2.45 | Clostridium | paraputrificum | 1.43 | A | 1 | 1 | 0 | Clostridium leptum | Clostridium paraputrificum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.23.22 | 3; sequenced [0207] | EFEx | 0 | Clostridium | nexile | Sellimonas | intestinalis | 1.27 | Aeromonas | caviae | 1.24 | B | 0 | 0 | 1 | Sellimonas intestinalis | Aeromonas caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium | nigeriense | Clostridium | nigeriense | 2.58 | Clostridium | tertium | 1.75 | A | 1 | 1 | 0 | Clostridium nigeriense | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium | novyi | Clostridium | novyi | 2.03 | Clostridium | novyi | 1.96 | A | 1 | 1 | 0 | Clostridium novyi | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium | perfringens | Clostridium | perfringens | 2.09 | Clostridium | perfringens | 1.9 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B | 3; case rerort | DT | 2013-08-06 | Clostridium | perfringens | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.4 | A | 1 | 1 | 0 | Clostridium perfringens | Clostridium perfringens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 10419 | 3; Public strain collection | eDT | 2021-10-26 | Clostridium | scindens | Clostridium | scindens | 2.23 | Clostridium | scindens | 2.15 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.1.26 | 3; sequenced [0207] | eDT | 0 | Clostridium | scindens | Clostridium | scindens | 2.2 | Clostridium | scindens | 2.06 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium | septicum | Clostridium | septicum | 2.44 | Clostridium | septicum | 2.26 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium | septicum | Clostridium | septicum | 2.27 | Clostridium | septicum | 2.23 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34091 | 3; sequenced | EFex | 2022-08-05 | Clostridium | sp-CVUAS-34091 | Clostridium | sp-CVUAS-34091 | 2.4 | Clostridium | nigeriense | 1.8 | A | 1 | 1 | 0 | Clostridium sp-CVUAS-34091 | Clostridium nigeriense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium | sp-DSM-1985 | Clostridium | botulinum | 2.26 | Campylobacter | jejuni | 1.33 | A | 0 | 0 | 1 | Clostridium botulinum | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium | sporogenes | Clostridium | sporogenes | 1.95 | Clostridium | sporogenes | 1.79 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium | sporogenes | Clostridium | sporogenes | 2.34 | Clostridium | sporogenes | 2.21 | A | 1 | 1 | 0 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.1.9 | 3; sequenced [0207] | eDT | 0 | Clostridium | symbiosum | Clostridium | symbiosum | 2.43 | Clostridium | symbiosum | 2.38 | A | 1 | 1 | 0 | Clostridium symbiosum | Clostridium symbiosum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WAL-14163 | 3; sequenced [0207] | eDT | 0 | Clostridium | symbiosum | Clostridium | symbiosum | 2.45 | Clostridium | symbiosum | 2.29 | A | 1 | 1 | 0 | Clostridium symbiosum | Clostridium symbiosum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141015223 | 2; 0 | DT | 2014-11-27 | Clostridium | tertium | Clostridium | tertium | 2.44 | Clostridium | tertium | 2.22 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium | tertium | Clostridium | tertium | 2.33 | Clostridium | tertium | 2.28 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium | tetani | Clostridium | tetani | 2.23 | Clostridium | tetani | 2.16 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium | tetani | Clostridium | tetani | 2.55 | Clostridium | tetani | 2.47 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya | limosa | Hathewaya | limosa | 2.29 | Hathewaya | limosa | 2.25 | A | 1 | 1 | 0 | Hathewaya limosa | Hathewaya limosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium | sordellii | Paeniclostridium | sordellii | 2.32 | Paeniclostridium | sordellii | 2.27 | A | 1 | 1 | 0 | Paeniclostridium sordellii | Paeniclostridium sordellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16296 | 3; Public strain collection | EFEx | 0 | Eubacterium | ramulus | Eubacterium | ramulus | 2.49 | Lacticaseibacillus | paracasei | 1.38 | A | 1 | 1 | 0 | Eubacterium ramulus | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H258 | 3; sequenced [0207] | eDT | 0 | Agathobacter | rectalis | Agathobacter | rectalis | 2.39 | Pseudomonas | putida | 1.45 | A | 1 | 1 | 0 | Agathobacter rectalis | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.9.13 | 3; sequenced [0207] | eDT | 0 | Agathobacter | rectalis | Agathobacter | rectalis | 2.14 | Pseudomonas | putida | 1.39 | A | 1 | 1 | 0 | Agathobacter rectalis | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Clinical isolate | 3; sequenced [0207] | EFEx | 0 | Anaerobutyricum | hallii | Anaerobutyricum | hallii | 2.36 | Anaerobutyricum | hallii | 2.08 | A | 1 | 1 | 0 | Anaerobutyricum hallii | Anaerobutyricum hallii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H53 | 3; sequenced [0207] | EFEx | 0 | Anaerostipes | hadrus | Anaerostipes | hadrus | 2.23 | Anaerostipes | hadrus | 2 | A | 1 | 1 | 0 | Anaerostipes hadrus | Anaerostipes hadrus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29492 | 3; type strain | EFEx | 0 | Blautia | caecimuris | Blautia | caecimuris | 2.72 | Roseburia | inulinvorans | 1.34 | A | 1 | 1 | 0 | Blautia caecimuris | Roseburia inulinvorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.18.51 | 3; sequenced [0207] | EFEx | 0 | Blautia | faecis | Blautia | faecis | 2.27 | Lacrimispora | celerecrescens_II | 1.39 | A | 1 | 1 | 0 | Blautia faecis | Lacrimispora celerecrescens_II | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22028 | 3; type strain | EFEx | 0 | Blautia | glucerasea | Blautia | caecimuris | 1.99 | Neisseria | oralis | 1.29 | B | 0 | 0 | 1 | Blautia caecimuris | Neisseria oralis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20583 | 3; type strain | EFEx | 0 | Blautia | hansenii | Aeromonas | molluscorum | 1.4 | Staphylococcus | simiae | 1.34 | B | 0 | 0 | 1 | Aeromonas molluscorum | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.19.60 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Blautia | schinkii_II | 2.16 | Blautia | faecis | 1.51 | A | 0 | 0 | 1 | Blautia schinkii_II | Blautia faecis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.9.19 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Blautia | schinkii_II | 2.27 | Weissella | viridescens | 1.44 | A | 0 | 0 | 1 | Blautia schinkii_II | Weissella viridescens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.20.66 | 3; sequenced [0207] | eDT | 0 | Blautia | obeum | Blautia | obeum | 2.13 | Blautia | obeum | 1.93 | A | 1 | 1 | 0 | Blautia obeum | Blautia obeum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9990 | 3; type strain | eDT | 0 | Blautia | producta | Blautia | coccoides | 2.33 | Blautia | coccoides | 2.03 | A | 0 | 0 | 1 | Blautia coccoides | Blautia coccoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.4.13 | 2; data not listed [0207] | eDT | 0 | Blautia | producta | Clostridium | haemolyticum | 1.43 | Clostridium | baratii | 1.4 | B | 0 | 0 | 1 | Clostridium haemolyticum | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26115 | 3; public strain collection | EFEx | 0 | Blautia | pseudococcoides | Blautia | coccoides | 1.61 | Blautia | coccoides | 1.6 | B | 0 | 0 | 1 | Blautia coccoides | Blautia coccoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.6.27 | 3; sequenced [0207] | EFEx | 2021-10-20 | Blautia | schinkii | Blautia | schinkii_II | 2.44 | Blautia | faecis | 1.3 | A | 0 | 0 | 1 | Blautia schinkii_II | Blautia faecis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.6.32 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia | schinkii | Blautia | schinkii_II | 2.44 | Blautia | faecis | 1.43 | A | 0 | 0 | 1 | Blautia schinkii_II | Blautia faecis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.18.16 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia | wexlerae | Blautia | wexlerae | 2.32 | Blautia | wexlerae | 2.18 | A | 1 | 1 | 0 | Blautia wexlerae | Blautia wexlerae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31264 | 3; Public strain collection | eDT | 2021-10-20 | Coprococcus | comes | Coprococcus | comes | 2.33 | Coprococcus | comes | 1.72 | A | 1 | 1 | 0 | Coprococcus comes | Coprococcus comes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.17.21 | 3; sequenced [0207] | eDT | 2021-10-20 | Coprococcus | comes | Coprococcus | comes | 2.28 | Coprococcus | comes | 1.72 | A | 1 | 1 | 0 | Coprococcus comes | Coprococcus comes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.10.6 | 3; sequenced [0207] | EFEx | 2021-10-20 | Coprococcus | eutactus | Coprococcus | eutactus | 1.66 | Staphylococcus | schleiferi | 1.4 | B | 0 | 0 | 1 | Coprococcus eutactus | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31256 | 3; Public strain collection | eDT | 0 | Dorea | formicigenerans | Dorea | formicigenerans | 2.26 | Dorea | phocaeensis | 1.79 | A | 1 | 1 | 0 | Dorea formicigenerans | Dorea phocaeensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.19.39 | 3; sequenced [0207] | eDT | 0 | Dorea | formicigenerans | Dorea | formicigenerans | 2.15 | Dorea | phocaeensis | 1.8 | A | 1 | 1 | 0 | Dorea formicigenerans | Dorea phocaeensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.18.49 | 3; sequenced [0207] | eDT | 0 | Dorea | longicatena | Dorea | longicatena | 2.06 | Lactobacillus | helveticus | 1.37 | A | 1 | 1 | 0 | Dorea longicatena | Lactobacillus helveticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108191 | 3; Public strain collection | EFEx | 0 | Enterocloster | clostridioformis | Enterocloster | clostridioformis | 2.34 | Enterocloster | clostridioformis | 2.32 | A | 1 | 1 | 0 | Enterocloster clostridioformis | Enterocloster clostridioformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.20.55 | 3; sequenced [0207] | eDT | 0 | Faecalicatena | fissicatena | Faecalicatena | fissicatena_II | 2.17 | Stenotrophomonas | maltophilia | 1.35 | A | 0 | 0 | 1 | Faecalicatena fissicatena_II | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.9.3 | 3; sequenced [0207] | eDT | 0 | Faecalicatena | fissicatena | Faecalicatena | fissicatena_II | 2.35 | Metabacillus | indicus | 1.25 | A | 0 | 0 | 1 | Faecalicatena fissicatena_II | Metabacillus indicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.14.54 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter | saccharivorans | Fusicatenibacter | saccharivorans | 2.45 | Serratia | rubidaea | 1.35 | A | 1 | 1 | 0 | Fusicatenibacter saccharivorans | Serratia rubidaea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.15.44 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter | saccharivorans | Fusicatenibacter | saccharivorans | 2.52 | Pseudomonas | antarctica | 1.38 | A | 1 | 1 | 0 | Fusicatenibacter saccharivorans | Pseudomonas antarctica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H260 | 3; sequenced [0207] | EFEx | 0 | Lachnospira | eligens | Lachnospira | eligens | 2.39 | Lachnospira | eligens | 2.18 | A | 1 | 1 | 0 | Lachnospira eligens | Lachnospira eligens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105336 | 3; Public strain collection | EFEx | 0 | Lacrimispora | celerecrescens | Lacrimispora | celerecrescens | 2.33 | Lacrimispora | sphenoides | 2.18 | B | 0 | 0 | 1 | Lacrimispora celerecrescens | Lacrimispora sphenoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H164 | 3; sequenced [0207] | EFEx | 0 | Roseburia | faecis | Roseburia | faecis | 2.17 | Roseburia | faecis | 1.86 | A | 1 | 1 | 0 | Roseburia faecis | Roseburia faecis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16839 | 3; type strain | EFEx | 0 | Roseburia | hominis | Roseburia | hominis | 2.36 | Roseburia | hominis | 2.14 | A | 1 | 1 | 0 | Roseburia hominis | Roseburia hominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108070 | 3; Public strain collection | EFEx | 0 | Roseburia | inulinivorans | Roseburia | inulinivorans | 2.3 | Roseburia | inulinvorans | 2.28 | B | 0 | 0 | 1 | Roseburia inulinivorans | Roseburia inulinvorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.1.25 | 3; sequenced [0207] | eDT | 0 | Sellimonas | intestinalis | Sellimonas | intestinalis | 2.11 | Magnusiomyces | capitatus | 1.4 | A | 1 | 1 | 0 | Sellimonas intestinalis | Magnusiomyces capitatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31255 | 3; Public strain collection | eDT | 0 | Anaerotruncus | colihominis | Anaerotruncus | colihominis | 2.38 | Anaerotruncus | colihominis | 2.24 | A | 1 | 1 | 0 | Anaerotruncus colihominis | Anaerotruncus colihominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H281 | 3; sequenced [0207] | EFEx | 0 | Faecalibacterium | prausnitzii | Faecalibacterium | prausnitzii | 2.06 | Faecalibacterium | prausnitzii | 1.86 | A | 1 | 1 | 0 | Faecalibacterium prausnitzii | Faecalibacterium prausnitzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H271 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | bromii | Ruminococcus | bromii | 1.98 | Clostridium | budayi | 1.31 | B | 0 | 0 | 1 | Ruminococcus bromii | Clostridium budayi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.22.53 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | gnavus | Ruminococcus | gnavus | 2.13 | Ruminococcus | gnavus | 1.97 | A | 1 | 1 | 0 | Ruminococcus gnavus | Ruminococcus gnavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK.23.60 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | gnavus | Ruminococcus | gnavus | 2.13 | Ruminococcus | gnavus | 2.07 | A | 1 | 1 | 0 | Ruminococcus gnavus | Ruminococcus gnavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H279 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus | lactaris | Ruminococcus | lactaris | 2.42 | Clostridioides | difficile | 1.32 | A | 1 | 1 | 0 | Ruminococcus lactaris | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H261 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus | torques | Ruminococcus | torques | 2.26 | Staphylococcus | arlettae | 1.46 | A | 1 | 1 | 0 | Ruminococcus torques | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium | bifermentans | Paraclostridium | bifermentans | 2.02 | Paraclostridium | bifermentans | 1.92 | A | 1 | 1 | 0 | Paraclostridium bifermentans | Paraclostridium bifermentans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix | anatis | Erysipelothrix | anatis | 2.41 | Erysipelothrix | anatis | 2.35 | A | 1 | 1 | 0 | Erysipelothrix anatis | Erysipelothrix anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix | anatis | Erysipelothrix | anatis | 2.65 | Erysipelothrix | anatis | 2.43 | A | 1 | 1 | 0 | Erysipelothrix anatis | Erysipelothrix anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix | aquatica | Erysipelothrix | aquatica | 2.48 | Erysipelothrix | aquatica | 2.43 | A | 1 | 1 | 0 | Erysipelothrix aquatica | Erysipelothrix aquatica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | VA92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | aquatica | Erysipelothrix | larvae | 2.68 | Erysipelothrix | larvae | 2.15 | A | 0 | 0 | 1 | Erysipelothrix larvae | Erysipelothrix larvae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | larvae | Erysipelothrix | larvae | 2.55 | Erysipelothrix | larvae | 2.18 | A | 1 | 1 | 0 | Erysipelothrix larvae | Erysipelothrix larvae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 110099 | 3; type strain | EFex [BRU03] | 2021-10-18 | Erysipelothrix | piscisicarius | Erysipelothrix | piscisicarius | 2.6 | Erysipelothrix | piscisicarius | 2.15 | A | 1 | 1 | 0 | Erysipelothrix piscisicarius | Erysipelothrix piscisicarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.7 | Erysipelothrix | rhusiopathiae | 2.54 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.25 | Erysipelothrix | rhusiopathiae | 2.22 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix | tonsillarum | Erysipelothrix | tonsillarum | 2.62 | Erysipelothrix | tonsillarum | 2.55 | A | 1 | 1 | 0 | Erysipelothrix tonsillarum | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-23 | Erysipelothrix | tonsillarum | Erysipelothrix | tonsillarum | 2.79 | Erysipelothrix | tonsillarum | 2.55 | A | 1 | 1 | 0 | Erysipelothrix tonsillarum | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus | sp-CVUAS-10169 | Anaerococcus | sp-CVUAS-10169 | 2.26 | Pseudomonas | putida | 1.34 | A | 1 | 1 | 0 | Anaerococcus sp-CVUAS-10169 | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32948 | 3; confirmed proficency test | EFEx | 2021-02-18 | Finegoldia | magna | Finegoldia | magna | 2.26 | Finegoldia | magna | 2.08 | A | 1 | 1 | 0 | Finegoldia magna | Finegoldia magna | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.25 | Helcococcus | kunzii | 2.25 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.45 | Helcococcus | kunzii | 2.4 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 433.2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus | ovis | Helcococcus | ovis | 2.41 | Helcococcus | ovis | 2.39 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus | ovis | Helcococcus | ovis | 2.28 | Helcococcus | ovis | 2.17 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus | sp-CVUAS-32321 | Helcococcus | sp-CVUAS-32321 | 2.67 | Helcococcus | sueciensis | 1.77 | A | 1 | 1 | 0 | Helcococcus sp-CVUAS-32321 | Helcococcus sueciensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 692.2 | 3; sequenced | DT | 2013-02-19 | Parvimonas | sp-CVUAS-692,2 | Parvimonas | sp-CVUAS-692,2 | 2.15 | Actinomyces | bovis | 1.44 | A | 1 | 1 | 0 | Parvimonas sp-CVUAS-692,2 | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides | finegoldii | Bacteroides | finegoldii | 2.6 | Bacteroides | finegoldii | 2.53 | A | 1 | 1 | 0 | Bacteroides finegoldii | Bacteroides finegoldii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides | fragilis | Bacteroides | fragilis | 2.53 | Bacteroides | fragilis | 2.42 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32947 | 3; confirmed proficency test | EFEx | 2021-02-18 | Bacteroides | fragilis | Bacteroides | fragilis | 2.35 | Bacteroides | fragilis | 2.22 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181001478 | 2; 0 | DT | 2018-02-02 | Bacteroides | pyogenes | Bacteroides | pyogenes | 2.48 | Bacteroides | pyogenes | 2.48 | A | 1 | 1 | 0 | Bacteroides pyogenes | Bacteroides pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella | heparinolytica | Prevotella | heparinolytica | 2.52 | Prevotella | heparinolytica | 2.49 | A | 1 | 1 | 0 | Prevotella heparinolytica | Prevotella heparinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium | chaponense | Chryseobacterium | chaponense | 2.46 | Chryseobacterium | chaponense | 1.81 | A | 1 | 1 | 0 | Chryseobacterium chaponense | Chryseobacterium chaponense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium | gleum | Chryseobacterium | gleum | 1.94 | Chryseobacterium | gleum | 1.9 | B | 0 | 0 | 1 | Chryseobacterium gleum | Chryseobacterium gleum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium | hominis | Chryseobacterium | hominis | 2.55 | Chryseobacterium | hominis | 2.08 | A | 1 | 1 | 0 | Chryseobacterium hominis | Chryseobacterium hominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 2.1 | Chryseobacterium | tructae | 2.07 | A | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium tructae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 1.85 | Chryseobacterium | indologenes | 1.8 | B | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium indologenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium | manosquense | Chryseobacterium | manosquense | 2.22 | Deinococcus | wulumuqiensis | 1.45 | A | 1 | 1 | 0 | Chryseobacterium manosquense | Deinococcus wulumuqiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 185-U1-6 | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium | manosquense | Chryseobacterium | manosquense | 2.13 | Staphylococcus | simiae | 1.28 | A | 1 | 1 | 0 | Chryseobacterium manosquense | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11450.4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium | sp-I-7 | Chryseobacterium | sp-I-7 | 2.7 | Chryseobacterium | geocarposphaerae | 1.42 | A | 1 | 1 | 0 | Chryseobacterium sp-I-7 | Chryseobacterium geocarposphaerae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | aquidurense | Flavobacterium | aquidurense | 2.54 | Flavobacterium | frigidimaris | 1.97 | A | 1 | 1 | 0 | Flavobacterium aquidurense | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5322.2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | branchiarum | Flavobacterium | branchiarum | 2.67 | Flavobacterium | pectinovorum | 1.94 | A | 1 | 1 | 0 | Flavobacterium branchiarum | Flavobacterium pectinovorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | chungangense | Flavobacterium | chungangense | 2.41 | Flavobacterium | plurextorum | 1.82 | A | 1 | 1 | 0 | Flavobacterium chungangense | Flavobacterium plurextorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium | columnare | Flavobacterium | columnare | 2.39 | Flavobacterium | columnare | 2.3 | A | 1 | 1 | 0 | Flavobacterium columnare | Flavobacterium columnare | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.41 | Flavobacterium | frigidimaris | 2.04 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4217.2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.6 | Flavobacterium | frigidimaris | 1.94 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium frigidimaris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium | hibernum | Flavobacterium | hibernum | 2 | Flavobacterium | saccharophilum | 1.97 | A | 1 | 1 | 0 | Flavobacterium hibernum | Flavobacterium saccharophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium | piscis | Flavobacterium | piscis | 2.78 | Flavobacterium | piscis | 2.29 | A | 1 | 1 | 0 | Flavobacterium piscis | Flavobacterium piscis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium | plurextorum | Flavobacterium | plurextorum | 2.62 | Flavobacterium | plurextorum | 2.35 | A | 1 | 1 | 0 | Flavobacterium plurextorum | Flavobacterium plurextorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5323.2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | psychoterrae | Flavobacterium | psychoterrae | 2.41 | Flavobacterium | hydatis | 1.81 | A | 1 | 1 | 0 | Flavobacterium psychoterrae | Flavobacterium hydatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11550.4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.52 | Flavobacterium | psychrophilum | 2.4 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.3 | Flavobacterium | psychrophilum | 2.23 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | reichenbachii | Flavobacterium | reichenbachii | 2 | Flavobacterium | reichenbachii | 1.9 | A | 1 | 1 | 0 | Flavobacterium reichenbachii | Flavobacterium reichenbachii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium | sp | Flavobacterium | saccharophilum | 2.04 | Flavobacterium | hibernum | 1.97 | A | 0 | 0 | 1 | Flavobacterium saccharophilum | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | sp-C2 | Flavobacterium | sp-C2 | 2.4 | Paraburkholderia | phenazinium | 1.57 | A | 1 | 1 | 0 | Flavobacterium sp-C2 | Paraburkholderia phenazinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-10323 | Flavobacterium | sp-CVUAS-10323 | 2.31 | Flavobacterium | chungbukense | 1.65 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-10323 | Flavobacterium chungbukense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium | sp-CVUAS-4146 | Flavobacterium | sp-CVUAS-4146 | 2.25 | Flavobacterium | branchiarum | 2.07 | B | 0 | 0 | 1 | Flavobacterium sp-CVUAS-4146 | Flavobacterium branchiarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4216.2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-4216,2 | Flavobacterium | sp-CVUAS-4216,2 | 2.07 | Flavobacterium | sp-JQMS | 1.47 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-4216,2 | Flavobacterium sp-JQMS | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-5506 | Flavobacterium | sp-CVUAS-5506 | 2.48 | Flavobacterium | johnsoniae | 1.82 | A | 1 | 1 | 0 | Flavobacterium sp-CVUAS-5506 | Flavobacterium johnsoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | sp-HMD2009 | Flavobacterium | sp-HMD2009 | 2.45 | Flavobacterium | succinicans | 1.69 | A | 1 | 1 | 0 | Flavobacterium sp-HMD2009 | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium | sp-JQMS | Flavobacterium | sp-JQMS | 2.59 | Flavobacterium | succinicans | 1.67 | A | 1 | 1 | 0 | Flavobacterium sp-JQMS | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33726 | 3; sequenced | EFEx | 2023-11-07 | Flavobacterium | sp-T157L | Flavobacterium | sp-T157L | 2.61 | Flavobacterium | succinicans | 1.6 | A | 1 | 1 | 0 | Flavobacterium sp-T157L | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium | succinicans | Flavobacterium | succinicans | 2.49 | Flavobacterium | tiangeerense | 1.78 | A | 1 | 1 | 0 | Flavobacterium succinicans | Flavobacterium tiangeerense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | tiangeerense | Flavobacterium | tiangeerense | 2.61 | Flavobacterium | succinicans | 2.14 | B | 0 | 0 | 1 | Flavobacterium tiangeerense | Flavobacterium succinicans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 82 | 2; confirmed proficency test | EFEx | 2021-09-10 | Bergeyella | zoohelcum | Bergeyella | zoohelcum | 2.14 | Bergeyella | zoohelcum | 2.14 | A | 1 | 1 | 0 | Bergeyella zoohelcum | Bergeyella zoohelcum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2589.2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas | tenax | Epilithonimonas | tenax | 2.34 | Chryseobacterium | aquaticum | 1.4 | A | 1 | 1 | 0 | Epilithonimonas tenax | Chryseobacterium aquaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.27 | Ornithobacterium | rhinotracheale | 2.24 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.21 | Ornithobacterium | rhinotracheale | 2.04 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella | columbina | Riemerella | columbina | 2.56 | Riemerella | columbina | 2.06 | A | 1 | 1 | 0 | Riemerella columbina | Riemerella columbina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34092 | 3; sequenced | EFex | 2022-08-05 | Weeksellaceae | sp-CVUAS-34092 | Weeksellaceae | sp-CVUAS-34092 | 2.47 | Empedobacter | falsenii | 1.39 | A | 1 | 1 | 0 | Weeksellaceae sp-CVUAS-34092 | Empedobacter falsenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9288.2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca | zopfii | Prototheca | zopfii | 2.24 | Prototheca | zopfii | 2.19 | A | 1 | 1 | 0 | Prototheca zopfii | Prototheca zopfii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9875.2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca | zopfii | Prototheca | zopfii | 2.54 | Prototheca | zopfii | 2.53 | A | 1 | 1 | 0 | Prototheca zopfii | Prototheca zopfii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus | sp-CVUAS-11266 | Deinococcus | sp-CVUAS-11266 | 2.76 | Acidovorax | facilis | 1.23 | A | 1 | 1 | 0 | Deinococcus sp-CVUAS-11266 | Acidovorax facilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium | necrophorum | Fusobacterium | necrophorum | 2.48 | Fusobacterium | necrophorum | 2.41 | A | 1 | 1 | 0 | Fusobacterium necrophorum | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium | varium | Fusobacterium | varium | 2.09 | Fusobacterium | varium | 2.04 | A | 1 | 1 | 0 | Fusobacterium varium | Fusobacterium varium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter | abscessus | Caviibacter | abscessus | 2.32 | Caviibacter | abscessus | 2.18 | A | 1 | 1 | 0 | Caviibacter abscessus | Caviibacter abscessus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga | salmonicida | Oceanivirga | salmonicida | 2.26 | Oceanivirga | salmonicida | 2.03 | A | 1 | 1 | 0 | Oceanivirga salmonicida | Oceanivirga salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus | hongkongensis | Pseudostreptobacillus | hongkongensis | 2.21 | Pseudostreptobacillus | hongkongensis | 1.81 | A | 1 | 1 | 0 | Pseudostreptobacillus hongkongensis | Pseudostreptobacillus hongkongensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella | termitidis | Sebaldella | termitidis | 2.16 | Sebaldella | termitidis | 1.65 | A | 1 | 1 | 0 | Sebaldella termitidis | Sebaldella termitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella | termitidis | Sebaldella | termitidis | 2.47 | Sebaldella | termitidis | 1.41 | A | 1 | 1 | 0 | Sebaldella termitidis | Sebaldella termitidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 38322 | 3; public strain collection | EFEx | 2015-12-16 | Sneathia | sanguinegens | Sneathia | sanguinegens | 2.76 | Sneathia | sanguinegens | 2.27 | A | 1 | 1 | 0 | Sneathia sanguinegens | Sneathia sanguinegens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCUG 41628 | 3; type strain | EFEx | 2020-10-08 | Sneathia | sanguinegens | Sneathia | sanguinegens | 2.49 | Sneathia | sanguinegens | 2.45 | A | 1 | 1 | 0 | Sneathia sanguinegens | Sneathia sanguinegens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Sneathia | vaginalis | 2.61 | Sneathia | vaginalis | 2.6 | A | 1 | 1 | 0 | Sneathia vaginalis | Sneathia vaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Sneathia | vaginalis | 2.57 | Sneathia | vaginalis | 2.56 | A | 1 | 1 | 0 | Sneathia vaginalis | Sneathia vaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus | felis | Streptobacillus | felis | 2.52 | Streptobacillus | felis | 2.5 | A | 1 | 1 | 0 | Streptobacillus felis | Streptobacillus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.58 | Streptobacillus | moniliformis | 2.57 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.47 | Streptobacillus | moniliformis | 2.43 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus | notomytis | Streptobacillus | notomytis | 2.46 | Streptobacillus | notomytis | 2.11 | A | 1 | 1 | 0 | Streptobacillus notomytis | Streptobacillus notomytis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus | ratti | Streptobacillus | ratti | 2.4 | Pseudostreptobacillus | hongkongensis | 1.43 | A | 1 | 1 | 0 | Streptobacillus ratti | Pseudostreptobacillus hongkongensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2280.3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas | glaciistagni | Aureimonas | glaciistagni | 2.53 | Bacillus | thuringiensis | 1.31 | A | 1 | 1 | 0 | Aureimonas glaciistagni | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 35685 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-23 | Bartonella | bacilliformis | Bartonella | bacilliformis | 2.69 | Bartonella | elizabethae | 2.1 | B | 0 | 0 | 1 | Bartonella bacilliformis | Bartonella elizabethae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 700132 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | grahamii | Bartonella | grahamii | 2.77 | Bartonella | queenslandensis | 1.9 | A | 1 | 1 | 0 | Bartonella grahamii | Bartonella queenslandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 49882 DSM 28221 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | henselae | Bartonella | henselae | 2.24 | Bartonella | henselae | 2.21 | A | 1 | 1 | 0 | Bartonella henselae | Bartonella henselae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BK8 | 2; strain collection of the RKI | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2022-05-10 | Bartonella | henselae | Bartonella | henselae | 2.32 | Schleiferilactobacillus | harbinensis | 1.34 | A | 1 | 1 | 0 | Bartonella henselae | Schleiferilactobacillus harbinensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CIP 103739 | 3; public strain collection | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | quintana | Bartonella | quintana | 2.67 | Lentilactobacillus | diolivorans | 1.55 | A | 1 | 1 | 0 | Bartonella quintana | Lentilactobacillus diolivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 700727 DSM 28218 | 3; type strain | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella | vinsonii | Bartonella | vinsonii | 2.16 | Bartonella | bacilliformis | 1.3 | A | 1 | 1 | 0 | Bartonella vinsonii | Bartonella bacilliformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM373 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | abortus | Brucella | abortus | 2.65 | Brucella | melitensis | 2.52 | B | 0 | 0 | 1 | Brucella abortus | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM369 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | abortus | Brucella | abortus | 2.62 | Brucella | melitensis | 2.49 | B | 0 | 0 | 1 | Brucella abortus | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-11 | Brucella | canis | Brucella | canis | 2.58 | Brucella | canis | 2.53 | A | 1 | 1 | 0 | Brucella canis | Brucella canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-03-05 | Brucella | canis | Brucella | canis | 2.56 | Brucella | melitensis | 2.47 | B | 0 | 0 | 1 | Brucella canis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 007 1 002 34 | 2; 0 | EFEx | 2013-06-20 | Brucella | ceti | Brucella | ceti | 2.83 | Brucella | ceti | 2.58 | A | 1 | 1 | 0 | Brucella ceti | Brucella ceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM389 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | ceti | Brucella | ceti | 2.75 | Brucella | ceti | 2.55 | A | 1 | 1 | 0 | Brucella ceti | Brucella ceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 10 1 001 44 | 2; 0 | EFEx | 2013-06-20 | Brucella | inopinata | Brucella | inopinata | 2.79 | Brucella | canis | 2.37 | B | 0 | 0 | 1 | Brucella inopinata | Brucella canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A1309 | 2; 0 | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 28.07.2021¦ 1:1 HCCA¦Andy 3.0 | 2021-07-28 | Brucella | intermedia | Brucella | intermedia | 2.87 | Ochrobactrum | intermedium | 2.43 | C | 0 | 0 | 1 | Brucella intermedia | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM374 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | melitensis | Brucella | melitensis | 2.77 | Brucella | melitensis | 2.7 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM407 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | melitensis | Brucella | melitensis | 2.75 | Brucella | melitensis | 2.72 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM408 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | melitensis | Brucella | melitensis | 2.73 | Brucella | melitensis | 2.62 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM410 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | melitensis | Brucella | melitensis | 2.69 | Brucella | melitensis | 2.66 | A | 1 | 1 | 0 | Brucella melitensis | Brucella melitensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 4915 CVUAS | 3; type strain | DT | 2013-11-26 | Brucella | microti | Brucella | microti | 2.67 | Brucella | sp. | 2.39 | B | 0 | 0 | 1 | Brucella microti | Brucella sp. | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 4915 RKI | 3; type strain | TFAextr | 2015-03-17 | Brucella | microti | Brucella | abortus | 2.57 | Brucella | microti | 2.52 | B | 0 | 0 | 1 | Brucella abortus | Brucella microti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-20 | Brucella | neotomae | Brucella | neotomae | 2.76 | Brucella | neotomae | 2.52 | A | 1 | 1 | 0 | Brucella neotomae | Brucella neotomae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM421 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | neotomae | Brucella | neotomae | 2.64 | Brucella | neotomae | 2.37 | A | 1 | 1 | 0 | Brucella neotomae | Brucella neotomae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 06 1 007 32 | 2; 0 | EFEx | 2013-06-20 | Brucella | ovis | Brucella | ovis | 2.8 | Brucella | ovis | 2.64 | A | 1 | 1 | 0 | Brucella ovis | Brucella ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-11 | Brucella | ovis | Brucella | ovis | 2.83 | Brucella | ovis | 2.53 | A | 1 | 1 | 0 | Brucella ovis | Brucella ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-26 | Brucella | pinnipedialis | Brucella | pinnipedialis | 2.76 | Brucella | pinnipedialis | 2.58 | A | 1 | 1 | 0 | Brucella pinnipedialis | Brucella pinnipedialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM405 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | pinnipedialis | Brucella | pinnipedialis | 2.71 | Brucella | ceti | 2.55 | B | 0 | 0 | 1 | Brucella pinnipedialis | Brucella ceti | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM431 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | suis | Brucella | suis | 2.7 | Brucella | suis | 2.54 | A | 1 | 1 | 0 | Brucella suis | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM427 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella | suis | Brucella | suis | 2.74 | Brucella | abortus | 2.68 | B | 0 | 0 | 1 | Brucella suis | Brucella abortus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella | suis | Brucella | suis | 2.62 | Brucella | suis | 2.04 | A | 1 | 1 | 0 | Brucella suis | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM428 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella | suis | Brucella | suis | 2.7 | Brucella | canis | 2.57 | B | 0 | 0 | 1 | Brucella suis | Brucella canis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM430 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-15 | Brucella | suis | Brucella | suis | 2.76 | Brucella | suis | 2.46 | A | 1 | 1 | 0 | Brucella suis | Brucella suis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2016-10-15 | Brucella | vulpis | Brucella | vulpis | 2.78 | Brucella | canis | 2.16 | B | 0 | 0 | 1 | Brucella vulpis | Brucella canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum | ovis | Falsochrobactrum | ovis | 2.48 | Falsochrobactrum | ovis | 2.41 | A | 1 | 1 | 0 | Falsochrobactrum ovis | Falsochrobactrum ovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_08 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2012-04-02 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.42 | Ochrobactrum | anthropi | 2.39 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2016-06-03 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.54 | Ochrobactrum | anthropi | 2.42 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.55 | Brucella | intermedia | 2.54 | C | 0 | 0 | 1 | Ochrobactrum intermedium | Brucella intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32748 | 3; sequenced | EFex | 2020-10-29 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.53 | Brucella | intermedia | 2.28 | C | 0 | 0 | 1 | Ochrobactrum intermedium | Brucella intermedia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum | sp | Ochrobactrum | anthropi | 2.46 | Ochrobactrum | anthropi | 2.41 | A | 0 | 0 | 1 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum | sp | Ochrobactrum | intermedium | 1.78 | Ochrobactrum | anthropi | 1.68 | B | 0 | 0 | 1 | Ochrobactrum intermedium | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012040 | 2; 0 | DT | 2017-10-04 | Ochrobactrum | sp | Ochrobactrum | tritici | 1.78 | Ochrobactrum | intermedium | 1.65 | B | 0 | 0 | 1 | Ochrobactrum tritici | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter | aganoensis | Aminobacter | aganoensis | 2.39 | Pseudomonas | mucidolens | 1.45 | A | 1 | 1 | 0 | Aminobacter aganoensis | Pseudomonas mucidolens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter | aminovorans | Aminobacter | aminovorans | 2.56 | Lactobacillus | curvatus | 1.5 | A | 1 | 1 | 0 | Aminobacter aminovorans | Lactobacillus curvatus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30768 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 16.09.2022¦ 1:1 HCCA¦Andy 3.1#EXT=Paracoccus yeei DSM 30768, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-10-07 | Paracoccus | yeei | Paracoccus | yeei | 2.37 | Paracoccus | yeei | 2.33 | A | 1 | 1 | 0 | Paracoccus yeei | Paracoccus yeei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-16 | Paracoccus | yeei | Paracoccus | yeei | 2.45 | Paracoccus | yeei | 2.24 | A | 1 | 1 | 0 | Paracoccus yeei | Paracoccus yeei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 103800 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 25.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Roseomonas gilardii DSM 103800, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-04-06 | Roseomonas | gilardii | Roseomonas | gilardii | 2.79 | Roseomonas | mucosa | 1.76 | A | 1 | 1 | 0 | Roseomonas gilardii | Roseomonas mucosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas | mucosa | Roseomonas | mucosa | 2.53 | Roseomonas | mucosa | 2.49 | A | 1 | 1 | 0 | Roseomonas mucosa | Roseomonas mucosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2007-01-24 | Inquilinus | limosus | Inquilinus | limosus | 2.63 | Inquilinus | limosus | 2.62 | A | 1 | 1 | 0 | Inquilinus limosus | Inquilinus limosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1098 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Sphingomonas paucimobilis DSM 30198, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Sphingomonas | paucimobilis | Sphingomonas | paucimobilis | 2.5 | Sphingomonas | paucimobilis | 2.42 | A | 1 | 1 | 0 | Sphingomonas paucimobilis | Sphingomonas paucimobilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30198 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Sphingomonas paucimobilis DSM 30198, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Sphingomonas | paucimobilis | Sphingomonas | paucimobilis | 2.53 | Sphingomonas | paucimobilis | 2.45 | A | 1 | 1 | 0 | Sphingomonas paucimobilis | Sphingomonas paucimobilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter | denitrificans | Achromobacter | denitrificans | 2.46 | Achromobacter | denitrificans | 2.22 | A | 1 | 1 | 0 | Achromobacter denitrificans | Achromobacter denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter | sp-DSM-30031 | Achromobacter | sp-DSM-30031 | 2.22 | Achromobacter | piechaudii | 1.95 | A | 1 | 1 | 0 | Achromobacter sp-DSM-30031 | Achromobacter piechaudii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-11-18 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.33 | Achromobacter | xylosoxidans | 2.16 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-24 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.26 | Achromobacter | xylosoxidans | 2.23 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30030 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Alcaligenes faecalis ssp. parafaecalis DSM 13975, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-29 | Alcaligenes | faecalis | Alcaligenes | faecalis | 2.5 | Alcaligenes | faecalis | 2.37 | A | 1 | 1 | 0 | Alcaligenes faecalis | Alcaligenes faecalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13975 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Alcaligenes faecalis subsp. Parafaecalis DSM 13975, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Alcaligenes | faecalis | Alcaligenes | faecalis | 2.5 | Alcaligenes | faecalis | 2.49 | A | 1 | 1 | 0 | Alcaligenes faecalis | Alcaligenes faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea | sp-CVUAS-31213 | Basilea | sp-CVUAS-31213 | 2.18 | Streptomyces | badius | 1.35 | A | 1 | 1 | 0 | Basilea sp-CVUAS-31213 | Streptomyces badius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella | avium | Bordetella | avium | 2.39 | Bordetella | avium | 2.29 | A | 1 | 1 | 0 | Bordetella avium | Bordetella avium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 21-2 | 2; confirmed proficency test | eDT | 2022-02-14 | Bordetella | trematum | Bordetella | trematum | 2.42 | Bordetella | trematum | 2.38 | A | 1 | 1 | 0 | Bordetella trematum | Bordetella trematum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32946 | 3; confirmed proficency test | EFEx | 2021-02-18 | Kerstersia | gyiorum | Kerstersia | gyiorum | 2.48 | Kerstersia | gyiorum | 2.48 | A | 1 | 1 | 0 | Kerstersia gyiorum | Kerstersia gyiorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 2.56 | Oligella | urethralis | 2.43 | A | 1 | 1 | 0 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 2.51 | Oligella | urethralis | 2.39 | A | 1 | 1 | 0 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga | kullae | Pigmentiphaga | kullae | 2.24 | Pigmentiphaga | daeguensis | 1.67 | A | 1 | 1 | 0 | Pigmentiphaga kullae | Pigmentiphaga daeguensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.12 | Taylorella | asinigenitalis | 2.09 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella asinigenitalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.32 | Taylorella | asinigenitalis | 2.01 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella asinigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.15 | Taylorella | equigenitalis | 2.1 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.35 | Taylorella | equigenitalis | 2.24 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.27 | Burkholderia | ambifaria | 2.24 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16087 LMG 19182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.36 | Burkholderia | diffusa | 2.32 | B | 0 | 0 | 1 | Burkholderia ambifaria | Burkholderia diffusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | anthina | Burkholderia | anthina | 2.42 | Burkholderia | anthina | 2.35 | A | 1 | 1 | 0 | Burkholderia anthina | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-08-06 | Burkholderia | anthina | Burkholderia | anthina | 2.79 | Burkholderia | pyrrocinia | 2.17 | B | 0 | 0 | 1 | Burkholderia anthina | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-18 | Burkholderia | arboris | Burkholderia | arboris | 2.81 | Burkholderia | puraquae | 2.37 | B | 0 | 0 | 1 | Burkholderia arboris | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | arboris | Burkholderia | arboris | 2.82 | Burkholderia | arboris | 2.33 | A | 1 | 1 | 0 | Burkholderia arboris | Burkholderia arboris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.5 | Burkholderia | cenocepacia | 2.48 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-07-02 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.52 | Burkholderia | cenocepacia | 2.42 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-05-04 | Burkholderia | cepacia | Burkholderia | cenocepacia | 2.71 | Burkholderia | cenocepacia | 2.46 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-09 | Burkholderia | cepacia | Burkholderia | cepacia | 2.73 | Burkholderia | cepacia | 2.4 | A | 1 | 1 | 0 | Burkholderia cepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-03 | Burkholderia | contaminans | Burkholderia | contaminans | 2.65 | Burkholderia | contaminans | 2.62 | A | 1 | 1 | 0 | Burkholderia contaminans | Burkholderia contaminans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 01 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-19 | Burkholderia | contaminans | Burkholderia | contaminans | 2.62 | Burkholderia | cepacia | 2.42 | B | 0 | 0 | 1 | Burkholderia contaminans | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-31 | Burkholderia | diffusa | Burkholderia | diffusa | 2.73 | Burkholderia | ambifaria | 2.35 | B | 0 | 0 | 1 | Burkholderia diffusa | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | diffusa | Burkholderia | diffusa | 2.75 | Burkholderia | territorii | 2.44 | B | 0 | 0 | 1 | Burkholderia diffusa | Burkholderia territorii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | dolosa | Burkholderia | dolosa | 2.46 | Burkholderia | dolosa | 2.29 | A | 1 | 1 | 0 | Burkholderia dolosa | Burkholderia dolosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-17 | Burkholderia | dolosa | Burkholderia | dolosa | 2.33 | Burkholderia | dolosa | 2.08 | A | 1 | 1 | 0 | Burkholderia dolosa | Burkholderia dolosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-24 | Burkholderia | gladioli | Burkholderia | gladioli | 2.45 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia | gladioli | Burkholderia | gladioli | 2.42 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | glumae | Burkholderia | glumae | 2.62 | Burkholderia | glumae | 2.33 | A | 1 | 1 | 0 | Burkholderia glumae | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-05-04 | Burkholderia | glumae | Burkholderia | glumae | 2.79 | Burkholderia | glumae | 2.33 | A | 1 | 1 | 0 | Burkholderia glumae | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-31 | Burkholderia | lata | Burkholderia | lata | 2.8 | Burkholderia | puraquae | 2.49 | B | 0 | 0 | 1 | Burkholderia lata | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-08-23 | Burkholderia | lata | Burkholderia | lata | 2.83 | Burkholderia | cepacia | 2.47 | B | 0 | 0 | 1 | Burkholderia lata | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-20 | Burkholderia | latens | Burkholderia | latens | 2.71 | Burkholderia | latens | 2.32 | A | 1 | 1 | 0 | Burkholderia latens | Burkholderia latens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-09-04 | Burkholderia | mallei | Burkholderia | mallei | 2.23 | Burkholderia | mallei | 2.19 | A | 1 | 1 | 0 | Burkholderia mallei | Burkholderia mallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-08-24 | Burkholderia | mallei | Burkholderia | pseudomallei | 2.12 | Burkholderia | pseudomallei | 2.01 | A | 0 | 0 | 1 | Burkholderia pseudomallei | Burkholderia pseudomallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-01 | Burkholderia | metallica | Burkholderia | metallica | 2.52 | Burkholderia | cepacia | 2.46 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | metallica | Burkholderia | metallica | 2.77 | Burkholderia | seminalis | 2.35 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia seminalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-13 | Burkholderia | multivorans | Burkholderia | multivorans | 2.22 | Burkholderia | multivorans | 2.21 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-08 | Burkholderia | multivorans | Burkholderia | multivorans | 2.13 | Burkholderia | multivorans | 2 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | oklahomensis | Burkholderia | oklahomensis | 2.72 | Burkholderia | oklahomensis | 2.49 | A | 1 | 1 | 0 | Burkholderia oklahomensis | Burkholderia oklahomensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | oklahomensis | Burkholderia | oklahomensis | 2.73 | Burkholderia | oklahomensis | 2.48 | A | 1 | 1 | 0 | Burkholderia oklahomensis | Burkholderia oklahomensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | plantarii | Burkholderia | plantarii | 2.49 | Burkholderia | glumae | 1.94 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia | plantarii | Burkholderia | plantarii | 2.5 | Burkholderia | glumae | 1.95 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-04 | Burkholderia | pseudomallei | Burkholderia | pseudomallei | 2.34 | Burkholderia | mallei | 2.28 | B | 0 | 0 | 1 | Burkholderia pseudomallei | Burkholderia mallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-01 | Burkholderia | pseudomallei | Burkholderia | pseudomallei | 2.49 | Burkholderia | pseudomallei | 2.47 | A | 1 | 1 | 0 | Burkholderia pseudomallei | Burkholderia pseudomallei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-21 | Burkholderia | pseudomultivorans | Burkholderia | pseudomultivorans | 2.75 | Burkholderia | ubonensis | 2.1 | B | 0 | 0 | 1 | Burkholderia pseudomultivorans | Burkholderia ubonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-21 | Burkholderia | pseudomultivorans | Burkholderia | pseudomultivorans | 2.79 | Burkholderia | pseudomultivorans | 2.28 | A | 1 | 1 | 0 | Burkholderia pseudomultivorans | Burkholderia pseudomultivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CAMPA 1043 (HST 249) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-01 | Burkholderia | puraquae | Burkholderia | puraquae | 2.73 | Burkholderia | puraquae | 2.7 | A | 1 | 1 | 0 | Burkholderia puraquae | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | R-50214 (BP 78) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-18 | Burkholderia | puraquae | Burkholderia | puraquae | 2.8 | Burkholderia | puraquae | 2.79 | A | 1 | 1 | 0 | Burkholderia puraquae | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | pyrrocinia | Burkholderia | pyrrocinia | 2.57 | Burkholderia | puraquae | 2.45 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia | pyrrocinia | Burkholderia | diffusa | 2.18 | Burkholderia | puraquae | 2.14 | B | 0 | 0 | 1 | Burkholderia diffusa | Burkholderia puraquae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia | seminalis | Burkholderia | seminalis | 2.79 | Burkholderia | seminalis | 2.41 | A | 1 | 1 | 0 | Burkholderia seminalis | Burkholderia seminalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-11 | Burkholderia | seminalis | Burkholderia | seminalis | 2.35 | Burkholderia | cepacia | 2.29 | B | 0 | 0 | 1 | Burkholderia seminalis | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia | stabilis | Burkholderia | stabilis | 2.51 | Burkholderia | stabilis | 2.44 | A | 1 | 1 | 0 | Burkholderia stabilis | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-19 | Burkholderia | stabilis | Burkholderia | stabilis | 2.28 | Burkholderia | pyrrocinia | 2.21 | B | 0 | 0 | 1 | Burkholderia stabilis | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-16 | Burkholderia | stagnalis | Burkholderia | stagnalis | 2.79 | Burkholderia | stagnalis | 2.65 | A | 1 | 1 | 0 | Burkholderia stagnalis | Burkholderia stagnalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia | stagnalis | Burkholderia | stagnalis | 2.82 | Burkholderia | stagnalis | 2.62 | A | 1 | 1 | 0 | Burkholderia stagnalis | Burkholderia stagnalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Burkholderia | taxonK | Burkholderia | contaminans | 2.7 | Burkholderia | contaminans | 2.5 | A | 0 | 0 | 1 | Burkholderia contaminans | Burkholderia contaminans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-26 | Burkholderia | taxonK | Burkholderia | contaminans | 2.57 | Burkholderia | contaminans | 2.55 | A | 0 | 0 | 1 | Burkholderia contaminans | Burkholderia contaminans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia | territorii | Burkholderia | territorii | 2.78 | Burkholderia | cenocepacia | 2.3 | B | 0 | 0 | 1 | Burkholderia territorii | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-16 | Burkholderia | territorii | Burkholderia | territorii | 2.81 | Burkholderia | diffusa | 2.51 | B | 0 | 0 | 1 | Burkholderia territorii | Burkholderia diffusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-02 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.51 | Burkholderia | thailandensis | 2.38 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13276 n | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2010-11-24 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.4 | Burkholderia | thailandensis | 2.34 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia | ubonensis | Burkholderia | ubonensis | 2.82 | Burkholderia | ubonensis | 2.31 | A | 1 | 1 | 0 | Burkholderia ubonensis | Burkholderia ubonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia | ubonensis | Burkholderia | ubonensis | 2.73 | Burkholderia | ubonensis | 2.47 | A | 1 | 1 | 0 | Burkholderia ubonensis | Burkholderia ubonensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9631 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2007-05-09 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.46 | Burkholderia | vietnamiensis | 2.33 | A | 1 | 1 | 0 | Burkholderia vietnamiensis | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10929 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-12-13 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.78 | Burkholderia | vietnamiensis | 2.43 | A | 1 | 1 | 0 | Burkholderia vietnamiensis | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-13 | Caballeronia | glathei | Caballeronia | glathei | 2.55 | Caballeronia | glathei | 2.33 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia glathei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-20 | Caballeronia | glathei | Caballeronia | glathei | 2.47 | Caballeronia | glathei | 2.23 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia glathei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea | apista | Pandoraea | apista | 2.36 | Pandoraea | apista | 2.28 | A | 1 | 1 | 0 | Pandoraea apista | Pandoraea apista | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea | pnomenusa | Pandoraea | pnomenusa | 2.4 | Pandoraea | pnomenusa | 2.33 | A | 1 | 1 | 0 | Pandoraea pnomenusa | Pandoraea pnomenusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea | pulmonicola | Pandoraea | pulmonicola | 2.49 | Pandoraea | pulmonicola | 2.25 | A | 1 | 1 | 0 | Pandoraea pulmonicola | Pandoraea pulmonicola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.52 | Paraburkholderia | caledonica | 2.45 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.46 | Paraburkholderia | caledonica | 2.37 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-13 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.71 | Paraburkholderia | caribensis | 2.25 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-13 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.8 | Paraburkholderia | caribensis | 2.23 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-20 | Paraburkholderia | hospita | Paraburkholderia | hospita | 2.42 | Paraburkholderia | azotifigens | 1.81 | A | 1 | 1 | 0 | Paraburkholderia hospita | Paraburkholderia azotifigens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-20 | Paraburkholderia | hospita | Paraburkholderia | hospita | 2.83 | Paraburkholderia | hospita | 2.07 | A | 1 | 1 | 0 | Paraburkholderia hospita | Paraburkholderia hospita | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-01 | Paraburkholderia | kururiensis | Paraburkholderia | kururiensis | 2.82 | Burkholderia | pyrrocinia | 1.64 | A | 1 | 1 | 0 | Paraburkholderia kururiensis | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-26 | Paraburkholderia | phymatum | Paraburkholderia | phymatum | 2.79 | Paraburkholderia | caribensis | 1.9 | A | 1 | 1 | 0 | Paraburkholderia phymatum | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-19 | Paraburkholderia | phytofirmans | Paraburkholderia | phytofirmans | 2.81 | Paraburkholderia | phytofirmans | 2.44 | A | 1 | 1 | 0 | Paraburkholderia phytofirmans | Paraburkholderia phytofirmans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-20 | Paraburkholderia | tuberum | Paraburkholderia | tuberum | 2.81 | Paraburkholderia | tuberum | 2.36 | A | 1 | 1 | 0 | Paraburkholderia tuberum | Paraburkholderia tuberum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2010-11-26 | Paraburkholderia | xenovorans | Paraburkholderia | xenovorans | 2.77 | Paraburkholderia | xenovorans | 2.49 | A | 1 | 1 | 0 | Paraburkholderia xenovorans | Paraburkholderia xenovorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2016-05-13 | Ralstonia | insidiosa | Ralstonia | insidiosa | 2.38 | Ralstonia | insidiosa | 2.35 | A | 1 | 1 | 0 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2016-05-13 | Ralstonia | mannitolilytica | Ralstonia | mannitolilytica | 2.46 | Ralstonia | mannitolilytica | 2.22 | A | 1 | 1 | 0 | Ralstonia mannitolilytica | Ralstonia mannitolilytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-22 | Ralstonia | pickettii | Ralstonia | pickettii | 2.15 | Ralstonia | pickettii | 2.12 | A | 1 | 1 | 0 | Ralstonia pickettii | Ralstonia pickettii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2009-05-07 | Ralstonia | pickettii | Ralstonia | insidiosa | 2.07 | Ralstonia | insidiosa | 2.07 | A | 0 | 0 | 1 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia | acidovorans | Delftia | acidovorans | 2.54 | Delftia | tsuruhatensis | 2.44 | B | 0 | 0 | 1 | Delftia acidovorans | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia | tsuruhatensis | Delftia | tsuruhatensis | 2.45 | Delftia | tsuruhatensis | 2.31 | A | 1 | 1 | 0 | Delftia tsuruhatensis | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga | intermedia | Hydrogenophaga | intermedia | 2.45 | Hydrogenophaga | intermedia | 2.4 | A | 1 | 1 | 0 | Hydrogenophaga intermedia | Hydrogenophaga intermedia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-07-25 | Xenophilus | sp | Aromatoleum | evansii | 1.43 | Ligilactobacillus | salivarius | 1.32 | B | 0 | 0 | 1 | Aromatoleum evansii | Ligilactobacillus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24713 extr | 3; type strain | Extr | 2016-02-02 | Deefgea | chitinilytica | Deefgea | chitinilytica | 2.72 | Mannheimia | haemolytica | 1.42 | A | 1 | 1 | 0 | Deefgea chitinilytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24713 eDT | 3; type strain | eDT | 2015-09-22 | Deefgea | chitinilytica | Deefgea | chitinilytica | 2.11 | Pichia | cactophila | 1.24 | A | 1 | 1 | 0 | Deefgea chitinilytica | Pichia cactophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea | rivuli | Deefgea | rivuli | 2.6 | Arthrobacter | woluwensis | 1.45 | A | 1 | 1 | 0 | Deefgea rivuli | Arthrobacter woluwensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 08-F17.2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Deefgea | sp-CVUAS-2703.2 | 2.03 | Deefgea | chitinilytica | 1.7 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-2703.2 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Deefgea | sp-CVUAS-2703.2 | 2.08 | Deefgea | chitinilytica | 1.65 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-2703.2 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6995.3 extr | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea | sp-CVUAS-6995,3 | Deefgea | sp-CVUAS-6995.3 | 2.61 | Deefgea | sp-CVUAS-9920 | 1.56 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-6995.3 | Deefgea sp-CVUAS-9920 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-6995,3 | Deefgea | sp-CVUAS-6995.3 | 2.38 | Pseudomonas | orientalis | 1.48 | A | 0 | 0 | 1 | Deefgea sp-CVUAS-6995.3 | Pseudomonas orientalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-9920 | Deefgea | sp-CVUAS-9920 | 2.42 | Deefgea | sp-CVUAS-2703.2 | 1.53 | A | 1 | 1 | 0 | Deefgea sp-CVUAS-9920 | Deefgea sp-CVUAS-2703.2 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea | sp-CVUAS-9920 | Deefgea | sp-CVUAS-9920 | 2.21 | Deefgea | chitinilytica | 1.37 | A | 1 | 1 | 0 | Deefgea sp-CVUAS-9920 | Deefgea chitinilytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter | limnosediminis | Iodobacter | limnosediminis | 2.6 | Iodobacter | fluviatilis | 1.45 | A | 1 | 1 | 0 | Iodobacter limnosediminis | Iodobacter fluviatilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella | kingae | Kingella | kingae | 2.49 | Kingella | kingae | 2.41 | A | 1 | 1 | 0 | Kingella kingae | Kingella kingae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10067.2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria | animaloris | Neisseria | animaloris | 2.41 | Neisseria | animaloris | 2.17 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria | animaloris | Neisseria | animaloris | 2.65 | Neisseria | animaloris | 2.42 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-4931 | Neisseria | sp-CVUAS-4931 | 2.56 | Neisseria | ornithorespiratoria' | 2.46 | B | 0 | 0 | 1 | Neisseria sp-CVUAS-4931 | Neisseria ornithorespiratoria' | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-8997 | Neisseria | sp-CVUAS-8997 | 2.27 | Neisseria_flavescens_subflava | group | 1.48 | A | 1 | 1 | 0 | Neisseria sp-CVUAS-8997 | Neisseria_flavescens_subflava group | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-LHL-1296 | Neisseria | sp-LHL-1296 | 2.59 | Neisseria | animaloris | 1.61 | A | 1 | 1 | 0 | Neisseria sp-LHL-1296 | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5699.3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria | sp-N95-16 | Neisseria | sp-N95-16 | 2.73 | Bergeriella | denitrificans | 1.99 | C | 0 | 0 | 1 | Neisseria sp-N95-16 | Bergeriella denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae | sp-CVUAS-8636 | Neisseriaceae | sp-CVUAS-8636 | 2.14 | Pseudomonas | taetrolens | 1.43 | A | 1 | 1 | 0 | Neisseriaceae sp-CVUAS-8636 | Pseudomonas taetrolens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter | cryaerophilus | Arcobacter | cryaerophilus | 2.03 | Arcobacter | cryaerophilus | 2.01 | A | 0 | 0 | 1 | Arcobacter cryaerophilus | Arcobacter cryaerophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.06 | Aliarcobacter | butzleri | 1.94 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.24 | Aliarcobacter | butzleri | 2.21 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter | coli | Campylobacter | coli | 2.18 | Campylobacter | coli | 1.97 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4689 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter | coli | Campylobacter | coli | 2.24 | Campylobacter | coli | 2.21 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter | fetus | Campylobacter | fetus | 2.34 | Campylobacter | fetus | 2.28 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5361 | 3; type strain | DT | 2013-09-24 | Campylobacter | fetus | Campylobacter | fetus | 2.51 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.37 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.19 | Campylobacter | hyointestinalis | 2.12 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6673.2 | 3; sequenced | DT | 2013-08-28 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.15 | Campylobacter | hyointestinalis | 2.06 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter | jejuni | Campylobacter | jejuni | 2.19 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter | jejuni | Campylobacter | jejuni | 2.39 | Campylobacter | jejuni | 2.32 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter | jejuni | Campylobacter | jejuni | 2.16 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | lari | Campylobacter | lari | 2.19 | Campylobacter | lari | 2.04 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.04 | Campylobacter | lari | 1.82 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.2 | Campylobacter | lari | 2.11 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 11375 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter | lari | Campylobacter | lari | 2.42 | Campylobacter | lari | 2.19 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Campylobacter | sp-W1 | 1.48 | A | 1 | 1 | 0 | Campylobacter peloridis | Campylobacter sp-W1 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-7048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Burkholderia | plantarii | 1.42 | A | 1 | 1 | 0 | Campylobacter peloridis | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter | sputorum | Campylobacter | sputorum | 2.01 | Campylobacter | sputorum | 1.9 | A | 1 | 1 | 0 | Campylobacter sputorum | Campylobacter sputorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter | sp-W1 | Campylobacter | sp-W1 | 2.61 | Campylobacter | jejuni | 1.49 | A | 1 | 1 | 0 | Campylobacter sp-W1 | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter | upsaliensis | Campylobacter | upsaliensis | 2.05 | Campylobacter | upsaliensis | 1.96 | A | 1 | 1 | 0 | Campylobacter upsaliensis | Campylobacter upsaliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11185.3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter | canadensis | Helicobacter | canadensis | 2.72 | Clostridium | cochlearium | 1.42 | A | 1 | 1 | 0 | Helicobacter canadensis | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas | bestiarum | Aeromonas | bestiarum | 2.44 | Aeromonas | salmonicida | 2.42 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas | bivalvium | Aeromonas | veronii | 2.4 | Aeromonas | bivalvium | 2.18 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas bivalvium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas | caviae | Aeromonas | caviae | 2.29 | Aeromonas | hydrophila | 2.28 | B | 0 | 0 | 1 | Aeromonas caviae | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | hydrophila | 2.03 | Aeromonas | dhakensis | 2 | B | 0 | 0 | 1 | Aeromonas hydrophila | Aeromonas dhakensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | dhakensis | 2.42 | Aeromonas | dhakensis | 2.41 | A | 1 | 1 | 0 | Aeromonas dhakensis | Aeromonas dhakensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30187 RKI | 4; type strain, WGS sequenced (BfR) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-14 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.57 | Aeromonas | hydrophila | 2.46 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.27 | Aeromonas | hydrophila | 2.25 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas | jandaei | Aeromonas | jandaei | 2.34 | Aeromonas | jandaei | 2.05 | A | 1 | 1 | 0 | Aeromonas jandaei | Aeromonas jandaei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | media | Aeromonas | media | 2.22 | Aeromonas | media | 2.2 | A | 1 | 1 | 0 | Aeromonas media | Aeromonas media | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas | molluscorum | Aeromonas | molluscorum | 2.28 | Aeromonas | molluscorum | 2.15 | A | 1 | 1 | 0 | Aeromonas molluscorum | Aeromonas molluscorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | popoffii | Aeromonas | popoffii | 2.5 | Aeromonas | salmonicida | 2.26 | B | 0 | 0 | 1 | Aeromonas popoffii | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.22 | Aeromonas | salmonicida | 2.07 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.31 | Aeromonas | salmonicida | 2.25 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.55 | Aeromonas | salmonicida | 2.38 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.53 | Aeromonas | salmonicida | 2.22 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8400.2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.19 | Aeromonas | salmonicida | 2.07 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.38 | Aeromonas | salmonicida | 2.36 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.47 | Aeromonas | salmonicida | 2.3 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.18 | Aeromonas | salmonicida | 2.11 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas | simiae | Aeromonas | simiae | 2.31 | Aeromonas | simiae | 2.07 | A | 1 | 1 | 0 | Aeromonas simiae | Aeromonas simiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas | veronii | Aeromonas | veronii | 2.46 | Aeromonas | ichthiosmia | 2.18 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas ichthiosmia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7386 CVUAS | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas | veronii | Aeromonas | veronii | 2.46 | Aeromonas | bivalvium | 2.26 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas bivalvium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas | pectinilytica | Pseudaeromonas | pectinilytica | 2.62 | Aeromonas | eucrenophila | 1.4 | A | 1 | 1 | 0 | Pseudaeromonas pectinilytica | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella | viscosa | Moritella | viscosa | 2.32 | Moritella | viscosa | 2.22 | A | 1 | 1 | 0 | Moritella viscosa | Moritella viscosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella | viscosa | Moritella | viscosa | 2.29 | Moritella | viscosa | 2.23 | A | 1 | 1 | 0 | Moritella viscosa | Moritella viscosa | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | sp-CVUAS-9366 | 2.64 | Shewanella | sp-CVUAS-9366 | 2.62 | A | 1 | 1 | 0 | Shewanella sp-CVUAS-9366 | Shewanella sp-CVUAS-9366 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | sp-CVUAS-9366 | 2.55 | Shewanella | sp-CVUAS-9366 | 2.5 | A | 1 | 1 | 0 | Shewanella sp-CVUAS-9366 | Shewanella sp-CVUAS-9366 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 81.2 | 3; sequenced | EFEx | 2013-11-28 | Shewanella | sp-RKVR22 | Shewanella | sp-RKVR22 | 2.4 | Shewanella | corallii | 1.59 | A | 1 | 1 | 0 | Shewanella sp-RKVR22 | Shewanella corallii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Citrobacter | amalonaticus | Citrobacter | amalonaticus | 2.41 | Citrobacter | amalonaticus | 2.3 | A | 1 | 1 | 0 | Citrobacter amalonaticus | Citrobacter amalonaticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17342 | 3; public strain collection | see [0185] | 2022-08-29 | Citrobacter | amalonaticus | Citrobacter | amalonaticus | 2.44 | Citrobacter | amalonaticus | 2.25 | A | 1 | 1 | 0 | Citrobacter amalonaticus | Citrobacter amalonaticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | freundii | 2.39 | Citrobacter | freundii | 2.37 | A | 1 | 1 | 0 | Citrobacter freundii | Citrobacter freundii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | braakii | 2.29 | Citrobacter | braakii | 2.23 | A | 0 | 0 | 1 | Citrobacter braakii | Citrobacter braakii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25408 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2008-04-11 | Citrobacter | koseri | Citrobacter | koseri | 2.52 | Citrobacter | koseri | 2.47 | A | 1 | 1 | 0 | Citrobacter koseri | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4570 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Citrobacter koseri DSM 4570, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-29 | Citrobacter | koseri | Citrobacter | koseri | 2.58 | Citrobacter | koseri | 2.49 | A | 1 | 1 | 0 | Citrobacter koseri | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter | condimenti | Cronobacter | condimenti | 2.49 | Cronobacter | malonaticus | 2.29 | B | 0 | 0 | 1 | Cronobacter condimenti | Cronobacter malonaticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter | dublinensis | Cronobacter | muytjensii | 2.22 | Cronobacter | sp | 2.2 | B | 0 | 0 | 1 | Cronobacter muytjensii | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-14 | Cronobacter | malonaticus | Cronobacter | sp | 2.29 | Cronobacter | malonaticus | 2.24 | B | 0 | 0 | 1 | Cronobacter sp | Cronobacter malonaticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18702 | 3; type strain | DT | 2019-03-05 | Cronobacter | malonaticus | Cronobacter | malonaticus | 2.56 | Cronobacter | sp | 2.46 | B | 0 | 0 | 1 | Cronobacter malonaticus | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32945 | 3; confirmed proficency test | EFEx | 2021-02-18 | Cronobacter | sakazakii | Cronobacter | sp | 2.14 | Cronobacter | sp | 1.99 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter | sakazakii | Cronobacter | sp | 2.44 | Cronobacter | sp | 2.33 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter | turicensis | Cronobacter | turicensis | 2.52 | Cronobacter | sp | 2.36 | B | 0 | 0 | 1 | Cronobacter turicensis | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30054 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Enterobacter | cloacae | Enterobacter | cloacae | 2.34 | Enterobacter | cloacae | 2.08 | A | 1 | 1 | 0 | Enterobacter cloacae | Enterobacter cloacae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2011-11-10 | Enterobacter | cloacae | Enterobacter | hormaechei | 2.4 | Enterobacter | hormaechei | 2.35 | A | 0 | 0 | 1 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14563 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter | hormaechei | Enterobacter | hormaechei | 2.93 | Enterobacter | hormaechei | 2.2 | A | 1 | 1 | 0 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12409 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter | hormaechei | Enterobacter | hormaechei | 2.84 | Enterobacter | hormaechei | 2.14 | A | 1 | 1 | 0 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16691 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter | hormaechei | Enterobacter | hormaechei | 2.83 | Enterobacter | hormaechei | 2.28 | A | 1 | 1 | 0 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2023000186 | 2; corresponds to expected result of PT | EFEx | 2023-09-08 | Enterobacter | sp | Enterobacter | bugandensis | 2.27 | Enterobacter | asburiae | 2.25 | B | 0 | 0 | 1 | Enterobacter bugandensis | Enterobacter asburiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Enterobacter | sp-ATCC-27988 | Kosakonia | radicincitans | 1.94 | Raoultella | ornithinolytica | 1.68 | B | 0 | 0 | 1 | Kosakonia radicincitans | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae | sp-CVUAS-30072 | Enterobacteriaceae | sp-CVUAS-30072 | 2.52 | Kosakonia | cowanii | 1.45 | A | 1 | 1 | 0 | Enterobacteriaceae sp-CVUAS-30072 | Kosakonia cowanii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia | coli | Escherichia | coli | 2.32 | Escherichia | coli | 2.29 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0500.2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia | coli | Escherichia | coli | 2.41 | Escherichia | coli | 2.34 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30053 | 3; type strain | EFex | 2019-02-12 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.53 | Klebsiella | aerogenes | 2.46 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NM20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.48 | Klebsiella | aerogenes | 2.29 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella | grimontii | Klebsiella | oxytoca | 2.42 | Klebsiella | grimontii | 2.39 | B | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella grimontii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2486.3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella | grimontii | Klebsiella | grimontii | 2.66 | Klebsiella | michiganensis | 2.35 | B | 0 | 0 | 1 | Klebsiella grimontii | Klebsiella michiganensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2404.3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.71 | Klebsiella | grimontii | 2.46 | B | 0 | 0 | 1 | Klebsiella michiganensis | Klebsiella grimontii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella | michiganensis | Klebsiella | michiganensis | 2.67 | Klebsiella | michiganensis | 2.37 | A | 1 | 1 | 0 | Klebsiella michiganensis | Klebsiella michiganensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.4 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-11-01 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.42 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 2008-04-11 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.33 | Klebsiella | pneumoniae | 2.33 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30104 CVUAS | 3; type strain | EFEx | 2013-09-17 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.52 | Klebsiella | pneumoniae | 2.4 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | quasipneumoniae | 2.7 | Klebsiella | quasipneumoniae | 2.3 | A | 1 | 1 | 0 | Klebsiella quasipneumoniae | Klebsiella quasipneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | quasipneumoniae | 2.49 | Klebsiella | quasipneumoniae | 2.41 | A | 1 | 1 | 0 | Klebsiella quasipneumoniae | Klebsiella quasipneumoniae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella | variicola | Klebsiella | variicola | 2.12 | Klebsiella | variicola | 2.08 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | variicola | Klebsiella | variicola | 2.58 | Klebsiella | variicola | 2.52 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia | sacchari | Kosakonia | sacchari | 2.29 | Salmonella | sp | 1.72 | C | 0 | 0 | 1 | Kosakonia sacchari | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.36 | Plesiomonas | shigelloides | 2.34 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.42 | Plesiomonas | shigelloides | 2.41 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Pluralibacter | gergoviae | Pluralibacter | gergoviae | 2.32 | Pluralibacter | gergoviae | 2.24 | A | 1 | 1 | 0 | Pluralibacter gergoviae | Pluralibacter gergoviae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7763 | 3; NRL Salmonella BfR | DT | 2012-10-16 | Salmonella | enterica | Salmonella | sp | 2.32 | Salmonella | sp | 2.23 | A | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Salmonella | enterica | Salmonella | sp | 2.36 | Salmonella | sp | 2.36 | A | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | II 06 06820 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Shigella | boydii | Escherichia | coli | 2.43 | Escherichia | coli | 2.39 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 3A 07 02503 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Shigella | flexneri | Shigella | flexneri | 2.46 | Escherichia | coli | 2.46 | C | 0 | 0 | 1 | Shigella flexneri | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 273 | 2; strain collection of the RKI | TFAextr | 2021-11-19 | Erwinia | amylovora | Erwinia | amylovora | 2.35 | Erwinia | amylovora | 2.07 | A | 1 | 1 | 0 | Erwinia amylovora | Erwinia amylovora | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3493 CVUAS | 3; type strain | eDT | 2015-07-09 | Pantoea | agglomerans | Pantoea | agglomerans | 2.35 | Pantoea | agglomerans | 2.31 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea | agglomerans | Pantoea | agglomerans | 2.52 | Pantoea | agglomerans | 2.47 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea | sp | Pantoea | sp | 1.99 | Pantoea | agglomerans | 1.93 | B | 0 | 0 | 1 | Pantoea sp | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4263.2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea | sp-FF5 | Pantoea | sp-FF5 | 2.72 | Pantoea | septica | 1.84 | A | 1 | 1 | 0 | Pantoea sp-FF5 | Pantoea septica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella | sp | Edwardsiella | tarda | 2.47 | Edwardsiella | tarda | 2.47 | A | 0 | 0 | 1 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Edwardsiella | tarda | Edwardsiella | tarda | 2.49 | Edwardsiella | tarda | 2.37 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-18 | Hafnia | alvei | Hafnia | alvei | 2.42 | Hafnia | alvei | 2.32 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia | alvei | Hafnia | alvei | 2.36 | Hafnia | alvei | 2.29 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea | myxofaciens | Cosenzaea | myxofaciens | 2.56 | Cosenzaea | myxofaciens | 2.54 | A | 1 | 1 | 0 | Cosenzaea myxofaciens | Cosenzaea myxofaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella | morganii | Morganella | morganii | 2.56 | Morganella | morganii | 2.55 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella | morganii | Morganella | morganii | 2.68 | Morganella | morganii | 2.61 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | hauseri | Proteus | hauseri | 2.6 | Proteus | hauseri | 2.55 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus | hauseri | Proteus | hauseri | 2.47 | Proteus | hauseri | 2.45 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus | mirabilis | Proteus | mirabilis | 2.56 | Proteus | mirabilis | 2.45 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4479 CVUAS | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | mirabilis | Proteus | mirabilis | 2.52 | Proteus | mirabilis | 2.46 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | penneri | Proteus | penneri | 2.48 | Proteus | penneri | 2.43 | A | 1 | 1 | 0 | Proteus penneri | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | cibarius | 2.51 | Proteus | terrae | 2.33 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | vulgaris | 2.45 | Proteus | penneri | 2.41 | B | 0 | 0 | 1 | Proteus vulgaris | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | cibarius | 2.42 | Proteus | terrae | 2.4 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus | terrae | Proteus | cibarius | 2.61 | Proteus | terrae | 2.5 | B | 0 | 0 | 1 | Proteus cibarius | Proteus terrae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | terrae | Proteus | terrae | 2.66 | Proteus | cibarius | 2.44 | B | 0 | 0 | 1 | Proteus terrae | Proteus cibarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Proteus | vulgaris | Proteus | vulgaris | 2.3 | Proteus | vulgaris | 2.28 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2140 CVUAS | 3; strain from public collection; sequenced (rpoB) | eDT | 2017-11-30 | Proteus | vulgaris | Proteus | vulgaris | 2.38 | Proteus | vulgaris | 2.33 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia | rettgeri | Providencia | rettgeri | 2.57 | Providencia | rettgeri | 2.52 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia | rettgeri | Providencia | rettgeri | 2.54 | Providencia | rettgeri | 2.31 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25827 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Providencia | stuartii | Providencia | stuartii | 2.32 | Providencia | stuartii | 2.31 | A | 1 | 1 | 0 | Providencia stuartii | Providencia stuartii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4539 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 25.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Providencia stuartii DSM 4539 (T), Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-04-06 | Providencia | stuartii | Providencia | stuartii | 2.4 | Providencia | stuartii | 2.24 | A | 1 | 1 | 0 | Providencia stuartii | Providencia stuartii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7849 | 2; sequenced | eDT | 2022-08-23 | Rahnella | aquatilis | Rahnella | aquatilis | 2.51 | Rahnella | aquatilis | 2.46 | A | 1 | 1 | 0 | Rahnella aquatilis | Rahnella aquatilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia | fonticola | Serratia | fonticola | 2.56 | Serratia | fonticola | 2.55 | A | 1 | 1 | 0 | Serratia fonticola | Serratia fonticola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-11-04 | Serratia | grimesii | Serratia | grimesii | 2.42 | Serratia | liquefaciens | 2.03 | B | 0 | 0 | 1 | Serratia grimesii | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.19 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.13 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia | quinivorans | Serratia | proteamaculans | 2.16 | Serratia | liquefaciens | 2.12 | B | 0 | 0 | 1 | Serratia proteamaculans | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6221 LT | 4; sequenced | DT | 2013-01-08 | Yersinia | aldovae | Yersinia | aldovae | 2.33 | Yersinia | aldovae | 2.32 | A | 1 | 1 | 0 | Yersinia aldovae | Yersinia aldovae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6319 sir | 2; sequenced (16S) | eDT | 2024-12-04 | Yersinia | aldovae | Yersinia | aldovae | 2.31 | Yersinia | aldovae | 2.16 | A | 1 | 1 | 0 | Yersinia aldovae | Yersinia aldovae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2007-03-09 | Yersinia | aleksiciae | Yersinia | aleksiciae | 2.28 | Yersinia | aleksiciae | 2.28 | A | 1 | 1 | 0 | Yersinia aleksiciae | Yersinia aleksiciae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMB 4355 | 2; 0 | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦ 1:1 HCCA-TA2 | 2007-07-13 | Yersinia | aleksiciae | Yersinia | aleksiciae | 2.23 | Yersinia | aleksiciae | 2.22 | A | 1 | 1 | 0 | Yersinia aleksiciae | Yersinia aleksiciae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia | bercovieri | Yersinia | bercovieri | 2.33 | Yersinia | bercovieri | 2.25 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.56 | Yersinia | bercovieri | 2.45 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4780 | 3; type strain | DT | 2013-05-08 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.49 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.48 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.43 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.34 | Yersinia | enterocolitica | 2.31 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.45 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.57 | Yersinia | enterocolitica | 2.53 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | entomophaga | Yersinia | entomophaga | 2.8 | Yersinia | nurmii | 2.36 | B | 0 | 0 | 1 | Yersinia entomophaga | Yersinia nurmii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 21 RKI | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-22 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.41 | Yersinia | frederiksenii | 2.36 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.24 | Yersinia | enterocolitica | 2.08 | B | 0 | 0 | 1 | Yersinia frederiksenii | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BP 76 sir | 2; 0 | eDT | 2024-12-04 | Yersinia | intermedia | Yersinia | intermedia | 2.47 | Yersinia | intermedia | 2.37 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia | intermedia | Yersinia | intermedia | 2.51 | Yersinia | intermedia | 2.45 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.57 | Yersinia | kristensenii | 2.53 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 302 LT | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.46 | Yersinia | kristensenii | 2.39 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-11-25 | Yersinia | massiliensis | Yersinia | massiliensis | 2.71 | Yersinia | massiliensis | 2.33 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | VB 925699 sir | 3; sequenced | eDT | 2024-12-04 | Yersinia | massiliensis | Yersinia | massiliensis | 2.39 | Yersinia | massiliensis | 2.18 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia | mollaretii | Yersinia | mollaretii | 2.32 | Yersinia | mollaretii | 2.25 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18520 | 3; type strain | DT | 16.01.2013 | Yersinia | mollaretii | Yersinia | mollaretii | 2.13 | Yersinia | mollaretii | 2.11 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | nurmii | Yersinia | nurmii | 2.75 | Yersinia | entomophaga | 2.33 | B | 0 | 0 | 1 | Yersinia nurmii | Yersinia entomophaga | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | pekkanenii | Yersinia | pekkanenii | 2.56 | Yersinia | pekkanenii | 2.36 | A | 1 | 1 | 0 | Yersinia pekkanenii | Yersinia pekkanenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 2028 | 3; public strain collection | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.48 | Yersinia | pseudotuberculosis | 2.48 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.57 | Yersinia | pseudotuberculosis | 2.43 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 304.2 sir | 4; WGS sequence (BfR) | EFEx | 2023-06-21 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.29 | Yersinia | pseudotuberculosis | 2.2 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8992 | 3; type strain | EFEx | 2023-09-06 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.69 | Yersinia | pseudotuberculosis | 2.67 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia | rohdei | Yersinia | rohdei | 2.46 | Yersinia | rohdei | 2.41 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2007-03-02 | Yersinia | rohdei | Yersinia | rohdei | 2.83 | Yersinia | rohdei | 2.31 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2772 CVUAS | 3; sequenced | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.61 | Yersinia | ruckeri | 2.56 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.56 | Yersinia | ruckeri | 2.56 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia | similis | Yersinia | similis | 2.83 | Yersinia | similis | 2.51 | A | 1 | 1 | 0 | Yersinia similis | Yersinia similis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia | wautersii | Yersinia | wautersii | 2.29 | Yersinia | pseudotuberculosis | 2.23 | B | 0 | 0 | 1 | Yersinia wautersii | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia | wautersii | Yersinia | wautersii | 2.69 | Yersinia | pseudotuberculosis | 2.39 | B | 0 | 0 | 1 | Yersinia wautersii | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella | londiniensis | Legionella | londiniensis | 2.27 | Legionella | londiniensis | 1.72 | A | 1 | 1 | 0 | Legionella londiniensis | Legionella londiniensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella | pneumophila | Legionella | pneumophila | 2.21 | Legionella | pneumophila | 1.73 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4122 | 3; 0 | eDT | 2016-01-20 | Legionella | pneumophila | Legionella | pneumophila | 2.33 | Legionella | pneumophila | 2.29 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella | taurinensis | Legionella | taurinensis | 2.41 | Legionella | sp | 1.86 | A | 1 | 1 | 0 | Legionella taurinensis | Legionella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10873.2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | stevensii | 2.33 | Halomonas | aquamarina | 1.54 | A | 1 | 1 | 0 | Halomonas stevensii | Halomonas aquamarina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10874.2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | stevensii | 2.43 | Dermacoccus | sp[3] | 1.35 | A | 1 | 1 | 0 | Halomonas stevensii | Dermacoccus sp[3] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) | caballi | Pasteurella | caballi | 2.22 | Pasteurella | caballi | 2.15 | A | 0 | 0 | 1 | Pasteurella caballi | Pasteurella caballi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) | caballi | Pasteurella | caballi | 2.54 | (Pasteurella) | caballi | 2.53 | C | 0 | 0 | 1 | Pasteurella caballi | (Pasteurella) caballi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32909 | 3; sequenced | eDT | 2021-02-05 | (Pasteurella) | mairii | (Pasteurella) | mairii | 2.46 | Actinobacillus | seminis | 1.99 | C | 0 | 0 | 1 | (Pasteurella) mairii | Actinobacillus seminis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) | mairii | (Pasteurella) | mairii | 2.28 | Pasteurella | mairii | 2.11 | C | 0 | 0 | 1 | (Pasteurella) mairii | Pasteurella mairii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) | testudinis | (Pasteurella) | testudinis | 2.33 | Pasteurella | testudinis | 2.01 | C | 0 | 0 | 1 | (Pasteurella) testudinis | Pasteurella testudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 957 | 3; sequenced | eDT | 2015-12-02 | (Pasteurella) | testudinis | (Pasteurella) | testudinis | 2.11 | Pasteurella | testudinis | 1.87 | C | 0 | 0 | 1 | (Pasteurella) testudinis | Pasteurella testudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | anseriformium | Actinobacillus | anseriformium | 2.48 | Actinobacillus | equuli | 2.05 | B | 0 | 0 | 1 | Actinobacillus anseriformium | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15UGF198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus | anseriformium | Actinobacillus | anseriformium | 2.53 | Actinobacillus | anseriformium | 2.22 | A | 1 | 1 | 0 | Actinobacillus anseriformium | Actinobacillus anseriformium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus | arthritidis | Actinobacillus | arthritidis | 2.33 | Actinobacillus | arthritidis_JRA_Pasturella-4 | 2 | B | 0 | 0 | 1 | Actinobacillus arthritidis | Actinobacillus arthritidis_JRA_Pasturella-4 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.42 | Actinobacillus | capsulatus | 2.2 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus capsulatus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.52 | Actinobacillus | capsulatus | 2.42 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus capsulatus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.4 | Actinobacillus | equuli | 2.38 | A | 1 | 1 | 0 | Actinobacillus equuli | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.28 | Actinobacillus | suis | 2.22 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.36 | Actinobacillus | equuli | 2.35 | A | 1 | 1 | 0 | Actinobacillus equuli | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus | equuli | Actinobacillus | equuli | 2.52 | Actinobacillus | suis | 2.52 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus | lignieresii | Actinobacillus | pleuropneumoniae | 2.34 | Actinobacillus | lignieresii | 2.33 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9945.2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | lignieresii | Actinobacillus | pleuropneumoniae | 2.45 | Actinobacillus | lignieresii | 2.37 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus | rossii | Actinobacillus | rossii | 2.36 | Actinobacillus | rossii | 2.08 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus | rossii | Actinobacillus | rossii | 2.67 | Actinobacillus | rossii | 2.64 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32505.2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus | sp-CVUAS-32505,2 | Actinobacillus | sp-CVUAS-32505,2 | 2.54 | Actinobacillus | arthritidis_JRA_Pasturella-4 | 2.18 | B | 0 | 0 | 1 | Actinobacillus sp-CVUAS-32505,2 | Actinobacillus arthritidis_JRA_Pasturella-4 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | sp-genomospecies-1 | 2.46 | Actinobacillus | pleuropneumoniae | 2.14 | B | 0 | 0 | 1 | Actinobacillus sp-genomospecies-1 | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | sp-genomospecies-1 | 2.53 | Actinobacillus | arthritidis | 2.09 | B | 0 | 0 | 1 | Actinobacillus sp-genomospecies-1 | Actinobacillus arthritidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus | sp-MK-2012 | Actinobacillus | sp-MK-2012 | 2.67 | Actinobacillus | sp-MK-2012 | 2.34 | A | 1 | 1 | 0 | Actinobacillus sp-MK-2012 | Actinobacillus sp-MK-2012 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-MK-2012 | Actinobacillus | sp-MK-2012 | 2.46 | Actinobacillus | sp-MK-2012 | 2.44 | A | 1 | 1 | 0 | Actinobacillus sp-MK-2012 | Actinobacillus sp-MK-2012 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus | suis | Actinobacillus | suis | 2.6 | Actinobacillus | suis | 2.55 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus | suis | Actinobacillus | suis | 2.5 | Actinobacillus | suis | 2.47 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium | gallinarum | Avibacterium | gallinarum | 2.19 | Avibacterium | gallinarum | 2.15 | A | 1 | 1 | 0 | Avibacterium gallinarum | Avibacterium gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium | gallinarum | Avibacterium | gallinarum | 2.72 | Avibacterium | gallinarum | 2.27 | A | 1 | 1 | 0 | Avibacterium gallinarum | Avibacterium gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1816.2 | 3; sequenced | DT | 2013-07-24 | Basfia | succiniciproducens | Basfia | succiniciproducens | 2.08 | Basfia | succiniciproducens | 2.03 | A | 1 | 1 | 0 | Basfia succiniciproducens | Basfia succiniciproducens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.52 | Bibersteinia | trehalosi | 2.48 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.47 | Bibersteinia | trehalosi | 2.39 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter | oris | Chelonobacter | oris | 2.42 | Chelonobacter | oris | 2.41 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter | oris | Chelonobacter | oris | 2.11 | Chelonobacter | oris | 2.05 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3802.2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter | osteomyelitidis | Cricetibacter | osteomyelitidis | 2.54 | Haemophilus | haemolyticus | 1.43 | A | 1 | 1 | 0 | Cricetibacter osteomyelitidis | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31212 | 3; sequenced | EFEx | 2018-10-16 | Exercitatus | varius | Exercitatus | varius | 2.66 | Exercitatus | varius | 2.3 | A | 1 | 1 | 0 | Exercitatus varius | Exercitatus varius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Exercitatus | varius | Exercitatus | varius | 2.36 | Exercitatus | varius | 2.34 | A | 1 | 1 | 0 | Exercitatus varius | Exercitatus varius | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia | canicola | Frederiksenia | canicola | 2.41 | Frederiksenia | canicola | 2.21 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia | canicola | Frederiksenia | canicola | 2.32 | Frederiksenia | canicola | 2.28 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium | anatis | Gallibacterium | anatis | 2.07 | Gallibacterium | anatis | 2.04 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium | anatis | Gallibacterium | anatis | 2.45 | Gallibacterium | anatis | 2.36 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium | sp-F151 | Gallibacterium | sp-genomospecies_3 | 2.49 | Gallibacterium | anatis | 1.44 | A | 0 | 0 | 1 | Gallibacterium sp-genomospecies_3 | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium | trehalosifermentans | Gallibacterium | trehalosifermentans | 2.65 | Gallibacterium | sp-genomospecies_3 | 1.88 | A | 1 | 1 | 0 | Gallibacterium trehalosifermentans | Gallibacterium sp-genomospecies_3 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus | haemoglobinophilus | Haemophilus | haemoglobinophilus | 2.29 | Haemophilus | sp-2269 | 1.91 | A | 1 | 1 | 0 | Haemophilus haemoglobinophilus | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4690 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 03.06.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Haemophilus influenzae DSM 4690, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-04 | Haemophilus | influenzae | Haemophilus | influenzae | 2.43 | Haemophilus | influenzae | 2.31 | A | 1 | 1 | 0 | Haemophilus influenzae | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | sp-2269 | 2.39 | Haemophilus | sp-2269 | 2.35 | A | 1 | 1 | 0 | Haemophilus sp-2269 | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | sp-2269 | 2.51 | Haemophilus | sp-2269 | 2.46 | A | 1 | 1 | 0 | Haemophilus sp-2269 | Haemophilus sp-2269 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P 1066 2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus | somni | Histophilus | somni | 2.2 | Histophilus | somni | 1.82 | A | 1 | 1 | 0 | Histophilus somni | Histophilus somni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia | caviae | Mannheimia | caviae | 2.62 | Mannheimia | caviae | 2.22 | A | 1 | 1 | 0 | Mannheimia caviae | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1064.4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia | glucosida | Mannheimia | glucosida | 2.35 | Mannheimia | glucosida | 2.22 | A | 1 | 1 | 0 | Mannheimia glucosida | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia | glucosida | Mannheimia | glucosida | 2.59 | Mannheimia | glucosida | 2.47 | A | 1 | 1 | 0 | Mannheimia glucosida | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia | granulomatis | Mannheimia | sp-14UCF195 | 2.3 | Mannheimia | granulomatis | 2.05 | B | 0 | 0 | 1 | Mannheimia sp-14UCF195 | Mannheimia granulomatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia | granulomatis | Mannheimia | granulomatis | 2.48 | Mannheimia | sp-14UCF195 | 2.26 | B | 0 | 0 | 1 | Mannheimia granulomatis | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1738 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.57 | Mannheimia | haemolytica | 2.57 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3952 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.64 | Mannheimia | haemolytica | 2.62 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.53 | Mannheimia | haemolytica | 2.47 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6197.2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.57 | Mannheimia | haemolytica | 2.52 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia | pernigra | Mannheimia | pernigra | 2.51 | Mannheimia | pernigra | 2.11 | A | 1 | 1 | 0 | Mannheimia pernigra | Mannheimia pernigra | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia | pernigra | Mannheimia | pernigra | 2.32 | Mannheimia | pernigra | 2.32 | A | 1 | 1 | 0 | Mannheimia pernigra | Mannheimia pernigra | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.57 | Mannheimia | ruminalis | 2.3 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.3 | Mannheimia | ruminalis | 2.19 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10396.2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia | sp-10-12559-1 | Mannheimia | sp-10-12559-1 | 2.46 | Mannheimia | caviae | 2.03 | B | 0 | 0 | 1 | Mannheimia sp-10-12559-1 | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10183.3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia | sp-14UCF195 | Mannheimia | sp-14UCF195 | 2.64 | Mannheimia | sp-14UCF195 | 2.32 | A | 1 | 1 | 0 | Mannheimia sp-14UCF195 | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia | sp-14UCF195 | Mannheimia | sp-14UCF195 | 2.68 | Mannheimia | sp-14UCF195 | 2.29 | A | 1 | 1 | 0 | Mannheimia sp-14UCF195 | Mannheimia sp-14UCF195 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia | sp-15USF15 | Mannheimia | sp-15USF15 | 2.68 | Mannheimia | indoligenes | 2.46 | B | 0 | 0 | 1 | Mannheimia sp-15USF15 | Mannheimia indoligenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia | sp-15USF15 | Mannheimia | sp-15USF15 | 2.49 | Mannheimia | caviae | 2.13 | B | 0 | 0 | 1 | Mannheimia sp-15USF15 | Mannheimia caviae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia | sp-CVUAS-11240 | Mannheimia | sp-CVUAS-11240 | 2.65 | Mannheimia | sp-15USF15 | 1.94 | A | 1 | 1 | 0 | Mannheimia sp-CVUAS-11240 | Mannheimia sp-15USF15 | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia | sp-CVUAS-1496 | Mannheimia | sp-CVUAS-1496 | 2.47 | Mannheimia | glucosida | 2.37 | B | 0 | 0 | 1 | Mannheimia sp-CVUAS-1496 | Mannheimia glucosida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.5 | Necropsobacter | rosorum | 2.43 | A | 1 | 1 | 0 | Necropsobacter rosorum | Necropsobacter rosorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.36 | Necropsobacter | rosorum | 2.35 | A | 1 | 1 | 0 | Necropsobacter rosorum | Necropsobacter rosorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella | semolina | Nicoletella | semolina | 2.55 | Nicoletella | semolina | 2.43 | A | 1 | 1 | 0 | Nicoletella semolina | Nicoletella semolina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella | semolina | Nicoletella | semolina | 2.46 | Nicoletella | semolina | 2.46 | A | 1 | 1 | 0 | Nicoletella semolina | Nicoletella semolina | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4060.3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella | aerogenes | Pasteurella | aerogenes | 2.47 | Pasteurella | aerogenes | 2.05 | A | 1 | 1 | 0 | Pasteurella aerogenes | Pasteurella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | canis | Pasteurella | canis | 2.07 | Pasteurella | canis | 2.04 | A | 1 | 1 | 0 | Pasteurella canis | Pasteurella canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella | multocida | Pasteurella | multocida | 2.48 | Pasteurella | multocida | 2.33 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32055 | 3; specific PCR | EFEx | 2024-09-25 | Pasteurella | multocida | Pasteurella | multocida | 2.67 | Pasteurella | multocida | 2.59 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella | multocida | Pasteurella | multocida | 2.62 | Pasteurella | multocida | 2.52 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.57 | Pasteurella | multocida | 2.56 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.51 | Pasteurella | multocida | 2.5 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella | multocida | Pasteurella | multocida | 2.57 | Pasteurella | multocida | 2.49 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | Bisgaard-Taxon-40 | Pasteurellaceae | Bisgaard-Taxon-40 | 2.45 | Haemophilus | parainfluenzae | 1.36 | A | 1 | 1 | 0 | Pasteurellaceae Bisgaard-Taxon-40 | Haemophilus parainfluenzae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9987.2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae | canine-oral-taxon-080 | Pasteurellaceae | canine-oral-taxon-080 | 2.56 | Frederiksenia | canicola | 1.92 | C | 0 | 0 | 1 | Pasteurellaceae canine-oral-taxon-080 | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 261.4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae | sp-CVUAS-261,4 | Psittacitella | sp-CVUAS-261.4 | 2.64 | Klebsiella | oxytoca | 1.47 | A | 0 | 0 | 1 | Psittacitella sp-CVUAS-261.4 | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae | sp-CVUAS-31988 | Pasteurellaceae | sp-CVUAS-31988 | 2.61 | Haemophilus | haemolyticus | 1.65 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-31988 | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34093 | 2; sequenced | EFex | 2022-08-05 | Pasteurellaceae | sp-CVUAS-34093 | Pasteurellaceae | sp-CVUAS-34093 | 2.59 | Mannheimia | indoligenes | 1.57 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-34093 | Mannheimia indoligenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | sp-CVUAS-4868 | Pasteurellaceae | sp-CVUAS-4868 | 2.54 | Enterococcus | faecalis | 1.34 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-4868 | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 724.3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae | sp-CVUAS-724,3 | Pasteurellaceae | sp-CVUAS-724,3 | 2.6 | Rodentibacter | heylii | 1.47 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-724,3 | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae | sp-CVUAS-9305 | Pasteurellaceae | sp-CVUAS-9305 | 2.65 | Actinobacillus | equuli | 1.41 | A | 1 | 1 | 0 | Pasteurellaceae sp-CVUAS-9305 | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9967.2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae | sp-CVUAS-9967,2 | Pasteurellaceae | sp-CVUAS-9967,2 | 2.45 | Haemophilus | haemolyticus | 1.74 | C | 0 | 0 | 1 | Pasteurellaceae sp-CVUAS-9967,2 | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae | sp-LHL-U-309 | Pasteurellaceae | sp-LHL-U-309 | 2.36 | Pasteurellaceae | sp-LHL-U-309 | 2.13 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-309 | Pasteurellaceae sp-LHL-U-309 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL U-309 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-309 | Pasteurellaceae | sp-LHL-U-309 | 2.46 | Pasteurellaceae | sp-LHL-U-309 | 2.22 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-309 | Pasteurellaceae sp-LHL-U-309 | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL U-3444 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-3444 | Pasteurellaceae | sp-LHL-U-3444 | 2.54 | Chelonobacter | oris | 1.54 | A | 1 | 1 | 0 | Pasteurellaceae sp-LHL-U-3444 | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae | sp-MCCM-2539 | Pasteurellaceae | sp-MCCM-2539 | 2.54 | Haemophilus | pittmaniae | 1.53 | A | 1 | 1 | 0 | Pasteurellaceae sp-MCCM-2539 | Haemophilus pittmaniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.48 | Rodentibacter | heylii | 2.4 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.55 | Rodentibacter | heylii | 2.5 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium | mucosae | Spirabiliibacterium | mucosae | 2.39 | Spirabiliibacterium | mucosae | 2.3 | A | 1 | 1 | 0 | Spirabiliibacterium mucosae | Spirabiliibacterium mucosae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium | mucosae | Spirabiliibacterium | mucosae | 2.49 | Spirabiliibacterium | mucosae | 2.45 | A | 1 | 1 | 0 | Spirabiliibacterium mucosae | Spirabiliibacterium mucosae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8966 | 3; sequenced | DT | 2013-07-23 | Volucribacter | amazonae | Volucribacter | amazonae | 2.43 | Volucribacter | psittacicida | 2.21 | B | 0 | 0 | 1 | Volucribacter amazonae | Volucribacter psittacicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LLBB 23UGF306 | 3; sequenced | OS Extr | 2023-10-11 | Volucribacter | psittacicida | Volucribacter | psittacicida | 2.27 | Volucribacter | amazonae | 2.14 | B | 0 | 0 | 1 | Volucribacter psittacicida | Volucribacter amazonae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 308.2 | 3; sequenced | DT | 2013-11-06 | Psittacicella | hinzii | Psittacicella | hinzii | 2.57 | Agromyces | cerinus | 1.43 | A | 1 | 1 | 0 | Psittacicella hinzii | Agromyces cerinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-14 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.51 | Acinetobacter | baumannii | 2.48 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-09-03 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.46 | Acinetobacter | baumannii | 2.41 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter | defluvii | Acinetobacter | defluvii | 2.42 | Acinetobacter | defluvii | 2.16 | A | 1 | 1 | 0 | Acinetobacter defluvii | Acinetobacter defluvii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | guillouiae | Acinetobacter | guillouiae | 2.34 | Acinetobacter | guillouiae | 2.24 | A | 1 | 1 | 0 | Acinetobacter guillouiae | Acinetobacter guillouiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | johnsonii | Acinetobacter | johnsonii | 2.3 | Acinetobacter | johnsonii | 2.26 | A | 1 | 1 | 0 | Acinetobacter johnsonii | Acinetobacter johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | junii | Acinetobacter | junii | 2.33 | Acinetobacter | junii | 2.31 | A | 1 | 1 | 0 | Acinetobacter junii | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7714 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | junii | Acinetobacter | junii | 2.49 | Acinetobacter | junii | 2.45 | A | 1 | 1 | 0 | Acinetobacter junii | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | lwoffii | Acinetobacter | lwoffii | 2.35 | Acinetobacter | lwoffii | 2.27 | A | 1 | 1 | 0 | Acinetobacter lwoffii | Acinetobacter lwoffii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30502 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter lwoffii DSM 30502, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Acinetobacter | lwoffii | Acinetobacter | lwoffii | 2.13 | Acinetobacter | lwoffii | 1.9 | A | 1 | 1 | 0 | Acinetobacter lwoffii | Acinetobacter lwoffii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9312 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 28.04.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter pittii DSM 9312, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-05-06 | Acinetobacter | pittii | Acinetobacter | pittii | 2.4 | Acinetobacter | pittii | 2.36 | A | 1 | 1 | 0 | Acinetobacter pittii | Acinetobacter pittii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9318 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter pittii DSM 9318, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Acinetobacter | pittii | Acinetobacter | pittii | 2.77 | Acinetobacter | lactucae | 2.28 | B | 0 | 0 | 1 | Acinetobacter pittii | Acinetobacter lactucae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7379.3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.29 | Acinetobacter | pseudolwoffii | 2.26 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.08 | Acinetobacter | pseudolwoffii | 2.07 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32366 | 3; sequenced | eDT | 2020-06-19 | Acinetobacter | sp-CVUAS-32366 | Acinetobacter | sp-CVUAS-32366 | 2.61 | Clostridium | baratii | 1.47 | A | 1 | 1 | 0 | Acinetobacter sp-CVUAS-32366 | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | variabilis | Acinetobacter | variabilis | 2.35 | Acinetobacter | variabilis | 2.17 | A | 1 | 1 | 0 | Acinetobacter variabilis | Acinetobacter variabilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | ovis | 2.44 | Moraxella | bovoculi | 2.16 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | ovis | 2.09 | Moraxella | bovoculi | 2.06 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella | catarrhalis | Moraxella | catarrhalis | 2.53 | Moraxella | catarrhalis | 2.51 | A | 1 | 1 | 0 | Moraxella catarrhalis | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11674.3 | 3; sequenced | eDT | 2022-03-22 | Moraxella | cuniculi | Moraxella | cuniculi | 2.55 | Moraxella | catarrhalis | 1.63 | A | 1 | 1 | 0 | Moraxella cuniculi | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella | osloensis | Moraxella | osloensis | 2.53 | Moraxella | osloensis | 1.94 | A | 1 | 1 | 0 | Moraxella osloensis | Moraxella osloensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella | ovis | Moraxella | ovis | 2.39 | Moraxella | bovoculi | 2.32 | B | 0 | 0 | 1 | Moraxella ovis | Moraxella bovoculi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | frigidicola | Psychrobacter | frigidicola | 2.49 | Psychrobacter | pulmonis | 1.47 | A | 1 | 1 | 0 | Psychrobacter frigidicola | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter | pasteurii | Psychrobacter | pasteurii | 2.52 | Psychrobacter | pasteurii | 2.04 | A | 1 | 1 | 0 | Psychrobacter pasteurii | Psychrobacter pasteurii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.58 | Psychrobacter | pulmonis | 2.35 | A | 1 | 1 | 0 | Psychrobacter pulmonis | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.3 | Psychrobacter | faecalis | 1.83 | A | 1 | 1 | 0 | Psychrobacter pulmonis | Psychrobacter faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas | (alkylphenolica) | Pseudomonas | alkylphenolica | 2.4 | Pseudomonas | japonica | 1.59 | A | 0 | 0 | 1 | Pseudomonas alkylphenolica | Pseudomonas japonica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas | [baetica] | Pseudomonas | baetica | 2.66 | Pseudomonas | koreensis | 2.06 | B | 0 | 0 | 1 | Pseudomonas baetica | Pseudomonas koreensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.33 | Pseudomonas | aeruginosa | 2.28 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.46 | Pseudomonas | aeruginosa | 2.38 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas | alcaligenes | Pseudomonas | alcaligenes | 2.44 | Pseudomonas | alcaligenes | 2.35 | A | 1 | 1 | 0 | Pseudomonas alcaligenes | Pseudomonas alcaligenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas | anguilliseptica | Pseudomonas | anguilliseptica | 2.02 | Pseudomonas | borbori | 1.95 | A | 1 | 1 | 0 | Pseudomonas anguilliseptica | Pseudomonas borbori | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas | brenneri | Pseudomonas | brenneri | 2.14 | Pseudomonas | veronii | 1.84 | A | 1 | 1 | 0 | Pseudomonas brenneri | Pseudomonas veronii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-02-20 | Pseudomonas | chlororaphis | Pseudomonas | chlororaphis | 2.44 | Pseudomonas | chlororaphis | 2.3 | A | 1 | 1 | 0 | Pseudomonas chlororaphis | Pseudomonas chlororaphis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas | fulva | Pseudomonas | fulva | 2.29 | Pseudomonas | fulva | 1.9 | A | 1 | 1 | 0 | Pseudomonas fulva | Pseudomonas fulva | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas | koreensis | Pseudomonas | koreensis | 2.26 | Pseudomonas | corrugata | 1.89 | A | 1 | 1 | 0 | Pseudomonas koreensis | Pseudomonas corrugata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas | monteilii | Pseudomonas | monteilii | 2.01 | Pseudomonas | putida | 1.92 | A | 1 | 1 | 0 | Pseudomonas monteilii | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1 A | 2; 0 | DT | 2017-08-31 | Pseudomonas | nitroreducens | Pseudomonas | nitroreducens | 2.07 | Pseudomonas | nitroreducens | 1.97 | A | 1 | 1 | 0 | Pseudomonas nitroreducens | Pseudomonas nitroreducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas | oryzihabitans | Pseudomonas | oryzihabitans | 2.12 | Pseudomonas | oryzihabitans | 2.09 | A | 1 | 1 | 0 | Pseudomonas oryzihabitans | Pseudomonas oryzihabitans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1 B | 2; 0 | DT | 2017-08-31 | Pseudomonas | putida | Pseudomonas | putida | 2.49 | Pseudomonas | putida | 2.31 | A | 1 | 1 | 0 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 23974 | 3; public strain collection | TFAextr | 2021-09-08 | Pseudomonas | putida | Pseudomonas | putida | 2.33 | Pseudomonas | putida | 2.22 | A | 1 | 1 | 0 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5190 | 3; type strain | TFAextr | 2022-06-04 | Pseudomonas | stutzeri | Pseudomonas | stutzeri | 2.61 | Pseudomonas | stutzeri | 2.27 | A | 1 | 1 | 0 | Pseudomonas stutzeri | Pseudomonas stutzeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas | taetrolens | Pseudomonas | taetrolens | 2 | Pseudomonas | lundensis | 1.95 | A | 1 | 1 | 0 | Pseudomonas taetrolens | Pseudomonas lundensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-12-12 | Allofrancisella | guangzhouensis | Allofrancisella | guangzhouensis | 2.69 | Paracoccus | denitrificans | 1.33 | A | 1 | 1 | 0 | Allofrancisella guangzhouensis | Paracoccus denitrificans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-11-18 | Francisella | hispaniensis | Francisella | hispaniensis | 2.75 | Francisella | tularensis | 1.99 | A | 1 | 1 | 0 | Francisella hispaniensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18777 | 3; strain collection of the RKI | EFEx | 2013-06-20 | Francisella | noatunensis | Francisella | noatunensis | 2.73 | Francisella | philomiragia | 1.78 | A | 1 | 1 | 0 | Francisella noatunensis | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-11-25 | Francisella | orientalis | Francisella | orientalis | 2.81 | Francisella | philomiragia | 1.57 | A | 1 | 1 | 0 | Francisella orientalis | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7535 | 3; type strain | EFEx | 2013-06-20 | Francisella | philomiragia | Francisella | philomiragia | 2.62 | Francisella | philomiragia | 1.9 | A | 1 | 1 | 0 | Francisella philomiragia | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4680.2 | 4; WGS (FLI) | DT | 2013-04-02 | Francisella | tularensis | Francisella | tularensis | 2.23 | Francisella | tularensis | 2.2 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 2013-04-23 | Francisella | tularensis | Francisella | tularensis | 2.36 | Francisella | tularensis | 2.34 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 2013-04-26 | Francisella | tularensis | Francisella | tularensis | 2.68 | Francisella | tularensis | 2.26 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | F064 | 3; strain collection of the FLI | EFEx | 2013-06-20 | Francisella | tularensis | Francisella | tularensis | 2.78 | Francisella | tularensis | 2.57 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2012-04-02 | Francisella | tularensis | Francisella | tularensis | 2.55 | Francisella | tularensis | 2.33 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 2017-06-01 | Francisella | tularensis | Francisella | tularensis | 2.8 | Francisella | tularensis | 2.44 | A | 1 | 1 | 0 | Francisella tularensis | Francisella tularensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33180 25C | 3; sequenced | DT | 2021-08-04 | Photobacterium | damselae | Photobacterium | damselae | 2.22 | Photobacterium | damselae | 2.02 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium | damselae | Photobacterium | damselae | 2.14 | Photobacterium | damselae | 2.07 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1266.2 | 3; sequenced | EFEx | 2019-09-25 | Photobacterium | piscicola | Photobacterium | piscicola | 2.6 | Photobacterium | iliopiscarium | 1.57 | A | 1 | 1 | 0 | Photobacterium piscicola | Photobacterium iliopiscarium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 842.2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium | swingsii | Photobacterium | swingsii | 2.63 | Providencia | alcalifaciens | 1.37 | A | 1 | 1 | 0 | Photobacterium swingsii | Providencia alcalifaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio | aestuarianus | Vibrio | aestuarianus | 2.52 | Vibrio | aestuarianus | 2.5 | A | 1 | 1 | 0 | Vibrio aestuarianus | Vibrio aestuarianus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7079 sir | 3; spezific PCR (CVFR) | eDT | 2022-06-23 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.18 | Vibrio | alginolyticus_17-VB00051_BfR | 2.17 | B | 0 | 0 | 1 | Vibrio alginolyticus | Vibrio alginolyticus_17-VB00051_BfR | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2171 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 2012-07-12 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.3 | Vibrio | alginolyticus | 2.29 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio | anguillarum | Vibrio | anguillarum | 2.34 | Vibrio | anguillarum | 2.3 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | A8-12 sir | 3; gyrB sequence | eDT | 2022-06-23 | Vibrio | anguillarum | Vibrio | anguillarum | 2.22 | Vibrio | anguillarum | 2.2 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio | campbelli | Vibrio | campbellii | 2.32 | Vibrio | campbellii | 2.28 | A | 0 | 0 | 1 | Vibrio campbellii | Vibrio campbellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio | campbelli | Vibrio | campbellii | 2.38 | Vibrio | campbellii | 2.37 | A | 0 | 0 | 1 | Vibrio campbellii | Vibrio campbellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16-VB00067-0 eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio | cholerae | Vibrio | cholerae | 2.37 | Vibrio | cholerae | 2.37 | A | 1 | 1 | 0 | Vibrio cholerae | Vibrio cholerae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10227 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio | cholerae | Vibrio | cholerae_23-VB00040_BfR | 2.27 | Vibrio | cholerae | 2.26 | B | 0 | 0 | 1 | Vibrio cholerae_23-VB00040_BfR | Vibrio cholerae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 45V066-10 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | cincinnatiensis | Vibrio | cincinnatiencis | 2.43 | Vibrio | cincinnatiensis | 2.41 | B | 0 | 0 | 1 | Vibrio cincinnatiencis | Vibrio cincinnatiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 2.55 | Vibrio | cincinnatiencis | 2.44 | B | 0 | 0 | 1 | Vibrio cincinnatiensis | Vibrio cincinnatiencis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.08 | Vibrio | diazotrophicus | 1.98 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8995 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.14 | Vibrio | diazotrophicus | 2.14 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio | fluvialis | Vibrio | fluvialis | 2.42 | Vibrio | fluvialis | 2.16 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio | fluvialis | Vibrio | fluvialis | 2.32 | Vibrio | fluvialis | 2.22 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1754 LT | 3; sequenced | eDT | 2013-04-16 | Vibrio | furnissii | Vibrio | furnissii | 2.2 | Vibrio | furnissii | 2.05 | A | 1 | 1 | 0 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14383 sir | 3; public strain collection, sequenced | eDT | 2022-06-23 | Vibrio | furnissii | Vibrio | furnissii | 2.21 | Vibrio | furnissii | 2.06 | A | 1 | 1 | 0 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 2012-08-16 | Vibrio | harveyi | Vibrio | harveyi | 2.28 | Vibrio | harveyi | 2.21 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio | harveyi | Vibrio | harveyi | 2.41 | Vibrio | harveyi | 2.4 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | P340-7 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio | jasicida | Vibrio | jasicida | 2.53 | Vibrio | harveyi | 2.49 | B | 0 | 0 | 1 | Vibrio jasicida | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | P340-9 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio | jasicida | Vibrio | jasicida | 2.41 | Vibrio | jasicida | 2.36 | A | 1 | 1 | 0 | Vibrio jasicida | Vibrio jasicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 12-VBH0498 | 4; WGS sequenced (BfR) | DT | 2021-02-02 | Vibrio | metoecus | Vibrio | mimicus | 2.48 | Vibrio | metoecus | 2.41 | B | 0 | 0 | 1 | Vibrio mimicus | Vibrio metoecus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00016 | 4; WGS sequenced (BfR) | DT | 2021-02-24 | Vibrio | metoecus | Vibrio | metoecus | 2.34 | Vibrio | metoecus | 2.32 | A | 1 | 1 | 0 | Vibrio metoecus | Vibrio metoecus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metschnikovii | Vibrio | metschnikovii | 2.18 | Vibrio | metschnikovii | 2.15 | A | 1 | 1 | 0 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 619 sir | 2; 0 | eDT | 2022-06-23 | Vibrio | metschnikovii | Vibrio | metschnikovii | 2.35 | Vibrio | metschnikovii | 2.3 | A | 1 | 1 | 0 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7872 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio | mimicus | Vibrio | mimicus | 2.41 | Vibrio | metoecus | 2.37 | B | 0 | 0 | 1 | Vibrio mimicus | Vibrio metoecus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19130 | 4; type strain; WGS sequenced | DT | 2021-02-24 | Vibrio | mimicus | Vibrio | mimicus | 2.31 | Vibrio | mimicus | 2.28 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00299 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | navarrensis | Vibrio | navarrensis | 2.33 | Vibrio | navarrensis | 2.21 | A | 1 | 1 | 0 | Vibrio navarrensis | Vibrio navarrensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | navarrensis | Vibrio | navarrensis | 2.16 | Vibrio | navarrensis | 1.99 | A | 1 | 1 | 0 | Vibrio navarrensis | Vibrio navarrensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio | owensii | Vibrio | owensii | 2.35 | Vibrio | jasicida | 2.27 | B | 0 | 0 | 1 | Vibrio owensii | Vibrio jasicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00254-0 | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.28 | Vibrio | parahaemolyticus | 2.26 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10027 sir | 3; type strain | eDT | 2022-06-23 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.21 | Vibrio | parahaemolyticus | 2.16 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 871.2 | 3; sequenced | eDT | 2014-08-21 | Vibrio | sp-CVUAS-871,2 | Vibrio | sp-CVUAS-871,2 | 2.28 | Vibrio | mimicus | 1.45 | A | 1 | 1 | 0 | Vibrio sp-CVUAS-871,2 | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778.2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio | tubiashii | Vibrio | tubiashii | 2.05 | Vibrio | europaeus | 2.03 | B | 0 | 0 | 1 | Vibrio tubiashii | Vibrio europaeus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778.2 | 3; sequenced | eDT | 2019-10-24 | Vibrio | tubiashii | Vibrio | tubiashii | 2.51 | Vibrio | brasiliensis | 2.18 | B | 0 | 0 | 1 | Vibrio tubiashii | Vibrio brasiliensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio | vulnificus | Vibrio | vulnificus | 2.45 | Vibrio | vulnificus | 2.35 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LVM05 LVU05 extr | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-14 | Vibrio | vulnificus | Vibrio | vulnificus | 2.46 | Vibrio | vulnificus | 2.45 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas | sp-G9077 | Wohlfahrtiimonas | sp-G9077 | 2.42 | Wohlfahrtiimonas | chitiniclastica | 2.35 | B | 0 | 0 | 1 | Wohlfahrtiimonas sp-G9077 | Wohlfahrtiimonas chitiniclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria | indica | Ignatzschineria | indica | 2.43 | Ignatzschineria | indica | 2.18 | A | 1 | 1 | 0 | Ignatzschineria indica | Ignatzschineria indica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria | ureiclastica | Ignatzschineria | ureiclastica | 2.58 | Ignatzschineria | ureiclastica | 2.42 | A | 1 | 1 | 0 | Ignatzschineria ureiclastica | Ignatzschineria ureiclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas | chelatiphaga | Stenotrophomonas | chelatiphaga | 2.14 | Stenotrophomonas | maltophilia | 1.61 | A | 1 | 1 | 0 | Stenotrophomonas chelatiphaga | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-09-03 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.26 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-05-14 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.22 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.42 | Stenotrophomonas | rhizophila | 1.8 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas rhizophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-18 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.77 | Stenotrophomonas | rhizophila | 2.4 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas rhizophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4239.2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas | tumulicola | Stenotrophomonas | tumulicola | 2.47 | Stenotrophomonas | chelatiphaga | 1.69 | A | 1 | 1 | 0 | Stenotrophomonas tumulicola | Stenotrophomonas chelatiphaga | 0 | 0 | 0 | 0

## #Parameter (Spezies)
Parameter: | Staphylococcus | haemolyticus-borealis | Codierung | Validität | Methode | Datum | verwendeter Organismus | Score | richtig-negativ | falsch-positiv
Test auf spektrale Nähe zum Parameter |  |  | A281_1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-05-31 | Alkalihalobacillus clausii | 1.03 | 1 | 0
Anzahl Isolate: | 2391 |  | CVUAS 31071.2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus clausii | 0.58 | 1 | 0
|  |  | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Alkalihalobacillus clausii | 0.87 | 1 | 0
richtig-negativ: | 2391 |  | Oman 24 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-06-08 | Alkalihalobacillus clausii | 0.67 | 1 | 0
|  |  | DSM 13666 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Alkalihalobacillus okuhidensis DSM 13666: preparation and measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-19 | Alkalihalobacillus okuhidensis | 1.08 | 1 | 0
falsch-positiv: | 0 |  | BT2.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-13 | Bacillus altitudinis | 0.51 | 1 | 0
|  |  | DSM 21631 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Strain obtained from DSMZ strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-05-09 | Bacillus altitudinis | 0.73 | 1 | 0
| 0% |  | DSM 26896 | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-05 | Bacillus altitudinis | 0.67 | 1 | 0
|  |  | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 2015-02-13 | Bacillus alveayuensis | 0.76 | 1 | 0
|  |  | DSM 7 B159 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2016-03-01 | Bacillus amyloliquefaciens | 0.86 | 1 | 0
Scorewertgrenzen: | 2 | Species | DSM 7 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2016-03-01 | Bacillus amyloliquefaciens | 0.87 | 1 | 0
| 1.7 | Genus | DSM 7 CVUAS | 3; type strain | DT | 2014-01-29 | Bacillus amyloliquefaciens | 0.98 | 1 | 0
Es wird gegen eine nur aus dem Parameter bestehende DB geprüft. |  |  | IGV 015 IGV | 3; sequenced | DT | 2016-11-30 | Bacillus amyloliquefaciens | 1.12 | 1 | 0
|  |  | IGV 016 IGV | 3; sequenced | DT | 2016-11-30 | Bacillus amyloliquefaciens | 0.77 | 1 | 0
|  |  | ABI02A | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-07-14 | Bacillus atrophaeus | 0.82 | 1 | 0
|  |  | ABI05A | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 21.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-11 | Bacillus atrophaeus | 0.41 | 1 | 0
|  |  | ALEQ 1 | 2; strain collection of the RKI | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2022-05-24 | Bacillus atrophaeus | 0.46 | 1 | 0
|  |  | ATCC 9372 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus atrophaeus | 0.86 | 1 | 0
|  |  | B 4 CVUAS | 2; 0 | DT | 2014-01-22 | Bacillus atrophaeus | 0.55 | 1 | 0
|  |  | B 4 RKI | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus atrophaeus | 0.66 | 1 | 0
|  |  | CCM 878 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Bacillus atrophaeus | 0.88 | 1 | 0
|  |  | CKL1 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-03-02 | Bacillus atrophaeus | 0.67 | 1 | 0
|  |  | CKL1 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 19.04.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-08 | Bacillus atrophaeus | 0.91 | 1 | 0
|  |  | DSM 2277 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-03-05 | Bacillus atrophaeus | 0.77 | 1 | 0
|  |  | DSM 675 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-04-13 | Bacillus atrophaeus | 0.96 | 1 | 0
|  |  | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-15 | Bacillus atrophaeus | 0.73 | 1 | 0
|  |  | EZ05-02 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2018-08-27 | Bacillus atrophaeus | 0.7 | 1 | 0
|  |  | EZ05-02 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 26.09.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2022-09-27 | Bacillus atrophaeus | 0.64 | 1 | 0
|  |  | GBSC56 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 22.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-09 | Bacillus atrophaeus | 1.06 | 1 | 0
|  |  | LSSC22 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 15.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-09 | Bacillus atrophaeus | 0.81 | 1 | 0
|  |  | LSSC3 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 02.08.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2019-01-09 | Bacillus atrophaeus | 0.81 | 1 | 0
|  |  | NMTD54 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 15.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-09 | Bacillus atrophaeus | 0.97 | 1 | 0
|  |  | TS1 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 26.09.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-09-27 | Bacillus atrophaeus | 0.95 | 1 | 0
|  |  | TS1 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 31.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-19 | Bacillus atrophaeus | 0.77 | 1 | 0
|  |  | TS1 (03) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 31.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-16 | Bacillus atrophaeus | 0.69 | 1 | 0
|  |  | DGL6 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 07.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-19 | Bacillus axarquiensis | 0.88 | 1 | 0
|  |  | LNXM37 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-16 | Bacillus axarquiensis | 0.97 | 1 | 0
|  |  | LNXM78 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 16.06.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2016-06-16 | Bacillus axarquiensis | 0.82 | 1 | 0
|  |  | CCEB 623 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-10 | Bacillus badius | 0.75 | 1 | 0
|  |  | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Bacillus badius | 0.98 | 1 | 0
|  |  | CIP 111456 | 3; type strain | eDT | 2025-03-18 | Bacillus albus | 0.72 | 1 | 0
|  |  | LMG 28875 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-01-25 | Bacillus albus | 1 | 1 | 0
|  |  | A1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-11 | Bacillus anthracis | 0.72 | 1 | 0
|  |  | A10 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus anthracis | 0.77 | 1 | 0
|  |  | A100 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus anthracis | 0.83 | 1 | 0
|  |  | A101 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus anthracis | 1.07 | 1 | 0
|  |  | A102 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus anthracis | 0.65 | 1 | 0
|  |  | A103 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus anthracis | 0.59 | 1 | 0
|  |  | A106 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-03 | Bacillus anthracis | 0.4 | 1 | 0
|  |  | A107 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-03 | Bacillus anthracis | 0.54 | 1 | 0
|  |  | A108 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus anthracis | 0.89 | 1 | 0
|  |  | A11 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-29 | Bacillus anthracis | 0.29 | 1 | 0
|  |  | A110 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-06 | Bacillus anthracis | 0.71 | 1 | 0
|  |  | A111 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus anthracis | 0.86 | 1 | 0
|  |  | A112 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-04 | Bacillus anthracis | 1.06 | 1 | 0
|  |  | A112-1 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-02-08 | Bacillus anthracis | 0.7 | 1 | 0
|  |  | A113 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-09 | Bacillus anthracis | 0.72 | 1 | 0
|  |  | A118 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus anthracis | 0.82 | 1 | 0
|  |  | A12 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus anthracis | 0.66 | 1 | 0
|  |  | A120 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus anthracis | 1.07 | 1 | 0
|  |  | A122 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus anthracis | 0.95 | 1 | 0
|  |  | A123 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus anthracis | 0.93 | 1 | 0
|  |  | A124 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus anthracis | 0.9 | 1 | 0
|  |  | A125 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-27 | Bacillus anthracis | 0.99 | 1 | 0
|  |  | A126 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus anthracis | 0.79 | 1 | 0
|  |  | A127 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-10 | Bacillus anthracis | 0.94 | 1 | 0
|  |  | A128 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-04 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A130 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-11 | Bacillus anthracis | 0.53 | 1 | 0
|  |  | A131 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.92 | 1 | 0
|  |  | A132 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-28 | Bacillus anthracis | 0.5 | 1 | 0
|  |  | A133 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-05 | Bacillus anthracis | 0.76 | 1 | 0
|  |  | A134 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-28 | Bacillus anthracis | 1 | 1 | 0
|  |  | A135 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.71 | 1 | 0
|  |  | A136 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-28 | Bacillus anthracis | 1 | 1 | 0
|  |  | A137 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-13 | Bacillus anthracis | 0.64 | 1 | 0
|  |  | A138-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-13 | Bacillus anthracis | 0.86 | 1 | 0
|  |  | A15 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-01 | Bacillus anthracis | 0.98 | 1 | 0
|  |  | A16 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-01 | Bacillus anthracis | 0.73 | 1 | 0
|  |  | A173-2 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-05-30 | Bacillus anthracis | 0.83 | 1 | 0
|  |  | A18 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-05 | Bacillus anthracis | 1.08 | 1 | 0
|  |  | A19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus anthracis | 0.6 | 1 | 0
|  |  | A22 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus anthracis | 0.94 | 1 | 0
|  |  | A23 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus anthracis | 0.9 | 1 | 0
|  |  | A24 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus anthracis | 0.58 | 1 | 0
|  |  | A25 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-05 | Bacillus anthracis | 0.75 | 1 | 0
|  |  | A269-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-02 | Bacillus anthracis | 0.96 | 1 | 0
|  |  | A27 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A28 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus anthracis | 0.57 | 1 | 0
|  |  | A29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus anthracis | 0.62 | 1 | 0
|  |  | A294 1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus anthracis | 0.74 | 1 | 0
|  |  | A3 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-19 | Bacillus anthracis | 0.49 | 1 | 0
|  |  | A30 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-05 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A300-1a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-07 | Bacillus anthracis | 0.9 | 1 | 0
|  |  | A300-1b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-14 | Bacillus anthracis | 0.73 | 1 | 0
|  |  | A300-2a | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-10-23 | Bacillus anthracis | 0.76 | 1 | 0
|  |  | A300-2b | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-13 | Bacillus anthracis | 0.41 | 1 | 0
|  |  | A302-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-11 | Bacillus anthracis | 0.81 | 1 | 0
|  |  | A306-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-12 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A315-1 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-12 | Bacillus anthracis | 0.72 | 1 | 0
|  |  | A315-2 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-03-12 | Bacillus anthracis | 0.78 | 1 | 0
|  |  | A32 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-08 | Bacillus anthracis | 0.83 | 1 | 0
|  |  | A33 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus anthracis | 0.5 | 1 | 0
|  |  | A34 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus anthracis | 0.72 | 1 | 0
|  |  | A35 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-06 | Bacillus anthracis | 0.68 | 1 | 0
|  |  | A36 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-06 | Bacillus anthracis | 0.62 | 1 | 0
|  |  | A37 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus anthracis | 0.63 | 1 | 0
|  |  | A38 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus anthracis | 0.73 | 1 | 0
|  |  | A39 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus anthracis | 0.71 | 1 | 0
|  |  | A4 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A40 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-06 | Bacillus anthracis | 0.57 | 1 | 0
|  |  | A41 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus anthracis | 0.81 | 1 | 0
|  |  | A42 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus anthracis | 1.01 | 1 | 0
|  |  | A43 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-09 | Bacillus anthracis | 0.82 | 1 | 0
|  |  | A44 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus anthracis | 0.42 | 1 | 0
|  |  | A45 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus anthracis | 0.6 | 1 | 0
|  |  | A46 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus anthracis | 0.74 | 1 | 0
|  |  | A49 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-06-22 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A5 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus anthracis | 0.85 | 1 | 0
|  |  | A58 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.89 | 1 | 0
|  |  | A59 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.84 | 1 | 0
|  |  | A6 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-19 | Bacillus anthracis | 1.03 | 1 | 0
|  |  | A60 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.58 | 1 | 0
|  |  | A61 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.69 | 1 | 0
|  |  | A62 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.75 | 1 | 0
|  |  | A63 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-10 | Bacillus anthracis | 0.78 | 1 | 0
|  |  | A64 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.7 | 1 | 0
|  |  | A65 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 1.26 | 1 | 0
|  |  | A66 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.69 | 1 | 0
|  |  | A67 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-11 | Bacillus anthracis | 0.91 | 1 | 0
|  |  | A68 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-12 | Bacillus anthracis | 0.51 | 1 | 0
|  |  | A69 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus anthracis | 1.09 | 1 | 0
|  |  | A7 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus anthracis | 0.58 | 1 | 0
|  |  | A72 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-12 | Bacillus anthracis | 0.77 | 1 | 0
|  |  | A73 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-12 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A74 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus anthracis | 0.72 | 1 | 0
|  |  | A8 ant | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-19 | Bacillus anthracis | 1.02 | 1 | 0
|  |  | A80 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus anthracis | 1.07 | 1 | 0
|  |  | A81 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus anthracis | 0.91 | 1 | 0
|  |  | A82 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus anthracis | 1.16 | 1 | 0
|  |  | A83 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus anthracis | 0.93 | 1 | 0
|  |  | A84 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-01 | Bacillus anthracis | 0.82 | 1 | 0
|  |  | A85 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus anthracis | 0.68 | 1 | 0
|  |  | A86 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-03 | Bacillus anthracis | 0.96 | 1 | 0
|  |  | A87 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus anthracis | 0.87 | 1 | 0
|  |  | A88 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus anthracis | 1.1 | 1 | 0
|  |  | A88 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus anthracis | 0.68 | 1 | 0
|  |  | A89 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-13 | Bacillus anthracis | 1 | 1 | 0
|  |  | A89 schwarz | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus anthracis | 0.71 | 1 | 0
|  |  | A9 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus anthracis | 0.7 | 1 | 0
|  |  | A90 blau | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus anthracis | 0.73 | 1 | 0
|  |  | A90 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-17 | Bacillus anthracis | 0.76 | 1 | 0
|  |  | A91 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-24 | Bacillus anthracis | 0.76 | 1 | 0
|  |  | A92 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus anthracis | 0.79 | 1 | 0
|  |  | A93 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-04 | Bacillus anthracis | 0.98 | 1 | 0
|  |  | AF 28 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.83 | 1 | 0
|  |  | AF 29 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.64 | 1 | 0
|  |  | AF 31 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.85 | 1 | 0
|  |  | AF 71 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.81 | 1 | 0
|  |  | AF 96 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.7 | 1 | 0
|  |  | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-11-09 | Bacillus anthracis | 0.66 | 1 | 0
|  |  | ATCC 14185 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-04-03 | Bacillus anthracis | 0.64 | 1 | 0
|  |  | ATCC 14578 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-08-13 | Bacillus anthracis | 0.72 | 1 | 0
|  |  | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-26 | Bacillus anthracis | 0.75 | 1 | 0
|  |  | SA 20 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-11 | Bacillus anthracis | 0.92 | 1 | 0
|  |  | Sterne | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus anthracis | 0.94 | 1 | 0
|  |  | vollum | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-03-06 | Bacillus anthracis | 0.52 | 1 | 0
|  |  | Wirt | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus anthracis | 1.03 | 1 | 0
|  |  | CIP 111020 | 3; public strain collection | eDT | 2025-03-18 | Bacillus bingmayongensis | 1.06 | 1 | 0
|  |  | DSM 25427 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 05.10.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus bingmayongensis DSM 25427, preparation and measurements: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-11-05 | Bacillus bingmayongensis | 0.84 | 1 | 0
|  |  | AaCol | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-03-06 | Bacillus cereus | 0.76 | 1 | 0
|  |  | SK1b | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-03-06 | Bacillus cereus | 0.6 | 1 | 0
|  |  | B194 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus cereus | 1.01 | 1 | 0
|  |  | B195 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus cereus | 0.89 | 1 | 0
|  |  | B196 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus cereus | 0.85 | 1 | 0
|  |  | CVUAS 34551 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus cereus | 0.94 | 1 | 0
|  |  | CVUAS 34552 | 2; 0 | EFex [BRU03] | 2022-12-14 | Bacillus cereus | 0.55 | 1 | 0
|  |  | B21 | 2; strain collection of the RKI | see [0185] | 2015-12-02 | Bacillus cereus | 0.78 | 1 | 0
|  |  | B22 | 2; strain collection of the RKI | see [0185] | 2015-12-02 | Bacillus cereus | 1.32 | 1 | 0
|  |  | B 340 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus cereus | 0.91 | 1 | 0
|  |  | B 341 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus cereus | 1.02 | 1 | 0
|  |  | B 342 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus cereus | 1.26 | 1 | 0
|  |  | B 343 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus cereus | 0.66 | 1 | 0
|  |  | A-071-2-A-2 Bac 25 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus sp. A-071; 2-A-2 Bac 25: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-26 | Bacillus cereus | 1.07 | 1 | 0
|  |  | A-082-2 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 29.02.2022 | Bacillus cereus | 0.76 | 1 | 0
|  |  | A-082-4 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 13.04.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 29.02.2022 | Bacillus cereus | 0.74 | 1 | 0
|  |  | A-175-2 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2021-03-31 | Bacillus cereus | 0.87 | 1 | 0
|  |  | A-550-5 (SC6) Bac 50 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus sp. A-550; 5 (SC6) Bac 50: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-12 | Bacillus cereus | 0.95 | 1 | 0
|  |  | BT250c Bac 33 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus sp. A-550; 5 (SC6) Bac 50: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-12 | Bacillus cereus | 0.63 | 1 | 0
|  |  | CD 3-1a | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-05-11 | Bacillus cereus | 1.03 | 1 | 0
|  |  | CD 3-1c | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-05-11 | Bacillus cereus | 0.84 | 1 | 0
|  |  | CD3.5 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-11-16 | Bacillus cereus | 0.74 | 1 | 0
|  |  | HD1.3 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-08-01 | Bacillus cereus | 0.8 | 1 | 0
|  |  | KN17-s3053 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=Source: AKRA Austria, Dr. Karner#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=Dr. Karger AKRA (Austria)# | 2022-05-11 | Bacillus cereus | 0.9 | 1 | 0
|  |  | DSM 4222 | 3; public strain collection | eDT | 2025-03-18 | Bacillus cereus | 0.51 | 1 | 0
|  |  | 161004439-1 | 2; 0 | DT | 2016-04-12 | Bacillus cereus | 0.97 | 1 | 0
|  |  | 161004439-2 | 2; 0 | DT | 2016-04-12 | Bacillus cereus | 1.03 | 1 | 0
|  |  | 161005557-1 | 2; 0 | DT | 2016-08-11 | Bacillus cereus | 0.83 | 1 | 0
|  |  | 171012589-5 | 2; 0 | DT | 2017-09-19 | Bacillus cereus | 0.93 | 1 | 0
|  |  | 181015909-7 | 2; 0 | DT | 2018-11-02 | Bacillus cereus | 0.47 | 1 | 0
|  |  | 2617 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-03-31 | Bacillus cereus | 0.49 | 1 | 0
|  |  | 301101RA0432 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-12 | Bacillus cereus | 0.86 | 1 | 0
|  |  | A-049 | 2; strain collection of the RKI | #CON=approx. 2 loops in 25 mkL#TRT=in ca. 80 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 30.03.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Strain from strain collection ZBS 3, Cultivation & preparation: M. Staemmler#LAS=LL43 pct. 1000, "LL70 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2022 02 28(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert220228.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2022-03-31 | Bacillus cereus | 0.95 | 1 | 0
|  |  | A-1607 12B | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus cereus A-1607; 12B: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Bacillus cereus | 0.58 | 1 | 0
|  |  | A22 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam I)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus cereus | 0.82 | 1 | 0
|  |  | A24 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus cereus | 0.79 | 1 | 0
|  |  | A31 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus cereus | 1.03 | 1 | 0
|  |  | A42 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam I)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus cereus | 0.84 | 1 | 0
|  |  | A8 cer | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus cereus | 0.78 | 1 | 0
|  |  | ATCC 10876 | 3; public strain collection | TFA Extr | 2015-04-29 | Bacillus cereus | 1.09 | 1 | 0
|  |  | ATCC 11778 IGV | 3; public strain collection | DT | 2016-02-03 | Bacillus cereus | 1.01 | 1 | 0
|  |  | ATCC 11778 LHL | 3; public strain collection | DT | 2015-06-23 | Bacillus cereus | 0.61 | 1 | 0
|  |  | ATCC 11778 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 2007-03-05 | Bacillus cereus | 0.8 | 1 | 0
|  |  | ATCC 14579 IGV | 3; type strain | DT | 2016-04-06 | Bacillus cereus | 1.06 | 1 | 0
|  |  | ATCC 14579 RKI | 3; type strain | in 80% TFA/ 1:1 HCCATA2 (A) | 2007-03-05 | Bacillus cereus | 0.53 | 1 | 0
|  |  | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus cereus | 0.7 | 1 | 0
|  |  | ATCC 33018 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-04 | Bacillus cereus | 0.92 | 1 | 0
|  |  | ATCC 33019 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-22 | Bacillus cereus | 0.84 | 1 | 0
|  |  | ATCC 6464 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus cereus | 0.93 | 1 | 0
|  |  | ATCC 7004 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus cereus | 0.99 | 1 | 0
|  |  | ATCC 9139 | 3; reference strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus cereus | 0.78 | 1 | 0
|  |  | B10 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-01 | Bacillus cereus | 0.83 | 1 | 0
|  |  | B108 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-15 | Bacillus cereus | 0.98 | 1 | 0
|  |  | B109 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-13 | Bacillus cereus | 0.94 | 1 | 0
|  |  | B110 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-15 | Bacillus cereus | 1.11 | 1 | 0
|  |  | B146 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-27 | Bacillus cereus | 0.82 | 1 | 0
|  |  | B16 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-02 | Bacillus cereus | 1.04 | 1 | 0
|  |  | B197 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus cereus | 0.85 | 1 | 0
|  |  | B198 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-14 | Bacillus cereus | 0.81 | 1 | 0
|  |  | B199 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus cereus | 1.14 | 1 | 0
|  |  | B200 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-17 | Bacillus cereus | 0.96 | 1 | 0
|  |  | B232 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-18 | Bacillus cereus | 1.11 | 1 | 0
|  |  | B 239 | 3; public strain collection | see [0185] | 2015-04-22 | Bacillus cereus | 1.05 | 1 | 0
|  |  | B252 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus cereus | 0.79 | 1 | 0
|  |  | B253 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus cereus | 1.05 | 1 | 0
|  |  | B254 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus cereus | 0.83 | 1 | 0
|  |  | B255 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus cereus | 0.92 | 1 | 0
|  |  | B257 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus cereus | 0.82 | 1 | 0
|  |  | B258 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus cereus | 0.86 | 1 | 0
|  |  | B259 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus cereus | 1.1 | 1 | 0
|  |  | B261 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus cereus | 1.16 | 1 | 0
|  |  | B262 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-14 | Bacillus cereus | 0.87 | 1 | 0
|  |  | B272 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus cereus | 0.99 | 1 | 0
|  |  | B273 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus cereus | 0.82 | 1 | 0
|  |  | B274 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus cereus | 0.88 | 1 | 0
|  |  | B275 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus cereus | 0.99 | 1 | 0
|  |  | B277 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus cereus | 0.86 | 1 | 0
|  |  | B278 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus cereus | 0.92 | 1 | 0
|  |  | B282 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-29 | Bacillus cereus | 0.67 | 1 | 0
|  |  | B293 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-19 | Bacillus cereus | 0.94 | 1 | 0
|  |  | B294 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus cereus | 0.74 | 1 | 0
|  |  | B295 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus cereus | 0.78 | 1 | 0
|  |  | B296 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-30 | Bacillus cereus | 0.74 | 1 | 0
|  |  | B297 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 0.86 | 1 | 0
|  |  | B298 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 0.79 | 1 | 0
|  |  | B299 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-19 | Bacillus cereus | 0.81 | 1 | 0
|  |  | B3 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-11 | Bacillus cereus | 0.8 | 1 | 0
|  |  | B300 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 0.76 | 1 | 0
|  |  | B302 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 0.77 | 1 | 0
|  |  | B303 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 1 | 1 | 0
|  |  | B304 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 0.74 | 1 | 0
|  |  | B305 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 0.88 | 1 | 0
|  |  | B306 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 1.05 | 1 | 0
|  |  | B308 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-05 | Bacillus cereus | 0.99 | 1 | 0
|  |  | B309 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-22 | Bacillus cereus | 0.82 | 1 | 0
|  |  | B310 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus cereus | 0.98 | 1 | 0
|  |  | B311 | 3; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus cereus | 0.87 | 1 | 0
|  |  | B316 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus cereus | 0.76 | 1 | 0
|  |  | B329 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 1:10 diluted , spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Cultivation & sample preparation by W. Beyer (University Hohenheim); Arrival: 09.08.2006; Preparation for MALDI-ToF MS#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Dr. Beyer (Uni Hohenheim)# | 2015-05-07 | Bacillus cereus | 0.76 | 1 | 0
|  |  | B33 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-10 | Bacillus cereus | 0.75 | 1 | 0
|  |  | B332 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-07 | Bacillus cereus | 0.91 | 1 | 0
|  |  | CCM 88 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-14 | Bacillus cereus | 0.84 | 1 | 0
|  |  | BW-A | 2; strain collection of the RKI | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:4 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2021-06-21 | Bacillus cereus | 0.66 | 1 | 0
|  |  | BW-B | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus cereus | 0.82 | 1 | 0
|  |  | CCM 99 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus cereus | 0.97 | 1 | 0
|  |  | CVUAS 1107 | 2; 0 | eDT | 2016-01-27 | Bacillus cereus | 1.07 | 1 | 0
|  |  | CVUAS 3901 | 3; 0 | eDT | 2016-04-26 | Bacillus cereus | 0.88 | 1 | 0
|  |  | DSM 1274 | 3; public strain collection ? | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-23 | Bacillus cereus | 0.87 | 1 | 0
|  |  | DSM 2299 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-16 | Bacillus cereus | 0.99 | 1 | 0
|  |  | DSM 2301 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus cereus | 0.55 | 1 | 0
|  |  | DSM 2301 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Bacillus cereus DSM 2301: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Bacillus cereus | 0.76 | 1 | 0
|  |  | DSM 2302 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus cereus | 0.95 | 1 | 0
|  |  | DSM 31 B181 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-30 | Bacillus cereus | 0.54 | 1 | 0
|  |  | DSM 31 B69 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-02-06 | Bacillus cereus | 0.55 | 1 | 0
|  |  | DSM 31 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus cereus | 1.11 | 1 | 0
|  |  | DSM 31 RKI | 3; type strain | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-09-21 | Bacillus cereus | 0.78 | 1 | 0
|  |  | DSM 3101 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-20 | Bacillus cereus | 1.01 | 1 | 0
|  |  | DSM 318 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-20 | Bacillus cereus | 0.62 | 1 | 0
|  |  | DSM 336 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-18 | Bacillus cereus | 0.66 | 1 | 0
|  |  | DSM 345 B170 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-30 | Bacillus cereus | 0.67 | 1 | 0
|  |  | DSM 345 B23 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-11 | Bacillus cereus | 1.1 | 1 | 0
|  |  | DSM 345 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus cereus | 1.05 | 1 | 0
|  |  | DSM 345 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 02.12.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus cereus DSM 345, preparation and measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-12-02 | Bacillus cereus | 0.89 | 1 | 0
|  |  | DSM 351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-18 | Bacillus cereus | 0.81 | 1 | 0
|  |  | DSM 3648 | 3; public strain collection | eDT | 2025-03-18 | Bacillus cereus | 0.49 | 1 | 0
|  |  | DSM 4218 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-20 | Bacillus cereus | 0.74 | 1 | 0
|  |  | DSM 4282 | 3; public strain collection | eDT | 2025-03-18 | Bacillus cereus | 0.54 | 1 | 0
|  |  | DSM 4312 B269 RKI | 3; public strain collection | TFA Extr | 2006-09-15 | Bacillus cereus | 0.93 | 1 | 0
|  |  | DSM 4312 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus cereus | 0.97 | 1 | 0
|  |  | DSM 4312 IGV | 3; public strain collection | DT | 2016-04-13 | Bacillus cereus | 1.1 | 1 | 0
|  |  | DSM 4312 RKI | 3; public strain collection | TFA Extr | 2016-06-30 | Bacillus cereus | 0.76 | 1 | 0
|  |  | DSM 4313 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus cereus | 0.85 | 1 | 0
|  |  | DSM 4313 IGV | 3; public strain collection | DT | 2016-04-06 | Bacillus cereus | 0.95 | 1 | 0
|  |  | DSM 4384 | 3; public strain collection | eDT | 2025-03-18 | Bacillus cereus | 0.71 | 1 | 0
|  |  | DSM 4490 | 3; public strain collection | in 80% TFA, Zentrifugation und Sterilfiltration, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-09-21 | Bacillus cereus | 0.81 | 1 | 0
|  |  | DSM 487 RKI | 3; public strain collection | see [0185] | 2021-12-03 | Bacillus cereus | 0.67 | 1 | 0
|  |  | DSM 487 B169 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus cereus | 0.73 | 1 | 0
|  |  | DSM 487 B185 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus cereus | 0.85 | 1 | 0
|  |  | DSM 508 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-04-07 | Bacillus cereus | 0.58 | 1 | 0
|  |  | DSM 609 | 3; strain from public collection | in 80% TFA, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-09-21 | Bacillus cereus | 0.94 | 1 | 0
|  |  | B241 | 3; public strain collection | TFA Extr | 2006-09-14 | Bacillus cereus | 0.81 | 1 | 0
|  |  | DSM 6127 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus cereus | 0.73 | 1 | 0
|  |  | DSM 6127 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 02.12.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus cereus DSM 6127, preparation and measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-12-02 | Bacillus cereus | 1.1 | 1 | 0
|  |  | DSM 626 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus cereus | 0.95 | 1 | 0
|  |  | DSM 679 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-14 | Bacillus cereus | 0.81 | 1 | 0
|  |  | DSM 689 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-09-14 | Bacillus cereus | 0.88 | 1 | 0
|  |  | DSM 8438 | 3; public strain collection | TFAextr | 2015-04-23 | Bacillus cereus | 0.78 | 1 | 0
|  |  | DSM 9378 | 3; public strain collection | eDT | 2025-03-18 | Bacillus cereus | 0.58 | 1 | 0
|  |  | HB3.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus cereus | 0.69 | 1 | 0
|  |  | HD1.4B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-16 | Bacillus cereus | 0.67 | 1 | 0
|  |  | Hohenheim | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-16 | Bacillus cereus | 1.12 | 1 | 0
|  |  | MS2.1B | 2; strain collection of the RKI | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2020-08-12 | Bacillus cereus | 1 | 1 | 0
|  |  | NCIMB 11925 | 3; reference strain collection | DT | 2015-07-22 | Bacillus cereus | 0.53 | 1 | 0
|  |  | NCTC 10320 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus cereus | 1.07 | 1 | 0
|  |  | NCTC 10989 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus cereus | 0.71 | 1 | 0
|  |  | NCTC 6474 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus cereus | 1.03 | 1 | 0
|  |  | NCTC 9939 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-28 | Bacillus cereus | 0.92 | 1 | 0
|  |  | NCTC 9945. ATCC 10876a (B264. Beyer) | 3; public strain collection | TFA Extr | 2015-04-28 | Bacillus cereus | 0.74 | 1 | 0
|  |  | NiMe_1 | 3; 16S, ces-pos | eDT | 2016-01-27 | Bacillus cereus | 0.88 | 1 | 0
|  |  | SL2009001164 | 3; 0 | eDT | 2016-01-27 | Bacillus cereus | 0.82 | 1 | 0
|  |  | SL2009005808 | 2; 0 | eDT | 2019-09-23 | Bacillus cereus | 0.94 | 1 | 0
|  |  | SL2009007190 | 2; 0 | eDT | 2019-09-25 | Bacillus cereus | 0.8 | 1 | 0
|  |  | SL2014002835 | 2; CytK2 pos, Krist neg | eDT | 2019-09-25 | Bacillus cereus | 0.48 | 1 | 0
|  |  | SN4-3 | 2; strain collection of the RKI | see [0185] | 2020-10-11 | Bacillus cereus | 0.84 | 1 | 0
|  |  | Wasserprobe Bac RKI | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus cereus | 0.6 | 1 | 0
|  |  | WIS-St. Nr. 29 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus cereus | 0.43 | 1 | 0
|  |  | WIS-St. Nr. 311 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus cereus | 0.68 | 1 | 0
|  |  | WSBC 10028 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus cereus | 0.94 | 1 | 0
|  |  | WSBC 10030 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus cereus | 1.12 | 1 | 0
|  |  | WSBC 10286 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus cereus | 0.69 | 1 | 0
|  |  | WSBC 10288 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus cereus | 1.05 | 1 | 0
|  |  | WSBC 10441 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus cereus | 1.26 | 1 | 0
|  |  | WSBC 10925 | 3; well defined isolat, ces-pos | eDT | 2016-01-27 | Bacillus cereus | 1.05 | 1 | 0
|  |  | 46776 CVUAS | 3; specific PCR, growth at 52°C | eDT | 2017-05-31 | Bacillus cytotoxicus | 0.92 | 1 | 0
|  |  | 46776 MVZ | 3; specific PCR, growth at 52°C | DT | 2017-05-23 | Bacillus cytotoxicus | 0.85 | 1 | 0
|  |  | 46778 CVUAS | 3; specific PCR, growth at 52°C | eDT | 2017-05-31 | Bacillus cytotoxicus | 0.96 | 1 | 0
|  |  | 46778 MVZ | 3; specific PCR, growth at 52°C | DT | 2017-05-23 | Bacillus cytotoxicus | 0.8 | 1 | 0
|  |  | Ad 2163 | 3; specific PCR | eDT | 2017-12-15 | Bacillus cytotoxicus | 0.81 | 1 | 0
|  |  | Ad 2164 | 3; specific PCR | eDT | 2017-12-15 | Bacillus cytotoxicus | 1.05 | 1 | 0
|  |  | Ad 2466 | 3; specific PCR | eDT | 2017-12-15 | Bacillus cytotoxicus | 0.96 | 1 | 0
|  |  | Ad 2469 | 3; specific PCR | eDT | 2017-12-15 | Bacillus cytotoxicus | 0.8 | 1 | 0
|  |  | BfR-BA-01360 | 3; sequenced | 0 | 2020-10-06 | Bacillus cytotoxicus | 1.03 | 1 | 0
|  |  | BfR-BA-01430 | 3; sequenced | 0 | 2020-10-06 | Bacillus cytotoxicus | 0.95 | 1 | 0
|  |  | BfR-BA-01431 | 3; sequenced | 0 | 2020-10-06 | Bacillus cytotoxicus | 1.1 | 1 | 0
|  |  | BfR-BA-01492 | 3; sequenced | 0 | 2020-10-06 | Bacillus cytotoxicus | 0.91 | 1 | 0
|  |  | BfR-BA-01522 | 3; sequenced | 0 | 2020-10-06 | Bacillus cytotoxicus | 1.02 | 1 | 0
|  |  | CH1 | 4; WGS | 0 | 0 | Bacillus cytotoxicus | 1.04 | 1 | 0
|  |  | CH13 | 4; WGS | 0 | 0 | Bacillus cytotoxicus | 1.07 | 1 | 0
|  |  | CH134 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.74 | 1 | 0
|  |  | CH15 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.02 | 1 | 0
|  |  | CH16 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.99 | 1 | 0
|  |  | CH2 | 4; WGS | 0 | 0 | Bacillus cytotoxicus | 1.01 | 1 | 0
|  |  | CH212 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.84 | 1 | 0
|  |  | CH213 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.73 | 1 | 0
|  |  | CH214 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.69 | 1 | 0
|  |  | CH216 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.88 | 1 | 0
|  |  | CH217 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.88 | 1 | 0
|  |  | CH218 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.89 | 1 | 0
|  |  | CH23 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.8 | 1 | 0
|  |  | CH25 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.16 | 1 | 0
|  |  | CH3 | 4; WGS | 0 | 0 | Bacillus cytotoxicus | 0.97 | 1 | 0
|  |  | CH348 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.02 | 1 | 0
|  |  | CH38 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.06 | 1 | 0
|  |  | CH39 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.12 | 1 | 0
|  |  | CH4 | 4; WGS | 0 | 0 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CH647 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.05 | 1 | 0
|  |  | CH755 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.08 | 1 | 0
|  |  | CH834 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.8 | 1 | 0
|  |  | CH835 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CH836 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.04 | 1 | 0
|  |  | CH837 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.99 | 1 | 0
|  |  | CH838 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.04 | 1 | 0
|  |  | CH839 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.97 | 1 | 0
|  |  | CH840 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.87 | 1 | 0
|  |  | CH841 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.05 | 1 | 0
|  |  | CH842 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.88 | 1 | 0
|  |  | CH843 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.77 | 1 | 0
|  |  | CH844 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.49 | 1 | 0
|  |  | CH845 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.95 | 1 | 0
|  |  | CH846 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.94 | 1 | 0
|  |  | CH847 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.92 | 1 | 0
|  |  | CH848 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.85 | 1 | 0
|  |  | CH849 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.89 | 1 | 0
|  |  | CH850 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CH851 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.86 | 1 | 0
|  |  | CH852 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.83 | 1 | 0
|  |  | CH853 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.07 | 1 | 0
|  |  | CH854 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.84 | 1 | 0
|  |  | CH855 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.6 | 1 | 0
|  |  | CH856 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.02 | 1 | 0
|  |  | CH857 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.78 | 1 | 0
|  |  | CH858 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.87 | 1 | 0
|  |  | CH859 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.07 | 1 | 0
|  |  | CH860 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.93 | 1 | 0
|  |  | CH861 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.81 | 1 | 0
|  |  | CH862 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.11 | 1 | 0
|  |  | CH863 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.5 | 1 | 0
|  |  | CH864 | 3; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.09 | 1 | 0
|  |  | CH865 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.89 | 1 | 0
|  |  | CH866 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.88 | 1 | 0
|  |  | CH867 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.9 | 1 | 0
|  |  | CH868 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.92 | 1 | 0
|  |  | CH869 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.91 | 1 | 0
|  |  | CH870 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.12 | 1 | 0
|  |  | CH871 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.82 | 1 | 0
|  |  | CH872 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.82 | 1 | 0
|  |  | CH873 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.87 | 1 | 0
|  |  | CH874 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.95 | 1 | 0
|  |  | CH875 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.93 | 1 | 0
|  |  | CH876 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.73 | 1 | 0
|  |  | CH877 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.83 | 1 | 0
|  |  | CH878 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.12 | 1 | 0
|  |  | CH879 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.99 | 1 | 0
|  |  | CH880 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.16 | 1 | 0
|  |  | CH881 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.87 | 1 | 0
|  |  | CH882 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 1.29 | 1 | 0
|  |  | CH883 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.85 | 1 | 0
|  |  | CH884 | 4; WGS | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.9 | 1 | 0
|  |  | CH886 | 3; PCR | 0 | 2023-01-26 | Bacillus cytotoxicus | 0.97 | 1 | 0
|  |  | CVUAS 0644 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.85 | 1 | 0
|  |  | CVUAS 0644 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 0.99 | 1 | 0
|  |  | CVUAS 10162 | 3; sequenced | eDT | 2015-11-30 | Bacillus cytotoxicus | 0.78 | 1 | 0
|  |  | CVUAS 10577 | 4; WGS | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.71 | 1 | 0
|  |  | CVUAS 11025 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 1.22 | 1 | 0
|  |  | CVUAS 11326.2 | 4; WGS | eDT | 2022-11-18 | Bacillus cytotoxicus | 1.12 | 1 | 0
|  |  | CVUAS 2831 | 3; sequenced; Wt 55°C, cytK1 | DT | 2013-01-18 | Bacillus cytotoxicus | 0.88 | 1 | 0
|  |  | CVUAS 2833 CVUAS | 4; 16S, draft genome [0009] | eDT | 2016-01-27 | Bacillus cytotoxicus | 0.94 | 1 | 0
|  |  | CVUAS 2833 RKI | 4; 16S, draft genome [0009] | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 0.82 | 1 | 0
|  |  | CVUAS 2894 | 3; sequenced; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 0.89 | 1 | 0
|  |  | CVUAS 2905 | 3; sequenced; Wt 55°C, cytK1 | DT | 2013-09-30 | Bacillus cytotoxicus | 0.96 | 1 | 0
|  |  | CVUAS 2979 CVUAS | 4; WGS, 16S | eDT | 2016-01-27 | Bacillus cytotoxicus | 0.95 | 1 | 0
|  |  | CVUAS 2979 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 1 | 1 | 0
|  |  | CVUAS 2982 | 3; sequenced; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 0.84 | 1 | 0
|  |  | CVUAS 3017.3 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.58 | 1 | 0
|  |  | CVUAS 30572 | 1; 0 | EFex [BRU03] | 2017-11-22 | Bacillus cytotoxicus | 1.04 | 1 | 0
|  |  | CVUAS 30605 | 1; 0 | eDT | 2017-12-01 | Bacillus cytotoxicus | 0.76 | 1 | 0
|  |  | CVUAS 30626 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.82 | 1 | 0
|  |  | CVUAS 31073 | 4; WGS, specific PCR | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.89 | 1 | 0
|  |  | CVUAS 31094 | 4; WGS, specific PCR | eDT | 2018-08-14 | Bacillus cytotoxicus | 0.93 | 1 | 0
|  |  | CVUAS 35239 | 4; WGS; specific PCR | eDT | 2024-09-24 | Bacillus cytotoxicus | 0.62 | 1 | 0
|  |  | CVUAS 35361 | 4; WGS; specific PCR | eDT | 2024-09-24 | Bacillus cytotoxicus | 1.04 | 1 | 0
|  |  | CVUAS 355.3 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.76 | 1 | 0
|  |  | CVUAS 356.3 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.92 | 1 | 0
|  |  | CVUAS 363.4 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 1 | 1 | 0
|  |  | CVUAS 3827 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.89 | 1 | 0
|  |  | CVUAS 3827 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 0.97 | 1 | 0
|  |  | CVUAS 3828 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.93 | 1 | 0
|  |  | CVUAS 3828 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 1.03 | 1 | 0
|  |  | CVUAS 3896 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-14 | Bacillus cytotoxicus | 1.03 | 1 | 0
|  |  | CVUAS 4117 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-14 | Bacillus cytotoxicus | 1.07 | 1 | 0
|  |  | CVUAS 4545 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 1.06 | 1 | 0
|  |  | CVUAS 4546.4 | 3; cyt K1 | eDT | 2015-05-20 | Bacillus cytotoxicus | 0.73 | 1 | 0
|  |  | CVUAS 5334 | 3; Wt 55°C, cytK1 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.59 | 1 | 0
|  |  | CVUAS 5335 | 3; Wt 55°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 0.93 | 1 | 0
|  |  | CVUAS 5403 | 2; Wt 55°C | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.6 | 1 | 0
|  |  | CVUAS 5404 | 2; Wt 55°C | eDT | 2016-11-22 | Bacillus cytotoxicus | 1.19 | 1 | 0
|  |  | CVUAS 5405 | 2; Wt 55°C | eDT | 2016-11-22 | Bacillus cytotoxicus | 1.2 | 1 | 0
|  |  | CVUAS 6081 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.64 | 1 | 0
|  |  | CVUAS 6082 CVUAS | 4; WGS; Wt 50°C, cytK1 | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.63 | 1 | 0
|  |  | CVUAS 6082 RKI | 4; WGS; Wt 50°C, cytK1 | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 1.01 | 1 | 0
|  |  | CVUAS 6083 | 3; Wt 50°C, cytK1 | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CVUAS 6092 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.95 | 1 | 0
|  |  | CVUAS 6128 | 3; Wt 50°C, cytK1 | EFEx | 2015-01-13 | Bacillus cytotoxicus | 1 | 1 | 0
|  |  | CVUAS 6129 | 3; Wt 50°C, cytK1 | EFEx | 2015-01-15 | Bacillus cytotoxicus | 1.02 | 1 | 0
|  |  | CVUAS 6130 | 3; Wt 50°C, cytK1 | EFEx | 2015-01-13 | Bacillus cytotoxicus | 0.66 | 1 | 0
|  |  | CVUAS 6183 CVUAS | 3; Wt 55°C, cytK1 | eDT | 2016-11-16 | Bacillus cytotoxicus | 1.08 | 1 | 0
|  |  | CVUAS 6183 RKI | 3; Wt 55°C, cytK1 | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 0.82 | 1 | 0
|  |  | CVUAS 6204 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CVUAS 6205 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-14 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CVUAS 6206 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 0.85 | 1 | 0
|  |  | CVUAS 6207 | 3; Wt 50°C, cytK1 | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.94 | 1 | 0
|  |  | CVUAS 6208 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 0.95 | 1 | 0
|  |  | CVUAS 6209 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-03-17 | Bacillus cytotoxicus | 1.08 | 1 | 0
|  |  | CVUAS 6210 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus cytotoxicus | 0.91 | 1 | 0
|  |  | CVUAS 6211 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.79 | 1 | 0
|  |  | CVUAS 6238 | 3; sequenced; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.86 | 1 | 0
|  |  | CVUAS 6239 | 3; sequenced; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.69 | 1 | 0
|  |  | CVUAS 6240 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.62 | 1 | 0
|  |  | CVUAS 6241 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus cytotoxicus | 1.29 | 1 | 0
|  |  | CVUAS 6242 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus cytotoxicus | 1.1 | 1 | 0
|  |  | CVUAS 6243 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus cytotoxicus | 1.06 | 1 | 0
|  |  | CVUAS 6244 CVUAS | 4; WGS, 16S | eDT | 2016-11-16 | Bacillus cytotoxicus | 1.09 | 1 | 0
|  |  | CVUAS 6244 RKI | 4; WGS, 16S | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 0.88 | 1 | 0
|  |  | CVUAS 6245 | 3; Wt 50°C, cytK1 | eDT | 2022-11-23 | Bacillus cytotoxicus | 0.69 | 1 | 0
|  |  | CVUAS 6292 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.95 | 1 | 0
|  |  | CVUAS 6336 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus cytotoxicus | 0.91 | 1 | 0
|  |  | CVUAS 6433.2 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.77 | 1 | 0
|  |  | CVUAS 6435 | 3; Wt 50°C, cytK1 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CVUAS 6475 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus cytotoxicus | 0.9 | 1 | 0
|  |  | CVUAS 6608 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus cytotoxicus | 0.69 | 1 | 0
|  |  | CVUAS 6655 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.62 | 1 | 0
|  |  | CVUAS 6656 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-03-11 | Bacillus cytotoxicus | 0.83 | 1 | 0
|  |  | CVUAS 6657 | 3; Wt 50°C, cytK1 | eDT | 2015-03-25 | Bacillus cytotoxicus | 0.69 | 1 | 0
|  |  | CVUAS 6658 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.73 | 1 | 0
|  |  | CVUAS 6659 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2016-04-26 | Bacillus cytotoxicus | 0.97 | 1 | 0
|  |  | CVUAS 6660 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 1.15 | 1 | 0
|  |  | CVUAS 6661 | 4; Wt 55°C, cytK1 | eDT | 2015-12-09 | Bacillus cytotoxicus | 0.84 | 1 | 0
|  |  | CVUAS 6688 CVUAS | 4; WGS; Wt 55°C, cytK1 | eDT | 2019-09-25 | Bacillus cytotoxicus | 1.11 | 1 | 0
|  |  | CVUAS 6688 RKI | 4; WGS; Wt 55°C, cytK1 | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 0.96 | 1 | 0
|  |  | CVUAS 6689 | 3; cytK1 | EFex [BRU03] | 2015-03-11 | Bacillus cytotoxicus | 1.12 | 1 | 0
|  |  | CVUAS 6690 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 1 | 1 | 0
|  |  | CVUAS 6691 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.6 | 1 | 0
|  |  | CVUAS 6692 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.7 | 1 | 0
|  |  | CVUAS 6693 | 3; Wt 50°C, cytK1 | eDT | 2015-01-21 | Bacillus cytotoxicus | 0.57 | 1 | 0
|  |  | CVUAS 6694 | 3; Wt 50°C, cytK1 | EFex [BRU03] | 2015-01-27 | Bacillus cytotoxicus | 1.05 | 1 | 0
|  |  | CVUAS 701.3 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.79 | 1 | 0
|  |  | CVUAS 7208 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 1.03 | 1 | 0
|  |  | CVUAS 7209 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.55 | 1 | 0
|  |  | CVUAS 7500 | 3; specific PCR | EFex [BRU03] | 2015-01-13 | Bacillus cytotoxicus | 1.02 | 1 | 0
|  |  | CVUAS 8754 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.93 | 1 | 0
|  |  | CVUAS 8755 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.63 | 1 | 0
|  |  | CVUAS 9297 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.39 | 1 | 0
|  |  | CVUAS 9298 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 1.05 | 1 | 0
|  |  | CVUAS 9299 | 1; 0 | eDT | 2022-11-18 | Bacillus cytotoxicus | 0.55 | 1 | 0
|  |  | DSM 22905 RKI | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-05-30 | Bacillus cytotoxicus | 1.03 | 1 | 0
|  |  | DSM 22905 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus cytotoxicus | 0.84 | 1 | 0
|  |  | IGV 003 | 3; sequenced | eDT | 2016-08-12 | Bacillus cytotoxicus | 0.55 | 1 | 0
|  |  | NVH 391-98 | 3; type strain | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.98 | 1 | 0
|  |  | CIP 111458 | 3; type strain | eDT | 2025-03-18 | Bacillus luti | 0.69 | 1 | 0
|  |  | LMG 28872 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-04 | Bacillus luti | 1.09 | 1 | 0
|  |  | CIP 111019 | 3; type strain | eDT | 2025-03-18 | Bacillus manliponensis | 1.15 | 1 | 0
|  |  | DSM 26473 CVUAS | 3; type strain | eDT | 2025-03-18 | Bacillus manliponensis | 0.93 | 1 | 0
|  |  | DSM 26473 RKI | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-07-25 | Bacillus manliponensis | 1.03 | 1 | 0
|  |  | CIP 111457 | 3; type strain | eDT | 2025-03-18 | Bacillus mobilis | 0.74 | 1 | 0
|  |  | LMG 28877 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-04 | Bacillus mobilis | 1.19 | 1 | 0
|  |  | ATCC 12826 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus mycoides | 0.47 | 1 | 0
|  |  | B 103 | 2; 0 | DT | 2015-07-22 | Bacillus mycoides | 0.5 | 1 | 0
|  |  | DSM 104109 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.99 | 1 | 0
|  |  | DSM 104111 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.42 | 1 | 0
|  |  | DSM 104112 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.81 | 1 | 0
|  |  | DSM 104113 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.71 | 1 | 0
|  |  | DSM 104118 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.48 | 1 | 0
|  |  | DSM 104119 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.74 | 1 | 0
|  |  | DSM 104119 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-05-09 | Bacillus mycoides | 0.88 | 1 | 0
|  |  | DSM 104121 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.69 | 1 | 0
|  |  | DSM 104122 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.56 | 1 | 0
|  |  | DSM 104135 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.71 | 1 | 0
|  |  | DSM 104137 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.86 | 1 | 0
|  |  | DSM 11821 CVUAS | 3; public strain collection | eDT | 2016-01-27 | Bacillus mycoides | 1.13 | 1 | 0
|  |  | DSM 11821 RKI | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-05-03 | Bacillus mycoides | 1.21 | 1 | 0
|  |  | DSM 12001 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.55 | 1 | 0
|  |  | DSM 12001 RKI | 3; public strain collection | see [0185] | 2021-12-02 | Bacillus mycoides | 1.12 | 1 | 0
|  |  | DSM 2048 CVUAS | 3; type strain | eDT | 09.12.2015 | Bacillus mycoides | 0.65 | 1 | 0
|  |  | DSM 2048 RKI | 3; type strain | EFex [BRU03] | 2022-01-13 | Bacillus mycoides | 1.17 | 1 | 0
|  |  | DSM 299 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.98 | 1 | 0
|  |  | DSM 299 IGV | 3; public strain collection | DT | 2016-04-08 | Bacillus mycoides | 0.59 | 1 | 0
|  |  | DSM 303 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 1.11 | 1 | 0
|  |  | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus mycoides | 0.64 | 1 | 0
|  |  | DSM 30963 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.67 | 1 | 0
|  |  | DSM 30966 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.85 | 1 | 0
|  |  | DSM 30969 | 3; public strain collection | eDT | 2025-03-18 | Bacillus mycoides | 0.84 | 1 | 0
|  |  | WS 2641 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA-TA2#EXT=Cultivation & preparation: 02.11.2006, preparation for MALDI MS 06.11.2006: M. Staemmler#LAS=LL26 pct. 600, "LL70 pct." 10 sh. ¦ Laser attenuator offset: 65 pct.¦Range: 20 pct.#CAL=ProteinStandard 1 (Bruker)#CUS=RKI ZBS6# | 2006-07-11 | Bacillus mycoides | 0.8 | 1 | 0
|  |  | WSBC 10204 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus mycoides | 1.26 | 1 | 0
|  |  | WSBC 10207 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus mycoides | 1.3 | 1 | 0
|  |  | WSBC 10277 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus mycoides | 0.83 | 1 | 0
|  |  | WSBC 10279 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus mycoides | 1.04 | 1 | 0
|  |  | WSBC 10291 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus mycoides | 0.99 | 1 | 0
|  |  | WSBC 10295 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus mycoides | 0.92 | 1 | 0
|  |  | WSBC 10360 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-07 | Bacillus mycoides | 1.01 | 1 | 0
|  |  | WSBC 10363 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus mycoides | 0.86 | 1 | 0
|  |  | WSBC 10387 | 3; well defined strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2006-07-11 | Bacillus mycoides | 0.83 | 1 | 0
|  |  | LMG 28871 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-01-25 | Bacillus nitratireducens | 1.3 | 1 | 0
|  |  | CIP 111460 | 3; type strain | eDT | 2025-03-18 | Bacillus nitratireduscens | 0.7 | 1 | 0
|  |  | ATCC 10987 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-10-23 | Bacillus pacificus | 0.5 | 1 | 0
|  |  | SN4-1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-04-11 | Bacillus pacificus | 0.81 | 1 | 0
|  |  | CIP 111542 | 3; type strain | eDT | 2025-03-18 | Bacillus pacificus | 0.41 | 1 | 0
|  |  | CIP 111461 | 3; type strain | eDT | 2025-03-18 | Bacillus paramycoides | 0.56 | 1 | 0
|  |  | LMG 28876 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 3.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Bacillus paramycoides LMG 28876, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-02-04 | Bacillus paramycoides | 0.68 | 1 | 0
|  |  | CIP 111454 | 3; type strain | eDT | 2025-03-18 | Bacillus paranthracis | 0.42 | 1 | 0
|  |  | CIP A30 | 3; public strain collection | eDT | 2025-03-18 | Bacillus paranthracis | 0.58 | 1 | 0
|  |  | LMG 28873 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-11 | Bacillus paranthracis | 1.25 | 1 | 0
|  |  | CIP 111459 | 3; type strain | eDT | 2025-03-18 | Bacillus proteolyticus | 0.75 | 1 | 0
|  |  | LMG 28870 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from LMG strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2019-12-01 | Bacillus proteolyticus | 1.26 | 1 | 0
|  |  | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus pseudomycoides | 1.02 | 1 | 0
|  |  | DSM 12442 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus pseudomycoides | 0.83 | 1 | 0
|  |  | DSM 12442 RKI | 3; type strain | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2007-05-03 | Bacillus pseudomycoides | 0.95 | 1 | 0
|  |  | DSM 12443 CVUAS | 3; public strain collection | eDT | 09.12.2015 | Bacillus pseudomycoides | 0.72 | 1 | 0
|  |  | DSM 12443 RKI | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-05 | Bacillus pseudomycoides | 0.85 | 1 | 0
|  |  | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.77 | 1 | 0
|  |  | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 1.1 | 1 | 0
|  |  | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus thuringiensis | 0.74 | 1 | 0
|  |  | DSM 5724 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.56 | 1 | 0
|  |  | DSM 5815 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.29 | 1 | 0
|  |  | DSM 6022 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.67 | 1 | 0
|  |  | DSM 5815 RKI | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 02.12.2021¦ 1:1 HCCA¦Andy 3.0#EXT=Bacillus thuringiensis DSM 5815, preparation and measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-12-02 | Bacillus thuringiensis | 1.21 | 1 | 0
|  |  | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 1.01 | 1 | 0
|  |  | DSM 5725 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.53 | 1 | 0
|  |  | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.82 | 1 | 0
|  |  | B 8 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2014-12-01 | Bacillus thuringiensis | 1.1 | 1 | 0
|  |  | CVUAS 10203 | 4; Cry-pos, Cyt-K2, Krist. Pos | eDT | 2016-01-27 | Bacillus thuringiensis | 0.67 | 1 | 0
|  |  | CVUAS 5760 | 4; Cry-pos, Krist. Pos | eDT | 2016-01-27 | Bacillus thuringiensis | 0.88 | 1 | 0
|  |  | CVUAS 5761 | 4; Cry-pos, Krist. Pos | eDT | 2016-01-27 | Bacillus thuringiensis | 1.08 | 1 | 0
|  |  | DSM 2046 (B188. Beyer) | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-07 | Bacillus thuringiensis | 0.71 | 1 | 0
|  |  | DSM 2046 CVUAS | 3; type strain | DT | 2014-01-22 | Bacillus thuringiensis | 0.91 | 1 | 0
|  |  | DSM 2046 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-09-21 | Bacillus thuringiensis | 0.95 | 1 | 0
|  |  | DSM 350 WIS-St. Nr. 315 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus thuringiensis | 0.74 | 1 | 0
|  |  | DSM 350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-03 | Bacillus thuringiensis | 0.51 | 1 | 0
|  |  | DSM 6022 IGV | 3; public strain collection | DT | 2016-04-06 | Bacillus thuringiensis | 1.2 | 1 | 0
|  |  | DSM 6024 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.74 | 1 | 0
|  |  | DSM 6030 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.87 | 1 | 0
|  |  | DSM 6033 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.44 | 1 | 0
|  |  | DSM 6073 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-12-04 | Bacillus thuringiensis | 0.96 | 1 | 0
|  |  | DSM 6087 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.74 | 1 | 0
|  |  | DSM 6099 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.5 | 1 | 0
|  |  | DSM 6100 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.63 | 1 | 0
|  |  | DSM 6101 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 1 | 1 | 0
|  |  | DSM 6102 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 1.03 | 1 | 0
|  |  | DSM 6106 | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.62 | 1 | 0
|  |  | DSM 6890 CVUAS | 3; public strain collection | eDT | 2025-03-18 | Bacillus thuringiensis | 0.53 | 1 | 0
|  |  | DSM 6890 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-12-02 | Bacillus thuringiensis | 0.64 | 1 | 0
|  |  | EZ05-05 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 05.10.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-10-09 | Bacillus thuringiensis | 0.92 | 1 | 0
|  |  | GBSC29 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 21.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-01 | Bacillus thuringiensis | 0.98 | 1 | 0
|  |  | Konkukian | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 1:10 diluted , spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Cells from solid cultures#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS2# | 2015-02-03 | Bacillus thuringiensis | 1.01 | 1 | 0
|  |  | NMTD81 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=inactivation date 15.03.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2021-08-21 | Bacillus thuringiensis | 0.94 | 1 | 0
|  |  | WS 2614 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus thuringiensis | 1.05 | 1 | 0
|  |  | WS 2621 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus thuringiensis | 1.04 | 1 | 0
|  |  | WS 2623 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-07-11 | Bacillus thuringiensis | 0.87 | 1 | 0
|  |  | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-11-07 | Bacillus thuringiensis | 0.83 | 1 | 0
|  |  | 43 Paciflor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus toyonensis | 1.02 | 1 | 0
|  |  | 8 Toyocerin | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2022-06-28 | Bacillus toyonensis | 1.16 | 1 | 0
|  |  | B260 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus toyonensis | 1.05 | 1 | 0
|  |  | CIP A31 | 3; public strain collection | eDT | 2025-03-18 | Bacillus toyonensis | 0.53 | 1 | 0
|  |  | NMTD92 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 03.11.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-02-01 | Bacillus toyonensis | 1.01 | 1 | 0
|  |  | Paciflor | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2022-06-21 | Bacillus toyonensis | 0.92 | 1 | 0
|  |  | Toyocerin | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-06-22 | Bacillus toyonensis | 0.88 | 1 | 0
|  |  | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-08 | Bacillus tropicus | 0.81 | 1 | 0
|  |  | CD 3-2 | 2; strain collection of the RKI | see [0185] | 2022-05-11 | Bacillus tropicus | 0.93 | 1 | 0
|  |  | CIP 111455 | 3; type-strain | eDT | 2025-03-18 | Bacillus tropicus | 0.66 | 1 | 0
|  |  | LMG 28874 | 3; type-strain | see [0185] | 2022-02-04 | Bacillus tropicus | 1.08 | 1 | 0
|  |  | SN1 | 2; strain collection of the RKI | see [0185] | 2020-12-15 | Bacillus tropicus | 0.76 | 1 | 0
|  |  | VM11 (01) | 2; strain collection of the RKI | see [0185] | 2018-11-14 | Bacillus tropicus | 0.64 | 1 | 0
|  |  | VM21 | 2; strain collection of the RKI | see [0185] | 2018-08-17 | Bacillus tropicus | 0.89 | 1 | 0
|  |  | B256 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus wiedmannii | 0.83 | 1 | 0
|  |  | DSM 102050 CVUAS | 3; type strain | eDT | 2025-03-18 | Bacillus wiedmannii | 0.91 | 1 | 0
|  |  | DSM 102050 RKI | 3; type strain | see [0185] | 2018-05-23 | Bacillus wiedmannii | 0.77 | 1 | 0
|  |  | GBSC29 (01) | 2; strain collection of the RKI | see [0185] | 2018-05-17 | Bacillus wiedmannii | 0.98 | 1 | 0
|  |  | GBSC45 | 2; strain collection of the RKI | see [0185] | 2018-08-23 | Bacillus wiedmannii | 0.66 | 1 | 0
|  |  | LLCG23 | 2; strain collection of the RKI | see [0185] | 2018-05-17 | Bacillus wiedmannii | 0.68 | 1 | 0
|  |  | NMSL88 | 2; strain collection of the RKI | see [0185] | 2018-08-21 | Bacillus wiedmannii | 0.78 | 1 | 0
|  |  | NMSW16 (02) | 2; strain collection of the RKI | see [0185] | 2018-01-10 | Bacillus wiedmannii | 1.06 | 1 | 0
|  |  | NMSW23 | 2; strain collection of the RKI | see [0185] | 2018-08-21 | Bacillus wiedmannii | 0.64 | 1 | 0
|  |  | B 338 | 2; strain collection of the RKI | see [0185] | 2015-05-08 | Bacillus cereus | 0.93 | 1 | 0
|  |  | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-31 | Bacillus ectoiniformans | 1.22 | 1 | 0
|  |  | CIP 111018 | 3; type strain | eDT | 2025-03-18 | Bacillus gaemokensis | 0.86 | 1 | 0
|  |  | DSM 29500 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Strain obtained from DSMZ strain collection#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2018-09-05 | Bacillus gobiensis | 0.71 | 1 | 0
|  |  | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-30 | Bacillus halotolerans | 0.83 | 1 | 0
|  |  | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-12 | Bacillus halotolerans | 0.61 | 1 | 0
|  |  | DSM 21200 | 3; public strain collection | #CON=not specified#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water, spotted 1:3 with matrix (12 g¦L HCCA in 70 pct. ACN, 3 pct. FA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-28 | Bacillus inaquosorum | 0.74 | 1 | 0
|  |  | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-12-02 | Bacillus inaquosorum | 0.89 | 1 | 0
|  |  | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-12 | Bacillus infantis | 0.74 | 1 | 0
|  |  | 141014442-3 | 3; VETQAS | DT | 2014-11-07 | Bacillus licheniformis | 0.87 | 1 | 0
|  |  | 141015909-3 | 2; 0 | DT | 2014-12-12 | Bacillus licheniformis | 0.66 | 1 | 0
|  |  | 161004445 | 2; 0 | DT | 2016-04-12 | Bacillus licheniformis | 0.74 | 1 | 0
|  |  | 161007249-3 | 2; 0 | DT | 2016-06-14 | Bacillus licheniformis | 0.8 | 1 | 0
|  |  | ATCC 14580 | 3; type strain | DT | 2015-06-23 | Bacillus licheniformis | 0.82 | 1 | 0
|  |  | B 084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus licheniformis | 0.47 | 1 | 0
|  |  | B 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-26 | Bacillus licheniformis | 0.81 | 1 | 0
|  |  | B 116 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-03-26 | Bacillus licheniformis | 1 | 1 | 0
|  |  | B 118 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-26 | Bacillus licheniformis | 0.5 | 1 | 0
|  |  | B 62 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus licheniformis | 0.92 | 1 | 0
|  |  | B 72 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-14 | Bacillus licheniformis | 0.74 | 1 | 0
|  |  | BRLK08_04 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2015-04-17 | Bacillus licheniformis | 1.09 | 1 | 0
|  |  | DSM 13 B178 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-08-16 | Bacillus licheniformis | 0.8 | 1 | 0
|  |  | DSM 13 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus licheniformis | 0.53 | 1 | 0
|  |  | DSM 13 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-12 | Bacillus licheniformis | 0.65 | 1 | 0
|  |  | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-10 | Bacillus mojavensis | 0.68 | 1 | 0
|  |  | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Bacillus mojavensis | 0.66 | 1 | 0
|  |  | ATCC 12759 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus paralicheniformis | 0.81 | 1 | 0
|  |  | EZ01-02 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-02-24 | Bacillus paralicheniformis | 0.49 | 1 | 0
|  |  | EZ01-05 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-02-27 | Bacillus paralicheniformis | 0.71 | 1 | 0
|  |  | EZ10-02 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2018-03-05 | Bacillus paralicheniformis | 0.74 | 1 | 0
|  |  | EZ15-08 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.07.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-07-23 | Bacillus paralicheniformis | 0.86 | 1 | 0
|  |  | NMSW12 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2027-04-10 | Bacillus paralicheniformis | 0.8 | 1 | 0
|  |  | CVUAS 34924 | 3; sequenced | EFEx | 2023-02-17 | Bacillus pumilus | 0.53 | 1 | 0
|  |  | LGA MB 2022000019 | 2; confirmed proficency test | eDT | 2022-02-14 | Bacillus pumilus | 0.56 | 1 | 0
|  |  | B 17 CVUAS | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus pumilus | 0.63 | 1 | 0
|  |  | B 17 RKI | 3; sequenced (CVUAS) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Bacillus pumilus | 1.01 | 1 | 0
|  |  | B 31 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-11 | Bacillus pumilus | 0.41 | 1 | 0
|  |  | B 91 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-14 | Bacillus pumilus | 0.92 | 1 | 0
|  |  | CCM 340 RKI | 3; reference strain collection | see [0185] | 2015-02-09 | Bacillus pumilus | 0.78 | 1 | 0
|  |  | CCM 340 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Bacillus pumilus | 0.37 | 1 | 0
|  |  | CJCL2 | 2; strain collection of the RKI | see [0185] | 2018-01-16 | Bacillus pumilus | 0.8 | 1 | 0
|  |  | DSM 13835 LMTG62N07 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-23 | Bacillus pumilus | 0.53 | 1 | 0
|  |  | DSM 13835 RKI | 3; public strain collection | see [0185] | 2022-08-05 | Bacillus pumilus | 0.92 | 1 | 0
|  |  | DSM 1794 | 3; public strain collection | see [0185] | 2022-08-05 | Bacillus pumilus | 0.47 | 1 | 0
|  |  | DSM 27 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Bacillus pumilus | 1.15 | 1 | 0
|  |  | DSM 27 Neukauf RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-12-15 | Bacillus pumilus | 1.26 | 1 | 0
|  |  | DSM 27 CVUAS | 3; type strain | DT | 2014-01-22 | Bacillus pumilus | 0.74 | 1 | 0
|  |  | DSM 355 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus pumilus | 0.6 | 1 | 0
|  |  | DSM 361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Bacillus pumilus | 1.14 | 1 | 0
|  |  | DSM 492 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-16 | Bacillus pumilus | 1.17 | 1 | 0
|  |  | DSM 8786 | 3; public strain collection | see [0185] | 2022-08-05 | Bacillus pumilus | 0.74 | 1 | 0
|  |  | EZ01-05 (01) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-02-07 |  | 0.68 | 1 | 0
|  |  | GBAC46 (01) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-01-10 |  | 0.76 | 1 | 0
|  |  | GBSC66 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-08-14 |  | 0.54 | 1 | 0
|  |  | GBSW19 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 |  | 0.81 | 1 | 0
|  |  | GBSW2 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-08 |  | 0.97 | 1 | 0
|  |  | LDZX38 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-08-15 |  | 0.74 | 1 | 0
|  |  | LLTC96 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-16 |  | 0.79 | 1 | 0
|  |  | LNTW65 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 |  | 0.71 | 1 | 0
|  |  | LNXM10 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 17.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 |  | 0.61 | 1 | 0
|  |  | LNXM12 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-07-12 | Bacillus pumilus | 0.88 | 1 | 0
|  |  | LNXM65 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-05-08 | Bacillus pumilus | 0.55 | 1 | 0
|  |  | LNXM70 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-03-09 | Bacillus pumilus | 0.34 | 1 | 0
|  |  | MM 171 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus pumilus | 0.87 | 1 | 0
|  |  | MZGC1 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-10 | Bacillus pumilus | 0.72 | 1 | 0
|  |  | NMCC4 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 26.07.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-11-22 | Bacillus pumilus | 0.97 | 1 | 0
|  |  | NMCC46 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Inactivation date 08.03.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-08-17 | Bacillus pumilus | 0.82 | 1 | 0
|  |  | NMCN1 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-08-15 | Bacillus pumilus | 0.88 | 1 | 0
|  |  | NMCN6 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 04.08.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-08-15 | Bacillus pumilus | 1.02 | 1 | 0
|  |  | NMSW10 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 01.09.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-01-09 | Bacillus pumilus | 1.03 | 1 | 0
|  |  | NMSX4 (02) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.02.2018#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2018-03-02 | Bacillus pumilus | 1.08 | 1 | 0
|  |  | NMTD17 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=Source: Prof. X.Gao via Prof. R.Borriss#LAS=Date of inactivation: 08.07.2017#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦Prof. Gao (Nanjing, China)# | 2017-07-18 | Bacillus pumilus | 0.89 | 1 | 0
|  |  | NVH-819-05 | 3; well described isolate [0017] | DT | 2015-04-21 | Bacillus pumilus | 0.63 | 1 | 0
|  |  | YBWC18 | 2; strain collection of the RKI | see [0185] | 2018-08-16 | Bacillus pumilus | 0.98 | 1 | 0
|  |  | DSM 103081 | 3; public strain collection | see [0185] | 2023-03-02 | Bacillus safensis | 0.45 | 1 | 0
|  |  | DSM 19292 | 3; type strain | see [0185] | 2023-03-02 | Bacillus safensis | 0.55 | 1 | 0
|  |  | GBSW22 | 2; strain collection of the RKI | see [0185] | 2023-03-02 | Bacillus safensis | 0.7 | 1 | 0
|  |  | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2018-05-24 | Bacillus siamensis | 0.8 | 1 | 0
|  |  | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-15 | Bacillus sonorensis | 0.6 | 1 | 0
|  |  | B 324 | 2; strain collection of the RKI | TFAextr | 2015-05-07 | Bacillus sp | 0.85 | 1 | 0
|  |  | DSM 12369 RKI | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Bacillus sp-DSM-12369 | 0.87 | 1 | 0
|  |  | ATCC 6633 RKI B | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-01 | Bacillus spizizenii | 0.7 | 1 | 0
|  |  | ATCC 6633 RKI A | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus spizizenii | 1.13 | 1 | 0
|  |  | ATCC 6633 CVUAS | 3; public strain collection | eDT | 2015-07-22 | Bacillus spizizenii | 0.59 | 1 | 0
|  |  | B 19 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-02 | Bacillus spizizenii | 1.06 | 1 | 0
|  |  | B 061 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-01-07 | Bacillus spizizenii | 0.56 | 1 | 0
|  |  | DSM 1087 (B156) | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2017-12-18 | Bacillus spizizenii | 0.65 | 1 | 0
|  |  | DSM 1087 CVUAS | 3; public strain collection | DT | 2014-01-22 | Bacillus spizizenii | 0.44 | 1 | 0
|  |  | DSM 1087 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus spizizenii | 0.47 | 1 | 0
|  |  | DSM 347 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2007-04-23 | Bacillus spizizenii | 1.11 | 1 | 0
|  |  | DSM 618 | 4; public strain collection | eDT | 2016-01-27 | Bacillus spizizenii | 0.37 | 1 | 0
|  |  | DSM 6395 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 17.11.2021¦ 1:1 HCCA¦Andy 3.0#EXT | 2021-11-17 | Bacillus spizizenii | 0.6 | 1 | 0
|  |  | WIS-St. Nr. 491 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-15 | Bacillus spizizenii | 0.58 | 1 | 0
|  |  | B 130 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus subtilis | 0.77 | 1 | 0
|  |  | DSM 402 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-28 | Bacillus subtilis | 0.67 | 1 | 0
|  |  | ATCC 6051 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus subtilis | 0.49 | 1 | 0
|  |  | B 014 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-08-11 | Bacillus subtilis | 0.78 | 1 | 0
|  |  | B 020 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-02 | Bacillus subtilis | 0.91 | 1 | 0
|  |  | B 129 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Bacillus subtilis | 1.01 | 1 | 0
|  |  | B rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-23 | Bacillus subtilis | 1.12 | 1 | 0
|  |  | BRLK08_02 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Bacillus subtilis | 0.77 | 1 | 0
|  |  | CCM 110 B068 | 3; reference strain collection | DT | 2014-01-22 | Bacillus subtilis | 0.69 | 1 | 0
|  |  | DSM 10 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-23 | Bacillus subtilis | 0.92 | 1 | 0
|  |  | DSM 1092 RKI | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-27 | Bacillus subtilis | 0.69 | 1 | 0
|  |  | DSM 1092 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Bacillus subtilis | 0.49 | 1 | 0
|  |  | E rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-23 | Bacillus subtilis | 0.85 | 1 | 0
|  |  | GR2.1 | 2; strain collection of the RKI | see [0185] | 2020-01-16 | Bacillus subtilis | 0.5 | 1 | 0
|  |  | HY 1152 | 2; strain collection of the RKI | see [0185] | 2022-05-24 | Bacillus subtilis | 0.95 | 1 | 0
|  |  | I rauh | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-11 | Bacillus subtilis | 0.95 | 1 | 0
|  |  | IGV 013 IGV | 3; sequenced | DT | 2016-11-09 | Bacillus subtilis | 0.59 | 1 | 0
|  |  | LVM03 | 3; sequenced (LGL) | eDT | 2015-07-09 | Bacillus subtilis | 0.72 | 1 | 0
|  |  | MM 508 subt RKI | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-03 | Bacillus subtilis | 0.94 | 1 | 0
|  |  | NMSX4 (01) | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-06-16 | Bacillus subtilis | 0.76 | 1 | 0
|  |  | RKI-w2gr | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-04-04 | Bacillus subtilis | 0.96 | 1 | 0
|  |  | SYST2 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-06-16 | Bacillus subtilis | 1 | 1 | 0
|  |  | B 12 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=in 80 pct. TFA, diluted by 1:1 HCCA:TA2 | 2014-12-01 | Bacillus subtilis | 0.74 | 1 | 0
|  |  | DSM 16466 | 3; public strain collection | see [0185] | 2022-08-19 | Bacillus taeanensis | 0.56 | 1 | 0
|  |  | WIS-St. Nr. 494 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦ 1:1 HCCA-TA2 (A)#EXT=Cultivation & sample preparation by WIS; Shipment: 20.12.2005; Preparation for MALDI-ToF Maren Staemmler#LAS=LL25 pct. 600, "LL70 pct." 10 sh. offset: 63 pct.¦Range: 20 pct. Laser neu Defok.#CAL=ProteinStandard 1 (Bruker)#CUS=RKI ZBS6 ¦WIS (Dr. Niederwoehrmeier, Munster)# | 2007-01-15 | Bacillus thuringiensis | 0.94 | 1 | 0
|  |  | CVUAS 2786.2 | 3; sequenced | DT | 2013-07-31 | Bacillus timorensis | 0.47 | 1 | 0
|  |  | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-07-20 | Bacillus vallismortis | 0.92 | 1 | 0
|  |  | JX-1 (01) | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2017-07-14 | Bacillus velezensis | 0.97 | 1 | 0
|  |  | JX-1 (02) | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2017-07-14 | Bacillus velezensis | 0.97 | 1 | 0
|  |  | A25 vel | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-10-15 | Bacillus velezensis | 0.66 | 1 | 0
|  |  | A35 vel | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 100 mkL#TRT=dissolved in 80 pct. TFA for 30 Min, diluted 1:10, mixed 1:1 by matrix (12 g¦L HCCA in TA2) and spotted#EXT=Source: ENDOBICA (Dr. Thanh Tam, Vietnam II)#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-10-19 | Bacillus velezensis | 0.91 | 1 | 0
|  |  | B 032 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-10 | Bacillus velezensis | 0.82 | 1 | 0
|  |  | BP1.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus velezensis | 0.86 | 1 | 0
|  |  | BP1.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus velezensis | 0.92 | 1 | 0
|  |  | BP1.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus velezensis | 0.86 | 1 | 0
|  |  | BP1.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2022-11-02 | Bacillus velezensis | 0.82 | 1 | 0
|  |  | BP2.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 1.07 | 1 | 0
|  |  | BP2.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 1.01 | 1 | 0
|  |  | BP2.3 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 0.86 | 1 | 0
|  |  | BP3.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 0.98 | 1 | 0
|  |  | BT2.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-13 | Bacillus velezensis | 0.61 | 1 | 0
|  |  | BT2.4 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-13 | Bacillus velezensis | 0.81 | 1 | 0
|  |  | CB1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-17 | Bacillus velezensis | 0.72 | 1 | 0
|  |  | CB1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus velezensis | 0.61 | 1 | 0
|  |  | CF1.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 1.15 | 1 | 0
|  |  | CF2.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 0.75 | 1 | 0
|  |  | CF2.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-23 | Bacillus velezensis | 0.86 | 1 | 0
|  |  | CF2.3 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-23 | Bacillus velezensis | 0.91 | 1 | 0
|  |  | CP1.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Bacillus velezensis | 0.7 | 1 | 0
|  |  | CP5.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-29 | Bacillus velezensis | 0.53 | 1 | 0
|  |  | CP5.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-29 | Bacillus velezensis | 0.62 | 1 | 0
|  |  | CP6 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-17 | Bacillus velezensis | 0.8 | 1 | 0
|  |  | CP7.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-02 | Bacillus velezensis | 0.52 | 1 | 0
|  |  | CP7.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-02 | Bacillus velezensis | 0.67 | 1 | 0
|  |  | CP7.1C | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-03 | Bacillus velezensis | 0.55 | 1 | 0
|  |  | CP7.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-03 | Bacillus velezensis | 0.97 | 1 | 0
|  |  | CP7.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-18 | Bacillus velezensis | 0.95 | 1 | 0
|  |  | CVUAS 35818 | 3; sequenced | eDT | 2023-11-02 | Bacillus velezensis | 0.96 | 1 | 0
|  |  | DJFZ40 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-03-07 | Bacillus velezensis | 1.01 | 1 | 0
|  |  | DL2.3A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 0.97 | 1 | 0
|  |  | DL2.3B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-22 | Bacillus velezensis | 0.99 | 1 | 0
|  |  | DP1.3B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-12-18 | Bacillus velezensis | 0.6 | 1 | 0
|  |  | DP2.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-08 | Bacillus velezensis | 0.97 | 1 | 0
|  |  | DSM 23117 | 3; public strain collection | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2015-03-10 | Bacillus velezensis | 1.13 | 1 | 0
|  |  | DSM 28326 | 3; public strain collection | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-05-24 | Bacillus velezensis | 0.52 | 1 | 0
|  |  | EG4.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus velezensis | 1.11 | 1 | 0
|  |  | EG4.2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-16 | Bacillus velezensis | 1.06 | 1 | 0
|  |  | EG5.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus velezensis | 0.78 | 1 | 0
|  |  | EG5.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus velezensis | 0.56 | 1 | 0
|  |  | EZ01-02 (01) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-08-31 | Bacillus velezensis | 1.2 | 1 | 0
|  |  | EZ05-04 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-08-31 | Bacillus velezensis | 0.76 | 1 | 0
|  |  | EZ10-02 (02) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-08-31 | Bacillus velezensis | 1.02 | 1 | 0
|  |  | EZ15-07 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-09-21 | Bacillus velezensis | 1.05 | 1 | 0
|  |  | EZ15-09 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-09-21 | Bacillus velezensis | 1.08 | 1 | 0
|  |  | FZB24 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2014-10-15 | Bacillus velezensis | 0.59 | 1 | 0
|  |  | FZB45 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2017-07-14 | Bacillus velezensis | 1 | 1 | 0
|  |  | GBSW11 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2018-03-07 | Bacillus velezensis | 0.62 | 1 | 0
|  |  | GS3.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-16 | Bacillus velezensis | 1.04 | 1 | 0
|  |  | HB1.4A | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2019-11-27 | Bacillus velezensis | 0.95 | 1 | 0
|  |  | HD2.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2019-12-18 | Bacillus velezensis | 0.88 | 1 | 0
|  |  | HD2.2 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2019-12-18 | Bacillus velezensis | 0.9 | 1 | 0
|  |  | HD2.3 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-08 | Bacillus velezensis | 0.6 | 1 | 0
|  |  | HD3.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus velezensis | 0.91 | 1 | 0
|  |  | HD3.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus velezensis | 0.45 | 1 | 0
|  |  | HD3.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus velezensis | 0.84 | 1 | 0
|  |  | HD3.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Bacillus velezensis | 0.67 | 1 | 0
|  |  | HD3.3B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus velezensis | 1.07 | 1 | 0
|  |  | HD3.3C | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus velezensis | 0.83 | 1 | 0
|  |  | HD5.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus velezensis | 1.13 | 1 | 0
|  |  | HD5.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus velezensis | 0.97 | 1 | 0
|  |  | HD5.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-10 | Bacillus velezensis | 1.11 | 1 | 0
|  |  | HD5 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-11-21 | Bacillus velezensis | 1.01 | 1 | 0
|  |  | JS1.1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-14 | Bacillus velezensis | 0.81 | 1 | 0
|  |  | KKD2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-01-23 | Bacillus velezensis | 0.96 | 1 | 0
|  |  | KT1 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2021-10-14 | Bacillus velezensis | 0.72 | 1 | 0
|  |  | LLCG43 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2018-08-17 | Bacillus velezensis | 1.09 | 1 | 0
|  |  | MM 508 vel RKI | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2007-03-05 | Bacillus velezensis | 0.54 | 1 | 0
|  |  | MR2.1B | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-17 | Bacillus velezensis | 0.55 | 1 | 0
|  |  | MS2.3A | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-17 | Bacillus velezensis | 1.12 | 1 | 0
|  |  | MS2.3B | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-17 | Bacillus velezensis | 0.86 | 1 | 0
|  |  | OL1.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-13 | Bacillus velezensis | 0.61 | 1 | 0
|  |  | OL2.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-14 | Bacillus velezensis | 1.19 | 1 | 0
|  |  | OR2.1 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-13 | Bacillus velezensis | 0.99 | 1 | 0
|  |  | OR2.2 | 2; strain collection of the RKI | #CON=approx. 2-3 loops in 25 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 1:1 HCCA¦TA2 (A)#EXT=Source: ZBS 6 strain collection#LAS=LL23 pct. 600, "LL60 pct." 10 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Isset: Autoflex20130808.isset) defocussed 4xunter bisher und ToM 200ns 20130611.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6, Prof, Borriss# | 2020-01-14 | Bacillus velezensis | 1.05 | 1 | 0
|  |  | RJGP41 (02) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2017-08-01 | Bacillus velezensis | 1.19 | 1 | 0
|  |  | RJGP41 (03) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-01-11 | Bacillus velezensis | 0.9 | 1 | 0
|  |  | RJGP41 (04) | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2017-08-01 | Bacillus velezensis | 0.65 | 1 | 0
|  |  | RS1.2A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus velezensis | 1.11 | 1 | 0
|  |  | RS1.2B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus velezensis | 0.57 | 1 | 0
|  |  | SC3.1A | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus velezensis | 0.95 | 1 | 0
|  |  | SC3.1B | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-20 | Bacillus velezensis | 0.99 | 1 | 0
|  |  | TK2 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2021-10-14 | Bacillus velezensis | 0.9 | 1 | 0
|  |  | VM10 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-11-14 | Bacillus velezensis | 1.09 | 1 | 0
|  |  | VM49 | 2; strain collection of the RKI | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2018-11-21 | Bacillus velezensis | 0.88 | 1 | 0
|  |  | DSM 29883 | 3; type strain | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2018-05-24 | Bacillus xiamenensis | 0.88 | 1 | 0
|  |  | LLTC93 | 2; strain collection of the RKI | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2) | 2017-07-18 | Bacillus zhangzhouensis | 0.87 | 1 | 0
|  |  | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Cytobacillus firmus | 0.84 | 1 | 0
|  |  | B 48 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-01-07 | Cytobacillus firmus | 0.81 | 1 | 0
|  |  | CCEB 627 DSM 12 | 3; type strain | DT | 2014-01-22 | Cytobacillus firmus | 0.83 | 1 | 0
|  |  | CCM 37 B 88 | 3; reference strain collection | 0 | 2015-02-09 | Cytobacillus firmus | 0.4 | 1 | 0
|  |  | CCM 37 | 3; reference strain collection | DT | 2014-01-22 | Cytobacillus firmus | 0.7 | 1 | 0
|  |  | DSM 359 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-09 | Cytobacillus firmus | 0.8 | 1 | 0
|  |  | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-15 | Cytobacillus horneckiae | 1.23 | 1 | 0
|  |  | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-15 | Cytobacillus horneckiae | 1.01 | 1 | 0
|  |  | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus kochii | 1.35 | 1 | 0
|  |  | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-02-19 | Domibacillus aminovorans | 0.64 | 1 | 0
|  |  | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-02-15 | Domibacillus aminovorans | 0.62 | 1 | 0
|  |  | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-19 | Evansella clarkii | 0.64 | 1 | 0
|  |  | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Evansella clarkii | 1.02 | 1 | 0
|  |  | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus stearothermophilus | 0.71 | 1 | 0
|  |  | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-25 | Halalkalihalobacillus halodurans | 1.17 | 1 | 0
|  |  | DSM 6939 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Halalkalihalobacillus halodurans | 0.96 | 1 | 0
|  |  | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Halalkalihalobacillus halodurans | 0.9 | 1 | 0
|  |  | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Heyndrickxia coagulans | 1.37 | 1 | 0
|  |  | DSM 2308 | 2; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Weizmannia coagulans DSM 2308: Preparation & measurement: N. Fischer#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-22 | Heyndrickxia coagulans | 1.39 | 1 | 0
|  |  | DSM 2311 | 2; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Weizmannia coagulans DSM 2308: Preparation & measurement: N. Fischer#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Heyndrickxia coagulans | 1.23 | 1 | 0
|  |  | DSM 30760 | 2; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Weizmannia coagulans DSM 2308: Preparation & measurement: N. Fischer#LAS=not specified#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-02 | Heyndrickxia coagulans | 1.45 | 1 | 0
|  |  | B 041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia sporothermodurans | 0.71 | 1 | 0
|  |  | B 106 | 3; sequenced | DT | 2015-07-22 | Heyndrickxia sporothermodurans | 0.7 | 1 | 0
|  |  | DSM 10599 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-06 | Heyndrickxia sporothermodurans | 0.77 | 1 | 0
|  |  | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-20 | Lederbergia lenta | 0.69 | 1 | 0
|  |  | DSM 9 CVUAS | 3; type strain | DT | 2014-01-22 | Lederbergia lenta | 0.57 | 1 | 0
|  |  | DSM 9 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-13 | Lederbergia lenta | 0.78 | 1 | 0
|  |  | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus bataviensis | 1.08 | 1 | 0
|  |  | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-20 | Neobacillus cucumis | 0.88 | 1 | 0
|  |  | B 216 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-25 | Niallia circulans | 1.15 | 1 | 0
|  |  | B46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-10 | Niallia circulans | 0.77 | 1 | 0
|  |  | DSM 11 | 3; type strain | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-04-13 | Niallia circulans | 0.98 | 1 | 0
|  |  | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 2007-03-05 | Niallia circulans | 0.83 | 1 | 0
|  |  | CVUAS 2582.2 | 3; sequenced | eDT | 2015-09-23 | Niallia nealsonii | 0.61 | 1 | 0
|  |  | CVUAS 33253 | 2; sequenced | EFex | 2023-01-12 | Niallia sp | 0.76 | 1 | 0
|  |  | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-10 | Peribacillus muralis | 0.58 | 1 | 0
|  |  | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-18 | Peribacillus psychrosaccharolyticus | 0.63 | 1 | 0
|  |  | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus simplex | 0.66 | 1 | 0
|  |  | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-09-03 | Peribacillus simplex | 0.38 | 1 | 0
|  |  | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-12-15 | Peribacillus simplex | 0.55 | 1 | 0
|  |  | DSM 110302 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2021-11-17 | Priestia endophytica | 0.83 | 1 | 0
|  |  | DSM 13796 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2021-11-17 | Priestia endophytica | 0.47 | 1 | 0
|  |  | DSM 1316 | 3; public strain collection | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from DSMZ strain collection#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Priestia flexa | 0.71 | 1 | 0
|  |  | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-18 | Priestia flexa | 0.9 | 1 | 0
|  |  | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-01-06 | Priestia flexa | 0.54 | 1 | 0
|  |  | DSM 30607 | 3; public strain collection | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦1:10 diluted¦1:1 HCCA:TA2#EXT=Strain obtained from DSMZ strain collection#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-22 | Priestia flexa | 0.64 | 1 | 0
|  |  | ATCC 14581 | 3; type strain | DT | 2014-01-29 | Priestia megaterium | 0.78 | 1 | 0
|  |  | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-26 | Priestia megaterium | 1.01 | 1 | 0
|  |  | B 2 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-12-01 | Priestia megaterium | 0.87 | 1 | 0
|  |  | B 333 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-11 | Priestia megaterium | 1.24 | 1 | 0
|  |  | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia megaterium | 0.45 | 1 | 0
|  |  | DSM 32 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Priestia megaterium | 1.07 | 1 | 0
|  |  | DSM 333 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Priestia megaterium | 0.85 | 1 | 0
|  |  | DSM 509 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Priestia megaterium | 1.07 | 1 | 0
|  |  | DSM 90 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-13 | Priestia megaterium | 1.07 | 1 | 0
|  |  | RRLK06_01 | 3; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-11 | Priestia megaterium | 1.01 | 1 | 0
|  |  | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-30 | Salibacterium salarium | 0.53 | 1 | 0
|  |  | DSM 6307 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Sutcliffiella cohnii | 0.77 | 1 | 0
|  |  | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-20 | Sutcliffiella cohnii | 0.59 | 1 | 0
|  |  | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-25 | Sutcliffiella cohnii | 0.68 | 1 | 0
|  |  | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus saccharophilus | 0.55 | 1 | 0
|  |  | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Virgibacillus pantothenticus | 0.62 | 1 | 0
|  |  | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus proomii | 0.56 | 1 | 0
|  |  | CVUAS 3616.2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus proomii | 0.78 | 1 | 0
|  |  | CVUAS 71.2 | 3; sequenced | DT | 2013-12-13 | Weizmannia ginsengihumi | 0.72 | 1 | 0
|  |  | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus halophilus | 0.73 | 1 | 0
|  |  | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus huakuii | 0.59 | 1 | 0
|  |  | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus psychrophilus | 0.95 | 1 | 0
|  |  | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 2013-03-14 | Macrococcus caseolyticus | 0.65 | 1 | 0
|  |  | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 0.85 | 1 | 0
|  |  | CVUAS 4448 | 3; 16S rDNA-sequencing | DT | 2013-03-14 | Macrococcus caseolyticus | 1 | 1 | 0
|  |  | CVUAS 4454 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 0.85 | 1 | 0
|  |  | CVUAS 4456 | 3; 16S rDNA-sequencing | DT | 2013-03-14 | Macrococcus caseolyticus | 0.75 | 1 | 0
|  |  | CVUAS 4503 | 3; sequenced | DT | 2013-03-14 | Macrococcus caseolyticus | 0.89 | 1 | 0
|  |  | CVUAS 4506 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 1.03 | 1 | 0
|  |  | CVUAS 4510 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 0.89 | 1 | 0
|  |  | CVUAS 4513 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 0.85 | 1 | 0
|  |  | CVUAS 4636 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 1.06 | 1 | 0
|  |  | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus carouselicus | 0.85 | 1 | 0
|  |  | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus lentus | 0.66 | 1 | 0
|  |  | 4363 | 2; API Staph ID32 | DT | 2013-04-30 | Mammaliicoccus sciuri | 0.65 | 1 | 0
|  |  | A 4778 | 2; API Staph ID32 | DT | 2013-04-30 | Mammaliicoccus sciuri | 0.67 | 1 | 0
|  |  | coa145 | 2; API Staph ID32 | DT | 2013-04-30 | Mammaliicoccus sciuri | 0.9 | 1 | 0
|  |  | CVUAS 3442 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus sciuri | 0.74 | 1 | 0
|  |  | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus sciuri | 0.74 | 1 | 0
|  |  | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus sciuri | 0.93 | 1 | 0
|  |  | CVUAS 9937.2 DT | 3; 16S rDNA-sequencing | DT | 2015-12-01 | Mammaliicoccus vitulinus | 1.02 | 1 | 0
|  |  | CVUAS 9937.2 extr | 3; 16S rDNA-sequencing | EFex | 2016-04-06 | Mammaliicoccus vitulinus | 1.06 | 1 | 0
|  |  | CVUAS 343.2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus vitulinus | 0.79 | 1 | 0
|  |  | CVUAS 5687 sir | 3; sequenced | EFex | 2025-07-08 | Staphylococcus agnetis | 0.98 | 1 | 0
|  |  | CVUAS 5687 LT | 3; sequenced | DT | 2013-04-16 | Staphylococcus agnetis | 1.26 | 1 | 0
|  |  | CVUAS 622.3 | 2; biochemical | EFex | 2025-07-08 | Staphylococcus agnetis | 1.25 | 1 | 0
|  |  | DSM 23656 LT | 3; type strain [0003] | DT | 2013-04-30 | Staphylococcus agnetis | 0.81 | 1 | 0
|  |  | DSM 23656 sir | 3; type strain [0003] | EFex | 2025-07-08 | Staphylococcus agnetis | 0.89 | 1 | 0
|  |  | TUM 7385 sir | 3; sequenced | EFex | 2025-07-08 | Staphylococcus agnetis | 1.12 | 1 | 0
|  |  | TUM 7385 LT | 3; sequenced | DT | 02.10.2013 | Staphylococcus agnetis | 1.11 | 1 | 0
|  |  | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus argenteus | 1.57 | 1 | 0
|  |  | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus argenteus | 1.08 | 1 | 0
|  |  | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus arlettae | 0.99 | 1 | 0
|  |  | 265U3-3 | 2; sequenced | DT | 2016-08-31 | Staphylococcus arlettae | 0.76 | 1 | 0
|  |  | coa031 | 2; API Staph ID32; 66342360 | DT | 2013-03-20 | Staphylococcus arlettae | 0.73 | 1 | 0
|  |  | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus arlettae | 0.91 | 1 | 0
|  |  | DSM 20672 | 3; type strain [0003] | DT | 04.10.2013 | Staphylococcus arlettae | 0.67 | 1 | 0
|  |  | KNS 35 | 2; API Staph ID32: 66340261 | DT | 2013-02-13 | Staphylococcus arlettae | 0.37 | 1 | 0
|  |  | TUM MG88 | 3; sequenced | DT | 02.10.2013 | Staphylococcus arlettae | 0.47 | 1 | 0
|  |  | AGES 001 | 3; 0 | EFEx | 2019-03-20 | Staphylococcus aureus | 0.8 | 1 | 0
|  |  | DSM 20231 CASO | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-02-25 | Staphylococcus aureus | 1.4 | 1 | 0
|  |  | 96-01678 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-09-28 | Staphylococcus aureus | 1.23 | 1 | 0
|  |  | 09-00981 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.3 | 1 | 0
|  |  | 09-01243 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.13 | 1 | 0
|  |  | 03-00397 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.35 | 1 | 0
|  |  | 08-00602 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-01-27 | Staphylococcus aureus | 1.42 | 1 | 0
|  |  | 04-03100 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-10-02 | Staphylococcus aureus | 1.26 | 1 | 0
|  |  | 07-03338 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.14 | 1 | 0
|  |  | 03-00220 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.57 | 1 | 0
|  |  | 06-00896 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.47 | 1 | 0
|  |  | 07-03347 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.61 | 1 | 0
|  |  | 09-00186 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus aureus | 1.65 | 1 | 0
|  |  | 09-00416 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-17 | Staphylococcus aureus | 1.45 | 1 | 0
|  |  | 09-01140 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus aureus | 1.35 | 1 | 0
|  |  | 09-02135 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-05-17 | Staphylococcus aureus | 1.46 | 1 | 0
|  |  | 08-01104 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Staphylococcus aureus | 1.43 | 1 | 0
|  |  | 07-01238 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-02 | Staphylococcus aureus | 1.27 | 1 | 0
|  |  | 07-01949 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-28 | Staphylococcus aureus | 1.52 | 1 | 0
|  |  | 07-01531 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-16 | Staphylococcus aureus | 1.5 | 1 | 0
|  |  | 09-02446 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2022-11-16 | Staphylococcus aureus | 1.22 | 1 | 0
|  |  | 09-02476 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-28 | Staphylococcus aureus | 1.47 | 1 | 0
|  |  | 09-02431 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus aureus | 1.32 | 1 | 0
|  |  | 09-02434 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus aureus | 1.52 | 1 | 0
|  |  | 09-02444 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus aureus | 1.47 | 1 | 0
|  |  | 08-00401 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus aureus | 1.39 | 1 | 0
|  |  | 08-00004 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus aureus | 1.24 | 1 | 0
|  |  | 08-02172 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus aureus | 1.35 | 1 | 0
|  |  | 08-02337 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus aureus | 1.57 | 1 | 0
|  |  | 08-02874 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus aureus | 1.37 | 1 | 0
|  |  | 08-02989 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-02 | Staphylococcus aureus | 1.46 | 1 | 0
|  |  | 08-02727 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-02 | Staphylococcus aureus | 1.38 | 1 | 0
|  |  | 09-00558 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-02 | Staphylococcus aureus | 1.38 | 1 | 0
|  |  | N315 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus aureus | 1.26 | 1 | 0
|  |  | Mu50 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus aureus | 1.33 | 1 | 0
|  |  | 07-00170 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus aureus | 1.28 | 1 | 0
|  |  | 07-02020 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-09 | Staphylococcus aureus | 1.4 | 1 | 0
|  |  | 06-00603 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-09 | Staphylococcus aureus | 1.41 | 1 | 0
|  |  | 02-02424 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2010-05-17 | Staphylococcus aureus | 1.33 | 1 | 0
|  |  | 06-00953-1 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-09 | Staphylococcus aureus | 1.28 | 1 | 0
|  |  | 04-02232 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-23 | Staphylococcus aureus | 1.41 | 1 | 0
|  |  | 06-01243 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus aureus | 1.43 | 1 | 0
|  |  | 02-03179 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.43 | 1 | 0
|  |  | 04-02981 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.44 | 1 | 0
|  |  | 08-00392 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2015-05-21 | Staphylococcus aureus | 1.41 | 1 | 0
|  |  | 09-00825 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-21 | Staphylococcus aureus | 1.26 | 1 | 0
|  |  | 08-00463 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-21 | Staphylococcus aureus | 1.33 | 1 | 0
|  |  | 07-00265 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-16 | Staphylococcus aureus | 1.52 | 1 | 0
|  |  | 07-00010 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-09 | Staphylococcus aureus | 1.44 | 1 | 0
|  |  | 08-00250 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus aureus | 1.34 | 1 | 0
|  |  | 08-00993 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-10-05 | Staphylococcus aureus | 1.38 | 1 | 0
|  |  | 08-01718 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-05 | Staphylococcus aureus | 1.45 | 1 | 0
|  |  | 09-00001 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-10-05 | Staphylococcus aureus | 1.17 | 1 | 0
|  |  | 09-01189 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus aureus | 1.21 | 1 | 0
|  |  | 09-02378 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus aureus | 1.39 | 1 | 0
|  |  | 09-00615 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus aureus | 1.33 | 1 | 0
|  |  | 09-02069 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus aureus | 1.35 | 1 | 0
|  |  | 09-01345 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-11-23 | Staphylococcus aureus | 1.63 | 1 | 0
|  |  | 09-00212 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-23 | Staphylococcus aureus | 1.4 | 1 | 0
|  |  | 09-01791-1 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-23 | Staphylococcus aureus | 1.43 | 1 | 0
|  |  | 07-02652 | 3; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2009-10-23 | Staphylococcus aureus | 1.44 | 1 | 0
|  |  | DSM 20231 TSA | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-02-25 | Staphylococcus aureus | 1.4 | 1 | 0
|  |  | 11-00783 | 2; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-04 | Staphylococcus aureus | 1.35 | 1 | 0
|  |  | 11-00902 | 2; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus aureus | 1.59 | 1 | 0
|  |  | 11-00961 | 2; well described strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-04 | Staphylococcus aureus | 1.5 | 1 | 0
|  |  | 11-01100 | 2; well described strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-05 | Staphylococcus aureus | 1.57 | 1 | 0
|  |  | 2017-4101-008269 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus aureus | 1.16 | 1 | 0
|  |  | 2018-4101-001305-003 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus aureus | 1.31 | 1 | 0
|  |  | 2018-4101-001376 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus aureus | 1.35 | 1 | 0
|  |  | 2018-4101-002922-005 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus aureus | 1.57 | 1 | 0
|  |  | 2018-4101-003040 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus aureus | 1.3 | 1 | 0
|  |  | 2018-4104-000006 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus aureus | 1.22 | 1 | 0
|  |  | 3321-2017-000527 | 2; Biochemistry | DT | 2018-06-27 | Staphylococcus aureus | 1.27 | 1 | 0
|  |  | 4201-2017-002411 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus aureus | 1.46 | 1 | 0
|  |  | 4201-2017-003006 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus aureus | 1.55 | 1 | 0
|  |  | 4201-2017-004220 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus aureus | 1.25 | 1 | 0
|  |  | 4201-2017-004348 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus aureus | 1.31 | 1 | 0
|  |  | 4201-2018-001040 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus aureus | 1.14 | 1 | 0
|  |  | 4201-2018-001081 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus aureus | 1.21 | 1 | 0
|  |  | 4201-2018-001446 | 3; confirmed by NRL BfR | DT | 2018-06-27 | Staphylococcus aureus | 1.28 | 1 | 0
|  |  | ATCC 29213 CVUAS | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus aureus | 1.46 | 1 | 0
|  |  | ATCC 29213 RKI | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-04 | Staphylococcus aureus | 1.51 | 1 | 0
|  |  | C3658-06 | 3; sequenced (LGL) | EFex | 2015-04-15 | Staphylococcus aureus | 1.2 | 1 | 0
|  |  | COL | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus aureus | 0.95 | 1 | 0
|  |  | CVUAS 10197.2 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus aureus | 1.28 | 1 | 0
|  |  | CVUAS 10507 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus aureus | 1.28 | 1 | 0
|  |  | CVUAS 1083 | 4; sequenztype | DT | 05.02.2013 | Staphylococcus aureus | 1.17 | 1 | 0
|  |  | CVUAS 1304.2 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus aureus | 1.44 | 1 | 0
|  |  | CVUAS 1305.2 | 4; sequenztype | DT | 2014-08-20 | Staphylococcus aureus | 1.3 | 1 | 0
|  |  | CVUAS 1894 | 4; sequenztype | DT | 2013-02-05 | Staphylococcus aureus | 1.05 | 1 | 0
|  |  | CVUAS 2645 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus aureus | 1.39 | 1 | 0
|  |  | CVUAS 3072 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus aureus | 1.25 | 1 | 0
|  |  | CVUAS 3194 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus aureus | 1.26 | 1 | 0
|  |  | CVUAS 3195 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus aureus | 1.12 | 1 | 0
|  |  | CVUAS 3196 | 4; sequenztype | DT | 13.01.2015 | Staphylococcus aureus | 1.18 | 1 | 0
|  |  | CVUAS 3199 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus aureus | 1.16 | 1 | 0
|  |  | CVUAS 3204 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus aureus | 1.34 | 1 | 0
|  |  | CVUAS 3206 | 4; sequenztype | DT | 31.10.2012 | Staphylococcus aureus | 1.22 | 1 | 0
|  |  | CVUAS 3260 | 4; sequenztype | DT | 31.10.2012 | Staphylococcus aureus | 1.13 | 1 | 0
|  |  | CVUAS 3320 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus aureus | 1.14 | 1 | 0
|  |  | CVUAS 3330 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus aureus | 1.3 | 1 | 0
|  |  | CVUAS 3433 | 4; sequenztype | DT | 31.10.2012 | Staphylococcus aureus | 1.36 | 1 | 0
|  |  | CVUAS 3452.2 | 4; sequenztype | DT | 10.09.2013 | Staphylococcus aureus | 1.33 | 1 | 0
|  |  | CVUAS 3740.2 | 4; sequenztype | DT | 2014-08-20 | Staphylococcus aureus | 1.32 | 1 | 0
|  |  | CVUAS 3771 | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus aureus | 1.31 | 1 | 0
|  |  | CVUAS 3776 | 3; 16S rDNA-sequencing | DT | 23.10.2012 | Staphylococcus aureus | 1.45 | 1 | 0
|  |  | CVUAS 4645 | 3; 16S rDNA-sequencing | DT | 2013-02-05 | Staphylococcus aureus | 1.07 | 1 | 0
|  |  | CVUAS 5756 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus aureus | 1.38 | 1 | 0
|  |  | CVUAS 6080 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus aureus | 1.28 | 1 | 0
|  |  | CVUAS 6459 | 3; 16S rDNA-sequencing | DT | 2014-05-14 | Staphylococcus aureus | 1.16 | 1 | 0
|  |  | CVUAS 6963 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus aureus | 1.12 | 1 | 0
|  |  | CVUAS 7112 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus aureus | 1.46 | 1 | 0
|  |  | CVUAS 7665 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus aureus | 1.13 | 1 | 0
|  |  | CVUAS 7779 | 3; 16S rDNA-sequencing | DT | 2012-09-12 | Staphylococcus aureus | 1.19 | 1 | 0
|  |  | CVUAS 7958.2 | 3; sequenztype | DT | 2013-04-30 | Staphylococcus aureus | 1.23 | 1 | 0
|  |  | CVUAS 7963.2 | 3; sequenztype | DT | 2013-04-30 | Staphylococcus aureus | 0.88 | 1 | 0
|  |  | CVUAS 8134 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus aureus | 1.33 | 1 | 0
|  |  | CVUAS 8135 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus aureus | 1.23 | 1 | 0
|  |  | CVUAS 8312 | 4; sequenztype | DT | 2012-10-23 | Staphylococcus aureus | 1.04 | 1 | 0
|  |  | CVUAS 8900 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus aureus | 1.08 | 1 | 0
|  |  | CVUAS 9019 | 4; sequenztype | DT | 2012-10-31 | Staphylococcus aureus | 1.17 | 1 | 0
|  |  | CVUAS 9197 | 4; sequenztype | DT | 2013-02-13 | Staphylococcus aureus | 1.07 | 1 | 0
|  |  | CVUAS 935 | 3; 16S rDNA-sequencing | DT | 2014-05-19 | Staphylococcus aureus | 1.42 | 1 | 0
|  |  | CVUAS 9373 | 4; sequenztype | DT | 2012-11-14 | Staphylococcus aureus | 1.1 | 1 | 0
|  |  | CVUAS 9388 | 4; sequenztype | DT | 2015-11-05 | Staphylococcus aureus | 1.23 | 1 | 0
|  |  | CVUAS 9642 | 2; biochemical tube tests | DT | 2012-12-04 | Staphylococcus aureus | 1.34 | 1 | 0
|  |  | CVUAS 9681 | 4; Sequencetyp (NRL-Staph, BfR) | DT | 2015-11-05 | Staphylococcus aureus | 1.13 | 1 | 0
|  |  | CVUAS 9682 | 4; Sequencetyp (NRL-Staph, BfR) | DT | 2015-11-05 | Staphylococcus aureus | 1.36 | 1 | 0
|  |  | CVUAS 9704 | 4; sequenztype | DT | 2015-01-13 | Staphylococcus aureus | 1.05 | 1 | 0
|  |  | CVUAS 9790 | 4; Sequencetyp (NRL-Staph, BfR) | eDT | 2015-11-05 | Staphylococcus aureus | 1.5 | 1 | 0
|  |  | CVUAS 9914.2 sir | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus aureus | 1.52 | 1 | 0
|  |  | CVUAS 9914.2 LT | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus aureus | 1.42 | 1 | 0
|  |  | DSM 1104 | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus aureus | 1.24 | 1 | 0
|  |  | DSM 11729 | 3; public strain collection | DT | 2013-05-14 | Staphylococcus aureus | 1.38 | 1 | 0
|  |  | DSM 12463 | 3; public strain collection | DT | 2019-11-21 | Staphylococcus aureus | 1.6 | 1 | 0
|  |  | DSM 20491 | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus aureus | 1.18 | 1 | 0
|  |  | DSM 346 | 3; public strain collection | eDT | 2013-03-12 | Staphylococcus aureus | 1.34 | 1 | 0
|  |  | DSM 4910 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-05-24 | Staphylococcus aureus | 1.37 | 1 | 0
|  |  | MSSA 129 | 4; 16S rDNA-sequencing | DT | 2013-03-12 | Staphylococcus aureus | 1.16 | 1 | 0
|  |  | RKI WG PS 29 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-04-09 | Staphylococcus aureus | 1.46 | 1 | 0
|  |  | RKI WG PS 42D | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-02-20 | Staphylococcus aureus | 1.4 | 1 | 0
|  |  | RKI WG PS 94 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-02-20 | Staphylococcus aureus | 1.41 | 1 | 0
|  |  | RKI WG PS 96 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-04-02 | Staphylococcus aureus | 1.34 | 1 | 0
|  |  | RKI4 | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus aureus | 1.18 | 1 | 0
|  |  | S6C | 3; 16S rDNA-sequencing | DT | 2012-10-23 | Staphylococcus aureus | 1.16 | 1 | 0
|  |  | WDCM 00131 | 3; public strain collection | DT | 2018-06-27 | Staphylococcus aureus | 1.33 | 1 | 0
|  |  | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus auricularis | 1.03 | 1 | 0
|  |  | MB 2023000185 | 2; corresponds to expected result of PT | EFEx | 2023-09-07 | Staphylococcus auricularis | 0.97 | 1 | 0
|  |  | DSM 6717 | 3; type strain [0003] | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0 | 2022-09-01 | Staphylococcus capitis | 1.36 | 1 | 0
|  |  | DSM 20326 | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus capitis | 1.39 | 1 | 0
|  |  | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus capitis | 0.99 | 1 | 0
|  |  | 21-1 | 2; confirmed proficency test | eDT | 2022-02-14 | Staphylococcus caprae | 1.08 | 1 | 0
|  |  | DSM 20608 RWTH | 3; type strain [0003] | EFEx | 2023-06-20 | Staphylococcus caprae | 1.22 | 1 | 0
|  |  | DSM 20608 CVUAS | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus caprae | 1.19 | 1 | 0
|  |  | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus caprae | 1.29 | 1 | 0
|  |  | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus carnosus | 1.05 | 1 | 0
|  |  | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus carnosus | 0.96 | 1 | 0
|  |  | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus chromogenes | 1.11 | 1 | 0
|  |  | coa009a sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus chromogenes | 1.39 | 1 | 0
|  |  | coa009a LT | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus chromogenes | 0.87 | 1 | 0
|  |  | coa051 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus chromogenes | 1.32 | 1 | 0
|  |  | coa051 LT | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus chromogenes | 1.17 | 1 | 0
|  |  | coa096 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus chromogenes | 0.97 | 1 | 0
|  |  | coa121 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus chromogenes | 1.19 | 1 | 0
|  |  | coa143 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus chromogenes | 0.92 | 1 | 0
|  |  | CVUAS 3348 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus chromogenes | 1.32 | 1 | 0
|  |  | CVUAS 3348 LT | 3; 16S rDNA-sequencing | DT | 2013-02-05 | Staphylococcus chromogenes | 1.2 | 1 | 0
|  |  | CVUAS 3349 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus chromogenes | 1.24 | 1 | 0
|  |  | CVUAS 3706 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus chromogenes | 1.27 | 1 | 0
|  |  | CVUAS 3706 LT | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus chromogenes | 1.13 | 1 | 0
|  |  | CVUAS 3784 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus chromogenes | 1.33 | 1 | 0
|  |  | CVUAS 3784 LT | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus chromogenes | 1.13 | 1 | 0
|  |  | KNS 44 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus chromogenes | 1.31 | 1 | 0
|  |  | KNS 60 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus chromogenes | 1.03 | 1 | 0
|  |  | KNS 73 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus chromogenes | 1.22 | 1 | 0
|  |  | MKNS 02 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus chromogenes | 0.94 | 1 | 0
|  |  | MKNS 26 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus chromogenes | 0.99 | 1 | 0
|  |  | TS 084 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus chromogenes | 1.51 | 1 | 0
|  |  | TS 084 LT | 3; 16S rDNA-sequencing | DT | 2013-03-20 | Staphylococcus chromogenes | 1.18 | 1 | 0
|  |  | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus chromogenes | 1.19 | 1 | 0
|  |  | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus cohnii | 0.96 | 1 | 0
|  |  | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus cohnii | 0.84 | 1 | 0
|  |  | 217b-3 | 3; 16S rDNA-sequencing | DT | 2016-08-31 | Staphylococcus cohnii | 0.88 | 1 | 0
|  |  | MKNS 18 | 2; API Staph ID32 | DT | 2013-03-20 | Staphylococcus cohnii | 0.93 | 1 | 0
|  |  | MKNS 24 | 3; sequenced, biochemical profile | DT | 20.03.2013 | Staphylococcus cohnii | 0.98 | 1 | 0
|  |  | TUM MG98 | 3; rpoB-sequencing | DT | 2013-10-24 | Staphylococcus cohnii | 0.97 | 1 | 0
|  |  | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus delphini | 0.84 | 1 | 0
|  |  | x Sta RKI | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-04 | Staphylococcus delphini | 0.99 | 1 | 0
|  |  | 11-00903 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2011-08-04 | Staphylococcus epidermidis | 1.08 | 1 | 0
|  |  | CCM 2124 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-05-19 | Staphylococcus epidermidis | 1.11 | 1 | 0
|  |  | coa076 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus epidermidis | 1.06 | 1 | 0
|  |  | CVUAS 3134 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 0.93 | 1 | 0
|  |  | CVUAS 3439 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 1.07 | 1 | 0
|  |  | CVUAS 3441 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 0.67 | 1 | 0
|  |  | CVUAS 3560 | 3; 16S rDNA-sequencing | DT | 2012-12-12 | Staphylococcus epidermidis | 1.13 | 1 | 0
|  |  | CVUAS 3701 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 0.8 | 1 | 0
|  |  | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 1.2 | 1 | 0
|  |  | CVUAS 3787 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 1.02 | 1 | 0
|  |  | CVUAS 5742 | 2; 0 | eDT | 2015-11-05 | Staphylococcus epidermidis | 1.29 | 1 | 0
|  |  | CVUAS 6867 | 2; biochemical tube tests | DT | 2012-12-04 | Staphylococcus epidermidis | 1.16 | 1 | 0
|  |  | DSM 1798 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2008-02-20 | Staphylococcus epidermidis | 0.88 | 1 | 0
|  |  | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus epidermidis | 1.1 | 1 | 0
|  |  | RRLK07_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Staphylococcus epidermidis | 0.89 | 1 | 0
|  |  | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus equorum | 1.04 | 1 | 0
|  |  | coa034 | 2; API Staph ID32 | DT | 2013-03-20 | Staphylococcus equorum | 0.74 | 1 | 0
|  |  | DSM 20674 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus equorum | 0.83 | 1 | 0
|  |  | TUM MG847 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus equorum | 0.62 | 1 | 0
|  |  | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus felis | 0.83 | 1 | 0
|  |  | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus felis | 0.89 | 1 | 0
|  |  | A269-3b | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-02 | Staphylococcus hominis | 1.2 | 1 | 0
|  |  | CVUAS 789.2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus hominis | 0.8 | 1 | 0
|  |  | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus hominis | 1.16 | 1 | 0
|  |  | DSM 20330 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus hominis | 0.98 | 1 | 0
|  |  | RRLK01_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-15 | Staphylococcus hominis | 1.07 | 1 | 0
|  |  | RRLK06_04 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Staphylococcus hominis | 1.14 | 1 | 0
|  |  | coa179 sir | 2; biochemical tests (API Staph ID32) | EFex | 2025-07-08 | Staphylococcus hyicus | 0.89 | 1 | 0
|  |  | coa179 LT | 2; biochemical tests (API Staph ID32) | DT | 2013-07-02 | Staphylococcus hyicus | 1.02 | 1 | 0
|  |  | CVUAS 1346 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus hyicus | 1.06 | 1 | 0
|  |  | CVUAS 1347 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus hyicus | 0.78 | 1 | 0
|  |  | CVUAS 1525 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus hyicus | 1.13 | 1 | 0
|  |  | CVUAS 1526 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus hyicus | 0.93 | 1 | 0
|  |  | CVUAS 1552 | 2; biochemical tube tests | DT | 2015-08-26 | Staphylococcus hyicus | 1.02 | 1 | 0
|  |  | CVUAS 1553 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus hyicus | 0.61 | 1 | 0
|  |  | CVUAS 1702 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus hyicus | 1.11 | 1 | 0
|  |  | CVUAS 1742 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus hyicus | 1.08 | 1 | 0
|  |  | CVUAS 1917 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus hyicus | 1.08 | 1 | 0
|  |  | CVUAS 1984 sir | 2; biochemical tube tests | EFex | 2025-07-08 | Staphylococcus hyicus | 1.35 | 1 | 0
|  |  | CVUAS 1984 LT | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus hyicus | 1.34 | 1 | 0
|  |  | CVUAS 212 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus hyicus | 1.19 | 1 | 0
|  |  | CVUAS 2148 | 2; biochemical tube tests | DT | 2013-04-16 | Staphylococcus hyicus | 1.09 | 1 | 0
|  |  | CVUAS 2152 | 2; biochemical tube tests | DT | 2013-04-10 | Staphylococcus hyicus | 1.18 | 1 | 0
|  |  | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus hyicus | 1.21 | 1 | 0
|  |  | CVUAS 32680 | 2; biochemical tube tests | EFex | 2025-07-08 | Staphylococcus hyicus | 1.18 | 1 | 0
|  |  | CVUAS 330 | 2; biochemical tube tests | DT | 2013-04-16 | Staphylococcus hyicus | 1.36 | 1 | 0
|  |  | CVUAS 3767 sir | 3; 16S rDNA-sequencing | EFex | 2025-07-08 | Staphylococcus hyicus | 1.22 | 1 | 0
|  |  | CVUAS 3767 LT | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus hyicus | 1.34 | 1 | 0
|  |  | CVUAS 593 | 2; biochemical tube tests | DT | 2013-02-05 | Staphylococcus hyicus | 1.13 | 1 | 0
|  |  | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus intermedius | 1.03 | 1 | 0
|  |  | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus intermedius | 0.71 | 1 | 0
|  |  | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus lugdunensis | 1.31 | 1 | 0
|  |  | 09-00272 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus lugdunensis | 1.41 | 1 | 0
|  |  | 09-00274 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus lugdunensis | 1.46 | 1 | 0
|  |  | 09-00276 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus lugdunensis | 1.37 | 1 | 0
|  |  | 09-00283 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-04-04 | Staphylococcus lugdunensis | 1.3 | 1 | 0
|  |  | 09-00285 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus lugdunensis | 1.34 | 1 | 0
|  |  | 09-00513 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus lugdunensis | 1.56 | 1 | 0
|  |  | 09-00515 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-08-04 | Staphylococcus lugdunensis | 1.43 | 1 | 0
|  |  | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2012-04-05 | Staphylococcus lugdunensis | 1.05 | 1 | 0
|  |  | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus microti | 0.96 | 1 | 0
|  |  | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus pasteuri | 0.86 | 1 | 0
|  |  | CVUAS 3086 | 3; 16S rDNA-sequencing | DT | 2013-04-10 | Staphylococcus pasteuri | 1 | 1 | 0
|  |  | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus pasteuri | 1.01 | 1 | 0
|  |  | MKNS 30 | 3; sequenced | DT | 2013-09-04 | Staphylococcus pasteuri | 0.98 | 1 | 0
|  |  | TUM 5823 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus pasteuri | 0.85 | 1 | 0
|  |  | VB969149 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus pseudintermedius | 0.88 | 1 | 0
|  |  | VB969390 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus pseudintermedius | 0.69 | 1 | 0
|  |  | VB971580 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus pseudintermedius | 1 | 1 | 0
|  |  | VB971904 | 4; sequenztype | DT | 2013-06-28 | Staphylococcus pseudintermedius | 0.94 | 1 | 0
|  |  | 11-0127-2 | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-03-02 | Staphylococcus pseudintermedius | 1.19 | 1 | 0
|  |  | CVUAS 2375 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus pseudintermedius | 0.8 | 1 | 0
|  |  | CVUAS 3688 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus pseudintermedius | 0.8 | 1 | 0
|  |  | CVUAS 3831 | 2; biochemical tube tests | DT | 2013-07-05 | Staphylococcus pseudintermedius | 0.93 | 1 | 0
|  |  | CVUAS 3856 | 3; 16S rDNA-sequencing (CVUAS) | eDT | 2015-04-21 | Staphylococcus pseudintermedius | 0.69 | 1 | 0
|  |  | CVUAS 5457 | 2; biochemical tube tests | DT | 2013-07-05 | Staphylococcus pseudintermedius | 1.08 | 1 | 0
|  |  | CVUAS 5594 | 2; biochemical tube tests | DT | 2013-07-05 | Staphylococcus pseudintermedius | 1.11 | 1 | 0
|  |  | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus pseudintermedius | 0.66 | 1 | 0
|  |  | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2016-03-02 | Staphylococcus pseudintermedius | 1.21 | 1 | 0
|  |  | CVUAS 6260.2 eDT | 3; sequenced | eDT | 2023-10-19 | Staphylococcus ratti | 1.08 | 1 | 0
|  |  | CVUAS 6260.2 extr | 3; sequenced | EFex | 2023-10-19 | Staphylococcus ratti | 1.51 | 1 | 0
|  |  | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus saprophyticus | 0.58 | 1 | 0
|  |  | A 436 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus saprophyticus | 1.02 | 1 | 0
|  |  | A 446 | 2; biochemical tube tests | DT | 2013-09-11 | Staphylococcus saprophyticus | 0.94 | 1 | 0
|  |  | coa045 | 2; biochemical tube tests | DT | 2013-03-20 | Staphylococcus saprophyticus | 0.91 | 1 | 0
|  |  | CVUAS 503.2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus saprophyticus | 0.99 | 1 | 0
|  |  | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus schleiferi | 1.23 | 1 | 0
|  |  | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus schweitzeri | 0.81 | 1 | 0
|  |  | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus simulans | 1.05 | 1 | 0
|  |  | 4919 | 2; API Staph ID32 | DT | 2013-07-02 | Staphylococcus simulans | 1.05 | 1 | 0
|  |  | coa037 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus simulans | 1.19 | 1 | 0
|  |  | coa117 | 2; API Staph ID32 | DT | 2013-02-13 | Staphylococcus simulans | 1.09 | 1 | 0
|  |  | coa126 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus simulans | 1.15 | 1 | 0
|  |  | coa150 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus simulans | 0.96 | 1 | 0
|  |  | DSM 20322 | 3; type strain | DT | 2013-10-02 | Staphylococcus simulans | 1.34 | 1 | 0
|  |  | KNS 05 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus simulans | 1.1 | 1 | 0
|  |  | KNS 27 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus simulans | 0.9 | 1 | 0
|  |  | MKNS 04 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus simulans | 1.22 | 1 | 0
|  |  | MKNS 08 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus simulans | 1.11 | 1 | 0
|  |  | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus simulans | 1.06 | 1 | 0
|  |  | CVUAS 2439.2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 0.69 | 1 | 0
|  |  | CVUAS 761.2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 0.69 | 1 | 0
|  |  | 4379 | 2; API Staph ID32 | DT | 2013-03-19 | Staphylococcus warneri | 1.04 | 1 | 0
|  |  | 4380 | 2; API Staph ID32 | DT | 2013-04-10 | Staphylococcus warneri | 0.6 | 1 | 0
|  |  | coa032 | 2; API Staph ID32 | DT | 2013-04-10 | Staphylococcus warneri | 0.95 | 1 | 0
|  |  | coa083 | 3; rpoB-sequencing | DT | 2013-04-10 | Staphylococcus warneri | 1.1 | 1 | 0
|  |  | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus warneri | 1.21 | 1 | 0
|  |  | CVUAS 687 | 3; 16S rDNA-sequencing | DT | 2013-04-10 | Staphylococcus warneri | 0.6 | 1 | 0
|  |  | DSM 20316 CVUAS | 3; type strain | DT | 2013-10-02 | Staphylococcus warneri | 0.6 | 1 | 0
|  |  | DSM 20316 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2011-09-08 | Staphylococcus warneri | 0.63 | 1 | 0
|  |  | TUM 6412 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus warneri | 0.66 | 1 | 0
|  |  | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.31 | 1 | 0
|  |  | 6124 | 2; biochemical tube tests | EFEx | 2016-09-22 | Staphylococcus xylosus | 1.17 | 1 | 0
|  |  | coa157 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus xylosus | 0.94 | 1 | 0
|  |  | CVUAS 3440 | 3; 16S, Xyl pos | DT | 2013-03-26 | Staphylococcus xylosus | 0.87 | 1 | 0
|  |  | CVUAS 3562 | 3; 16S, Xyl pos | DT | 2013-03-26 | Staphylococcus xylosus | 1.36 | 1 | 0
|  |  | CVUAS 3563 | 3; sequenced | DT | 2012-11-14 | Staphylococcus xylosus | 1.22 | 1 | 0
|  |  | CVUAS 3566 | 3; 16S, Xyl pos | DT | 2013-03-26 | Staphylococcus xylosus | 0.78 | 1 | 0
|  |  | CVUAS 3692 | 3; 16S, Xyl pos | DT | 2012-10-31 | Staphylococcus xylosus | 1.11 | 1 | 0
|  |  | CVUAS 3699 | 3; 16S, Xyl pos | DT | 2013-07-24 | Staphylococcus xylosus | 0.53 | 1 | 0
|  |  | CVUAS 974.2 | 3; 16S, Xyl pos | eDT | 2015-08-25 | Staphylococcus xylosus | 0.84 | 1 | 0
|  |  | FLI-ING 4922 | 3; 16S, Xyl pos | EFEx | 2016-06-01 | Staphylococcus xylosus | 0.77 | 1 | 0
|  |  | KNS 61 | 2; biochemical tube tests | DT | 2013-03-20 | Staphylococcus xylosus | 0.44 | 1 | 0
|  |  | MKNS 06 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus xylosus | 0.67 | 1 | 0
|  |  | TS 227 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.79 | 1 | 0
|  |  | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.62 | 1 | 0
|  |  | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus devriesei | 1.38 | 1 | 0
|  |  | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces bovis | 0.87 | 1 | 0
|  |  | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces denticolens | 0.27 | 1 | 0
|  |  | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces sp-CVUAS-31303 | 0.61 | 1 | 0
|  |  | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces viscosus | 0.9 | 1 | 0
|  |  | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces weissii | 1.05 | 1 | 0
|  |  | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium bovis | 0.76 | 1 | 0
|  |  | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium canis | 0.92 | 1 | 0
|  |  | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium haemolyticum | 1.21 | 1 | 0
|  |  | 17-2056 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium hippocoleae | 0.71 | 1 | 0
|  |  | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium hippocoleae | 0.58 | 1 | 0
|  |  | DSM 10002 | 3; type strain | EFex | 2017-10-04 | Arcanobacterium phocae | 1.12 | 1 | 0
|  |  | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium phocae | 1.06 | 1 | 0
|  |  | 2698 | 3; type strain | DT | 2013-08-07 | Arcanobacterium phocisimile | 0.45 | 1 | 0
|  |  | 14418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium pinnipediorum | 0.57 | 1 | 0
|  |  | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium pinnipediorum | 0.65 | 1 | 0
|  |  | DSM 13483 | 3; type strain | EFex | 2017-10-04 | Arcanobacterium pluranimalium | 1.07 | 1 | 0
|  |  | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 0.8 | 1 | 0
|  |  | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 0.6 | 1 | 0
|  |  | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium sp-2701 | 1.07 | 1 | 0
|  |  | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium sp-2701 | 0.73 | 1 | 0
|  |  | CVUAS 1060.2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium wilhelmae | 0.88 | 1 | 0
|  |  | CVUAS 647 | 4; type strain | DT | 2014-09-04 | Arcanobacterium wilhelmae | 0.8 | 1 | 0
|  |  | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia sp-CVUAS-30561 | 0.75 | 1 | 0
|  |  | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia canis | 1.06 | 1 | 0
|  |  | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia cardiffensis | 0.8 | 1 | 0
|  |  | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia hyovaginalis | 0.87 | 1 | 0
|  |  | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia hyovaginalis | 0.94 | 1 | 0
|  |  | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia sp-CVUAS-32183 | 0.78 | 1 | 0
|  |  | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia sp-CVUAS-8688 | 0.98 | 1 | 0
|  |  | CVUAS 31845.2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia sp-CVUAS-8688 | 0.99 | 1 | 0
|  |  | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella abortisuis | 0.36 | 1 | 0
|  |  | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella abortisuis | 0.8 | 1 | 0
|  |  | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella bernardiae | 1 | 1 | 0
|  |  | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella bernardiae | 0.77 | 1 | 0
|  |  | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella bialowiezensis | 0.78 | 1 | 0
|  |  | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella bonasi | 0.76 | 1 | 0
|  |  | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.67 | 1 | 0
|  |  | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.13 | 1 | 0
|  |  | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella pyogenes | 0.32 | 1 | 0
|  |  | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella pyogenes | 1.15 | 1 | 0
|  |  | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium amycolatum | 0.99 | 1 | 0
|  |  | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-17 | Corynebacterium amycolatum | 0.82 | 1 | 0
|  |  | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium aurimucosum | 0.44 | 1 | 0
|  |  | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium belfantii | 0.59 | 1 | 0
|  |  | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium belfantii | 0.02 | 1 | 0
|  |  | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium bovis | 0.69 | 1 | 0
|  |  | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium camporealensis | 0.95 | 1 | 0
|  |  | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium casei | 0.89 | 1 | 0
|  |  | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium confusum | 0.22 | 1 | 0
|  |  | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium confusum | 0.44 | 1 | 0
|  |  | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium cystitidis | 0.41 | 1 | 0
|  |  | KL0163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.35 | 1 | 0
|  |  | KL0167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.16 | 1 | 0
|  |  | KL0173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.54 | 1 | 0
|  |  | KL0956 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 0.76 | 1 | 0
|  |  | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium frankenforstense | 0.96 | 1 | 0
|  |  | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium frankenforstense | 0.97 | 1 | 0
|  |  | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium freiburgense | 0.58 | 1 | 0
|  |  | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium hansenii | 0.45 | 1 | 0
|  |  | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium hansenii | 0.36 | 1 | 0
|  |  | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium humireducens | 0.81 | 1 | 0
|  |  | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium kutscheri | 1.14 | 1 | 0
|  |  | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium lactis | 0.93 | 1 | 0
|  |  | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium lactis | 0.5 | 1 | 0
|  |  | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium phoceense | 0.54 | 1 | 0
|  |  | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium pilosum | 0.57 | 1 | 0
|  |  | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium pseudotuberculosis | 0.85 | 1 | 0
|  |  | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium pseudotuberculosis | 0.89 | 1 | 0
|  |  | DSM 20689 CVUAS | 3; type strain | DT | 2012-12-18 | Corynebacterium pseudotuberculosis | 0.94 | 1 | 0
|  |  | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium pseudotuberculosis | 0.48 | 1 | 0
|  |  | C104 | 4; sequenced (rpoB) UGent; NGS (CVUAS) | DT | 2020-06-18 | Corynebacterium ramonii | 0.53 | 1 | 0
|  |  | CIP 112226 EFex | 3; type strain | EFEx | 2024-08-13 | Corynebacterium ramonii | 0.7 | 1 | 0
|  |  | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium rouxii | 0.15 | 1 | 0
|  |  | CVUAS 3559.2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium rouxii | 0.27 | 1 | 0
|  |  | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium rouxii | 0.3 | 1 | 0
|  |  | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium silvaticum | 0.76 | 1 | 0
|  |  | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium silvaticum | 0.64 | 1 | 0
|  |  | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium sp-208-U2-6 | 0.85 | 1 | 0
|  |  | CVUAS 32165.2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium sp-2300500 | 0.4 | 1 | 0
|  |  | CVUAS 5775.2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium sp-2300500 | 0.84 | 1 | 0
|  |  | CVUAS 2554.2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium sp-2396 | 0.78 | 1 | 0
|  |  | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium sp-243-z-3 | 0.87 | 1 | 0
|  |  | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium sp-257-y-4 | 0.5 | 1 | 0
|  |  | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium sp-284z-2 | 0.17 | 1 | 0
|  |  | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium sp-CVUAS-11664 | 0.56 | 1 | 0
|  |  | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium sp-CVUAS-1888 | 0.89 | 1 | 0
|  |  | CVUAS 1721.3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium sp-CVUAS-2484.4 | 0.8 | 1 | 0
|  |  | CVUAS 2484.4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388; 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium sp-CVUAS-2484.4 | 0.81 | 1 | 0
|  |  | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium sp-CVUAS-30404 | 0.72 | 1 | 0
|  |  | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-CVUAS-30404 | 0.92 | 1 | 0
|  |  | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium sp-CVUAS-31533 | 1.03 | 1 | 0
|  |  | CVUAS 32278.2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium sp-CVUAS-32278.2 | 0.82 | 1 | 0
|  |  | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium sp-CVUAS-874,4 | 0.76 | 1 | 0
|  |  | CVUAS 874.4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium sp-CVUAS-874,4 | 0.85 | 1 | 0
|  |  | LALLF 21TRD0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-LALLF-21TRD0553 | 0.61 | 1 | 0
|  |  | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium sp-LHL-0497 | 0.64 | 1 | 0
|  |  | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium sp-LHL-0558 | 0.7 | 1 | 0
|  |  | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium sp-LHL-1322 | 0.5 | 1 | 0
|  |  | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium stationis | 0.45 | 1 | 0
|  |  | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium testudinoris | 0.79 | 1 | 0
|  |  | CCM 2823 | 4; reference strain collection, WGS (ANI) CVUAS | DT | 2012-10-18 | Corynebacterium ulcerans | 0.84 | 1 | 0
|  |  | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium ulcerans | 0.27 | 1 | 0
|  |  | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium ulcerans | 0.71 | 1 | 0
|  |  | JB 362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium ulcerans | 0.78 | 1 | 0
|  |  | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 0.6 | 1 | 0
|  |  | CVUAS 6522.4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium ulcerans | 0.9 | 1 | 0
|  |  | CVUAS 5437.3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium variabile | 0.98 | 1 | 0
|  |  | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium xerosis | 1.08 | 1 | 0
|  |  | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium xerosis-group | 0.83 | 1 | 0
|  |  | CVUAS 2379.2 | 3; sequenced | EFex | 2021-04-08 | Gordonia paraffinivorans | 0.83 | 1 | 0
|  |  | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia paraffinivorans | 0.59 | 1 | 0
|  |  | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia rubripertincta | 1.05 | 1 | 0
|  |  | CVUAS 10185.3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium smegmatis | 0.96 | 1 | 0
|  |  | CVUAS 31296.3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia cyriacigeorgica | 0.33 | 1 | 0
|  |  | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia farcinica | 0.77 | 1 | 0
|  |  | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus equi | 0.98 | 1 | 0
|  |  | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 0.79 | 1 | 0
|  |  | CVUAS 1447.3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella sp | 0.39 | 1 | 0
|  |  | CVUAS 2895.3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium luteolum | 0.59 | 1 | 0
|  |  | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas flavigena | 0.33 | 1 | 0
|  |  | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium sp-234z-2 | 0.71 | 1 | 0
|  |  | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium sp-CVUAS-30560 | 0.82 | 1 | 0
|  |  | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus sedentarius | 0.38 | 1 | 0
|  |  | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus congolensis | 0.64 | 1 | 0
|  |  | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus congolensis | 0.64 | 1 | 0
|  |  | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium flaccumfaciens | 0.49 | 1 | 0
|  |  | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia sp-MUVA-A-508948 | 0 | 1 | 0
|  |  | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium sp | 1.14 | 1 | 0
|  |  | CVUAS 34837 | 3; sequenced | EFEx | 2023-01-19 | Microbacterium sp-CVUAS-34837 | 0.9 | 1 | 0
|  |  | CVUAS 34838 | 3; sequenced | EFEx | 2023-01-19 | Microbacterium sp-G5-3 | 0.35 | 1 | 0
|  |  | DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2011-09-01 | Glutamicibacter nicotianae | 0.76 | 1 | 0
|  |  | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter protophormiae | 1.13 | 1 | 0
|  |  | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria atrinae | 0.53 | 1 | 0
|  |  | 56a-5 | 3; sequenced | EFEx | 2020-05-20 | Kocuria gwangalensis | 1 | 1 | 0
|  |  | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-04-14 | Kocuria palustris | 0.52 | 1 | 0
|  |  | DSM 20030 | 3; type strain | #CON=3-5 loops in 100 mkL#TRT=in 80 pct. TFA¦ 30 Min treatment¦1:10 diluted¦ 2:2 HCCA¦TA2 (A) | 2011-08-25 | Micrococcus luteus | 0.53 | 1 | 0
|  |  | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-17 | Micrococcus luteus | 0.72 | 1 | 0
|  |  | DSM 20032 CVUAS | 3; type strain | eDT | 2015-07-09 | Rothia kristinae | 0.99 | 1 | 0
|  |  | IMET 10432 DSM 20127 | 3; type strain | unknown | 2011-09-08 | Sinomonas atrocyanea | 0.62 | 1 | 0
|  |  | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium cellulans | 0.73 | 1 | 0
|  |  | CVUAS 749.2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter inulinus | 1.05 | 1 | 0
|  |  | DSM 108415 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=100 pct. TFA, treatment 30 min, 1:10 diluted (water), spotted 1:3 with matrix matrix (12g¦L HCCA in 70 pct. ACN, 3 pct. TFA, 17 pct. H2O v¦v)#EXT=Cutibacterium acnes DSM 108415: Preparation & measurement: N. Fischer#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-23 | Cutibacterium acnes | 0.74 | 1 | 0
|  |  | DSM 1079 | 3; public strain collection | EFEx | 2017-01-20 | Penicillium roqueforti | 0.65 | 1 | 0
|  |  | CVUAS 10354.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton verrucosum | 0.89 | 1 | 0
|  |  | CVUAS 10357.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton verrucosum | 0.67 | 1 | 0
|  |  | DSM 1386 | 3; public strain collection | Fextr | 2016-03-10 | Candida albicans | 0.88 | 1 | 0
|  |  | CVUAS 31652 | 3; standard identification | EFEx | 2019-03-25 | Saccharomyces cerevisiae | 0.79 | 1 | 0
|  |  | IGV 011 | 3; sequenced | EFEx | 2016-12-08 | Sarocladium strictum | 0.86 | 1 | 0
|  |  | CVUAS 2520.4 | 3; sequenced (ITS) | F-Extr | 2018-07-26 | Fusarium sambucinum | 0.06 | 1 | 0
|  |  | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.65 | 1 | 0
|  |  | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.66 | 1 | 0
|  |  | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia zopfii | 1.18 | 1 | 0
|  |  | DSM 100506 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 03.11.2021¦ 1:1 HCCA | 2021-11-03 | Lysinibacillus endophyticus | 0.69 | 1 | 0
|  |  | B 092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus fusiformis | 0.99 | 1 | 0
|  |  | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus fusiformis | 0.73 | 1 | 0
|  |  | CVUAS 7192.2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus macroides | 0.77 | 1 | 0
|  |  | DSM 54 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-03 | Lysinibacillus macroides | 1.22 | 1 | 0
|  |  | DSM 24784 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-15 | Lysinibacillus pakistanensis | 0.64 | 1 | 0
|  |  | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus sphaericus | 0.68 | 1 | 0
|  |  | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus sphaericus | 0.95 | 1 | 0
|  |  | DSM 23493 | 3; type strain | approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted | 2021-11-15 | Lysinibacillus xylanilyticus | 0.71 | 1 | 0
|  |  | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus saudimassiliensis | 0.76 | 1 | 0
|  |  | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus odysseyi | 0.91 | 1 | 0
|  |  | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium acetylicum | 0.88 | 1 | 0
|  |  | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium enclense | 0.55 | 1 | 0
|  |  | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium profundum | 0.73 | 1 | 0
|  |  | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix thermosphacta | 0.73 | 1 | 0
|  |  | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix thermosphacta | 0.48 | 1 | 0
|  |  | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria aquatica | 1.13 | 1 | 0
|  |  | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria aquatica | 0.9 | 1 | 0
|  |  | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria fleischmannii | 1 | 1 | 0
|  |  | LU2006-3 | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria fleischmannii | 0.6 | 1 | 0
|  |  | CIP 105447 | 3; type strain | EFEx | 2017-02-15 | Listeria grayi | 0.54 | 1 | 0
|  |  | DSM 20596 | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria grayi | 0.68 | 1 | 0
|  |  | CVUAS 10132.2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria innocua | 0.64 | 1 | 0
|  |  | CVUAS 8569 | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria innocua | 0.87 | 1 | 0
|  |  | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria ivanovii | 0.91 | 1 | 0
|  |  | CVUAS 32100 | 3; sequenced | EFex | 2019-12-13 | Listeria ivanovii | 0.86 | 1 | 0
|  |  | DSM 23813 | 3; type strain | EFEx | 2017-02-15 | Listeria marthii | 0.7 | 1 | 0
|  |  | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.55 | 1 | 0
|  |  | CVUAS 873.2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.68 | 1 | 0
|  |  | CVUAS 240.2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.6 | 1 | 0
|  |  | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.68 | 1 | 0
|  |  | CLIP 70676 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.9 | 1 | 0
|  |  | SLCC 2479 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 1.13 | 1 | 0
|  |  | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.45 | 1 | 0
|  |  | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.64 | 1 | 0
|  |  | CVUAS 7366.2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 1.06 | 1 | 0
|  |  | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria newyorkensis | 0.61 | 1 | 0
|  |  | CVUAS 36192 | 3; sequenced | EFex | 2024-03-06 | Listeria newyorkensis | 0.97 | 1 | 0
|  |  | DSM 22097 | 3; type strain | EFEx | 2017-02-15 | Listeria rocourtiae | 0.94 | 1 | 0
|  |  | CVUAS 2763 | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria seeligeri | 1.07 | 1 | 0
|  |  | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria seeligeri | 0.9 | 1 | 0
|  |  | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria sp | 0.57 | 1 | 0
|  |  | CVUAS 2412.3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria welshimeri | 0.46 | 1 | 0
|  |  | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria welshimeri | 0.74 | 1 | 0
|  |  | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus laterosporus | 0.97 | 1 | 0
|  |  | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus laterosporus | 0.91 | 1 | 0
|  |  | HD1.4A | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-08 | Brevibacillus parabrevis | 1.2 | 1 | 0
|  |  | HD3.3A | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2020-01-09 | Brevibacillus parabrevis | 1.01 | 1 | 0
|  |  | HB1.1 | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Brevibacillus porteri | 0.31 | 1 | 0
|  |  | HB1.2 | 3; sequenced | #CON=1 loop in 50 mkL#TRT=in 80 pct. TFA, 30 Min, 1:10 diluted, spotted 1:1 with matrix (12 g¦L HCCA in TA2)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6 ¦ Dr. Thanh Tam (IPPR, Hanoi, Vietnam)# | 2019-11-27 | Brevibacillus porteri | 0.7 | 1 | 0
|  |  | B 210 | 3; eudo | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-01 | Paenibacillus alvei | 0.95 | 1 | 0
|  |  | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus amylolyticus | 1 | 1 | 0
|  |  | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus amylolyticus | 0.82 | 1 | 0
|  |  | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus barengoltzii | 1.01 | 1 | 0
|  |  | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus barengoltzii | 1.18 | 1 | 0
|  |  | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus chibensis | 0.98 | 1 | 0
|  |  | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus glucanolyticus | 0.67 | 1 | 0
|  |  | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus glucanolyticus | 0.52 | 1 | 0
|  |  | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus larvae | 0.85 | 1 | 0
|  |  | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus larvae | 1.03 | 1 | 0
|  |  | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 1.03 | 1 | 0
|  |  | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 0.75 | 1 | 0
|  |  | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus macerans | 0.83 | 1 | 0
|  |  | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus nuruki | 0.62 | 1 | 0
|  |  | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus pabuli | 0.91 | 1 | 0
|  |  | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus pabuli | 0.83 | 1 | 0
|  |  | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus pasadenensis | 0.99 | 1 | 0
|  |  | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-09-12 | Paenibacillus polymyxa | 0.62 | 1 | 0
|  |  | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-11-24 | Paenibacillus polymyxa | 0.37 | 1 | 0
|  |  | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus sp-CVUAS-6369 | 0.89 | 1 | 0
|  |  | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus sp-E8a | 0.86 | 1 | 0
|  |  | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus taichungensis | 0.47 | 1 | 0
|  |  | FLI 003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus taichungensis | 0.74 | 1 | 0
|  |  | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus tundrae | 0.76 | 1 | 0
|  |  | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus validus | 0.64 | 1 | 0
|  |  | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus sp-185-U1-2 | 1.01 | 1 | 0
|  |  | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus suis | 0.83 | 1 | 0
|  |  | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus urinaeequi | 0.48 | 1 | 0
|  |  | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 2016-09-27 | Aerococcus viridans | 0.63 | 1 | 0
|  |  | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter phocae | 0.42 | 1 | 0
|  |  | CVUAS 2442.2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium divergens | 0.95 | 1 | 0
|  |  | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium divergens | 1.28 | 1 | 0
|  |  | CVUAS 6273 | 3; sequenced | eDT | 2016-01-13 | Carnobacterium maltaromaticum | 0.71 | 1 | 0
|  |  | CVUAS 1519.3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella adiacens | 0.76 | 1 | 0
|  |  | CVUAS 2616 | 3; sequenced | DT | 2013-09-25 | Marinilactibacillus sp-CVUAS-2616 | 0.91 | 1 | 0
|  |  | CVUAS 8980.3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus sp-CVUAS-2616 | 1.13 | 1 | 0
|  |  | 16 Ec | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus casseliflavus | 0.37 | 1 | 0
|  |  | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus casseliflavus | 1.05 | 1 | 0
|  |  | CVUAS 5030 | 3; sequenced | eDT | 2014-07-23 | Enterococcus cecorum | 0.77 | 1 | 0
|  |  | CVUAS 8897 | 3; sequenced | eDT | 2014-07-23 | Enterococcus cecorum | 0.99 | 1 | 0
|  |  | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus durans | 0.52 | 1 | 0
|  |  | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-16 | Enterococcus durans | 1.03 | 1 | 0
|  |  | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus faecalis | 1.04 | 1 | 0
|  |  | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus faecalis | 0.72 | 1 | 0
|  |  | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2011-07-11 | Enterococcus faecium | 0.87 | 1 | 0
|  |  | MR11-207 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus faecium | 0.59 | 1 | 0
|  |  | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-23 | Enterococcus gallinarum | 0.92 | 1 | 0
|  |  | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus gallinarum | 0.76 | 1 | 0
|  |  | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 2011-09-23 | Enterococcus hirae | 0.89 | 1 | 0
|  |  | CVUAS 34954 | 3; sequenced | eDT | 2023-11-30 | Enterococcus hirae | 0.93 | 1 | 0
|  |  | CVUAS 34219 | 3; sequenced | eDT | 2023-11-30 | Enterococcus mundtii | 0.99 | 1 | 0
|  |  | CVUAS 11490 | 3; sequenced | eDT | 2023-11-30 | Enterococcus phoeniculicola | 0.54 | 1 | 0
|  |  | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus solitarius | 0.98 | 1 | 0
|  |  | CVUAS 1549.2 | 3; sequenced | DT | 2016-09-27 | Vagococcus salmoninarum | 0.98 | 1 | 0
|  |  | MB 2023000188 | 2; corresponds to expected result of PT | EFEx | 2023-09-07 | Lacticaseibacillus rhamnosus | 0.5 | 1 | 0
|  |  | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus acidophilus | 1.22 | 1 | 0
|  |  | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus johnsonii | 0.75 | 1 | 0
|  |  | CVUAS 32277.2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus sp-CVUAS-32277,2 | 0.73 | 1 | 0
|  |  | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus claussenii | 0.5 | 1 | 0
|  |  | LTH 1213 | 2; 0 | DT | 2013-07-09 | Leuconostoc mesenteroides | 0.56 | 1 | 0
|  |  | LTH 1146 | 2; 0 | DT | 2013-07-09 | Lactococcus lactis | 0.71 | 1 | 0
|  |  | DSM 20622 | 3; type strain | EFEx | 2023-06-20 | Streptococcus acidominimus | 0.96 | 1 | 0
|  |  | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus agalactiae | 0.52 | 1 | 0
|  |  | DSM 6784 CVUAS | 3; public strain collection | eDT | 2019-07-09 | Streptococcus agalactiae | 0.72 | 1 | 0
|  |  | DSM 20728 | 3; type strain | EFEx | 2023-09-05 | Streptococcus alactolyticus | 0.54 | 1 | 0
|  |  | DSM 20563 | 3; type strain | EFEx | 2023-08-29 | Streptococcus anginosus | 0.62 | 1 | 0
|  |  | DSM 25818 | 3; type strain | EFEx | 2023-09-07 | Streptococcus anginosus | 0.61 | 1 | 0
|  |  | DSM 15627 | 3; type strain | EFEx | 2023-08-29 | Streptococcus australis | 0.48 | 1 | 0
|  |  | DSM 103678 | 3; type strain | EFEx | 2023-09-07 | Streptococcus azizii | 0.86 | 1 | 0
|  |  | CCUG 69277 | 3; type strain | EFEx | 2023-06-20 | Streptococcus bovimastitidis | 0.87 | 1 | 0
|  |  | DSM 19004 | 3; type strain | EFEx | 2023-09-07 | Streptococcus caballi | 0.26 | 1 | 0
|  |  | CCUG 73951 | 3; type strain | EFEx | 2023-06-20 | Streptococcus caledonicus | 1.01 | 1 | 0
|  |  | DSM 105751 | 3; type strain | EFEx | 2023-09-07 | Streptococcus cameli | 0.49 | 1 | 0
|  |  | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus canis | 0.83 | 1 | 0
|  |  | LHL Mi-845-17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus canis | 0.41 | 1 | 0
|  |  | CCUG 67170 | 3; type strain | EFEx | 2023-06-20 | Streptococcus caprae | 0.71 | 1 | 0
|  |  | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus castoreus | 0.47 | 1 | 0
|  |  | DSM 17536 CVUAS | 3; type strain | EFEx | 2018-05-03 | Streptococcus castoreus | 0.32 | 1 | 0
|  |  | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus catagoni | 1.2 | 1 | 0
|  |  | 99-1-2017 | 3; type strain | EFEx | 2018-05-03 | Streptococcus catagoni | 1.2 | 1 | 0
|  |  | DSM 102819 | 3; type strain | EFEx | 2023-09-07 | Streptococcus caviae | 0.97 | 1 | 0
|  |  | DSM 106182 | 3; type strain | EFEx | 2023-09-07 | Streptococcus chenjunshii | 0.79 | 1 | 0
|  |  | LMG 14507 | 3; type strain | EFEx | 2023-09-07 | Streptococcus constellatus | 0.26 | 1 | 0
|  |  | DSM 17475 | 3; type strain | EFEx | 2023-09-06 | Streptococcus constellatus | 0.68 | 1 | 0
|  |  | DSM 25819 | 3; type strain | EFEx | 2023-09-07 | Streptococcus constellatus | 0.53 | 1 | 0
|  |  | DSM 20562 | 3; type strain | EFEx | 2023-08-29 | Streptococcus criceti | 0.43 | 1 | 0
|  |  | DSM 8249 | 3; type strain | EFEx | 2023-08-29 | Streptococcus cristatus | 0.78 | 1 | 0
|  |  | CCUG 65085 | 3; type strain | EFEx | 2023-06-20 | Streptococcus cuniculi | 0.82 | 1 | 0
|  |  | CCUG 66496 | 3; type strain | EFEx | 2023-06-20 | Streptococcus cuniculipharyngis | 1.09 | 1 | 0
|  |  | DSM 22233 | 3; type strain | EFEx | 2023-09-07 | Streptococcus danieliae | 0.69 | 1 | 0
|  |  | DSM 21999 | 3; type strain | EFEx | 2023-09-07 | Streptococcus dentapri | 0.63 | 1 | 0
|  |  | DSM 27381 | 3; type strain | EFEx | 2023-09-07 | Streptococcus dentiloxodontae | 0.84 | 1 | 0
|  |  | DSM 18963 | 3; type strain | EFEx | 2023-06-20 | Streptococcus dentirousetti | 0.6 | 1 | 0
|  |  | DSM 25137 | 3; type strain | EFEx | 2023-08-29 | Streptococcus dentisani | 0.77 | 1 | 0
|  |  | CCUG 47155 | 3; type strain | EFEx | 2023-09-06 | Streptococcus devriesei | 0.67 | 1 | 0
|  |  | DSM 15616 | 3; type strain | EFEx | 2023-09-06 | Streptococcus didelphis | 0.57 | 1 | 0
|  |  | DSM 5635 | 3; type strain | EFEx | 2023-08-29 | Streptococcus downei | 0.47 | 1 | 0
|  |  | CCUG 73139 | 3; type strain | EFEx | 2023-06-20 | Streptococcus downii | 0.52 | 1 | 0
|  |  | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus dysgalactiae | 0.27 | 1 | 0
|  |  | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0.71 | 1 | 0
|  |  | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus dysgalactiae | 0.57 | 1 | 0
|  |  | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0.75 | 1 | 0
|  |  | DSM 14446 | 3; type strain | EFEx | 2023-09-06 | Streptococcus entericus | 0.71 | 1 | 0
|  |  | CVUAS 32462.2 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus equi | 0.84 | 1 | 0
|  |  | CVUAS 32770 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus equi | 0.79 | 1 | 0
|  |  | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus equi | 0.42 | 1 | 0
|  |  | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus equi | 0.44 | 1 | 0
|  |  | DSM 20558 RWTH | 3; type strain | EFEx | 2023-09-05 | Streptococcus equinus | 0.2 | 1 | 0
|  |  | DSM 20646 | 3; type strain | EFEx | 2023-09-07 | Streptococcus ferus | 0.67 | 1 | 0
|  |  | DSM 15349 | 3; type strain | EFEx | 2023-08-30 | Streptococcus gallinaceus | 0.55 | 1 | 0
|  |  | DSM 16831 | 3; type strain | EFEx | 2023-09-07 | Streptococcus gallolyticus | 0.69 | 1 | 0
|  |  | DSM 15879 | 3; type strain | EFEx | 2023-09-05 | Streptococcus gallolyticus | 0.5 | 1 | 0
|  |  | DSM 15351 RWTH | 3; type strain | EFEx | 2023-09-06 | Streptococcus gallolyticus | 0.25 | 1 | 0
|  |  | DSM 6777 | 3; type strain | EFEx | 2023-08-29 | Streptococcus gordonii | 0.33 | 1 | 0
|  |  | DSM 17028 | 3; type strain | EFEx | 2023-09-05 | Streptococcus halichoeri | 0.76 | 1 | 0
|  |  | DSM 101996 | 3; type strain | EFEx | 2023-09-07 | Streptococcus halotolerans | 0.28 | 1 | 0
|  |  | DSM 19005 | 3; type strain | EFEx | 2023-09-07 | Streptococcus henryi | 0.58 | 1 | 0
|  |  | DSM 107951 | 3; type strain | EFEx | 2023-06-20 | Streptococcus hillyeri | 0.45 | 1 | 0
|  |  | DSM 101997 | 3; type strain | EFEx | 2023-09-07 | Streptococcus himalayensis | 1 | 1 | 0
|  |  | DSM 26014 | 3; type strain | EFEx | 2023-06-20 | Streptococcus hongkongensis | 0.43 | 1 | 0
|  |  | DSM 20770 | 3; type strain | EFEx | 2023-09-05 | Streptococcus hyointestinalis | 0.68 | 1 | 0
|  |  | DSM 12219 | 3; type strain | EFEx | 2023-09-05 | Streptococcus hyovaginalis | 0.75 | 1 | 0
|  |  | CCUG 52536 | 3; type strain | EFEx | 2023-09-07 | Streptococcus ictaluri | 0.51 | 1 | 0
|  |  | DSM 22960 | 3; type strain | EFEx | 2023-08-30 | Streptococcus infantarius | 0.31 | 1 | 0
|  |  | CCUG 39817 | 3; type strain | EFEx | 2023-09-05 | Streptococcus infantis | 0.84 | 1 | 0
|  |  | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus iniae | 0.55 | 1 | 0
|  |  | DSM 20576 | 3; type strain | EFEx | 2023-08-29 | Streptococcus iniae | 0.63 | 1 | 0
|  |  | DSM 20573 | 3; type strain | EFEx | 2023-09-06 | Streptococcus intermedius | 0.57 | 1 | 0
|  |  | DSM 23027 | 3; type strain | EFEx | 2023-09-07 | Streptococcus lactarius | 0.93 | 1 | 0
|  |  | DSM 27382 | 3; type strain | EFEx | 2023-06-20 | Streptococcus loxodontisalivarius | 0.65 | 1 | 0
|  |  | DSM 15350 | 3; type strain | EFEx | 2023-08-29 | Streptococcus lutetiensis | 0.65 | 1 | 0
|  |  | DSM 20724 | 3; type strain | EFEx | 2023-08-29 | Streptococcus macacae | 0.43 | 1 | 0
|  |  | DSM 18627 | 3; type strain | EFEx | 2023-09-07 | Streptococcus marimammalium | 1.24 | 1 | 0
|  |  | DSM 101995 | 3; type strain | EFEx | 2023-06-20 | Streptococcus marmotae | 0.82 | 1 | 0
|  |  | CCUG 49690 | 3; type strain | EFEx | 2023-09-07 | Streptococcus massiliensis | 0.49 | 1 | 0
|  |  | DSM 19192 | 3; type strain | EFEx | 2023-09-07 | Streptococcus merionis | 0.75 | 1 | 0
|  |  | DSM 17118 | 3; type strain | EFEx | 2023-09-06 | Streptococcus minor | 0.57 | 1 | 0
|  |  | DSM 12643 RWTH | 3; type strain | EFEx | 2023-09-05 | Streptococcus mitis | 0.89 | 1 | 0
|  |  | DSM 12643 RKI | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 11.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus agalactiae DSM 2134, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-02-18 | Streptococcus mitis | 0.7 | 1 | 0
|  |  | DSM 105143 | 3; type strain | EFEx | 2023-06-20 | Streptococcus moroccensis | 0.65 | 1 | 0
|  |  | DSM 20523 | 3; type strain | EFEx | 2023-08-29 | Streptococcus mutans | 0.57 | 1 | 0
|  |  | DSM 27088 | 3; type strain | EFEx | 2023-06-20 | Streptococcus oralis | 0.52 | 1 | 0
|  |  | DSM 20627 | 3; type strain | EFEx | 2023-08-30 | Streptococcus oralis | 0.53 | 1 | 0
|  |  | DSM 24864 | 3; type strain | EFEx | 2023-06-20 | Streptococcus oralis | 0.56 | 1 | 0
|  |  | DSM 100101 | 3; type strain | EFEx | 2023-06-20 | Streptococcus oricebi | 1.07 | 1 | 0
|  |  | DSM 27377 | 3; type strain | EFEx | 2023-06-20 | Streptococcus oriloxodontae | 0.47 | 1 | 0
|  |  | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus orisasini | 0.5 | 1 | 0
|  |  | DSM 25193 | 3; type strain | EFEx | 2023-09-07 | Streptococcus orisasini | 0.73 | 1 | 0
|  |  | CCUG 43577 | 3; type strain | EFEx | 2023-08-30 | Streptococcus orisratti | 0.61 | 1 | 0
|  |  | DSM 18307 | 3; type strain | EFEx | 2023-09-07 | Streptococcus orisuis | 0.58 | 1 | 0
|  |  | CCUG 39485 | 3; type strain | EFEx | 2023-08-29 | Streptococcus ovis | 1.3 | 1 | 0
|  |  | CCUG 69612 | 3; type strain | EFEx | 2023-06-20 | Streptococcus ovuberis | 0.74 | 1 | 0
|  |  | DSM 111148 | 3; type strain | EFEx | 2023-06-20 | Streptococcus pacificus | 0.6 | 1 | 0
|  |  | DSM 29921 | 3; type strain | EFEx | 2023-09-07 | Streptococcus panodentis | 0.83 | 1 | 0
|  |  | DSM 102135 | 3; type strain | EFEx | 2023-09-07 | Streptococcus pantholopis | 0.61 | 1 | 0
|  |  | B 9056 | 2; strain collection of the RKI | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 11.2.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus agalactiae DSM 2134, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-09-01 | Streptococcus parasanguinis | 0 | 1 | 0
|  |  | DSM 6778 | 3; type strain | EFEx | 2023-08-30 | Streptococcus parasanguinis | 0.48 | 1 | 0
|  |  | DSM 29126 | 3; type strain | EFEx | 2023-09-07 | Streptococcus parasuis | 1.04 | 1 | 0
|  |  | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus parauberis | 0.68 | 1 | 0
|  |  | DSM 6631 | 3; type strain | EFEx | 2023-09-05 | Streptococcus parauberis | 0.92 | 1 | 0
|  |  | DSM 26545 | 3; type strain | EFEx | 2023-09-07 | Streptococcus penaeicida | 0.75 | 1 | 0
|  |  | CCUG 39814 | 3; type strain | EFEx | 2023-09-05 | Streptococcus peroris | 0.89 | 1 | 0
|  |  | LMG 16735 | 3; type strain | EFEx | 2023-09-05 | Streptococcus phocae | 0.76 | 1 | 0
|  |  | DSM 24768 | 3; type strain | EFEx | 2023-09-07 | Streptococcus phocae | 0.81 | 1 | 0
|  |  | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus phocae | 0.52 | 1 | 0
|  |  | 12-13862-001-01 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus pluranimalium | 0.67 | 1 | 0
|  |  | LMG 14177 | 3; type strain | EFEx | 2023-09-05 | Streptococcus pluranimalium | 0.73 | 1 | 0
|  |  | DSM 22810 | 3; type strain | EFEx | 2023-09-07 | Streptococcus plurextorum | 1.02 | 1 | 0
|  |  | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 1.07 | 1 | 0
|  |  | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.55 | 1 | 0
|  |  | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.71 | 1 | 0
|  |  | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus pneumoniae | 0.55 | 1 | 0
|  |  | DSM 23759 | 3; type strain | EFEx | 2023-06-20 | Streptococcus porci | 1.02 | 1 | 0
|  |  | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus porcinus | 0.47 | 1 | 0
|  |  | DSM 20725 | 3; type strain | EFEx | 2023-09-05 | Streptococcus porcinus | 0.48 | 1 | 0
|  |  | DSM 28302 | 3; type strain | EFEx | 2023-09-07 | Streptococcus porcorum | 0.8 | 1 | 0
|  |  | CCUG 49455 | 3; type strain | EFEx | 2023-08-30 | Streptococcus pseudopneumoniae | 0.88 | 1 | 0
|  |  | DSM 18513 | 3; type strain | EFEx | 2023-09-07 | Streptococcus pseudoporcinus | 1.04 | 1 | 0
|  |  | B 1567 | 2; strain collection of the RKI | #CON=approx. 3 loops per Eppend. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, Inactivation date: 19.08.2021¦ 1:1 HCCA¦ Andy 3.0#EXT=Streptococcus pneumoniae B2046, preparation: C. Hartl, measurements: C. Hartl#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2021-09-02 | Streptococcus pyogenes | 0.85 | 1 | 0
|  |  | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-04 | Streptococcus pyogenes | 0.21 | 1 | 0
|  |  | DSM 20564 | 3; type strain | EFEx | 2023-08-29 | Streptococcus ratti | 0.46 | 1 | 0
|  |  | DSM 101998 | 3; type strain | EFEx | 2023-09-07 | Streptococcus respiraculi | 1.05 | 1 | 0
|  |  | DSM 105174 | 3; type strain | EFEx | 2023-09-07 | Streptococcus rifensis | 0.89 | 1 | 0
|  |  | DSM 26920 | 3; type strain | EFEx | 2023-08-30 | Streptococcus rubneri | 0.6 | 1 | 0
|  |  | DSM 104980 | 3; type strain | EFEx | 2023-09-07 | Streptococcus ruminantium | 1 | 1 | 0
|  |  | DSM 28303 | 3; type strain | EFEx | 2023-09-07 | Streptococcus rupicaprae | 0.59 | 1 | 0
|  |  | DSM 20560 | 3; type strain | EFEx | 2023-09-05 | Streptococcus salivarius | 0.8 | 1 | 0
|  |  | DSM 20617 | 3; type strain | EFEx | 2023-08-29 | Streptococcus salivarius | 0.71 | 1 | 0
|  |  | CVUAS 10145.3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus salivarius | 1.06 | 1 | 0
|  |  | DSM 27513 | 3; type strain | EFEx | 2023-09-07 | Streptococcus saliviloxodontae | 0.1 | 1 | 0
|  |  | DSM 20068 | 3; public strain collection | TFAextr | 2022-02-18 | Streptococcus sanguinis | 0.74 | 1 | 0
|  |  | DSM 20567 RWTH | 3; type strain | EFEx | 2023-08-30 | Streptococcus sanguinis | 0.69 | 1 | 0
|  |  | DSM 14990 | 3; type strain | EFEx | 2023-08-30 | Streptococcus sinensis | 0.6 | 1 | 0
|  |  | ATCC 33478 | 3; type strain | EFEx | 2023-08-29 | Streptococcus sobrinus | 0.64 | 1 | 0
|  |  | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus sp-CVUAS-32349 | 1.03 | 1 | 0
|  |  | DSM 9682 | 3; type strain | EFEx | 2023-09-05 | Streptococcus suis | 0.72 | 1 | 0
|  |  | DSM 105173 | 3; type strain | EFEx | 2023-09-07 | Streptococcus tangierensis | 0.65 | 1 | 0
|  |  | CCUG 32906 | 3; type strain | EFEx | 2023-09-05 | Streptococcus thoraltensis | 0.78 | 1 | 0
|  |  | DSM 25324 | 3; type strain | EFEx | 2023-09-07 | Streptococcus troglodytae | 0.49 | 1 | 0
|  |  | CVUAS 2260.4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus troglodytidis | 0.77 | 1 | 0
|  |  | KCTC 33006 | 3; type strain | EFEx | 2023-06-20 | Streptococcus troglodytidis | 0.89 | 1 | 0
|  |  | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0.49 | 1 | 0
|  |  | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0.35 | 1 | 0
|  |  | CCUG 41590 | 3; type strain | EFEx | 2023-09-06 | Streptococcus urinalis | 1.13 | 1 | 0
|  |  | DSM 22768 | 3; type strain | EFEx | 2023-09-07 | Streptococcus ursoris | 0.35 | 1 | 0
|  |  | DSM 5636 | 3; type strain | EFEx | 2023-08-29 | Streptococcus vestibularis | 1.13 | 1 | 0
|  |  | DSM 110741 | 3; type strain; WGS (PMID: 33999792) | EFEx | 2023-06-20 | Streptococcus vicugnae | 0.39 | 1 | 0
|  |  | DSM 110742 | 3; type strain, WGS (PMID: 33999792) | EFEx | 2023-06-20 | Streptococcus zalophi | 0.86 | 1 | 0
|  |  | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 1.14 | 1 | 0
|  |  | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.57 | 1 | 0
|  |  | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.83 | 1 | 0
|  |  | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.45 | 1 | 0
|  |  | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.87 | 1 | 0
|  |  | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.35 | 1 | 0
|  |  | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.56 | 1 | 0
|  |  | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.77 | 1 | 0
|  |  | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium barati | 0.62 | 1 | 0
|  |  | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium botulinum | 0.97 | 1 | 0
|  |  | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium butyricum | 0.78 | 1 | 0
|  |  | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium chauvoei | 0.74 | 1 | 0
|  |  | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium chauvoei | 0.65 | 1 | 0
|  |  | 171002729 | 2; 0 | DT | 2017-02-28 | Clostridium cochlearium | 0.61 | 1 | 0
|  |  | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium colicanis | 0.46 | 1 | 0
|  |  | CVUAS 849.2 | 3; sequenced | DT | 2014-12-17 | Clostridium isatidis | 0.47 | 1 | 0
|  |  | DSM 108256 | 3; Public strain collection | EFEx | 0 | Clostridium leptum | 0.69 | 1 | 0
|  |  | MSK.23.22 | 3; sequenced [0207] | EFEx | 0 | Clostridium nexile | 0.9 | 1 | 0
|  |  | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium nigeriense | 0.46 | 1 | 0
|  |  | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium novyi | 0.62 | 1 | 0
|  |  | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium perfringens | 0.53 | 1 | 0
|  |  | B | 3; case rerort | DT | 2013-08-06 | Clostridium perfringens | 0.45 | 1 | 0
|  |  | JCM 10419 | 3; Public strain collection | eDT | 2021-10-26 | Clostridium scindens | 0.62 | 1 | 0
|  |  | MSK.1.26 | 3; sequenced [0207] | eDT | 0 | Clostridium scindens | 0.54 | 1 | 0
|  |  | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium septicum | 0.68 | 1 | 0
|  |  | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium septicum | 1.02 | 1 | 0
|  |  | CVUAS 34091 | 3; sequenced | EFex | 2022-08-05 | Clostridium sp-CVUAS-34091 | 0.6 | 1 | 0
|  |  | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium sp-DSM-1985 | 0.69 | 1 | 0
|  |  | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium sporogenes | 0.75 | 1 | 0
|  |  | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium sporogenes | 0.77 | 1 | 0
|  |  | MSK.1.9 | 3; sequenced [0207] | eDT | 0 | Clostridium symbiosum | 0.88 | 1 | 0
|  |  | WAL-14163 | 3; sequenced [0207] | eDT | 0 | Clostridium symbiosum | 0.78 | 1 | 0
|  |  | 141015223 | 2; 0 | DT | 2014-11-27 | Clostridium tertium | 0.57 | 1 | 0
|  |  | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium tertium | 0.81 | 1 | 0
|  |  | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium tetani | 0.63 | 1 | 0
|  |  | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium tetani | 0.53 | 1 | 0
|  |  | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya limosa | 0.78 | 1 | 0
|  |  | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium sordellii | 0.9 | 1 | 0
|  |  | DSM 16296 | 3; Public strain collection | EFEx | 0 | Eubacterium ramulus | 0.7 | 1 | 0
|  |  | CLA-AA-H258 | 3; sequenced [0207] | eDT | 0 | Agathobacter rectalis | 0.39 | 1 | 0
|  |  | MSK.9.13 | 3; sequenced [0207] | eDT | 0 | Agathobacter rectalis | 0.76 | 1 | 0
|  |  | Clinical isolate | 3; sequenced [0207] | EFEx | 0 | Anaerobutyricum hallii | 0.86 | 1 | 0
|  |  | CLA-AA-H53 | 3; sequenced [0207] | EFEx | 0 | Anaerostipes hadrus | 0.7 | 1 | 0
|  |  | DSM 29492 | 3; type strain | EFEx | 0 | Blautia caecimuris | 0.83 | 1 | 0
|  |  | MSK.18.51 | 3; sequenced [0207] | EFEx | 0 | Blautia faecis | 0.68 | 1 | 0
|  |  | DSM 22028 | 3; type strain | EFEx | 0 | Blautia glucerasea | 1.15 | 1 | 0
|  |  | DSM 20583 | 3; type strain | EFEx | 0 | Blautia hansenii | 1.08 | 1 | 0
|  |  | MSK.19.60 | 3; sequenced [0207] | eDT | 0 | Blautia massiliensis | 0.74 | 1 | 0
|  |  | MSK.9.19 | 3; sequenced [0207] | eDT | 0 | Blautia massiliensis | 0.85 | 1 | 0
|  |  | MSK.20.66 | 3; sequenced [0207] | eDT | 0 | Blautia obeum | 0.78 | 1 | 0
|  |  | CCUG 9990 | 3; type strain | eDT | 0 | Blautia producta | 0.49 | 1 | 0
|  |  | MSK.4.13 | 2; data not listed [0207] | eDT | 0 | Blautia producta | 0.95 | 1 | 0
|  |  | DSM 26115 | 3; public strain collection | EFEx | 0 | Blautia pseudococcoides | 0.82 | 1 | 0
|  |  | MSK.6.27 | 3; sequenced [0207] | EFEx | 2021-10-20 | Blautia schinkii | 0.89 | 1 | 0
|  |  | MSK.6.32 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia schinkii | 0.71 | 1 | 0
|  |  | MSK.18.16 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia wexlerae | 0.92 | 1 | 0
|  |  | JCM 31264 | 3; Public strain collection | eDT | 2021-10-20 | Coprococcus comes | 1.17 | 1 | 0
|  |  | MSK.17.21 | 3; sequenced [0207] | eDT | 2021-10-20 | Coprococcus comes | 0.99 | 1 | 0
|  |  | MSK.10.6 | 3; sequenced [0207] | EFEx | 2021-10-20 | Coprococcus eutactus | 0.95 | 1 | 0
|  |  | JCM 31256 | 3; Public strain collection | eDT | 0 | Dorea formicigenerans | 0.78 | 1 | 0
|  |  | MSK.19.39 | 3; sequenced [0207] | eDT | 0 | Dorea formicigenerans | 0.67 | 1 | 0
|  |  | MSK.18.49 | 3; sequenced [0207] | eDT | 0 | Dorea longicatena | 1.25 | 1 | 0
|  |  | DSM 108191 | 3; Public strain collection | EFEx | 0 | Enterocloster clostridioformis | 0.8 | 1 | 0
|  |  | MSK.20.55 | 3; sequenced [0207] | eDT | 0 | Faecalicatena fissicatena | 1.15 | 1 | 0
|  |  | MSK.9.3 | 3; sequenced [0207] | eDT | 0 | Faecalicatena fissicatena | 1.01 | 1 | 0
|  |  | MSK.14.54 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter saccharivorans | 0.43 | 1 | 0
|  |  | MSK.15.44 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter saccharivorans | 0.38 | 1 | 0
|  |  | CLA-AA-H260 | 3; sequenced [0207] | EFEx | 0 | Lachnospira eligens | 0.78 | 1 | 0
|  |  | DSM 105336 | 3; Public strain collection | EFEx | 0 | Lacrimispora celerecrescens | 0.89 | 1 | 0
|  |  | CLA-AA-H164 | 3; sequenced [0207] | EFEx | 0 | Roseburia faecis | 0.73 | 1 | 0
|  |  | DSM 16839 | 3; type strain | EFEx | 0 | Roseburia hominis | 0.68 | 1 | 0
|  |  | DSM 108070 | 3; Public strain collection | EFEx | 0 | Roseburia inulinivorans | 0.92 | 1 | 0
|  |  | MSK.1.25 | 3; sequenced [0207] | eDT | 0 | Sellimonas intestinalis | 1.09 | 1 | 0
|  |  | JCM 31255 | 3; Public strain collection | eDT | 0 | Anaerotruncus colihominis | 0.65 | 1 | 0
|  |  | CLA-AA-H281 | 3; sequenced [0207] | EFEx | 0 | Faecalibacterium prausnitzii | 0.96 | 1 | 0
|  |  | CLA-AA-H271 | 3; sequenced [0207] | eDT | 0 | Ruminococcus bromii | 0.62 | 1 | 0
|  |  | MSK.22.53 | 3; sequenced [0207] | eDT | 0 | Ruminococcus gnavus | 0.62 | 1 | 0
|  |  | MSK.23.60 | 3; sequenced [0207] | eDT | 0 | Ruminococcus gnavus | 0.8 | 1 | 0
|  |  | CLA-AA-H279 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus lactaris | 0.53 | 1 | 0
|  |  | CLA-AA-H261 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus torques | 0.44 | 1 | 0
|  |  | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium bifermentans | 0.44 | 1 | 0
|  |  | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix anatis | 0.37 | 1 | 0
|  |  | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix anatis | 0.78 | 1 | 0
|  |  | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix aquatica | 0.57 | 1 | 0
|  |  | VA92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix aquatica | 0.86 | 1 | 0
|  |  | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix larvae | 0.73 | 1 | 0
|  |  | DSM 110099 | 3; type strain | EFex [BRU03] | 2021-10-18 | Erysipelothrix piscisicarius | 0.94 | 1 | 0
|  |  | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix rhusiopathiae | 0.51 | 1 | 0
|  |  | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix rhusiopathiae | 0.94 | 1 | 0
|  |  | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix tonsillarum | 0.32 | 1 | 0
|  |  | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-23 | Erysipelothrix tonsillarum | 0.69 | 1 | 0
|  |  | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus sp-CVUAS-10169 | 0.49 | 1 | 0
|  |  | CVUAS 32948 | 3; confirmed proficency test | EFEx | 2021-02-18 | Finegoldia magna | 0 | 1 | 0
|  |  | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus kunzii | 0.69 | 1 | 0
|  |  | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus kunzii | 0.73 | 1 | 0
|  |  | CVUAS 433.2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus ovis | 0.4 | 1 | 0
|  |  | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus ovis | 0.44 | 1 | 0
|  |  | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus sp-CVUAS-32321 | 0.84 | 1 | 0
|  |  | CVUAS 692.2 | 3; sequenced | DT | 2013-02-19 | Parvimonas sp-CVUAS-692,2 | 0.97 | 1 | 0
|  |  | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides finegoldii | 0.96 | 1 | 0
|  |  | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides fragilis | 0.34 | 1 | 0
|  |  | CVUAS 32947 | 3; confirmed proficency test | EFEx | 2021-02-18 | Bacteroides fragilis | 1 | 1 | 0
|  |  | 181001478 | 2; 0 | DT | 2018-02-02 | Bacteroides pyogenes | 0.79 | 1 | 0
|  |  | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella heparinolytica | 0.69 | 1 | 0
|  |  | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium chaponense | 0.85 | 1 | 0
|  |  | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium gleum | 0.3 | 1 | 0
|  |  | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium hominis | 0.75 | 1 | 0
|  |  | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium indologenes | 0.64 | 1 | 0
|  |  | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium indologenes | 0.75 | 1 | 0
|  |  | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium manosquense | 0.88 | 1 | 0
|  |  | 185-U1-6 | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium manosquense | 1.18 | 1 | 0
|  |  | CVUAS 11450.4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium sp-I-7 | 0.67 | 1 | 0
|  |  | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium aquidurense | 0.77 | 1 | 0
|  |  | CVUAS 5322.2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium branchiarum | 0.43 | 1 | 0
|  |  | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium chungangense | 0.62 | 1 | 0
|  |  | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium columnare | 0.75 | 1 | 0
|  |  | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium frigidimaris | 0.21 | 1 | 0
|  |  | CVUAS 4217.2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium frigidimaris | 0.4 | 1 | 0
|  |  | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium hibernum | 0.64 | 1 | 0
|  |  | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium piscis | 0.75 | 1 | 0
|  |  | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium plurextorum | 0.58 | 1 | 0
|  |  | CVUAS 5323.2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium psychoterrae | 0.73 | 1 | 0
|  |  | CVUAS 11550.4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium psychrophilum | 1.08 | 1 | 0
|  |  | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium psychrophilum | 0.9 | 1 | 0
|  |  | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium reichenbachii | 0.65 | 1 | 0
|  |  | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium sp | 0.75 | 1 | 0
|  |  | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium sp-C2 | 1.1 | 1 | 0
|  |  | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-10323 | 0.36 | 1 | 0
|  |  | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium sp-CVUAS-4146 | 0.58 | 1 | 0
|  |  | CVUAS 4216.2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-4216,2 | 0.52 | 1 | 0
|  |  | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-5506 | 0.17 | 1 | 0
|  |  | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium sp-HMD2009 | 0.87 | 1 | 0
|  |  | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium sp-JQMS | 0.64 | 1 | 0
|  |  | CVUAS 33726 | 3; sequenced | EFEx | 2023-11-07 | Flavobacterium sp-T157L | 0.53 | 1 | 0
|  |  | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium succinicans | 0.51 | 1 | 0
|  |  | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium tiangeerense | 0.73 | 1 | 0
|  |  | 82 | 2; confirmed proficency test | EFEx | 2021-09-10 | Bergeyella zoohelcum | 0.85 | 1 | 0
|  |  | CVUAS 2589.2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas tenax | 1.05 | 1 | 0
|  |  | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium rhinotracheale | 0.94 | 1 | 0
|  |  | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium rhinotracheale | 0.74 | 1 | 0
|  |  | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella columbina | 0.78 | 1 | 0
|  |  | CVUAS 34092 | 3; sequenced | EFex | 2022-08-05 | Weeksellaceae sp-CVUAS-34092 | 1.15 | 1 | 0
|  |  | CVUAS 9288.2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca zopfii | 0.99 | 1 | 0
|  |  | CVUAS 9875.2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca zopfii | 0.7 | 1 | 0
|  |  | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus sp-CVUAS-11266 | 0.31 | 1 | 0
|  |  | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium necrophorum | 0.68 | 1 | 0
|  |  | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium varium | 0.11 | 1 | 0
|  |  | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter abscessus | 1.02 | 1 | 0
|  |  | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga salmonicida | 0.5 | 1 | 0
|  |  | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus hongkongensis | 0.87 | 1 | 0
|  |  | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella termitidis | 0.23 | 1 | 0
|  |  | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella termitidis | 0.2 | 1 | 0
|  |  | CCUG 38322 | 3; public strain collection | EFEx | 2015-12-16 | Sneathia sanguinegens | 0.99 | 1 | 0
|  |  | CCUG 41628 | 3; type strain | EFEx | 2020-10-08 | Sneathia sanguinegens | 0.23 | 1 | 0
|  |  | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.69 | 1 | 0
|  |  | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.67 | 1 | 0
|  |  | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus felis | 0.42 | 1 | 0
|  |  | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus moniliformis | 0.99 | 1 | 0
|  |  | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus moniliformis | 0.61 | 1 | 0
|  |  | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus notomytis | 0.79 | 1 | 0
|  |  | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus ratti | 1.02 | 1 | 0
|  |  | CVUAS 2280.3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas glaciistagni | 0.56 | 1 | 0
|  |  | ATCC 35685 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-23 | Bartonella bacilliformis | 0.85 | 1 | 0
|  |  | ATCC 700132 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella grahamii | 0.99 | 1 | 0
|  |  | ATCC 49882 DSM 28221 | 3; type strain | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella henselae | 0.5 | 1 | 0
|  |  | BK8 | 2; strain collection of the RKI | #CON=3 loops in 150 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2022-05-10 | Bartonella henselae | 0.54 | 1 | 0
|  |  | CIP 103739 | 3; public strain collection | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella quintana | 0.65 | 1 | 0
|  |  | ATCC 700727 DSM 28218 | 3; type strain | #CON=3 loops in 100 mkL#TRT=100 pct. TFA, 30 Min, 1:10 diluted in water , spotted 1:1 with matrix (12 g¦L HCCA in 70 pct.ACN, 3 pct. TFA, 17 pct. water)#EXT=not specified#LAS=#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS 2¦ ZBS6# | 2021-11-12 | Bartonella vinsonii | 0.44 | 1 | 0
|  |  | BM373 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella abortus | 0.53 | 1 | 0
|  |  | BM369 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella abortus | 0.41 | 1 | 0
|  |  | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-11 | Brucella canis | 0.58 | 1 | 0
|  |  | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-03-05 | Brucella canis | 0.58 | 1 | 0
|  |  | 007 1 002 34 | 2; 0 | EFEx | 2013-06-20 | Brucella ceti | 0.55 | 1 | 0
|  |  | BM389 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella ceti | 0.62 | 1 | 0
|  |  | 10 1 001 44 | 2; 0 | EFEx | 2013-06-20 | Brucella inopinata | 0.61 | 1 | 0
|  |  | A1309 | 2; 0 | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted; inactivation date: 28.07.2021¦ 1:1 HCCA¦Andy 3.0 | 2021-07-28 | Brucella intermedia | 0.45 | 1 | 0
|  |  | BM374 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella melitensis | 0.62 | 1 | 0
|  |  | BM407 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella melitensis | 0.55 | 1 | 0
|  |  | BM408 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella melitensis | 0.67 | 1 | 0
|  |  | BM410 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella melitensis | 0.5 | 1 | 0
|  |  | CCM 4915 CVUAS | 3; type strain | DT | 2013-11-26 | Brucella microti | 0.3 | 1 | 0
|  |  | CCM 4915 RKI | 3; type strain | TFAextr | 2015-03-17 | Brucella microti | 0.57 | 1 | 0
|  |  | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-03-20 | Brucella neotomae | 0.69 | 1 | 0
|  |  | BM421 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella neotomae | 0.7 | 1 | 0
|  |  | 06 1 007 32 | 2; 0 | EFEx | 2013-06-20 | Brucella ovis | 0.55 | 1 | 0
|  |  | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-02-11 | Brucella ovis | 0.86 | 1 | 0
|  |  | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-26 | Brucella pinnipedialis | 0.76 | 1 | 0
|  |  | BM405 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella pinnipedialis | 1.06 | 1 | 0
|  |  | BM431 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella suis | 0.55 | 1 | 0
|  |  | BM427 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-08 | Brucella suis | 0.48 | 1 | 0
|  |  | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella suis | 0.47 | 1 | 0
|  |  | BM428 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-22 | Brucella suis | 0.6 | 1 | 0
|  |  | BM430 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-15 | Brucella suis | 0.5 | 1 | 0
|  |  | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2016-10-15 | Brucella vulpis | 0.35 | 1 | 0
|  |  | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum ovis | 0.62 | 1 | 0
|  |  | A-269_08 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2012-04-02 | Ochrobactrum anthropi | 0.75 | 1 | 0
|  |  | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 2016-06-03 | Ochrobactrum anthropi | 0.72 | 1 | 0
|  |  | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum intermedium | 0.08 | 1 | 0
|  |  | CVUAS 32748 | 3; sequenced | EFex | 2020-10-29 | Ochrobactrum intermedium | 0.76 | 1 | 0
|  |  | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum sp | 0.53 | 1 | 0
|  |  | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum sp | 0.41 | 1 | 0
|  |  | 171012040 | 2; 0 | DT | 2017-10-04 | Ochrobactrum sp | 0.54 | 1 | 0
|  |  | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter aganoensis | 0.6 | 1 | 0
|  |  | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter aminovorans | 0.7 | 1 | 0
|  |  | DSM 30768 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 16.09.2022¦ 1:1 HCCA¦Andy 3.1#EXT=Paracoccus yeei DSM 30768, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-10-07 | Paracoccus yeei | 0.42 | 1 | 0
|  |  | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-04-16 | Paracoccus yeei | 1.34 | 1 | 0
|  |  | DSM 103800 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 25.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Roseomonas gilardii DSM 103800, Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-04-06 | Roseomonas gilardii | 0.86 | 1 | 0
|  |  | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas mucosa | 0.78 | 1 | 0
|  |  | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 2007-01-24 | Inquilinus limosus | 0.8 | 1 | 0
|  |  | DSM 1098 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Sphingomonas paucimobilis DSM 30198, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Sphingomonas paucimobilis | 0.9 | 1 | 0
|  |  | DSM 30198 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Sphingomonas paucimobilis DSM 30198, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-09-01 | Sphingomonas paucimobilis | 0.96 | 1 | 0
|  |  | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter denitrificans | 0.71 | 1 | 0
|  |  | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter sp-DSM-30031 | 0.72 | 1 | 0
|  |  | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2008-11-18 | Achromobacter xylosoxidans | 0.5 | 1 | 0
|  |  | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-02-24 | Achromobacter xylosoxidans | 0.77 | 1 | 0
|  |  | DSM 30030 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Alcaligenes faecalis ssp. parafaecalis DSM 13975, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-29 | Alcaligenes faecalis | 0.93 | 1 | 0
|  |  | DSM 13975 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Alcaligenes faecalis subsp. Parafaecalis DSM 13975, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Alcaligenes faecalis | 0.41 | 1 | 0
|  |  | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea sp-CVUAS-31213 | 1.03 | 1 | 0
|  |  | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella avium | 0.83 | 1 | 0
|  |  | 21-2 | 2; confirmed proficency test | eDT | 2022-02-14 | Bordetella trematum | 0.59 | 1 | 0
|  |  | CVUAS 32946 | 3; confirmed proficency test | EFEx | 2021-02-18 | Kerstersia gyiorum | 0.2 | 1 | 0
|  |  | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.56 | 1 | 0
|  |  | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.53 | 1 | 0
|  |  | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga kullae | 0.66 | 1 | 0
|  |  | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella asinigenitalis | 0.7 | 1 | 0
|  |  | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella asinigenitalis | 0.87 | 1 | 0
|  |  | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella equigenitalis | 0.34 | 1 | 0
|  |  | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella equigenitalis | 0.35 | 1 | 0
|  |  | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Burkholderia ambifaria | 0.97 | 1 | 0
|  |  | DSM 16087 LMG 19182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia ambifaria | 0.57 | 1 | 0
|  |  | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia anthina | 0.75 | 1 | 0
|  |  | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-08-06 | Burkholderia anthina | 0.65 | 1 | 0
|  |  | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-18 | Burkholderia arboris | 0.57 | 1 | 0
|  |  | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia arboris | 0.76 | 1 | 0
|  |  | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia cenocepacia | 0.43 | 1 | 0
|  |  | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-07-02 | Burkholderia cenocepacia | 0.79 | 1 | 0
|  |  | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-05-04 | Burkholderia cepacia | 0.72 | 1 | 0
|  |  | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-05-09 | Burkholderia cepacia | 0.54 | 1 | 0
|  |  | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-03 | Burkholderia contaminans | 0.96 | 1 | 0
|  |  | BC 01 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-19 | Burkholderia contaminans | 0.8 | 1 | 0
|  |  | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-31 | Burkholderia diffusa | 0.67 | 1 | 0
|  |  | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia diffusa | 0.82 | 1 | 0
|  |  | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia dolosa | 0.58 | 1 | 0
|  |  | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-17 | Burkholderia dolosa | 0.81 | 1 | 0
|  |  | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-24 | Burkholderia gladioli | 0.55 | 1 | 0
|  |  | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia gladioli | 0.87 | 1 | 0
|  |  | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia glumae | 0.55 | 1 | 0
|  |  | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-05-04 | Burkholderia glumae | 0.51 | 1 | 0
|  |  | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-31 | Burkholderia lata | 0.56 | 1 | 0
|  |  | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-08-23 | Burkholderia lata | 0.58 | 1 | 0
|  |  | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-05-20 | Burkholderia latens | 0.55 | 1 | 0
|  |  | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-09-04 | Burkholderia mallei | 0.74 | 1 | 0
|  |  | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-08-24 | Burkholderia mallei | 0.82 | 1 | 0
|  |  | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-01 | Burkholderia metallica | 0.76 | 1 | 0
|  |  | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia metallica | 0.76 | 1 | 0
|  |  | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-13 | Burkholderia multivorans | 0.78 | 1 | 0
|  |  | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-11-08 | Burkholderia multivorans | 1.28 | 1 | 0
|  |  | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia oklahomensis | 0.84 | 1 | 0
|  |  | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia oklahomensis | 0.83 | 1 | 0
|  |  | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia plantarii | 0.38 | 1 | 0
|  |  | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Burkholderia plantarii | 0.44 | 1 | 0
|  |  | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-04 | Burkholderia pseudomallei | 1.01 | 1 | 0
|  |  | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-01 | Burkholderia pseudomallei | 0.89 | 1 | 0
|  |  | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-21 | Burkholderia pseudomultivorans | 0.88 | 1 | 0
|  |  | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-21 | Burkholderia pseudomultivorans | 0.43 | 1 | 0
|  |  | CAMPA 1043 (HST 249) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-04-01 | Burkholderia puraquae | 0.81 | 1 | 0
|  |  | R-50214 (BP 78) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-03-18 | Burkholderia puraquae | 0.99 | 1 | 0
|  |  | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia pyrrocinia | 0.89 | 1 | 0
|  |  | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-04-03 | Burkholderia pyrrocinia | 0.57 | 1 | 0
|  |  | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia seminalis | 0.83 | 1 | 0
|  |  | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-11 | Burkholderia seminalis | 0.99 | 1 | 0
|  |  | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-25 | Burkholderia stabilis | 0.4 | 1 | 0
|  |  | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2006-12-19 | Burkholderia stabilis | 0.6 | 1 | 0
|  |  | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-16 | Burkholderia stagnalis | 0.38 | 1 | 0
|  |  | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia stagnalis | 0.56 | 1 | 0
|  |  | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Burkholderia taxonK | 0.69 | 1 | 0
|  |  | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-06-26 | Burkholderia taxonK | 0.5 | 1 | 0
|  |  | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia territorii | 0.43 | 1 | 0
|  |  | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-16 | Burkholderia territorii | 0.81 | 1 | 0
|  |  | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-02 | Burkholderia thailandensis | 0.76 | 1 | 0
|  |  | DSM 13276 n | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2010-11-24 | Burkholderia thailandensis | 0.56 | 1 | 0
|  |  | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-12 | Burkholderia ubonensis | 1 | 1 | 0
|  |  | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2017-03-27 | Burkholderia ubonensis | 0.28 | 1 | 0
|  |  | CCUG 9631 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2007-05-09 | Burkholderia vietnamiensis | 0.62 | 1 | 0
|  |  | LMG 10929 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2006-12-13 | Burkholderia vietnamiensis | 0.59 | 1 | 0
|  |  | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-13 | Caballeronia glathei | 0.2 | 1 | 0
|  |  | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-20 | Caballeronia glathei | 0 | 1 | 0
|  |  | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea apista | 1.01 | 1 | 0
|  |  | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea pnomenusa | 0.68 | 1 | 0
|  |  | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-24 | Pandoraea pulmonicola | 0.73 | 1 | 0
|  |  | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Paraburkholderia caledonica | 0.94 | 1 | 0
|  |  | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-27 | Paraburkholderia caledonica | 0.76 | 1 | 0
|  |  | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-13 | Paraburkholderia caribensis | 1.29 | 1 | 0
|  |  | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-13 | Paraburkholderia caribensis | 1.22 | 1 | 0
|  |  | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-20 | Paraburkholderia hospita | 1.02 | 1 | 0
|  |  | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2012-04-20 | Paraburkholderia hospita | 0.56 | 1 | 0
|  |  | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-12-01 | Paraburkholderia kururiensis | 0.69 | 1 | 0
|  |  | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-26 | Paraburkholderia phymatum | 0.76 | 1 | 0
|  |  | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2010-11-19 | Paraburkholderia phytofirmans | 1.24 | 1 | 0
|  |  | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2012-04-20 | Paraburkholderia tuberum | 0.23 | 1 | 0
|  |  | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2010-11-26 | Paraburkholderia xenovorans | 0.86 | 1 | 0
|  |  | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2016-05-13 | Ralstonia insidiosa | 0.34 | 1 | 0
|  |  | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2016-05-13 | Ralstonia mannitolilytica | 0.39 | 1 | 0
|  |  | DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2007-01-22 | Ralstonia pickettii | 0.57 | 1 | 0
|  |  | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2009-05-07 | Ralstonia pickettii | 0.95 | 1 | 0
|  |  | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia acidovorans | 0.59 | 1 | 0
|  |  | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia tsuruhatensis | 0.67 | 1 | 0
|  |  | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga intermedia | 1.17 | 1 | 0
|  |  | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-07-25 | Xenophilus sp | 0.48 | 1 | 0
|  |  | DSM 24713 extr | 3; type strain | Extr | 2016-02-02 | Deefgea chitinilytica | 0.83 | 1 | 0
|  |  | DSM 24713 eDT | 3; type strain | eDT | 2015-09-22 | Deefgea chitinilytica | 0.76 | 1 | 0
|  |  | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea rivuli | 1.29 | 1 | 0
|  |  | 08-F17.2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.8 | 1 | 0
|  |  | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.79 | 1 | 0
|  |  | CVUAS 6995.3 extr | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea sp-CVUAS-6995,3 | 0.67 | 1 | 0
|  |  | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-6995,3 | 0.39 | 1 | 0
|  |  | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-9920 | 0.76 | 1 | 0
|  |  | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea sp-CVUAS-9920 | 1.04 | 1 | 0
|  |  | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter limnosediminis | 0.6 | 1 | 0
|  |  | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella kingae | 0.57 | 1 | 0
|  |  | CVUAS 10067.2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria animaloris | 0.64 | 1 | 0
|  |  | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria animaloris | 0.31 | 1 | 0
|  |  | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-4931 | 0.79 | 1 | 0
|  |  | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-8997 | 0.69 | 1 | 0
|  |  | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-LHL-1296 | 0.68 | 1 | 0
|  |  | CVUAS 5699.3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria sp-N95-16 | 0.92 | 1 | 0
|  |  | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae sp-CVUAS-8636 | 0.38 | 1 | 0
|  |  | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter cryaerophilus | 0.72 | 1 | 0
|  |  | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter butzleri | 0.78 | 1 | 0
|  |  | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter butzleri | 1.08 | 1 | 0
|  |  | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter coli | 0.61 | 1 | 0
|  |  | DSM 4689 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter coli | 0.81 | 1 | 0
|  |  | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter fetus | 0.91 | 1 | 0
|  |  | DSM 5361 | 3; type strain | DT | 2013-09-24 | Campylobacter fetus | 0.68 | 1 | 0
|  |  | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter fetus | 0.96 | 1 | 0
|  |  | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter fetus | 0.81 | 1 | 0
|  |  | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter hyointestinalis | 0.97 | 1 | 0
|  |  | CVUAS 6673.2 | 3; sequenced | DT | 2013-08-28 | Campylobacter hyointestinalis | 0.4 | 1 | 0
|  |  | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter jejuni | 0.78 | 1 | 0
|  |  | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter jejuni | 1.07 | 1 | 0
|  |  | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter jejuni | 0.56 | 1 | 0
|  |  | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter lari | 0.81 | 1 | 0
|  |  | BfR-CA-6556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.52 | 1 | 0
|  |  | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.72 | 1 | 0
|  |  | DSM 11375 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter lari | 0.77 | 1 | 0
|  |  | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter peloridis | 0.75 | 1 | 0
|  |  | BfR-CA-7048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter peloridis | 0.61 | 1 | 0
|  |  | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter sputorum | 1.23 | 1 | 0
|  |  | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter sp-W1 | 0.29 | 1 | 0
|  |  | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter upsaliensis | 0.47 | 1 | 0
|  |  | CVUAS 11185.3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter canadensis | 1.07 | 1 | 0
|  |  | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas bestiarum | 0.94 | 1 | 0
|  |  | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas bivalvium | 0.78 | 1 | 0
|  |  | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas caviae | 0.32 | 1 | 0
|  |  | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.45 | 1 | 0
|  |  | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.97 | 1 | 0
|  |  | DSM 30187 RKI | 4; type strain, WGS sequenced (BfR) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-14 | Aeromonas hydrophila | 0.85 | 1 | 0
|  |  | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas hydrophila | 0.8 | 1 | 0
|  |  | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas jandaei | 0.53 | 1 | 0
|  |  | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas media | 0.63 | 1 | 0
|  |  | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas molluscorum | 0.83 | 1 | 0
|  |  | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas popoffii | 0.69 | 1 | 0
|  |  | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas salmonicida | 0.71 | 1 | 0
|  |  | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas salmonicida | 0.62 | 1 | 0
|  |  | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas salmonicida | 0.63 | 1 | 0
|  |  | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas salmonicida | 0.48 | 1 | 0
|  |  | CVUAS 8400.2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas salmonicida | 0.8 | 1 | 0
|  |  | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas salmonicida | 0.57 | 1 | 0
|  |  | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas salmonicida | 0.44 | 1 | 0
|  |  | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas salmonicida | 0.76 | 1 | 0
|  |  | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas simiae | 0.79 | 1 | 0
|  |  | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas veronii | 0.71 | 1 | 0
|  |  | DSM 7386 CVUAS | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas veronii | 0.83 | 1 | 0
|  |  | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas pectinilytica | 0.71 | 1 | 0
|  |  | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella viscosa | 0.69 | 1 | 0
|  |  | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella viscosa | 0.62 | 1 | 0
|  |  | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.81 | 1 | 0
|  |  | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.78 | 1 | 0
|  |  | CVUAS 81.2 | 3; sequenced | EFEx | 2013-11-28 | Shewanella sp-RKVR22 | 0.76 | 1 | 0
|  |  | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Citrobacter amalonaticus | 0.84 | 1 | 0
|  |  | DSM 17342 | 3; public strain collection | see [0185] | 2022-08-29 | Citrobacter amalonaticus | 0.64 | 1 | 0
|  |  | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter freundii | 0.36 | 1 | 0
|  |  | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter freundii | 0.41 | 1 | 0
|  |  | ATCC 25408 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2008-04-11 | Citrobacter koseri | 0.52 | 1 | 0
|  |  | DSM 4570 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=+ 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation: 26.08.2022¦ 1:1 HCCA¦Andy 3.0#EXT=Citrobacter koseri DSM 4570, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-08-29 | Citrobacter koseri | 0.49 | 1 | 0
|  |  | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter condimenti | 0.41 | 1 | 0
|  |  | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter dublinensis | 0.54 | 1 | 0
|  |  | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-14 | Cronobacter malonaticus | 0.6 | 1 | 0
|  |  | DSM 18702 | 3; type strain | DT | 2019-03-05 | Cronobacter malonaticus | 0.09 | 1 | 0
|  |  | CVUAS 32945 | 3; confirmed proficency test | EFEx | 2021-02-18 | Cronobacter sakazakii | 0.56 | 1 | 0
|  |  | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter sakazakii | 0.48 | 1 | 0
|  |  | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter turicensis | 0.33 | 1 | 0
|  |  | DSM 30054 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Enterobacter cloacae | 0.68 | 1 | 0
|  |  | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2011-11-10 | Enterobacter cloacae | 0.93 | 1 | 0
|  |  | DSM 14563 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter hormaechei | 0.82 | 1 | 0
|  |  | DSM 12409 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter hormaechei | 1.02 | 1 | 0
|  |  | DSM 16691 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 04.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Enterobacter hormaechei subsp. hoffmannii DSM 14563 (T), Präparation:Anne und Chris, Messung: Anne#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2023-03-16 | Enterobacter hormaechei | 1.03 | 1 | 0
|  |  | MB 2023000186 | 2; corresponds to expected result of PT | EFEx | 2023-09-08 | Enterobacter sp | 0.6 | 1 | 0
|  |  | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Enterobacter sp-ATCC-27988 | 0.52 | 1 | 0
|  |  | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae sp-CVUAS-30072 | 0.58 | 1 | 0
|  |  | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia coli | 0.75 | 1 | 0
|  |  | CVUAS 0500.2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia coli | 0.91 | 1 | 0
|  |  | DSM 30053 | 3; type strain | EFex | 2019-02-12 | Klebsiella aerogenes | 0.47 | 1 | 0
|  |  | NM20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Klebsiella aerogenes | 0.52 | 1 | 0
|  |  | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella grimontii | 0.7 | 1 | 0
|  |  | CVUAS 2486.3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella grimontii | 0.78 | 1 | 0
|  |  | CVUAS 2404.3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella michiganensis | 0.52 | 1 | 0
|  |  | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella michiganensis | 0.52 | 1 | 0
|  |  | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella oxytoca | 0.56 | 1 | 0
|  |  | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2013-11-01 | Klebsiella oxytoca | 0.74 | 1 | 0
|  |  | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 2008-04-11 | Klebsiella pneumoniae | 0.39 | 1 | 0
|  |  | DSM 30104 CVUAS | 3; type strain | EFEx | 2013-09-17 | Klebsiella pneumoniae | 0.62 | 1 | 0
|  |  | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.64 | 1 | 0
|  |  | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.49 | 1 | 0
|  |  | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella variicola | 0.78 | 1 | 0
|  |  | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella variicola | 0.42 | 1 | 0
|  |  | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia sacchari | 0.63 | 1 | 0
|  |  | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.82 | 1 | 0
|  |  | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.87 | 1 | 0
|  |  | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Pluralibacter gergoviae | 0.92 | 1 | 0
|  |  | CVUAS 7763 | 3; NRL Salmonella BfR | DT | 2012-10-16 | Salmonella enterica | 0.88 | 1 | 0
|  |  | 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Salmonella enterica | 0.6 | 1 | 0
|  |  | II 06 06820 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Shigella boydii | 0.79 | 1 | 0
|  |  | 3A 07 02503 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-06-09 | Shigella flexneri | 0.91 | 1 | 0
|  |  | 273 | 2; strain collection of the RKI | TFAextr | 2021-11-19 | Erwinia amylovora | 0.9 | 1 | 0
|  |  | DSM 3493 CVUAS | 3; type strain | eDT | 2015-07-09 | Pantoea agglomerans | 0.43 | 1 | 0
|  |  | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea agglomerans | 0.93 | 1 | 0
|  |  | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea sp | 0.73 | 1 | 0
|  |  | CVUAS 4263.2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea sp-FF5 | 0.7 | 1 | 0
|  |  | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella sp | 0.59 | 1 | 0
|  |  | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Edwardsiella tarda | 1.12 | 1 | 0
|  |  | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-18 | Hafnia alvei | 0.66 | 1 | 0
|  |  | LHL 1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia alvei | 0.63 | 1 | 0
|  |  | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea myxofaciens | 0.66 | 1 | 0
|  |  | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella morganii | 0.31 | 1 | 0
|  |  | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella morganii | 0.61 | 1 | 0
|  |  | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus hauseri | 0.5 | 1 | 0
|  |  | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus hauseri | 0.56 | 1 | 0
|  |  | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus mirabilis | 0.43 | 1 | 0
|  |  | DSM 4479 CVUAS | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus mirabilis | 0.99 | 1 | 0
|  |  | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus penneri | 0.7 | 1 | 0
|  |  | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.78 | 1 | 0
|  |  | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 1.08 | 1 | 0
|  |  | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.63 | 1 | 0
|  |  | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus terrae | 0.89 | 1 | 0
|  |  | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus terrae | 0.45 | 1 | 0
|  |  | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Proteus vulgaris | 0.75 | 1 | 0
|  |  | DSM 2140 CVUAS | 3; strain from public collection; sequenced (rpoB) | eDT | 2017-11-30 | Proteus vulgaris | 0.82 | 1 | 0
|  |  | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia rettgeri | 0.63 | 1 | 0
|  |  | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia rettgeri | 0.53 | 1 | 0
|  |  | ATCC 25827 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-04-11 | Providencia stuartii | 1 | 1 | 0
|  |  | DSM 4539 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 25.03.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Providencia stuartii DSM 4539 (T), Preparation & measurement: A. Gries und Chris#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-04-06 | Providencia stuartii | 0.66 | 1 | 0
|  |  | CVUAS 7849 | 2; sequenced | eDT | 2022-08-23 | Rahnella aquatilis | 0.94 | 1 | 0
|  |  | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia fonticola | 0.73 | 1 | 0
|  |  | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-11-04 | Serratia grimesii | 1 | 1 | 0
|  |  | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia marcescens | 0.63 | 1 | 0
|  |  | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia marcescens | 0.48 | 1 | 0
|  |  | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia quinivorans | 0.09 | 1 | 0
|  |  | CVUAS 6221 LT | 4; sequenced | DT | 2013-01-08 | Yersinia aldovae | 0.75 | 1 | 0
|  |  | CVUAS 6319 sir | 2; sequenced (16S) | eDT | 2024-12-04 | Yersinia aldovae | 0.72 | 1 | 0
|  |  | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2007-03-09 | Yersinia aleksiciae | 0.54 | 1 | 0
|  |  | IMB 4355 | 2; 0 | #CON=3 loops in 100 mkL#TRT=in 80 pct. TFA¦ 1:1 HCCA-TA2 | 2007-07-13 | Yersinia aleksiciae | 0.46 | 1 | 0
|  |  | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia bercovieri | 0.56 | 1 | 0
|  |  | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia bercovieri | 0.89 | 1 | 0
|  |  | DSM 4780 | 3; type strain | DT | 2013-05-08 | Yersinia enterocolitica | 0.62 | 1 | 0
|  |  | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.71 | 1 | 0
|  |  | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia enterocolitica | 0.6 | 1 | 0
|  |  | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.53 | 1 | 0
|  |  | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia enterocolitica | 0.91 | 1 | 0
|  |  | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia enterocolitica | 0.81 | 1 | 0
|  |  | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia entomophaga | 0.54 | 1 | 0
|  |  | 21 RKI | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-22 | Yersinia frederiksenii | 0.32 | 1 | 0
|  |  | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia frederiksenii | 0.64 | 1 | 0
|  |  | BP 76 sir | 2; 0 | eDT | 2024-12-04 | Yersinia intermedia | 0.84 | 1 | 0
|  |  | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia intermedia | 0.42 | 1 | 0
|  |  | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia kristensenii | 0.71 | 1 | 0
|  |  | CVUAS 302 LT | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia kristensenii | 0.68 | 1 | 0
|  |  | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-11-25 | Yersinia massiliensis | 0.96 | 1 | 0
|  |  | VB 925699 sir | 3; sequenced | eDT | 2024-12-04 | Yersinia massiliensis | 0.83 | 1 | 0
|  |  | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia mollaretii | 0.89 | 1 | 0
|  |  | DSM 18520 | 3; type strain | DT | 16.01.2013 | Yersinia mollaretii | 0.49 | 1 | 0
|  |  | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia nurmii | 0.61 | 1 | 0
|  |  | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia pekkanenii | 0.82 | 1 | 0
|  |  | NCTC 2028 | 3; public strain collection | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Yersinia pestis | 0.95 | 1 | 0
|  |  | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 2009-11-27 | Yersinia pestis | 0.7 | 1 | 0
|  |  | CVUAS 304.2 sir | 4; WGS sequence (BfR) | EFEx | 2023-06-21 | Yersinia pseudotuberculosis | 0.64 | 1 | 0
|  |  | DSM 8992 | 3; type strain | EFEx | 2023-09-06 | Yersinia pseudotuberculosis | 0.58 | 1 | 0
|  |  | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia rohdei | 1.04 | 1 | 0
|  |  | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2007-03-02 | Yersinia rohdei | 0.5 | 1 | 0
|  |  | CVUAS 2772 CVUAS | 3; sequenced | DT | 2013-01-30 | Yersinia ruckeri | 0.32 | 1 | 0
|  |  | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia ruckeri | 0.6 | 1 | 0
|  |  | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia similis | 0.18 | 1 | 0
|  |  | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia wautersii | 0.67 | 1 | 0
|  |  | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 2015-12-03 | Yersinia wautersii | 0.62 | 1 | 0
|  |  | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella londiniensis | 0.82 | 1 | 0
|  |  | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella pneumophila | 0.8 | 1 | 0
|  |  | CVUAS 4122 | 3; 0 | eDT | 2016-01-20 | Legionella pneumophila | 0.77 | 1 | 0
|  |  | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella taurinensis | 0.7 | 1 | 0
|  |  | CVUAS 10873.2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0.9 | 1 | 0
|  |  | CVUAS 10874.2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0.85 | 1 | 0
|  |  | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) caballi | 0.49 | 1 | 0
|  |  | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) caballi | 0.59 | 1 | 0
|  |  | CVUAS 32909 | 3; sequenced | eDT | 2021-02-05 | (Pasteurella) mairii | 0.69 | 1 | 0
|  |  | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) mairii | 0.82 | 1 | 0
|  |  | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) testudinis | 0.91 | 1 | 0
|  |  | LHL 957 | 3; sequenced | eDT | 2015-12-02 | (Pasteurella) testudinis | 0.58 | 1 | 0
|  |  | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus anseriformium | 0.77 | 1 | 0
|  |  | 15UGF198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus anseriformium | 0.61 | 1 | 0
|  |  | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus arthritidis | 0.75 | 1 | 0
|  |  | LLBB 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus capsulatus | 0.96 | 1 | 0
|  |  | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus capsulatus | 0.77 | 1 | 0
|  |  | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.77 | 1 | 0
|  |  | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.43 | 1 | 0
|  |  | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.57 | 1 | 0
|  |  | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus equuli | 0.51 | 1 | 0
|  |  | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus lignieresii | 0.64 | 1 | 0
|  |  | CVUAS 9945.2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus lignieresii | 0.43 | 1 | 0
|  |  | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus rossii | 0.93 | 1 | 0
|  |  | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus rossii | 0.86 | 1 | 0
|  |  | CVUAS 32505.2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus sp-CVUAS-32505,2 | 0.51 | 1 | 0
|  |  | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus sp-genomospecies-1 | 0.74 | 1 | 0
|  |  | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-genomospecies-1 | 0.96 | 1 | 0
|  |  | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus sp-MK-2012 | 0.84 | 1 | 0
|  |  | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-MK-2012 | 0.96 | 1 | 0
|  |  | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus suis | 0.61 | 1 | 0
|  |  | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus suis | 0.56 | 1 | 0
|  |  | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium gallinarum | 0.53 | 1 | 0
|  |  | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium gallinarum | 1.06 | 1 | 0
|  |  | CVUAS 1816.2 | 3; sequenced | DT | 2013-07-24 | Basfia succiniciproducens | 0.96 | 1 | 0
|  |  | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 1.22 | 1 | 0
|  |  | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 0.92 | 1 | 0
|  |  | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter oris | 1.04 | 1 | 0
|  |  | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter oris | 0.92 | 1 | 0
|  |  | CVUAS 3802.2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter osteomyelitidis | 0.65 | 1 | 0
|  |  | CVUAS 31212 | 3; sequenced | EFEx | 2018-10-16 | Exercitatus varius | 0.58 | 1 | 0
|  |  | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Exercitatus varius | 0.46 | 1 | 0
|  |  | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia canicola | 0.61 | 1 | 0
|  |  | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia canicola | 0.5 | 1 | 0
|  |  | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium anatis | 1.3 | 1 | 0
|  |  | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium anatis | 1.1 | 1 | 0
|  |  | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium sp-F151 | 0.83 | 1 | 0
|  |  | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium trehalosifermentans | 0.53 | 1 | 0
|  |  | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus haemoglobinophilus | 0.72 | 1 | 0
|  |  | DSM 4690 | 3; type strain | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 03.06.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Haemophilus influenzae DSM 4690, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-04 | Haemophilus influenzae | 0.81 | 1 | 0
|  |  | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus sp-2269 | 0.67 | 1 | 0
|  |  | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus sp-2269 | 0.88 | 1 | 0
|  |  | P 1066 2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus somni | 0.79 | 1 | 0
|  |  | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia caviae | 0.62 | 1 | 0
|  |  | CVUAS 1064.4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia glucosida | 0.75 | 1 | 0
|  |  | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia glucosida | 0.79 | 1 | 0
|  |  | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia granulomatis | 0.74 | 1 | 0
|  |  | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia granulomatis | 0.71 | 1 | 0
|  |  | CVUAS 1738 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia haemolytica | 0.42 | 1 | 0
|  |  | CVUAS 3952 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia haemolytica | 0.21 | 1 | 0
|  |  | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia haemolytica | 0.57 | 1 | 0
|  |  | CVUAS 6197.2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia haemolytica | 0.48 | 1 | 0
|  |  | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia pernigra | 0.22 | 1 | 0
|  |  | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia pernigra | 0.96 | 1 | 0
|  |  | LLBB 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia ruminalis | 0.63 | 1 | 0
|  |  | LLBB 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia ruminalis | 0.54 | 1 | 0
|  |  | CVUAS 10396.2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia sp-10-12559-1 | 0.52 | 1 | 0
|  |  | CVUAS 10183.3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia sp-14UCF195 | 0.57 | 1 | 0
|  |  | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia sp-14UCF195 | 0.55 | 1 | 0
|  |  | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia sp-15USF15 | 0.24 | 1 | 0
|  |  | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia sp-15USF15 | 0.51 | 1 | 0
|  |  | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia sp-CVUAS-11240 | 0.74 | 1 | 0
|  |  | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia sp-CVUAS-1496 | 0.44 | 1 | 0
|  |  | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter rosorum | 0.57 | 1 | 0
|  |  | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter rosorum | 0.6 | 1 | 0
|  |  | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella semolina | 0.47 | 1 | 0
|  |  | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella semolina | 0.67 | 1 | 0
|  |  | CVUAS 4060.3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella aerogenes | 0.93 | 1 | 0
|  |  | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella canis | 0.61 | 1 | 0
|  |  | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella multocida | 0.59 | 1 | 0
|  |  | CVUAS 32055 | 3; specific PCR | EFEx | 2024-09-25 | Pasteurella multocida | 0.78 | 1 | 0
|  |  | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella multocida | 0.75 | 1 | 0
|  |  | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella multocida | 0.31 | 1 | 0
|  |  | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella multocida | 0.64 | 1 | 0
|  |  | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella multocida | 0.66 | 1 | 0
|  |  | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae Bisgaard-Taxon-40 | 0.85 | 1 | 0
|  |  | CVUAS 9987.2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae canine-oral-taxon-080 | 0.58 | 1 | 0
|  |  | CVUAS 261.4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae sp-CVUAS-261,4 | 0.92 | 1 | 0
|  |  | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae sp-CVUAS-31988 | 0.91 | 1 | 0
|  |  | CVUAS 34093 | 2; sequenced | EFex | 2022-08-05 | Pasteurellaceae sp-CVUAS-34093 | 0.46 | 1 | 0
|  |  | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae sp-CVUAS-4868 | 0.49 | 1 | 0
|  |  | CVUAS 724.3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae sp-CVUAS-724,3 | 0.73 | 1 | 0
|  |  | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae sp-CVUAS-9305 | 0.32 | 1 | 0
|  |  | CVUAS 9967.2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae sp-CVUAS-9967,2 | 1.06 | 1 | 0
|  |  | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae sp-LHL-U-309 | 0.88 | 1 | 0
|  |  | LHL U-309 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-309 | 0.57 | 1 | 0
|  |  | LHL U-3444 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-3444 | 0.66 | 1 | 0
|  |  | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae sp-MCCM-2539 | 0.82 | 1 | 0
|  |  | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.61 | 1 | 0
|  |  | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.69 | 1 | 0
|  |  | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium mucosae | 0.58 | 1 | 0
|  |  | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium mucosae | 0.58 | 1 | 0
|  |  | CVUAS 8966 | 3; sequenced | DT | 2013-07-23 | Volucribacter amazonae | 0.51 | 1 | 0
|  |  | LLBB 23UGF306 | 3; sequenced | OS Extr | 2023-10-11 | Volucribacter psittacicida | 0.52 | 1 | 0
|  |  | CVUAS 308.2 | 3; sequenced | DT | 2013-11-06 | Psittacicella hinzii | 0.8 | 1 | 0
|  |  | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2016-12-14 | Acinetobacter baumannii | 0.59 | 1 | 0
|  |  | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2014-09-03 | Acinetobacter baumannii | 0.37 | 1 | 0
|  |  | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter defluvii | 0.16 | 1 | 0
|  |  | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter guillouiae | 0.32 | 1 | 0
|  |  | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter johnsonii | 0.74 | 1 | 0
|  |  | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter junii | 0.7 | 1 | 0
|  |  | CVUAS 7714 | 3; sequenced | DT | 2013-09-18 | Acinetobacter junii | 0.6 | 1 | 0
|  |  | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter lwoffii | 0.03 | 1 | 0
|  |  | DSM 30502 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter lwoffii DSM 30502, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Acinetobacter lwoffii | 0.6 | 1 | 0
|  |  | DSM 9312 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 28.04.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter pittii DSM 9312, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-05-06 | Acinetobacter pittii | 0.49 | 1 | 0
|  |  | DSM 9318 | 3; public strain collection | #CON=approx. 3 loops in Epp. tube#TRT=plus 50-100 mkL 100 pct. TFA¦ 30 Min treatment¦1:10 diluted, inactivation date: 01.07.2022¦ 1:1 HCCA¦ Andy 3.0#EXT=Acinetobacter pittii DSM 9318, Preparation & measurement: A. Gries#LAS=LL30 pct. 1000, "LL60 pct." 10-20 sh. ¦ Laser attenuator offset: 52 pct.¦Range: 20 pct. (Autoflex 2019 04 04(PeterSetup).isset) defocussed 5xunter Strichueberlappung und ToM 200ns 20190404 gesichert2.par#CAL=E. coli K12 (DSM 3871), linear calibration#CUS=RKI ZBS6# | 2022-07-20 | Acinetobacter pittii | 0.51 | 1 | 0
|  |  | CVUAS 7379.3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter pseudolwoffii | 0.51 | 1 | 0
|  |  | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter pseudolwoffii | 0.82 | 1 | 0
|  |  | CVUAS 32366 | 3; sequenced | eDT | 2020-06-19 | Acinetobacter sp-CVUAS-32366 | 0.69 | 1 | 0
|  |  | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter variabilis | 0.49 | 1 | 0
|  |  | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 1.05 | 1 | 0
|  |  | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 0.91 | 1 | 0
|  |  | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella catarrhalis | 1.17 | 1 | 0
|  |  | CVUAS 11674.3 | 3; sequenced | eDT | 2022-03-22 | Moraxella cuniculi | 0.74 | 1 | 0
|  |  | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella osloensis | 0.88 | 1 | 0
|  |  | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella ovis | 0.65 | 1 | 0
|  |  | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter frigidicola | 0.75 | 1 | 0
|  |  | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter pasteurii | 1.13 | 1 | 0
|  |  | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.87 | 1 | 0
|  |  | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.77 | 1 | 0
|  |  | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas (alkylphenolica) | 0.37 | 1 | 0
|  |  | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas [baetica] | 0.64 | 1 | 0
|  |  | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas aeruginosa | 0.34 | 1 | 0
|  |  | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-01-10 | Pseudomonas aeruginosa | 0.64 | 1 | 0
|  |  | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas alcaligenes | 0.92 | 1 | 0
|  |  | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas anguilliseptica | 0.7 | 1 | 0
|  |  | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas brenneri | 0.8 | 1 | 0
|  |  | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2008-02-20 | Pseudomonas chlororaphis | 0.83 | 1 | 0
|  |  | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas fulva | 0.66 | 1 | 0
|  |  | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas koreensis | 0.61 | 1 | 0
|  |  | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas monteilii | 0.5 | 1 | 0
|  |  | 171011722-1 A | 2; 0 | DT | 2017-08-31 | Pseudomonas nitroreducens | 0.57 | 1 | 0
|  |  | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas oryzihabitans | 0.88 | 1 | 0
|  |  | 171011722-1 B | 2; 0 | DT | 2017-08-31 | Pseudomonas putida | 0.57 | 1 | 0
|  |  | ATCC 23974 | 3; public strain collection | TFAextr | 2021-09-08 | Pseudomonas putida | 0.36 | 1 | 0
|  |  | DSM 5190 | 3; type strain | TFAextr | 2022-06-04 | Pseudomonas stutzeri | 0.84 | 1 | 0
|  |  | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas taetrolens | 0.61 | 1 | 0
|  |  | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2016-12-12 | Allofrancisella guangzhouensis | 0.86 | 1 | 0
|  |  | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-11-18 | Francisella hispaniensis | 0.78 | 1 | 0
|  |  | DSM 18777 | 3; strain collection of the RKI | EFEx | 2013-06-20 | Francisella noatunensis | 0.57 | 1 | 0
|  |  | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 2015-11-25 | Francisella orientalis | 0.54 | 1 | 0
|  |  | DSM 7535 | 3; type strain | EFEx | 2013-06-20 | Francisella philomiragia | 0.7 | 1 | 0
|  |  | CVUAS 4680.2 | 4; WGS (FLI) | DT | 2013-04-02 | Francisella tularensis | 0.3 | 1 | 0
|  |  | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 2013-04-23 | Francisella tularensis | 0.91 | 1 | 0
|  |  | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 2013-04-26 | Francisella tularensis | 0.67 | 1 | 0
|  |  | F064 | 3; strain collection of the FLI | EFEx | 2013-06-20 | Francisella tularensis | 0.84 | 1 | 0
|  |  | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 2012-04-02 | Francisella tularensis | 0.81 | 1 | 0
|  |  | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 2017-06-01 | Francisella tularensis | 0.55 | 1 | 0
|  |  | CVUAS 33180 25C | 3; sequenced | DT | 2021-08-04 | Photobacterium damselae | 0.51 | 1 | 0
|  |  | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium damselae | 0.23 | 1 | 0
|  |  | CVUAS 1266.2 | 3; sequenced | EFEx | 2019-09-25 | Photobacterium piscicola | 0.43 | 1 | 0
|  |  | CVUAS 842.2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium swingsii | 1.11 | 1 | 0
|  |  | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio aestuarianus | 0.61 | 1 | 0
|  |  | CVUAS 7079 sir | 3; spezific PCR (CVFR) | eDT | 2022-06-23 | Vibrio alginolyticus | 0.81 | 1 | 0
|  |  | DSM 2171 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 2012-07-12 | Vibrio alginolyticus | 0.58 | 1 | 0
|  |  | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio anguillarum | 0.65 | 1 | 0
|  |  | A8-12 sir | 3; gyrB sequence | eDT | 2022-06-23 | Vibrio anguillarum | 0.65 | 1 | 0
|  |  | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio campbelli | 0.43 | 1 | 0
|  |  | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio campbelli | 0.44 | 1 | 0
|  |  | 16-VB00067-0 eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio cholerae | 0.78 | 1 | 0
|  |  | CVUAS 10227 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio cholerae | 0.74 | 1 | 0
|  |  | 45V066-10 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio cincinnatiensis | 0.92 | 1 | 0
|  |  | LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio cincinnatiensis | 0.83 | 1 | 0
|  |  | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio diazotrophicus | 0.85 | 1 | 0
|  |  | CVUAS 8995 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio diazotrophicus | 0.24 | 1 | 0
|  |  | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio fluvialis | 0.33 | 1 | 0
|  |  | LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio fluvialis | 0.35 | 1 | 0
|  |  | CVUAS 1754 LT | 3; sequenced | eDT | 2013-04-16 | Vibrio furnissii | 0.79 | 1 | 0
|  |  | DSM 14383 sir | 3; public strain collection, sequenced | eDT | 2022-06-23 | Vibrio furnissii | 0.3 | 1 | 0
|  |  | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 2012-08-16 | Vibrio harveyi | 0.33 | 1 | 0
|  |  | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio harveyi | 0.32 | 1 | 0
|  |  | P340-7 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio jasicida | 0.27 | 1 | 0
|  |  | P340-9 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio jasicida | 0.45 | 1 | 0
|  |  | 12-VBH0498 | 4; WGS sequenced (BfR) | DT | 2021-02-02 | Vibrio metoecus | 0.47 | 1 | 0
|  |  | 20-VB00016 | 4; WGS sequenced (BfR) | DT | 2021-02-24 | Vibrio metoecus | 0.83 | 1 | 0
|  |  | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio metschnikovii | 0.72 | 1 | 0
|  |  | CVUAS 619 sir | 2; 0 | eDT | 2022-06-23 | Vibrio metschnikovii | 0.67 | 1 | 0
|  |  | CVUAS 7872 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio mimicus | 0.47 | 1 | 0
|  |  | DSM 19130 | 4; type strain; WGS sequenced | DT | 2021-02-24 | Vibrio mimicus | 0.62 | 1 | 0
|  |  | 20-VB00299 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio navarrensis | 0.61 | 1 | 0
|  |  | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio navarrensis | 1 | 1 | 0
|  |  | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio owensii | 0.91 | 1 | 0
|  |  | 17-VB00254-0 | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio parahaemolyticus | 0.32 | 1 | 0
|  |  | DSM 10027 sir | 3; type strain | eDT | 2022-06-23 | Vibrio parahaemolyticus | 0.51 | 1 | 0
|  |  | CVUAS 871.2 | 3; sequenced | eDT | 2014-08-21 | Vibrio sp-CVUAS-871,2 | 0.49 | 1 | 0
|  |  | CVUAS 2778.2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio tubiashii | 0.61 | 1 | 0
|  |  | CVUAS 2778.2 | 3; sequenced | eDT | 2019-10-24 | Vibrio tubiashii | 0.78 | 1 | 0
|  |  | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio vulnificus | 0.62 | 1 | 0
|  |  | LVM05 LVU05 extr | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-14 | Vibrio vulnificus | 0.87 | 1 | 0
|  |  | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas sp-G9077 | 0.27 | 1 | 0
|  |  | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria indica | 0.97 | 1 | 0
|  |  | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria ureiclastica | 0.6 | 1 | 0
|  |  | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas chelatiphaga | 0.64 | 1 | 0
|  |  | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-09-03 | Stenotrophomonas maltophilia | 0.42 | 1 | 0
|  |  | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2009-05-14 | Stenotrophomonas maltophilia | 0.57 | 1 | 0
|  |  | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas rhizophila | 0.71 | 1 | 0
|  |  | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2015-05-18 | Stenotrophomonas rhizophila | 0.39 | 1 | 0
|  |  | CVUAS 4239.2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas tumulicola | 0.54 | 1 | 0

## Report (Spezies)
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Parameter: Staphylococcus haemolyticus-borealis
Version: | 4
erstellt von: | Dyk | geprüft von: | Rau
erstellt am: | 2025-11-26 | geprüft am: | 2026-01-20
Anwendung | Bakterienisolate aus Lebensmitteln & der veterinärmedizinischen Diagnostik
Umfeld | Lebensmittel-Mikrobiologie, Mastitis, Veterinärmedizinische Diagnostik
Verwendete Datenbank / Programmversion |  | BRU 12438 + UA BW 1667 | BT 4.x / 5.x
Validierungsset:
Validierungsisolate/-materialien (Parameter): 23 |  | Vergleichsisolate/-materialien (#Parameter): 2391
Kultivierungsbedingungen bei Mikroorganismen: SBA, 37° C, 24h, aerob/anaerob/CO2/mikroaerophil
Zustand der Probe bei sonstigen Materialien:
Aufreinigungsmethode: DT, eDT, EFex, 80%TFA
Es werden nur identifizierte Proben für die Ermittlung der In- und Exklusivität herangezogen.
Richtig-positiv: 1. Hit =Ziel-Parameter mit Score ≥ 2 und Score 2. Hit < 1.7 oder Score 2. Hit ≥ 1.7 & Gattung gleich 1. Hit oder Score 2. Hit ≥ 2 & Art gleich 1. Hit |  | Richtig-negativ: 1. Hit ohne Zuordnung zum Ziel-Parameter mit score ≥ 2
Falsch-negativ: Kriterien für richtig-positiv werden nicht erfüllt |  | Falsch-positiv: Kriterien für richtig-negativ werden nicht erfüllt
Validierungsergebnis | Ziel-Parameter: | Staphylococcus haemolyticus-borealis |  | SOLL | erfüllt
Zahl aller Spektren | 23
davon mit ID-Ergebnis | 22 | Identifiziert Rate | 95.65% | 20 | ja | 1
davon Eigenspektren | 20 |  |  | 4.4 | ja | 1
In Bezug auf Identifizierte: |  | Inklusivität
Positive Übereinstimmungen (PA) | 22 | Richtig-Positiv Rate | 100% | 95% | ja | 1
Negative Abweichungen (ND) | 0 | Falsch-Negativ Rate | 0% | 1% | ja | 1
Nicht-Ziel-Parameter
Zahl aller Spektren | 2391
davon mit ID-Ergebnis | 1599 | Identifiziert Rate | 66.88% | 30 | ja | 1
davon Eigenspektren | 872 |  |  | 319.8 | ja | 1
In Bezug auf Identifizierte: |  | Exklusivität
Negative Übereinstimmungen (NA) | 1599 | Richtig-Negativ Rate | 100% | 99% | ja | 1
Positive Abweichungen (PD) | 0 | Falsch-Positiv Rate | 0% | 1% | ja | 1
Test auf spektrale Nähe | 2391 | Anteil nah verwandter Sp | 0%
Von 22 identifizierten Proben des Parameters wurden unter Verwendung der vollständigen Datenbank 22 (=100%) richtig erkannt (Inklusivität). 0 (=0%) der identifizierten Proben des Parameters wurden falsch einer anderen Spezies zugeordnet.
In den 1599 identifizierten Proben der Nicht-Ziel-Parameter wurden unter Verwendung der vollständigen Datenbank 100% richtig-negativ (Exklusivität) gezählt. Von diesen 1599 Proben wurden 0 (=0%) fehlerhaft als Parameter identifiziert.
In den 2391 Proben der Nicht-Ziel-Parameter wurden unter Verwendung einer nur den Parameter enthaltenden Datenbank bei 0 (=0%) Einträgen ein score > 2 erreicht.
Die Kriterien für die | Zielgerichtete Identifizierung | sind somit | erfüllt
Kommentar:
S. haemolyticus und S. borealis sind auch per 16S Sequenzierung nicht zu unterscheiden. Daher wird der Parameter zusammenfassend validiert.
Freigabe an Standort | für (Abteilung / Bereich) | am (Datum) | von (Hdz.)
CVUA S | T, MT, D | 2026-01-20 | Rau
Das Verfahren ist für den vorgesehenen Einsatzzweck geeignet und zur Verwendung freigegeben.
als Prüfmethode für Parameter in LIMS eingepflegt
Eine unterschriebene Version dieses Validierungsreports befindet sich im erstellenden Labor.
Andere Keime, die in dieser Matrix vorkommen können.
Andere Keime, die auf diesem Agar wachsen können.
Vergleichsisolate/-materialien, die ähnliche Spektren liefern können (siehe Blatt "Ähnliche zu Parameter").
ja
nein