## Settings
Settings für die Validierung
Bestätigungsprüfung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Institution / Labor | CVUA S
Parameter | Gattung (Genus) |  | Art (Spezies)
|  | Clostridium | perfringens
Unterscheidungsgrenzen
| Score zur Unterscheidung der Gattung | 1.7 | Score zur Unterscheidung der Spezies | 2

## Parameter (Spezies) DB
Parameter: | Clostridium | perfringens | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | nicht zugeordnet | richtig-positiv | falsch-negativ | falsch zugeordnet
|  |  | 1 | 1 | #105003 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.29 | Clostridium | perfringens | 2.27 | A | 1 | 0 | 1 | 0 | 0
Zahl aller Spektren | 101 |  | 1 | 1 | #2012 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.01 | Clostridium | perfringens | 1.99 | A | 1 | 0 | 1 | 0 | 0
davon Identifiziert | 101 |  | 1 | 1 | 03YSZ32 - 99 | 3; 0 | DT | 2013-08-06 | Clostridium | perfringens | 2.09 | Clostridium | perfringens | 2.09 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | 04 28 Balanca | 3; 0 | DT | 2013-01-09 | Clostridium | perfringens | 2.48 | Clostridium | perfringens | 2.38 | A | 1 | 0 | 1 | 0 | 0
von allen Isolaten |  |  | 1 | 1 | 04UCF87 - 150 | 3; 0 | DT | 2013-08-06 | Clostridium | perfringens | 2.44 | Clostridium | perfringens | 2.33 | A | 1 | 0 | 1 | 0 | 0
richtig-positiv: | 101 |  | 1 | 1 | 07BA 35790 001 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.27 | Clostridium | perfringens | 2.26 | A | 1 | 0 | 1 | 0 | 0
nicht zugeordnet | 0 |  | 1 | 1 | 07BA 35790 002 | 3; 0 | DT | 2013-08-06 | Clostridium | perfringens | 2.03 | Clostridium | perfringens | 2.02 | A | 1 | 0 | 1 | 0 | 0
falsch zugeordnet | 0 |  | 1 | 1 | 07BA 35790 003 | 3; 0 | DT | 2013-08-06 | Clostridium | perfringens | 2.06 | Clostridium | perfringens | 2.02 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | 09UG Typ C | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.22 | Clostridium | perfringens | 2.05 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | 09UG Typ D 1 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.42 | Clostridium | perfringens | 2.4 | A | 1 | 0 | 1 | 0 | 0
Scorewertgrenzen: | 2 | Spezies | 1 | 1 | 09UG Typ D 2 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.14 | Clostridium | perfringens | 2.06 | A | 1 | 0 | 1 | 0 | 0
| 1.7 | Genus | 1 | 1 | 09UG Typ Ee+ | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.27 | Clostridium | perfringens | 2.24 | A | 1 | 0 | 1 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 0 | 0 | 151003186 | 2; 0 | DT | 2015-03-16 | Clostridium | perfringens | 2.22 | Clostridium | perfringens | 2.22 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 151003678-2 | 3; clostridium perfringens PCR | DT | 2015-03-25 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 151010872 | 3; clostridium perfringens PCR | DT | 2015-08-17 | Clostridium | perfringens | 2.3 | Clostridium | perfringens | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 151015524-2 | 2; 0 | DT | 2015-12-10 | Clostridium | perfringens | 2.53 | Clostridium | perfringens | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 171002729_2 perf | 3; clostridium perfringens PCR | DT | 2017-02-28 | Clostridium | perfringens | 2.57 | Clostridium | perfringens | 2.53 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 171011776-8 | 2; 0 | DT | 2017-09-01 | Clostridium | perfringens | 2.39 | Clostridium | perfringens | 2.34 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 171013066 | 3; clostridium perfringens PCR | DT | 2017-10-02 | Clostridium | perfringens | 2.54 | Clostridium | perfringens | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 171013196-2 | 2; 0 | DT | 2017-10-02 | Clostridium | perfringens | 2.56 | Clostridium | perfringens | 2.45 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 181001399 | 2; 0 | DT | 2018-02-01 | Clostridium | perfringens | 2.47 | Clostridium | perfringens | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 181001479 | 2; 0 | DT | 2018-02-02 | Clostridium | perfringens | 2.3 | Clostridium | perfringens | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 181002100 | 2; 0 | DT | 2018-02-14 | Clostridium | perfringens | 2.55 | Clostridium | perfringens | 2.44 | A | 1 | 0 | 1 | 0 | 0
Positive Übereinstimmungen (PA) | 101 |  | 0 | 0 | 181002933 | 3; clostridium perfringens PCR | DT | 2018-02-03 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.34 | A | 1 | 0 | 1 | 0 | 0
Negative Abweichungen (ND) | 0 |  | 0 | 0 | 181008082-6 | 3; clostridium perfringens PCR | DT | 2018-02-14 | Clostridium | perfringens | 2.52 | Clostridium | perfringens | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 181008314 Cl | 2; 0 | DT | 2018-05-06 | Clostridium | perfringens | 2.47 | Clostridium | perfringens | 2.44 | A | 1 | 0 | 1 | 0 | 0
nicht identifiziert in % | 0% |  | 0 | 0 | 181008379 | 2; 0 | DT | 2018-06-07 | Clostridium | perfringens | 2.32 | Clostridium | perfringens | 2.29 | A | 1 | 0 | 1 | 0 | 0
falsch zugeordnert in % | 0% |  | 0 | 0 | 181010904 | 2; 0 | DT | 2018-07-23 | Clostridium | perfringens | 2.35 | Clostridium | perfringens | 2.27 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 181010929 | 2; 0 | DT | 2018-07-23 | Clostridium | perfringens | 2.41 | Clostridium | perfringens | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 181010930 | 2; 0 | DT | 2018-07-23 | Clostridium | perfringens | 2.52 | Clostridium | perfringens | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009517-002_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009517-002_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.27 | Clostridium | perfringens | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009544-002_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.42 | Clostridium | perfringens | 2.36 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009544-002_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009550_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.52 | Clostridium | perfringens | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009550_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.47 | Clostridium | perfringens | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009566_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.26 | Clostridium | perfringens | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009566_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.32 | Clostridium | perfringens | 2.25 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009567-001_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.41 | Clostridium | perfringens | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009567-001_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.33 | Clostridium | perfringens | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009604_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.25 | Clostridium | perfringens | 2.21 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4101-009604_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.21 | Clostridium | perfringens | 2.19 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001041_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.07 | Clostridium | perfringens | 2.05 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001041_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.35 | Clostridium | perfringens | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001047_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.26 | Clostridium | perfringens | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001047_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.39 | Clostridium | perfringens | 2.38 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001051_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001051_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.31 | Clostridium | perfringens | 2.25 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001052_DT | 3; PCR | DT | 2017-11-23 | Clostridium | perfringens | 2.21 | Clostridium | perfringens | 2.19 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | 2017-4104-001052_eDT | 3; PCR | eDT | 2017-11-23 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | A | 3; case rerort | DT | 2012-12-19 | Clostridium | perfringens | 2.38 | Clostridium | perfringens | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | ATCC 13124 | 3; type strain | DT | 2013-08-06 | Clostridium | perfringens | 2.44 | Clostridium | perfringens | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | ATCC 27324 | 3; public strain collection | DT | 2012-12-19 | Clostridium | perfringens | 2.09 | Clostridium | perfringens | 1.9 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | B | 3; case rerort | DT | 2013-08-06 | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 0079 | 3; spec. PCR | DT | 2012-12-19 | Clostridium | perfringens | 2.07 | Clostridium | perfringens | 2.05 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 0088,2 | 3; sequenced | eDT | 2014-01-09 | Clostridium | perfringens | 2.16 | Clostridium | perfringens | 2.14 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 10313 | 2; specific PCR | eDT | 2021-05-11 | Clostridium | perfringens | 2.46 | Clostridium | perfringens | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 10314 | 2; specific PCR | eDT | 2021-05-11 | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 1655,2 | 3; 0 | eDT | 2017-02-16 | Clostridium | perfringens | 2.15 | Clostridium | perfringens | 2.15 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2711 | 3; case rerort [0112] | eDT | 2015-03-26 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2712 | 3; 0 | DT | 2013-08-06 | Clostridium | perfringens | 2.32 | Clostridium | perfringens | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2713 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.03 | Clostridium | perfringens | 1.99 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 2891 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.36 | Clostridium | perfringens | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 30066 | 2; specific PCR | eDT | 2024-06-06 | Clostridium | perfringens | 2.31 | Clostridium | perfringens | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 30090 | 2; 0 | eDT | 2024-06-06 | Clostridium | perfringens | 2.26 | Clostridium | perfringens | 2.24 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 31008 | 2; 0 | eDT | 2018-06-21 | Clostridium | perfringens | 2.23 | Clostridium | perfringens | 2.11 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 32280 | 2; LVU | eDT | 2020-04-17 | Clostridium | perfringens | 2.33 | Clostridium | perfringens | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34064 | 2; specific PCR | eDT | 2022-07-21 | Clostridium | perfringens | 2.36 | Clostridium | perfringens | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34280 | 2; specific PCR | eDT | 2022-09-09 | Clostridium | perfringens | 2.47 | Clostridium | perfringens | 2.47 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34443 | 1; standard DEV-Verfahren | eDT | 2022-10-21 | Clostridium | perfringens | 2.24 | Clostridium | perfringens | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34592 | 3; well described strain | eDT | 2022-12-05 | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.4 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34593 | 3; well described strain | eDT | 2022-12-05 | Clostridium | perfringens | 2.48 | Clostridium | perfringens | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34594 | 3; well described strain | eDT | 2022-12-05 | Clostridium | perfringens | 2.53 | Clostridium | perfringens | 2.49 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34595 | 3; well described strain | eDT | 2022-12-05 | Clostridium | perfringens | 2.39 | Clostridium | perfringens | 2.39 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 34879 | 3; well described strain | eDT | 2023-01-31 | Clostridium | perfringens | 2.5 | Clostridium | perfringens | 2.46 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 36477 | 2; acc. DIN EN ISO 14189:2016-11 | EFEx | 2024-06-14 | Clostridium | perfringens | 2.14 | Clostridium | perfringens | 2.02 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 36478 | 2; acc. DIN EN ISO 14189:2016-11 | EFEx | 2024-06-14 | Clostridium | perfringens | 2.1 | Clostridium | perfringens | 1.89 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 4670,2 | 3; well described strain | eDT | 2021-06-18 | Clostridium | perfringens | 2.46 | Clostridium | perfringens | 2.44 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 5543,4 | 2; specific PCR | eDT | 2019-09-25 | Clostridium | perfringens | 2.47 | Clostridium | perfringens | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 6338 | 3; spez. PCR | DT | 2013-08-06 | Clostridium | perfringens | 2.15 | Clostridium | perfringens | 2.04 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | CVUAS 7825 | 3; spec. PCR | DT | 2012-12-19 | Clostridium | perfringens | 2.29 | Clostridium | perfringens | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | D24_0736-1-1-11 Ripac | 2; 0 | eDT | 2024-06-19 | Clostridium | perfringens | 2.07 | Clostridium | perfringens | 2.01 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | DSM 756 SK22 | 3; type strain | DT | 2021-11-12 | Clostridium | perfringens | 2.49 | Clostridium | perfringens | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | DSM 756 CVUAS | 3; type strain | EFEx | 2013-05-22 | Clostridium | perfringens | 2.2 | Clostridium | perfringens | 2.16 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | DSM 756 DT LUA RLP | 3; type strain | DT | 2017-11-23 | Clostridium | perfringens | 2.26 | Clostridium | perfringens | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | DSM 756 eDT LUA RLP | 3; type strain | eDT | 2017-11-23 | Clostridium | perfringens | 2.42 | Clostridium | perfringens | 2.42 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | DSM 756-Ref.139 | 3; type strain | DT | 2017-12-06 | Clostridium | perfringens | 2.31 | Clostridium | perfringens | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | K 287 | 3; spec. PCR | eDT | 2024-06-06 | Clostridium | perfringens | 2.2 | Clostridium | perfringens | 2.14 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF 18LCP0391-01#0002 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.35 | Clostridium | perfringens | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF 18LMT1451-01#0001 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.59 | Clostridium | perfringens | 2.43 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF_18LMT1624-01#0003 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.51 | Clostridium | perfringens | 2.37 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF_20LMT0086-01#0003 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.43 | Clostridium | perfringens | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF_20LMT0087-01#0004 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF_20LMT0088-01#0001 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.37 | Clostridium | perfringens | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF_20LMT1312-01#0001 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.42 | Clostridium | perfringens | 2.32 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF_21LMT1377-01#0009 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.46 | Clostridium | perfringens | 2.33 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | LALLF_23LMT0581-02#0003 | 1; biochemistry | DT | 2021-11-12 | Clostridium | perfringens | 2.36 | Clostridium | perfringens | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | MB_0002 | 3; confirmed in laboratory ring trial | DT | 2019-09-26 | Clostridium | perfringens | 2.4 | Clostridium | perfringens | 2.3 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | Mü 738 | 3; spec. PCR | DT | 2012-12-19 | Clostridium | perfringens | 2.1 | Clostridium | perfringens | 2.08 | A | 1 | 0 | 1 | 0 | 0
|  |  | 0 | 0 | NCTC 8748 | 3; public strain collection | DT | 2015-06-23 | Clostridium | perfringens | 2.27 | Clostridium | perfringens | 2.2 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | UG Typ A | 3; spec. PCR | DT | 2012-12-19 | Clostridium | perfringens | 2.06 | Clostridium | perfringens | 1.94 | A | 1 | 0 | 1 | 0 | 0

## #Parameter (Spezies) DB
Parameter: | Clostridium perfringens |  | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | verwendete Gattung | verwendete Art | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | richtige ID | keine richtige ID | 1. Hit | 2. Hit | 1. Hit Treffer Parameter | 2. Hit Treffer Parameter | falsch als Parameter identifiziert | PM unter den ersten 2 Hits
|  |  | 1 | 0 | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | bovis | Actinomyces | bovis | 1.92 | Actinomyces | bovis | 1.68 | B | 0 | 0 | 1 | Actinomyces bovis | Actinomyces bovis | 0 | 0 | 0 | 0
Anzahl Isolate: | 1144 |  | 1 | 1 | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces | denticolens | Actinomyces | denticolens | 2.07 | Actinomyces | denticolens | 2.05 | A | 1 | 1 | 0 | Actinomyces denticolens | Actinomyces denticolens | 0 | 0 | 0 | 0
davon Identifiziert | 543 |  | 1 | 0 | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces | sp-CVUAS-31303 | Campylobacter | jejuni | 1.26 | Staphylococcus | carnosus | 1.22 | B | 0 | 0 | 1 | Campylobacter jejuni | Staphylococcus carnosus | 0 | 0 | 0 | 0
richtige ID | 543 |  | 1 | 1 | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces | viscosus | Actinomyces | viscosus | 2.04 | Actinomyces | viscosus | 1.87 | A | 1 | 1 | 0 | Actinomyces viscosus | Actinomyces viscosus | 0 | 0 | 0 | 0
keine richtige ID | 601 |  | 1 | 1 | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces | weissii | Actinomyces | weissii | 2.35 | Actinomyces | weissii | 2.35 | A | 1 | 1 | 0 | Actinomyces weissii | Actinomyces weissii | 0 | 0 | 0 | 0
falsch als Parameter identifiziert | 0 |  | 1 | 0 | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium | bovis | Arcanobacterium | pluranimalium | 1.77 | Arcanobacterium | pluranimalium | 1.52 | B | 0 | 0 | 1 | Arcanobacterium pluranimalium | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium | canis | Arcanobacterium | canis | 2.19 | Legionella | adelaidensis | 1.26 | A | 1 | 1 | 0 | Arcanobacterium canis | Legionella adelaidensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium | haemolyticum | Arcanobacterium | haemolyticum | 2.3 | Arcanobacterium | haemolyticum | 2.18 | A | 1 | 1 | 0 | Arcanobacterium haemolyticum | Arcanobacterium haemolyticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 17-6203 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.29 | Arcanobacterium | hippocoleae | 1.34 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium | hippocoleae | Arcanobacterium | hippocoleae | 2.1 | Arcanobacterium | hippocoleae | 1.49 | A | 1 | 1 | 0 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
Parameter wird getroffen mit: |  |  | 1 | 1 | DSM 10002 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.5 | Arcanobacterium | phocae | 2.4 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium | phocae | Arcanobacterium | phocae | 2.4 | Arcanobacterium | phocae | 2.29 | A | 1 | 1 | 0 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
1. Hit Treffer Parameter | 0 |  | 1 | 1 | 2698T | 3; type strain | DT | 2013-08-07 | Arcanobacterium | phocisimile | Arcanobacterium | phocisimile | 2.15 | Arcanobacterium | phocisimile | 1.84 | A | 1 | 1 | 0 | Arcanobacterium phocisimile | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
2. Hit Treffer Parameter | 0 |  | 1 | 1 | 014418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.12 | Arcanobacterium | phocisimile | 1.79 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
PM unter den ersten 2 Hits | 0 |  | 1 | 1 | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium | pinnipediorum | Arcanobacterium | pinnipediorum | 2.34 | Arcanobacterium | phocisimile | 1.92 | A | 1 | 1 | 0 | Arcanobacterium pinnipediorum | Arcanobacterium phocisimile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13483 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium | pluranimalium | Arcanobacterium | pluranimalium | 2.1 | Arcanobacterium | pluranimalium | 1.57 | A | 1 | 1 | 0 | Arcanobacterium pluranimalium | Arcanobacterium pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | hippocoleae | 2.2 | Arcanobacterium | hippocoleae | 2.17 | A | 0 | 0 | 1 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium | sp-17-2065.02 | Arcanobacterium | hippocoleae | 2.28 | Arcanobacterium | hippocoleae | 2.25 | A | 0 | 0 | 1 | Arcanobacterium hippocoleae | Arcanobacterium hippocoleae | 0 | 0 | 0 | 0
Scorewertgrenzen: | 2 | Spezies | 1 | 0 | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium | sp-2701 | Arcanobacterium | phocae | 2.49 | Arcanobacterium | phocae | 2.36 | A | 0 | 0 | 1 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
| 1.7 | Genus | 1 | 0 | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium | sp-2701 | Arcanobacterium | phocae | 2.46 | Arcanobacterium | phocae | 2.37 | A | 0 | 0 | 1 | Arcanobacterium phocae | Arcanobacterium phocae | 0 | 0 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 1 | 0 | CVUAS 1060,2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium | wilhelmae | Staphylococcus | schleiferi | 1.33 | Staphylococcus | schleiferi | 1.3 | B | 0 | 0 | 1 | Staphylococcus schleiferi | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 647 | 3; type strain | DT | 2014-09-04 | Arcanobacterium | wilhelmae | Lactobacillus | amylovorus | 1.34 | Torulaspora | delbrueckii | 1.25 | B | 0 | 0 | 1 | Lactobacillus amylovorus | Torulaspora delbrueckii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia | sp-CVUAS-30561 | Streptomyces | polychromogenes | 1.39 | Streptomyces | racemochromogenes | 1.38 | B | 0 | 0 | 1 | Streptomyces polychromogenes | Streptomyces racemochromogenes | 0 | 0 | 0 | 0
Negative Übereinstimmungen (NA) | 1144 |  | 1 | 1 | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia | canis | Schaalia | canis | 2.26 | Schaalia | canis | 2.24 | A | 1 | 1 | 0 | Schaalia canis | Schaalia canis | 0 | 0 | 0 | 0
Positive Abweichungen (PD) | 0 |  | 1 | 0 | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia | cardiffensis | Schaalia | cardiffensis | 1.4 | Veillonella | atypica | 1.29 | B | 0 | 0 | 1 | Schaalia cardiffensis | Veillonella atypica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.12 | Schaalia | hyovaginalis | 2.06 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
PD in % | 0% |  | 1 | 1 | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia | hyovaginalis | Schaalia | hyovaginalis | 2.4 | Schaalia | hyovaginalis | 2.29 | A | 1 | 1 | 0 | Schaalia hyovaginalis | Schaalia hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia | sp-CVUAS-32183 | Schaalia | funkei | 1.46 | Schaalia | funkei | 1.42 | B | 0 | 0 | 1 | Schaalia funkei | Schaalia funkei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia | sp-CVUAS-8688 | Schaalia | odontolytica | 1.39 | Schaalia | odontolytica | 1.3 | B | 0 | 0 | 1 | Schaalia odontolytica | Schaalia odontolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31845,2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia | sp-CVUAS-8688 | Schaalia | odontolytica | 1.44 | Schaalia | odontolytica | 1.37 | B | 0 | 0 | 1 | Schaalia odontolytica | Schaalia odontolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella | abortisuis | Trueperella | abortisuis | 2.27 | Trueperella | abortisuis | 2.2 | A | 1 | 1 | 0 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | abortisuis | Trueperella | abortisuis | 1.78 | Trueperella | abortisuis | 1.75 | B | 0 | 0 | 1 | Trueperella abortisuis | Trueperella abortisuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella | bernardiae | Trueperella | bernardiae | 2.11 | Trueperella | bernardiae | 2.03 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella | bernardiae | Trueperella | bernardiae | 2.23 | Trueperella | bernardiae | 2.15 | A | 1 | 1 | 0 | Trueperella bernardiae | Trueperella bernardiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella | bialowiezensis | Trueperella | bialowiezensis | 2.02 | Oceanobacillus | picturae | 1.3 | A | 1 | 1 | 0 | Trueperella bialowiezensis | Oceanobacillus picturae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella | bonasi | Trueperella | bonasi | 2.47 | Trueperella | bonasi | 2.32 | A | 1 | 1 | 0 | Trueperella bonasi | Trueperella bonasi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Trueperella | pyogenes | 1.6 | Trueperella | pyogenes | 1.5 | B | 0 | 0 | 1 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella | pecoris | Bergeyella | zoohelcum | 1.44 | Levilactobacillus | brevis | 1.34 | B | 0 | 0 | 1 | Bergeyella zoohelcum | Levilactobacillus brevis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.42 | Trueperella | pyogenes | 2.39 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella | pyogenes | Trueperella | pyogenes | 2.32 | Trueperella | pyogenes | 2.3 | A | 1 | 1 | 0 | Trueperella pyogenes | Trueperella pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.19 | Corynebacterium | amycolatum | 2.15 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2014 E8 | Corynebacterium | amycolatum | Corynebacterium | amycolatum | 2.3 | Corynebacterium | amycolatum | 2.29 | A | 1 | 1 | 0 | Corynebacterium amycolatum | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium | aurimucosum | Corynebacterium | aurimucosum | 2.07 | Corynebacterium | aurimucosum | 2.05 | A | 1 | 1 | 0 | Corynebacterium aurimucosum | Corynebacterium aurimucosum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.3 | Corynebacterium | diphtheriae | 2.26 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | belfantii | Corynebacterium | diphtheriae | 2.28 | Corynebacterium | diphtheriae | 2.28 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium | bovis | Corynebacterium | bovis | 2.35 | Corynebacterium | bovis | 2.33 | A | 1 | 1 | 0 | Corynebacterium bovis | Corynebacterium bovis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium | camporealensis | Corynebacterium | camporealensis | 1.96 | Corynebacterium | flavescens | 1.55 | B | 0 | 0 | 1 | Corynebacterium camporealensis | Corynebacterium flavescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium | casei | Corynebacterium | casei | 1.98 | Corynebacterium | stationis | 1.64 | B | 0 | 0 | 1 | Corynebacterium casei | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium | confusum | Corynebacterium | confusum | 1.78 | Corynebacterium | confusum | 1.68 | B | 0 | 0 | 1 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | confusum | Corynebacterium | confusum | 1.64 | Corynebacterium | confusum | 1.51 | B | 0 | 0 | 1 | Corynebacterium confusum | Corynebacterium confusum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium | cystitidis | Corynebacterium | cystitidis | 2.36 | Pseudomonas | abietaniphila | 1.37 | A | 1 | 1 | 0 | Corynebacterium cystitidis | Pseudomonas abietaniphila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.37 | Corynebacterium | diphtheriae | 2.36 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.45 | Corynebacterium | diphtheriae | 2.37 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | KL0173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.5 | Corynebacterium | diphtheriae | 2.41 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | KL0956 | 3; MLST | 0 | 0 | Corynebacterium | diphtheriae | Corynebacterium | diphtheriae | 2.04 | Corynebacterium | diphtheriae | 2 | A | 1 | 1 | 0 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 1.97 | Candida | tropicalis | 1.57 | B | 0 | 0 | 1 | Corynebacterium frankenforstense | Candida tropicalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | frankenforstense | Corynebacterium | frankenforstense | 2.48 | Streptococcus | pneumoniae | 1.29 | A | 1 | 1 | 0 | Corynebacterium frankenforstense | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium | freiburgense | Pseudomonas | gessardii | 1.27 | Inquilinus | limosus | 1.24 | B | 0 | 0 | 1 | Pseudomonas gessardii | Inquilinus limosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | hansenii | Corynebacterium | xerosis | 1.42 | Corynebacterium | amycolatum | 1.35 | B | 0 | 0 | 1 | Corynebacterium xerosis | Corynebacterium amycolatum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | hansenii | Corynebacterium | xerosis | 1.58 | Corynebacterium | hansenii | 1.53 | B | 0 | 0 | 1 | Corynebacterium xerosis | Corynebacterium hansenii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | humireducens | Corynebacterium | glutamicum | 1.44 | Gluconobacter | oxydans | 1.3 | B | 0 | 0 | 1 | Corynebacterium glutamicum | Gluconobacter oxydans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium | kutscheri | Corynebacterium | kutscheri | 2.56 | Corynebacterium | kutscheri | 2.37 | A | 1 | 1 | 0 | Corynebacterium kutscheri | Corynebacterium kutscheri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium | lactis | Corynebacterium | amycolatum | 1.46 | Tsukamurella | spumae | 1.3 | B | 0 | 0 | 1 | Corynebacterium amycolatum | Tsukamurella spumae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium | lactis | Phocaeicola | vulgatus | 1.33 | Pseudomonas | veronii | 1.31 | B | 0 | 0 | 1 | Phocaeicola vulgatus | Pseudomonas veronii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium | phoceense | Corynebacterium | phoceense | 2.19 | Corynebacterium | phoceense | 1.89 | A | 1 | 1 | 0 | Corynebacterium phoceense | Corynebacterium phoceense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | pilosum | Corynebacterium | pilosum | 1.78 | Burkholderia | cenocepacia | 1.32 | B | 0 | 0 | 1 | Corynebacterium pilosum | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.48 | Corynebacterium | pseudotuberculosis | 2.34 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.41 | Corynebacterium | pseudotuberculosis | 2.33 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20689 CVUAS | 3; type strain | DT | 2012-12-18 | Corynebacterium | pseudotuberculosis | Corynebacterium | pseudotuberculosis | 2.39 | Corynebacterium | pseudotuberculosis | 2.38 | A | 1 | 1 | 0 | Corynebacterium pseudotuberculosis | Corynebacterium pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C138 | 3; sequenced (rpoB) UGent | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.32 | Corynebacterium | diphtheriae | 2.29 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.09 | Corynebacterium | diphtheriae | 2.01 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3559,2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.16 | Corynebacterium | diphtheriae | 2.15 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium | rouxii | Corynebacterium | diphtheriae | 2.24 | Corynebacterium | diphtheriae | 2.07 | A | 0 | 0 | 1 | Corynebacterium diphtheriae | Corynebacterium diphtheriae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.28 | Corynebacterium | ulcerans | 1.94 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium | silvaticum | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.04 | A | 0 | 0 | 1 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium | sp-208-U2-6 | Sphingomonas | yabuuchiae | 1.39 | Acidiphilium | acidophilum | 1.38 | B | 0 | 0 | 1 | Sphingomonas yabuuchiae | Acidiphilium acidophilum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32165,2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium | sp-2300500 | Acinetobacter | baumannii | 1.26 | Lysinibacillus | endophyticus | 1.22 | B | 0 | 0 | 1 | Acinetobacter baumannii | Lysinibacillus endophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5775,2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium | sp-2300500 | Acinetobacter | baumannii | 1.29 | Agromyces | fucosus | 1.28 | B | 0 | 0 | 1 | Acinetobacter baumannii | Agromyces fucosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2554,2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium | sp-2396 | Corynebacterium | sp | 1.92 | Arthrobacter | pyridinolis | 1.35 | B | 0 | 0 | 1 | Corynebacterium sp | Arthrobacter pyridinolis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium | sp-243-z-3 | Corynebacterium | resistens | 1.78 | Corynebacterium | resistens | 1.76 | B | 0 | 0 | 1 | Corynebacterium resistens | Corynebacterium resistens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium | sp-257-y-4 | Proteus | mirabilis | 1.36 | Ligilactobacillus | agilis | 1.27 | B | 0 | 0 | 1 | Proteus mirabilis | Ligilactobacillus agilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium | sp-284z-2 | Proteus | mirabilis | 1.3 | Loigolactobacillus | bifermentans | 1.3 | B | 0 | 0 | 1 | Proteus mirabilis | Loigolactobacillus bifermentans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium | sp-CVUAS-11664 | Pseudomonas | aeruginosa | 1.32 | Liquorilactobacillus | satsumensis | 1.24 | B | 0 | 0 | 1 | Pseudomonas aeruginosa | Liquorilactobacillus satsumensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium | sp-CVUAS-1888 | Corynebacterium | striatum | 1.3 | Pseudomonas | aeruginosa | 1.28 | B | 0 | 0 | 1 | Corynebacterium striatum | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1721,3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium | sp-CVUAS-2484.4 | Streptococcus | gallolyticus | 1.5 | Corynebacterium | felinum | 1.41 | B | 0 | 0 | 1 | Streptococcus gallolyticus | Corynebacterium felinum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2484,4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388; 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium | sp-CVUAS-2484.4 | Pseudomonas | citronellolis | 1.29 | Lacticaseibacillus | paracasei | 1.27 | B | 0 | 0 | 1 | Pseudomonas citronellolis | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | falsenii | 1.42 | Corynebacterium | sp | 1.32 | B | 0 | 0 | 1 | Corynebacterium falsenii | Corynebacterium sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-CVUAS-30404 | Corynebacterium | sp | 1.57 | Arthrobacter | psychrolactophilus | 1.36 | B | 0 | 0 | 1 | Corynebacterium sp | Arthrobacter psychrolactophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium | sp-CVUAS-31533 | Corynebacterium | amycolatum | 1.45 | Paeniglutamicibacter | sulfureus | 1.38 | B | 0 | 0 | 1 | Corynebacterium amycolatum | Paeniglutamicibacter sulfureus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32278,2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium | sp-CVUAS-32278.2 | Agromyces | cerinus | 1.2 | Listeria | grayi | 1.17 | B | 0 | 0 | 1 | Agromyces cerinus | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium | sp-CVUAS-874,4 | Clostridium | beijerinckii | 1.38 | Staphylococcus | carnosus | 1.34 | B | 0 | 0 | 1 | Clostridium beijerinckii | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 874,4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium | sp-CVUAS-874,4 | Neobacillus | bataviensis | 1.3 | Paucilactobacillus | suebicus | 1.29 | B | 0 | 0 | 1 | Neobacillus bataviensis | Paucilactobacillus suebicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 21TRB0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium | sp-LALLF-21TRB0553 | Cryptococcus | neoformans | 1.38 | Corynebacterium | flavescens | 1.37 | B | 0 | 0 | 1 | Cryptococcus neoformans | Corynebacterium flavescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium | sp-LHL-0497 | Haemophilus | influenzae | 1.26 | Staphylococcus | epidermidis | 1.17 | B | 0 | 0 | 1 | Haemophilus influenzae | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium | sp-LHL-0558 | Nocardia | cyriacigeorgica | 1.32 | Glutamicibacter | protophormiae | 1.3 | B | 0 | 0 | 1 | Nocardia cyriacigeorgica | Glutamicibacter protophormiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium | sp-LHL-1322 | Agromyces | humatus | 1.41 | Corynebacterium | xerosis | 1.29 | B | 0 | 0 | 1 | Agromyces humatus | Corynebacterium xerosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium | stationis | Corynebacterium | stationis | 2.27 | Corynebacterium | stationis | 2.21 | A | 1 | 1 | 0 | Corynebacterium stationis | Corynebacterium stationis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium | testudinoris | Bacteroides | fragilis | 1.29 | Citrobacter | koseri | 1.28 | B | 0 | 0 | 1 | Bacteroides fragilis | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 2823 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.35 | Corynebacterium | ulcerans | 2.31 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.49 | Corynebacterium | ulcerans | 2.36 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.34 | Corynebacterium | ulcerans | 2.29 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.25 | Corynebacterium | ulcerans | 2.2 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.42 | Corynebacterium | ulcerans | 2.28 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6522,4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium | ulcerans | Corynebacterium | ulcerans | 2.26 | Corynebacterium | ulcerans | 2.23 | A | 1 | 1 | 0 | Corynebacterium ulcerans | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5437,3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium | variabile | Corynebacterium | variabile | 1.62 | Blastomonas | ursincola | 1.34 | B | 0 | 0 | 1 | Corynebacterium variabile | Blastomonas ursincola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium | xerosis | Corynebacterium | xerosis | 1.57 | Levilactobacillus | brevis | 1.35 | B | 0 | 0 | 1 | Corynebacterium xerosis | Levilactobacillus brevis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium | xerosis-group | Corynebacterium | xerosis | 1.58 | Aromatoleum | bremensis | 1.45 | B | 0 | 0 | 1 | Corynebacterium xerosis | Aromatoleum bremensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2379,2 | 3; sequenced | EFex | 2021-04-08 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.04 | Gordonia | paraffinivorans | 2 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia | paraffinivorans | Gordonia | paraffinivorans | 2.14 | Gordonia | paraffinivorans | 1.92 | A | 1 | 1 | 0 | Gordonia paraffinivorans | Gordonia paraffinivorans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia | rubripertincta | Gordonia | rubripertincta | 2.25 | Gordonia | rubripertincta | 2.17 | A | 1 | 1 | 0 | Gordonia rubripertincta | Gordonia rubripertincta | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10185.3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium | smegmatis | Mycobacterium | smegmatis | 1.48 | Mycobacterium | smegmatis | 1.42 | B | 0 | 0 | 1 | Mycobacterium smegmatis | Mycobacterium smegmatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31296,3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia | cyriacigeorgica | Nocardia | cyriacigeorgica | 2.45 | Nocardia | cyriacigeorgica | 2.28 | A | 1 | 1 | 0 | Nocardia cyriacigeorgica | Nocardia cyriacigeorgica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia | farcinica | Nocardia | farcinica | 2.36 | Nocardia | kroppenstedtii | 2.26 | B | 0 | 0 | 1 | Nocardia farcinica | Nocardia kroppenstedtii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.25 | Rhodococcus | hoagii | 2.22 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus | equi | Rhodococcus | hoagii | 2.19 | Rhodococcus | hoagii | 2.04 | A | 0 | 0 | 1 | Rhodococcus hoagii | Rhodococcus hoagii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1447,3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella | sp | Tsukamurella | tyrosinosolvens | 1.54 | Tsukamurella | paurometabola | 1.54 | B | 0 | 0 | 1 | Tsukamurella tyrosinosolvens | Tsukamurella paurometabola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2895,3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium | luteolum | Brevibacterium | luteolum | 1.92 | Brevibacterium | luteolum | 1.71 | B | 0 | 0 | 1 | Brevibacterium luteolum | Brevibacterium luteolum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas | flavigena | Bacteroides | fragilis | 1.43 | Neisseria | meningitidis | 1.35 | B | 0 | 0 | 1 | Bacteroides fragilis | Neisseria meningitidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium | sp-234z-2 | Glutamicibacter | protophormiae | 1.31 | Phocaeicola | vulgatus | 1.29 | B | 0 | 0 | 1 | Glutamicibacter protophormiae | Phocaeicola vulgatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium | sp-CVUAS-30560 | Raoultella | ornithinolytica | 1.39 | Glutamicibacter | ardleyensis | 1.31 | B | 0 | 0 | 1 | Raoultella ornithinolytica | Glutamicibacter ardleyensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus | sedentarius | Kytococcus | sedentarius | 2.1 | Kytococcus | sedentarius | 1.84 | A | 1 | 1 | 0 | Kytococcus sedentarius | Kytococcus sedentarius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus | congolensis | Dermatophilus | congolensis | 1.62 | Dermatophilus | congolensis | 1.4 | B | 0 | 0 | 1 | Dermatophilus congolensis | Dermatophilus congolensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus | congolensis | Dermatophilus | congolensis | 1.58 | Chryseobacterium | arthrosphaerae | 1.39 | B | 0 | 0 | 1 | Dermatophilus congolensis | Chryseobacterium arthrosphaerae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium | flaccumfaciens | Curtobacterium | flaccumfaciens | 2.42 | Curtobacterium | flaccumfaciens | 2.31 | A | 1 | 1 | 0 | Curtobacterium flaccumfaciens | Curtobacterium flaccumfaciens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia | sp-MUVA-A-508948 | Staphylococcus | aureus | 1.46 | Clostridium | chauvoei | 1.45 | B | 0 | 0 | 1 | Staphylococcus aureus | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium | sp | Microbacterium | phyllosphaerae | 1.84 | Microbacterium | foliorum | 1.84 | B | 0 | 0 | 1 | Microbacterium phyllosphaerae | Microbacterium foliorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10353 DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.09.2011 L8 | Glutamicibacter | nicotianae | Lactobacillus | pentosus | 1.33 | Burkholderia | cepacia | 1.26 | B | 0 | 0 | 1 | Lactobacillus pentosus | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMET 10432 DSM 20127 | 3; type strain | unknown | 08.09.2011, L13, [0034] | Sinomonas | atrocyanea | Sinomonas | atrocyanea | 2.52 | Sinomonas | atrocyanea | 2.24 | A | 1 | 1 | 0 | Sinomonas atrocyanea | Sinomonas atrocyanea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter | protophormiae | Lactobacillus | sakei | 1.37 | Loigolactobacillus | coryniformis | 1.31 | B | 0 | 0 | 1 | Lactobacillus sakei | Loigolactobacillus coryniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria | atrinae | Kocuria | atrinae | 2 | Kocuria | carniphila | 1.83 | A | 1 | 1 | 0 | Kocuria atrinae | Kocuria carniphila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 56a-5 | 3; sequenced | EFEx | 2020-05-20 | Kocuria | gwangalensis | Kocuria | carniphila | 1.64 | Kocuria | carniphila | 1.59 | B | 0 | 0 | 1 | Kocuria carniphila | Kocuria carniphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2014 E22 | Kocuria | palustris | Kocuria | palustris | 2.38 | Kocuria | palustris | 2.33 | A | 1 | 1 | 0 | Kocuria palustris | Kocuria palustris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015 K12 | Micrococcus | luteus | Micrococcus | luteus | 2.37 | Micrococcus | luteus | 2.24 | A | 1 | 1 | 0 | Micrococcus luteus | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20032 | 3; type strain | eDT | 2015-07-09 | Rothia | kristinae | Rothia | kristinae | 2.26 | Rothia | kristinae | 2.25 | A | 1 | 1 | 0 | Rothia kristinae | Rothia kristinae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium | cellulans | Cellulosimicrobium | cellulans | 2.05 | Cellulosimicrobium | cellulans | 1.97 | A | 1 | 1 | 0 | Cellulosimicrobium cellulans | Cellulosimicrobium cellulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 749,2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter | inulinus | Sanguibacter | inulinus | 2.46 | Sanguibacter | keddieii | 1.76 | A | 1 | 1 | 0 | Sanguibacter inulinus | Sanguibacter keddieii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1079 | 3; public strain collection | EFEx | 2017-01-20 | Penicillium | roqueforti | Penicillium | camemberti | 1.55 | Staphylococcus | warneri | 1.49 | B | 0 | 0 | 1 | Penicillium camemberti | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10354,2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | tonsurans | 1.38 | Limosilactobacillus | fermentum | 1.25 | B | 0 | 0 | 1 | Trichophyton tonsurans | Limosilactobacillus fermentum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10357.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton | verrucosum | Trichophyton | interdigitale | 1.38 | Trichophyton | mentagrophytes | 1.33 | B | 0 | 0 | 1 | Trichophyton interdigitale | Trichophyton mentagrophytes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1386 | 3; public strain collection | Fextr | 2016-03-10 | Candida | albicans | Candida | albicans | 2.2 | Candida | albicans | 2.14 | A | 1 | 1 | 0 | Candida albicans | Candida albicans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IGV 011 IGV | 3; sequenced | EFEx | 2016-12-08 | Sarocladium | strictum | Sarocladium | strictum | 1.54 | Sarocladium | strictum | 1.41 | B | 0 | 0 | 1 | Sarocladium strictum | Sarocladium strictum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2520.4 | 3; sequenced (ITS) | F-Extr | 2018-07-26 | Fusarium | sambucinum | Kodamaea | ohmeri | 1.23 | Lentilactobacillus | parabuchneri | 1.2 | B | 0 | 0 | 1 | Kodamaea ohmeri | Lentilactobacillus parabuchneri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus | amylolyticus | Paenibacillus | amylolyticus | 2.26 | Paenibacillus | amylolyticus | 2.25 | A | 1 | 1 | 0 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus | amylolyticus | Paenibacillus | amylolyticus | 2.17 | Paenibacillus | amylolyticus | 2.12 | A | 1 | 1 | 0 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.53 | Paenibacillus | barengoltzii | 2.5 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus | barengoltzii | Paenibacillus | barengoltzii | 2.43 | Paenibacillus | barengoltzii | 2.4 | A | 1 | 1 | 0 | Paenibacillus barengoltzii | Paenibacillus barengoltzii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.18 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus | glucanolyticus | Paenibacillus | glucanolyticus | 2.25 | Paenibacillus | lautus | 1.78 | A | 1 | 1 | 0 | Paenibacillus glucanolyticus | Paenibacillus lautus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus | larvae | Paenibacillus | larvae | 2.32 | Paenibacillus | larvae | 2.28 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.33 | Paenibacillus | larvae | 2.24 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.33 | Paenibacillus | larvae | 2.24 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus | larvae | Paenibacillus | larvae | 2.18 | Paenibacillus | larvae | 2.12 | A | 1 | 1 | 0 | Paenibacillus larvae | Paenibacillus larvae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus | macerans | Paenibacillus | macerans | 2.28 | Paenibacillus | macerans | 2.07 | A | 1 | 1 | 0 | Paenibacillus macerans | Paenibacillus macerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.12 | Paenibacillus | pabuli | 2.02 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus | pabuli | Paenibacillus | tylopili | 2.22 | Paenibacillus | pabuli | 2.15 | B | 0 | 0 | 1 | Paenibacillus tylopili | Paenibacillus pabuli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus | pasadenensis | Paenibacillus | pasadenensis | 2.52 | Paenibacillus | pasadenensis | 2.05 | A | 1 | 1 | 0 | Paenibacillus pasadenensis | Paenibacillus pasadenensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus | validus | Paenibacillus | validus | 2.26 | Paenibacillus | validus | 2.16 | A | 1 | 1 | 0 | Paenibacillus validus | Paenibacillus validus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CDH 260-11 | 3; public strain collection | DT | 2017-12-06 | Clostridioides | difficile | Clostridioides | difficile | 2.14 | Clostridioides | difficile | 1.99 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium | barati | Clostridium | budayi | 2.31 | Clostridium | baratii | 2.28 | B | 0 | 0 | 1 | Clostridium budayi | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium | butyricum | Clostridium | butyricum | 2.61 | Clostridium | butyricum | 2.48 | A | 1 | 1 | 0 | Clostridium butyricum | Clostridium butyricum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium | chauvoei | Clostridium | chauvoei | 2.13 | Clostridium | chauvoei | 2.01 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017156-1 | 3; FLI | DT | 2018-11-27 | Clostridium | chauvoei | Clostridium | chauvoei | 2.28 | Clostridium | chauvoei | 2.17 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017156-2 | 3; FLI | DT | 2018-11-27 | Clostridium | chauvoei | Clostridium | chauvoei | 2.2 | Clostridium | chauvoei | 2.14 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017156-3 | 3; FLI | DT | 2018-11-27 | Clostridium | chauvoei | Clostridium | chauvoei | 2.14 | Clostridium | chauvoei | 2.13 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017156-4 | 3; FLI | DT | 2018-11-27 | Clostridium | chauvoei | Clostridium | chauvoei | 2.11 | Clostridium | chauvoei | 2.09 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002729_1 coch | 2; 0 | DT | 2017-02-28 | Clostridium | cochlearium | Clostridium | cochlearium | 2.51 | Clostridium | cochlearium | 2.5 | A | 1 | 1 | 0 | Clostridium cochlearium | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium | colicanis | Clostridium | colicanis | 2.29 | Bifidobacterium | boum | 1.42 | A | 1 | 1 | 0 | Clostridium colicanis | Bifidobacterium boum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium | novyi | Clostridium | novyi | 2.03 | Clostridium | novyi | 1.96 | A | 1 | 1 | 0 | Clostridium novyi | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium | septicum | Clostridium | septicum | 2.44 | Clostridium | septicum | 2.26 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium | sporogenes | Clostridium | sporogenes | 2.34 | Clostridium | sporogenes | 2.21 | A | 1 | 1 | 0 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141015223 aerob | 2; 0 | DT | 2014-11-27 | Clostridium | tertium | Clostridium | tertium | 2.44 | Clostridium | tertium | 2.22 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161002486-2 aerob | 2; 0 | DT | 2016-03-01 | Clostridium | tertium | Clostridium | tertium | 2.47 | Clostridium | tertium | 2.34 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003180 aerob | 2; 0 | DT | 2018-06-30 | Clostridium | tertium | Clostridium | tertium | 2.36 | Clostridium | tertium | 2.3 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium | tetani | Clostridium | tetani | 2.23 | Clostridium | tetani | 2.16 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya | limosa | Hathewaya | limosa | 2.29 | Hathewaya | limosa | 2.25 | A | 1 | 1 | 0 | Hathewaya limosa | Hathewaya limosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium | sordellii | Paeniclostridium | sordellii | 2.32 | Paeniclostridium | sordellii | 2.27 | A | 1 | 1 | 0 | Paeniclostridium sordellii | Paeniclostridium sordellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium | bifermentans | Paraclostridium | bifermentans | 2.02 | Paraclostridium | bifermentans | 1.92 | A | 1 | 1 | 0 | Paraclostridium bifermentans | Paraclostridium bifermentans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 44 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N11 | Bacillus | badius | Bacillus | badius | 2.06 | Bacillus | badius | 1.99 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.03.2017 I18 | Bacillus | ectoiniformans | Streptococcus | dysgalactiae | 1.49 | Bacillus | licheniformis | 1.28 | B | 0 | 0 | 1 | Streptococcus dysgalactiae | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F22 | Cytobacillus | firmus | Cytobacillus | firmus | 2.45 | Cytobacillus | firmus | 2.24 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 37 B 88 | 3; reference strain collection | 0 | 09.02.2015, L16 | Cytobacillus | firmus | Cytobacillus | firmus | 2.06 | Cytobacillus | firmus | 1.98 | A | 1 | 1 | 0 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N14 | Niallia | circulans | Niallia | circulans | 1.98 | Niallia | circulans | 1.88 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 N5 | Virgibacillus | pantothenticus | Virgibacillus | pantothenticus | 2.01 | Virgibacillus | pantothenticus | 1.84 | A | 1 | 1 | 0 | Virgibacillus pantothenticus | Virgibacillus pantothenticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31071,2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 1.69 | Alkalihalobacillus | clausii | 1.58 | B | 0 | 0 | 1 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015; H10 | Alkalihalobacillus | clausii | Alkalihalobacillus | clausii | 1.71 | Alkalihalobacillus | clausii | 1.7 | B | 0 | 0 | 1 | Alkalihalobacillus clausii | Alkalihalobacillus clausii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 13.02.2015, K10, [0034] | Bacillus | alveayuensis | Bacillus | alveayuensis | 2.52 | Peribacillus | psychrosaccharolyticus | 1.35 | A | 1 | 1 | 0 | Bacillus alveayuensis | Peribacillus psychrosaccharolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7 RKI 2015 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCA/TA2 (A) | 16.02.2015, K16, [0034] | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.24 | Bacillus | amyloliquefaciens | 1.99 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7 CVUAS | 3; type strain | DT | 2014-01-29 | Bacillus | amyloliquefaciens | Bacillus | amyloliquefaciens | 2.15 | Bacillus | amyloliquefaciens | 2.15 | A | 1 | 1 | 0 | Bacillus amyloliquefaciens | Bacillus amyloliquefaciens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E06 | Bacillus | anthracis | Bacillus | cereus | 2.01 | Bacillus | cereus | 1.96 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; E04 | Bacillus | anthracis | Bacillus | cereus | 2.19 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 004 | 2; 0 | DT | 2014-01-22 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.14 | Bacillus | atrophaeus | 2.14 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 L8 | Bacillus | atrophaeus | Bacillus | atrophaeus | 2.43 | Bacillus | atrophaeus | 2.36 | A | 1 | 1 | 0 | Bacillus atrophaeus | Bacillus atrophaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H20 | Bacillus | badius | Bacillus | badius | 2.45 | Bacillus | badius | 2.26 | A | 1 | 1 | 0 | Bacillus badius | Bacillus badius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus | cereus | Bacillus | cereus | 2.05 | Bacillus | thuringiensis | 2.02 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus thuringiensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1107 | 2; 0 | eDT | 27.01.2016 | Bacillus | cereus | Bacillus | cereus | 2.41 | Bacillus | cereus | 2.38 | A | 1 | 1 | 0 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.05.2015; E09 | Bacillus | tropicus | Bacillus | cereus | 2.28 | Bacillus | cereus | 2.19 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2833 CVUAS | 4; 16S, draft genome [0009] | eDT | 2016-01-27 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.48 | Bacillus | cytotoxicus | 2.45 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NVH 391-98 | 3; type strain | eDT | 2016-11-16 | Bacillus | cytotoxicus | Bacillus | cytotoxicus | 2.45 | Bacillus | cytotoxicus | 2.26 | A | 1 | 1 | 0 | Bacillus cytotoxicus | Bacillus cytotoxicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCEB 627 | 3; type strain | DT | 2014-01-22 | Bacillus | firmus | Cytobacillus | firmus | 2.19 | Cytobacillus | firmus | 2.13 | A | 0 | 0 | 1 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 37 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Bacillus | firmus | Cytobacillus | firmus | 2.33 | Cytobacillus | firmus | 2.13 | A | 0 | 0 | 1 | Cytobacillus firmus | Cytobacillus firmus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.11.2017 F14 | Bacillus | halotolerans | Bacillus | mojavensis | 1.96 | Bacillus | mojavensis | 1.83 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.12.2017 I8 | Bacillus | halotolerans | Bacillus | mojavensis | 2.23 | Bacillus | subtilis | 2.1 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.12.2016 [H3] | Bacillus | inaquosorum | Bacillus | subtilis | 2.27 | Bacillus | subtilis | 2.12 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.02.2016 C20 | Bacillus | infantis | Bacillus | infantis | 2.46 | Bacillus | infantis | 2.39 | A | 1 | 1 | 0 | Bacillus infantis | Bacillus infantis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus | licheniformis | Bacillus | licheniformis | 2.17 | Bacillus | licheniformis | 2.15 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 13 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus | licheniformis | Bacillus | licheniformis | 2.27 | Bacillus | licheniformis | 1.81 | A | 1 | 1 | 0 | Bacillus licheniformis | Bacillus licheniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K15 | Bacillus | mojavensis | Bacillus | subtilis | 2.13 | Bacillus | subtilis | 2.08 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F4 | Bacillus | mojavensis | Bacillus | mojavensis | 2.33 | Bacillus | subtilis | 2.27 | B | 0 | 0 | 1 | Bacillus mojavensis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2048 CVUAS | 3; type strain | eDT | 09.12.2015 | Bacillus | mycoides | Bacillus | mycoides | 2.22 | Bacillus | mycoides | 2.08 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus | mycoides | Bacillus | mycoides | 2.04 | Bacillus | mycoides | 2.01 | A | 1 | 1 | 0 | Bacillus mycoides | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 11778 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E8, [0034] | Bacillus | pacificus | Bacillus | cereus | 1.92 | Bacillus | mycoides | 1.62 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus | pseudomycoides | Bacillus | cereus | 2.43 | Bacillus | cereus | 2.4 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 12443 CVUAS | 3; public strain collection | eDT | 09.12.2015 | Bacillus | pseudomycoides | Bacillus | mycoides | 2.13 | Bacillus | cereus | 2.13 | B | 0 | 0 | 1 | Bacillus mycoides | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B017 CVUAS | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus | pumilus | Bacillus | pumilus | 2.15 | Bacillus | pumilus | 1.96 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 340 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | pumilus | Bacillus | pumilus | 2.04 | Bacillus | pumilus | 1.95 | A | 1 | 1 | 0 | Bacillus pumilus | Bacillus pumilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.05.2018, N12 | Bacillus | siamensis | Bacillus | subtilis | 1.84 | Bacillus | mojavensis | 1.8 | B | 0 | 0 | 1 | Bacillus subtilis | Bacillus mojavensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2016, K18 | Bacillus | sonorensis | Bacillus | sonorensis | 1.95 | Listeria | grayi | 1.56 | B | 0 | 0 | 1 | Bacillus sonorensis | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 12369 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Bacillus | sp-DSM-12369 | Bacillus | sonorensis | 1.53 | Thauera | mechernichensis | 1.39 | B | 0 | 0 | 1 | Bacillus sonorensis | Thauera mechernichensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ATCC 6633 CVUAS | 3; public strain collection | eDT | 2015-07-22 | Bacillus | spizizenii | Bacillus | subtilis | 2.08 | Bacillus | subtilis | 1.96 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 618 | 4; public strain collection | eDT | 27.01.2016 | Bacillus | spizizenii | Bacillus | subtilis | 2.09 | Bacillus | subtilis | 2.04 | A | 0 | 0 | 1 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 110 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.33 | Bacillus | subtilis | 2.15 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 1092 | 3; reference strain collection | DT | 2014-01-22 | Bacillus | subtilis | Bacillus | subtilis | 2.23 | Bacillus | subtilis | 2.2 | A | 1 | 1 | 0 | Bacillus subtilis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 2.17 | Bacillus | cereus | 2.14 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 2.12 | Bacillus | cereus | 2.04 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus | thuringiensis | Bacillus | cereus | 2.03 | Bacillus | cereus | 1.97 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 2.24 | Bacillus | cereus | 2.19 | A | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus | thuringiensis | Bacillus | cereus | 1.99 | Bacillus | cereus | 1.95 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6890 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2015, L4 | Bacillus | thuringiensis | Bacillus | thuringiensis | 2.18 | Bacillus | cereus | 2.04 | B | 0 | 0 | 1 | Bacillus thuringiensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 07.11.2006 K4 | Bacillus | thuringiensis | Bacillus | cereus | 1.75 | Bacillus | cereus | 1.7 | B | 0 | 0 | 1 | Bacillus cereus | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2786,2 | 3; sequenced | DT | 2013-07-31 | Bacillus | timorensis | Bacillus | timonensis | 1.49 | Bacillus | smithii | 1.38 | B | 0 | 0 | 1 | Bacillus timonensis | Bacillus smithii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.07.2017, H17 | Bacillus | vallismortis | Bacillus | vallismortis | 2.37 | Bacillus | subtilis | 1.89 | A | 1 | 1 | 0 | Bacillus vallismortis | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.02.2016 K19 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.11 | Cytobacillus | horneckiae | 2.01 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.04.2014 H5 | Cytobacillus | horneckiae | Cytobacillus | horneckiae | 2.32 | Cytobacillus | horneckiae | 2.18 | A | 1 | 1 | 0 | Cytobacillus horneckiae | Cytobacillus horneckiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus | kochii | Cytobacillus | kochii | 1.86 | Cytobacillus | kochii | 1.83 | B | 0 | 0 | 1 | Cytobacillus kochii | Cytobacillus kochii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 19.02.2015, E18 | Domibacillus | aminovorans | Pseudomonas | aeruginosa | 1.37 | Micrococcus | luteus | 1.35 | B | 0 | 0 | 1 | Pseudomonas aeruginosa | Micrococcus luteus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 15.02.2015, E8 | Domibacillus | aminovorans | Staphylococcus | lugdunensis | 1.46 | Lodderomyces | elongisporus | 1.45 | B | 0 | 0 | 1 | Staphylococcus lugdunensis | Lodderomyces elongisporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.02.2015 H5 | Evansella | clarkii | Bacillus | clarkii | 2.35 | Bacillus | mycoides | 1.45 | A | 0 | 0 | 1 | Bacillus clarkii | Bacillus mycoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F7 | Evansella | clarkii | Bacillus | clarkii | 2.04 | Alkalihalobacillus | krulwichiae | 1.46 | A | 0 | 0 | 1 | Bacillus clarkii | Alkalihalobacillus krulwichiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus | stearothermophilus | Geobacillus | stearothermophilus | 2.39 | Geobacillus | stearothermophilus | 2.29 | A | 1 | 1 | 0 | Geobacillus stearothermophilus | Geobacillus stearothermophilus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 I10; (single sp is included in the corresponding msp) | Halalkalihalobacillus | halodurans | Alkalihalobacillus | halodurans | 2.19 | Alkalihalobacillus | okuhidensis | 1.98 | A | 0 | 0 | 1 | Alkalihalobacillus halodurans | Alkalihalobacillus okuhidensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 E19; (single sp is included in the corresponding msp) | Halalkalihalobacillus | halodurans | Alkalihalobacillus | okuhidensis | 2.14 | Alkalihalobacillus | halodurans | 1.92 | A | 0 | 0 | 1 | Alkalihalobacillus okuhidensis | Alkalihalobacillus halodurans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia | sporothermodurans | Heyndrickxia | sporothermodurans | 1.34 | Acidovorax | citrulli | 1.3 | B | 0 | 0 | 1 | Heyndrickxia sporothermodurans | Acidovorax citrulli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 10599 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015, H9 | Heyndrickxia | sporothermodurans | Heyndrickxia | sporothermodurans | 2.35 | Paracoccus | versutus | 1.32 | A | 1 | 1 | 0 | Heyndrickxia sporothermodurans | Paracoccus versutus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017; N12 | Lederbergia | lenta | Lederbergia | lenta | 2.43 | Lederbergia | lenta | 2.33 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 9 CVUAS | 3; type strain | DT | 2014-01-22 | Lederbergia | lenta | Lederbergia | lenta | 2.22 | Lederbergia | lenta | 2.08 | A | 1 | 1 | 0 | Lederbergia lenta | Lederbergia lenta | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus | bataviensis | Neobacillus | bataviensis | 2.29 | Staphylococcus | cohnii | 1.41 | A | 1 | 1 | 0 | Neobacillus bataviensis | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017 N13 | Neobacillus | cucumis | Neobacillus | bataviensis | 2.02 | Bacillus | cereus | 1.54 | A | 0 | 0 | 1 | Neobacillus bataviensis | Bacillus cereus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E14, [0034] | Niallia | circulans | Niallia | circulans | 1.9 | Niallia | circulans | 1.65 | B | 0 | 0 | 1 | Niallia circulans | Niallia circulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2582,2 | 3; sequenced | eDT | 2015-09-23 | Niallia | nealsonii | Niallia | taxi | 2.04 | Citrobacter | freundii | 1.42 | A | 0 | 0 | 1 | Niallia taxi | Citrobacter freundii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K20 | Peribacillus | muralis | Peribacillus | muralis | 2.38 | Peribacillus | simplex | 1.81 | A | 1 | 1 | 0 | Peribacillus muralis | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016; K20 | Peribacillus | psychrosaccharolyticus | Peribacillus | psychrosaccharolyticus | 2.21 | Staphylococcus | simiae | 1.44 | A | 1 | 1 | 0 | Peribacillus psychrosaccharolyticus | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus | simplex | Peribacillus | simplex | 1.67 | Schleiferilactobacillus | perolens | 1.45 | B | 0 | 0 | 1 | Peribacillus simplex | Schleiferilactobacillus perolens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2016 H13 | Peribacillus | simplex | Peribacillus | simplex | 1.78 | Peribacillus | simplex | 1.76 | B | 0 | 0 | 1 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 15.12.2016, K15, [0096] | Peribacillus | simplex | Peribacillus | simplex | 2.19 | Peribacillus | simplex | 1.64 | A | 1 | 1 | 0 | Peribacillus simplex | Peribacillus simplex | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016 K15 | Priestia | flexa | Priestia | flexa | 2.54 | Priestia | flexa | 2.17 | A | 1 | 1 | 0 | Priestia flexa | Priestia flexa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H18 | Priestia | flexa | Priestia | megaterium | 2.01 | Priestia | megaterium | 1.8 | A | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015 H21 | Priestia | megaterium | Priestia | megaterium | 1.92 | Priestia | megaterium | 1.61 | B | 0 | 0 | 1 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia | megaterium | Priestia | megaterium | 2.06 | Priestia | megaterium | 2.06 | A | 1 | 1 | 0 | Priestia megaterium | Priestia megaterium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E3 | Salibacillus | salarius | Salibacterium | salarium | 2.05 | Aeromonas | salmonicida | 1.35 | A | 0 | 0 | 1 | Salibacterium salarium | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F13 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.13 | Sutcliffiella | cohnii | 1.86 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 K8 | Sutcliffiella | cohnii | Sutcliffiella | cohnii | 2.03 | Sutcliffiella | cohnii | 1.79 | A | 1 | 1 | 0 | Sutcliffiella cohnii | Sutcliffiella cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus | saccharophilus | Lodderomyces | elongisporus | 1.45 | Liquorilactobacillus | nagelii | 1.34 | B | 0 | 0 | 1 | Lodderomyces elongisporus | Liquorilactobacillus nagelii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus | proomii | Virgibacillus | proomii | 2.02 | Virgibacillus | proomii | 1.99 | A | 1 | 1 | 0 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3616,2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus | proomii | Virgibacillus | proomii | 1.82 | Virgibacillus | proomii | 1.71 | B | 0 | 0 | 1 | Virgibacillus proomii | Virgibacillus proomii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H15 | Weizmannia | coagulans | Weizmannia | coagulans | 2.45 | Weizmannia | coagulans | 2.12 | A | 1 | 1 | 0 | Weizmannia coagulans | Weizmannia coagulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 71,2 | 3; sequenced | DT | 2013-12-13 | Weizmannia | ginsengihumi | Weizmannia | ginsengihumi | 2.19 | Weizmannia | ginsengihumi | 2.06 | A | 1 | 1 | 0 | Weizmannia ginsengihumi | Weizmannia ginsengihumi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.47 | Kurthia | gibsonii | 2.39 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia | gibsonii | Kurthia | gibsonii | 2.27 | Kurthia | gibsonii | 2.26 | A | 1 | 1 | 0 | Kurthia gibsonii | Kurthia gibsonii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia | zopfii | Kurthia | zopfii | 1.66 | Kurthia | zopfii | 1.61 | B | 0 | 0 | 1 | Kurthia zopfii | Kurthia zopfii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus | fusiformis | Lysinibacillus | fusiformis | 2.34 | Lysinibacillus | fusiformis | 2.23 | A | 1 | 1 | 0 | Lysinibacillus fusiformis | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus | fusiformis | Lysinibacillus | sphaericus | 2.36 | Lysinibacillus | fusiformis | 2.09 | B | 0 | 0 | 1 | Lysinibacillus sphaericus | Lysinibacillus fusiformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7192,2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus | macroides | Lysinibacillus | xylanilyticus | 2.23 | Lysinibacillus | boronitolerans | 2.18 | B | 0 | 0 | 1 | Lysinibacillus xylanilyticus | Lysinibacillus boronitolerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.34 | Lysinibacillus | sphaericus | 2.31 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus | sphaericus | Lysinibacillus | sphaericus | 2.11 | Lysinibacillus | sphaericus | 1.96 | A | 1 | 1 | 0 | Lysinibacillus sphaericus | Lysinibacillus sphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus | saudimassiliensis | Lysinibacillus | fusiformis | 1.48 | Ureibacillus | thermosphaericus | 1.37 | B | 0 | 0 | 1 | Lysinibacillus fusiformis | Ureibacillus thermosphaericus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus | odysseyi | Ureibacillus | sinduriensis | 1.47 | Lysinibacillus | sp[2] | 1.4 | B | 0 | 0 | 1 | Ureibacillus sinduriensis | Lysinibacillus sp[2] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium | acetylicum | Exiguobacterium | indicum | 1.82 | Exiguobacterium | soli | 1.6 | B | 0 | 0 | 1 | Exiguobacterium indicum | Exiguobacterium soli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium | enclense | Exiguobacterium | soli | 1.68 | Exiguobacterium | sp[2] | 1.33 | B | 0 | 0 | 1 | Exiguobacterium soli | Exiguobacterium sp[2] | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium | profundum | Exiguobacterium | sp[3] | 2.41 | Exiguobacterium | sp[3] | 2.36 | A | 0 | 0 | 1 | Exiguobacterium sp[3] | Exiguobacterium sp[3] | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.32 | Brochothrix | thermosphacta | 2.24 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix | thermosphacta | Brochothrix | thermosphacta | 2.39 | Brochothrix | thermosphacta | 2.37 | A | 1 | 1 | 0 | Brochothrix thermosphacta | Brochothrix thermosphacta | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria | aquatica | Listeria | aquatica | 2.26 | Listeria | grayi | 1.48 | A | 1 | 1 | 0 | Listeria aquatica | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria | aquatica | Listeria | aquatica | 2.29 | Listeria | grayi | 1.44 | A | 1 | 1 | 0 | Listeria aquatica | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.11 | Listeria | fleischmannii | 2.1 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LU2006-3 EFex | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria | fleischmannii | Listeria | fleischmannii | 2.03 | Listeria | fleischmannii | 1.77 | A | 1 | 1 | 0 | Listeria fleischmannii | Listeria fleischmannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105447 EFex | 3; type strain | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.32 | Listeria | grayi | 2.32 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20596 EFex | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria | grayi | Listeria | grayi | 2.25 | Listeria | grayi | 2.12 | A | 1 | 1 | 0 | Listeria grayi | Listeria grayi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10132,2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | innocua | Listeria | innocua | 2.45 | Listeria | innocua | 2.44 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8569 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | innocua | Listeria | innocua | 2.48 | Listeria | innocua | 2.46 | A | 1 | 1 | 0 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria | ivanovii | Listeria | ivanovii | 2.33 | Listeria | ivanovii | 2.32 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32100 | 2; sequenced | EFex | 2019-12-13 | Listeria | ivanovii | Listeria | ivanovii | 2.41 | Listeria | ivanovii | 2.35 | A | 1 | 1 | 0 | Listeria ivanovii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 23813 EFEx | 3; type strain | EFEx | 2017-02-15 | Listeria | marthii | Listeria | monocytogenes | 2.28 | Listeria | monocytogenes | 2.19 | A | 0 | 0 | 1 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.42 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 873,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.43 | Listeria | monocytogenes | 2.41 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 240,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.29 | Listeria | monocytogenes | 2.24 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0206 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.41 | Listeria | monocytogenes | 2.37 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0066 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.19 | Listeria | monocytogenes | 2.19 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.22 | Listeria | monocytogenes | 2.21 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria | monocytogenes | Listeria | monocytogenes | 2.37 | Listeria | monocytogenes | 2.36 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7366,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria | monocytogenes | Listeria | monocytogenes | 2.54 | Listeria | monocytogenes | 2.51 | A | 1 | 1 | 0 | Listeria monocytogenes | Listeria monocytogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria | newyorkensis | Listeria | newyorkensis | 2.07 | Listeria | cornellensis | 2.06 | B | 0 | 0 | 1 | Listeria newyorkensis | Listeria cornellensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 22097 EFex | 3; type strain | EFEx | 2017-02-15 | Listeria | rocourtiae | Listeria | rocourtiae | 2.1 | Listeria | cornellensis | 1.94 | A | 1 | 1 | 0 | Listeria rocourtiae | Listeria cornellensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2763 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria | seeligeri | Listeria | seeligeri | 2.45 | Listeria | seeligeri | 2.45 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | seeligeri | Listeria | seeligeri | 2.36 | Listeria | seeligeri | 2.3 | A | 1 | 1 | 0 | Listeria seeligeri | Listeria seeligeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria | sp | Listeria | newyorkensis | 2.13 | Listeria | cornellensis | 1.89 | A | 0 | 0 | 1 | Listeria newyorkensis | Listeria cornellensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2412,3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria | welshimeri | Listeria | welshimeri | 2.38 | Listeria | welshimeri | 2.28 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria | welshimeri | Listeria | welshimeri | 2.44 | Listeria | welshimeri | 2.41 | A | 1 | 1 | 0 | Listeria welshimeri | Listeria welshimeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 1.95 | Brevibacillus | laterosporus | 1.91 | B | 0 | 0 | 1 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus | laterosporus | Brevibacillus | laterosporus | 2.42 | Brevibacillus | laterosporus | 2.25 | A | 1 | 1 | 0 | Brevibacillus laterosporus | Brevibacillus laterosporus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | B 210 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 K19 | Paenibacillus | alvei | Paenibacillus | alvei | 1.82 | Paenibacillus | curdlanolyticus | 1.45 | B | 0 | 0 | 1 | Paenibacillus alvei | Paenibacillus curdlanolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus | chibensis | Paenibacillus | chibensis | 2.28 | Paenibacillus | chibensis | 2.15 | A | 1 | 1 | 0 | Paenibacillus chibensis | Paenibacillus chibensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus | nuruki | Ochrobactrum | anthropi | 1.29 | Paucilactobacillus | suebicus | 1.23 | B | 0 | 0 | 1 | Ochrobactrum anthropi | Paucilactobacillus suebicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.09.2013 C8 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 1.9 | Paenibacillus | polymyxa | 1.78 | B | 0 | 0 | 1 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.11.2017 C18 | Paenibacillus | polymyxa | Paenibacillus | polymyxa | 2.1 | Paenibacillus | polymyxa | 2.02 | A | 1 | 1 | 0 | Paenibacillus polymyxa | Paenibacillus polymyxa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus | sp-CVUAS-6369 | Lentilactobacillus | parabuchneri | 1.35 | Lactobacillus | amylovorus | 1.3 | B | 0 | 0 | 1 | Lentilactobacillus parabuchneri | Lactobacillus amylovorus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus | sp-E8a | Paenibacillus | amylolyticus | 2.24 | Paenibacillus | amylolyticus | 2.17 | A | 0 | 0 | 1 | Paenibacillus amylolyticus | Paenibacillus amylolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus | taichungensis | Paenibacillus | illinoisensis | 1.77 | Paenibacillus | tylopili | 1.76 | B | 0 | 0 | 1 | Paenibacillus illinoisensis | Paenibacillus tylopili | 0 | 0 | 0 | 0
|  |  | 1 | 0 | FLI003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus | taichungensis | Paenibacillus | pabuli | 1.83 | Paenibacillus | illinoisensis | 1.75 | B | 0 | 0 | 1 | Paenibacillus pabuli | Paenibacillus illinoisensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus | tundrae | Paenibacillus | pabuli | 1.78 | Paenibacillus | tylopili | 1.77 | B | 0 | 0 | 1 | Paenibacillus pabuli | Paenibacillus tylopili | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | halophilus | Listeria | innocua | 1.68 | Listeria | innocua | 1.48 | B | 0 | 0 | 1 | Listeria innocua | Listeria innocua | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | huakuii | Macrococcus | canis | 1.37 | Neisseria | meningitidis | 1.36 | B | 0 | 0 | 1 | Macrococcus canis | Neisseria meningitidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus | psychrophilus | Jeotgalicoccus | psychrophilus | 1.84 | Legionella | geestiana | 1.41 | B | 0 | 0 | 1 | Jeotgalicoccus psychrophilus | Legionella geestiana | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus | carouselicus | Staphylococcus | simulans | 1.4 | Fredinandcohnia | humi | 1.39 | B | 0 | 0 | 1 | Staphylococcus simulans | Fredinandcohnia humi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 14.03.2013 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 2.08 | Macrococcus | caseolyticus | 1.67 | A | 1 | 1 | 0 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus | caseolyticus | Macrococcus | caseolyticus | 1.87 | Macrococcus | caseolyticus | 1.77 | B | 0 | 0 | 1 | Macrococcus caseolyticus | Macrococcus caseolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus | lentus | Staphylococcus | lentus | 1.8 | Staphylococcus | lentus | 1.51 | B | 0 | 0 | 1 | Staphylococcus lentus | Staphylococcus lentus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus | sciuri | Staphylococcus | sciuri | 2.37 | Staphylococcus | sciuri | 2.27 | A | 0 | 0 | 1 | Staphylococcus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus | sciuri | Staphylococcus | sciuri | 1.99 | Staphylococcus | sciuri | 1.68 | B | 0 | 0 | 1 | Staphylococcus sciuri | Staphylococcus sciuri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9937,2 EFex | 3; 16S rDNA-sequencing | EFex | 2016-04-06 | Mammaliicoccus | vitulinus | Staphylococcus | vitulinus | 1.87 | Staphylococcus | vitulinus | 1.81 | B | 0 | 0 | 1 | Staphylococcus vitulinus | Staphylococcus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 343,2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus | vitulinus | Staphylococcus | vitulinus | 2.12 | Staphylococcus | vitulinus | 2.05 | A | 0 | 0 | 1 | Staphylococcus vitulinus | Staphylococcus vitulinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5687 | 3; sequenced | DT | 2013-04-16 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | hyicus | 1.91 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TUM 7385 | 3; sequenced | DT | 02.10.2013 | Staphylococcus | agnetis | Staphylococcus | hyicus | 2.19 | Staphylococcus | hyicus | 1.96 | A | 0 | 0 | 1 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.39 | Staphylococcus | argenteus | 2.38 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus | argenteus | Staphylococcus | argenteus | 2.28 | Staphylococcus | argenteus | 2.2 | A | 1 | 1 | 0 | Staphylococcus argenteus | Staphylococcus argenteus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus | arlettae | Staphylococcus | arlettae | 1.72 | Staphylococcus | arlettae | 1.51 | B | 0 | 0 | 1 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus | arlettae | Staphylococcus | arlettae | 1.6 | Staphylococcus | arlettae | 1.58 | B | 0 | 0 | 1 | Staphylococcus arlettae | Staphylococcus arlettae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 29213 CVUAS | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.39 | Staphylococcus | aureus | 2.38 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9914,2 sir | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus | aureus | Staphylococcus | aureus | 2.47 | Staphylococcus | aureus | 2.43 | A | 1 | 1 | 0 | Staphylococcus aureus | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus | auricularis | Staphylococcus | auricularis | 2.14 | Staphylococcus | auricularis | 1.85 | A | 1 | 1 | 0 | Staphylococcus auricularis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20326 (T) | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.25 | Staphylococcus | capitis | 2.13 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | capitis | Staphylococcus | capitis | 2.2 | Staphylococcus | capitis | 2.19 | A | 1 | 1 | 0 | Staphylococcus capitis | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20608 (T) | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.16 | Staphylococcus | caprae | 2.06 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus | caprae | Staphylococcus | caprae | 2.06 | Staphylococcus | caprae | 2.03 | A | 1 | 1 | 0 | Staphylococcus caprae | Staphylococcus caprae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.18 | Staphylococcus | carnosus | 1.55 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | carnosus | Staphylococcus | carnosus | 2.24 | Staphylococcus | carnosus | 1.46 | A | 1 | 1 | 0 | Staphylococcus carnosus | Staphylococcus carnosus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.26 | Staphylococcus | hyicus | 1.85 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus | chromogenes | Staphylococcus | chromogenes | 2.28 | Staphylococcus | chromogenes | 1.91 | A | 1 | 1 | 0 | Staphylococcus chromogenes | Staphylococcus chromogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | cohnii | Staphylococcus | xylosus | 1.36 | Staphylococcus | schleiferi | 1.36 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus | cohnii | Staphylococcus | cohnii | 1.88 | Staphylococcus | ureilyticus | 1.84 | B | 0 | 0 | 1 | Staphylococcus cohnii | Staphylococcus ureilyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | x Sta | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012 K10 | Staphylococcus | delphini | Staphylococcus | delphini | 2.32 | Staphylococcus | delphini | 2.01 | A | 1 | 1 | 0 | Staphylococcus delphini | Staphylococcus delphini | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus | delphini | Staphylococcus | delphini | 1.9 | Staphylococcus | intermedius | 1.83 | B | 0 | 0 | 1 | Staphylococcus delphini | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus | devriesei | Staphylococcus | capitis | 1.46 | Staphylococcus | devriesei | 1.45 | B | 0 | 0 | 1 | Staphylococcus capitis | Staphylococcus devriesei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.09 | Staphylococcus | epidermidis | 2.08 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus | epidermidis | Staphylococcus | epidermidis | 2.31 | Staphylococcus | epidermidis | 1.99 | A | 1 | 1 | 0 | Staphylococcus epidermidis | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | equorum | Staphylococcus | equorum | 2.16 | Staphylococcus | equorum | 1.8 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20674 T | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus | equorum | Staphylococcus | equorum | 2.11 | Staphylococcus | equorum | 1.88 | A | 1 | 1 | 0 | Staphylococcus equorum | Staphylococcus equorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | felis | Staphylococcus | felis | 2.13 | Staphylococcus | felis | 1.96 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | felis | Staphylococcus | felis | 2.14 | Staphylococcus | felis | 2 | A | 1 | 1 | 0 | Staphylococcus felis | Staphylococcus felis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.08.2011 H9 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.46 | Staphylococcus | haemolyticus | 2.43 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus | haemolyticus | Staphylococcus | haemolyticus | 2.12 | Staphylococcus | haemolyticus | 2.04 | A | 1 | 1 | 0 | Staphylococcus haemolyticus | Staphylococcus haemolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N6 [0034] | Staphylococcus | hominis | Staphylococcus | hominis | 2.27 | Staphylococcus | hominis | 2.23 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 789,2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus | hominis | Staphylococcus | hominis | 2.13 | Staphylococcus | hominis | 1.96 | A | 1 | 1 | 0 | Staphylococcus hominis | Staphylococcus hominis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.2 | Staphylococcus | hyicus | 2.14 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3767 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus | hyicus | Staphylococcus | hyicus | 2.38 | Staphylococcus | hyicus | 2.12 | A | 1 | 1 | 0 | Staphylococcus hyicus | Staphylococcus hyicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, N8, [0034] | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.15 | Staphylococcus | intermedius | 2.02 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N7, [0034] | Staphylococcus | intermedius | Staphylococcus | intermedius | 2.07 | Staphylococcus | intermedius | 1.97 | A | 1 | 1 | 0 | Staphylococcus intermedius | Staphylococcus intermedius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 05.04.2012, N5, [0034] | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.4 | Staphylococcus | lugdunensis | 2.4 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, H16 | Staphylococcus | lugdunensis | Staphylococcus | lugdunensis | 2.5 | Staphylococcus | lugdunensis | 2.43 | A | 1 | 1 | 0 | Staphylococcus lugdunensis | Staphylococcus lugdunensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus | microti | Staphylococcus | microti | 1.85 | Staphylococcus | microti | 1.85 | B | 0 | 0 | 1 | Staphylococcus microti | Staphylococcus microti | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus | pasteuri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 2.17 | A | 0 | 0 | 1 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus | pasteuri | Staphylococcus | pasteuri | 2 | Staphylococcus | warneri | 1.81 | A | 1 | 1 | 0 | Staphylococcus pasteuri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.03.2016, H05, [0034] | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 2.05 | Staphylococcus | pseudintermedius | 1.97 | A | 1 | 1 | 0 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus | pseudintermedius | Staphylococcus | pseudintermedius | 1.84 | Staphylococcus | pseudintermedius | 1.82 | B | 0 | 0 | 1 | Staphylococcus pseudintermedius | Staphylococcus pseudintermedius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus | saprophyticus | Staphylococcus | kloosii | 1.43 | Staphylococcus | cohnii | 1.26 | B | 0 | 0 | 1 | Staphylococcus kloosii | Staphylococcus cohnii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 503,2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus | saprophyticus | Staphylococcus | saprophyticus | 2.18 | Staphylococcus | saprophyticus | 2.06 | A | 1 | 1 | 0 | Staphylococcus saprophyticus | Staphylococcus saprophyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus | schleiferi | Staphylococcus | schleiferi | 2.15 | Staphylococcus | schleiferi | 2.02 | A | 1 | 1 | 0 | Staphylococcus schleiferi | Staphylococcus schleiferi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus | schweitzeri | Staphylococcus | schweitzeri | 2.17 | Staphylococcus | schweitzeri | 2.17 | A | 1 | 1 | 0 | Staphylococcus schweitzeri | Staphylococcus schweitzeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus | simulans | Staphylococcus | simulans | 2.13 | Staphylococcus | simulans | 2.11 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus | simulans | Staphylococcus | simulans | 2.17 | Staphylococcus | simulans | 2.04 | A | 1 | 1 | 0 | Staphylococcus simulans | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2439,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.26 | Staphylococcus | succinus | 2.23 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 761,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus | succinus | Staphylococcus | succinus | 2.28 | Staphylococcus | succinus | 2.21 | A | 1 | 1 | 0 | Staphylococcus succinus | Staphylococcus succinus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus | warneri | Staphylococcus | warneri | 2.07 | Staphylococcus | warneri | 2.02 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 20316 CVUAS | 3; type strain | DT | 2013-10-02 | Staphylococcus | warneri | Staphylococcus | warneri | 2.18 | Staphylococcus | warneri | 1.96 | A | 1 | 1 | 0 | Staphylococcus warneri | Staphylococcus warneri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 2.04 | Staphylococcus | xylosus | 2.02 | A | 1 | 1 | 0 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus | xylosus | Staphylococcus | xylosus | 1.88 | Staphylococcus | xylosus | 1.76 | B | 0 | 0 | 1 | Staphylococcus xylosus | Staphylococcus xylosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus | sp-185-U1-2 | Xanthomonas | hortorum | 1.38 | Xanthomonas | hortorum | 1.37 | B | 0 | 0 | 1 | Xanthomonas hortorum | Xanthomonas hortorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus | suis | Aerococcus | suis | 2.3 | Levilactobacillus | brevis | 1.42 | A | 1 | 1 | 0 | Aerococcus suis | Levilactobacillus brevis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus | urinaeequi | Aerococcus | viridans | 1.72 | Aerococcus | viridans | 1.65 | B | 0 | 0 | 1 | Aerococcus viridans | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 27.09.2016 | Aerococcus | viridans | Aerococcus | viridans | 2.04 | Aerococcus | viridans | 1.94 | A | 1 | 1 | 0 | Aerococcus viridans | Aerococcus viridans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter | phocae | Atopobacter | phocae | 2.33 | Atopobacter | phocae | 2.31 | A | 1 | 1 | 0 | Atopobacter phocae | Atopobacter phocae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2442,2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium | divergens | Carnobacterium | divergens | 1.99 | Carnobacterium | divergens | 1.96 | B | 0 | 0 | 1 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium | divergens | Carnobacterium | divergens | 2.37 | Carnobacterium | divergens | 2.27 | A | 1 | 1 | 0 | Carnobacterium divergens | Carnobacterium divergens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6273 | 3; sequenced | eDT | 13.01.2016 | Carnobacterium | maltaromaticum | Carnobacterium | maltaromaticum | 2.55 | Carnobacterium | maltaromaticum | 2.54 | A | 1 | 1 | 0 | Carnobacterium maltaromaticum | Carnobacterium maltaromaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1519,3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella | adiacens | Granulicatella | adiacens | 2.3 | Granulicatella | adiacens | 2.24 | A | 1 | 1 | 0 | Granulicatella adiacens | Granulicatella adiacens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2616 | 3; sequenced | DT | 25.09.2013 | Marinilactibacillus | sp-CVUAS-2616 | Lacticaseibacillus | sharpeae | 1.42 | Staphylococcus | simulans | 1.34 | B | 0 | 0 | 1 | Lacticaseibacillus sharpeae | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8980,3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus | sp-CVUAS-2616 | Enterococcus | cecorum | 1.53 | Ligilactobacillus | murinus | 1.5 | B | 0 | 0 | 1 | Enterococcus cecorum | Ligilactobacillus murinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16 Ec RKI | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011 H20 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.3 | Enterococcus | casseliflavus | 2.2 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H12 | Enterococcus | durans | Enterococcus | hirae | 2.09 | Enterococcus | hirae | 1.8 | A | 0 | 0 | 1 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H13 | Enterococcus | durans | Enterococcus | durans | 2 | Enterococcus | durans | 1.94 | A | 1 | 1 | 0 | Enterococcus durans | Enterococcus durans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.07.2011 N8 | Enterococcus | faecium | Enterococcus | faecium | 2.09 | Enterococcus | faecium | 2.05 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N13 | Enterococcus | gallinaru | Enterococcus | gallinarum | 2.04 | Enterococcus | gallinarum | 1.98 | A | 0 | 0 | 1 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N3 | Enterococcus | hirae | Enterococcus | hirae | 2.33 | Enterococcus | hirae | 2.16 | A | 1 | 1 | 0 | Enterococcus hirae | Enterococcus hirae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus | casseliflavus | Enterococcus | casseliflavus | 2.33 | Enterococcus | casseliflavus | 2.32 | A | 1 | 1 | 0 | Enterococcus casseliflavus | Enterococcus casseliflavus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5030 | 3; sequenced | eDT | 23.07.2014 | Enterococcus | cecorum | Enterococcus | cecorum | 2.34 | Enterococcus | cecorum | 2.29 | A | 1 | 1 | 0 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8897 | 3; sequenced | eDT | 23.07.2014 | Enterococcus | cecorum | Enterococcus | cecorum | 1.41 | Enterococcus | cecorum | 1.38 | B | 0 | 0 | 1 | Enterococcus cecorum | Enterococcus cecorum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus | faecalis | Enterococcus | faecalis | 2.51 | Enterococcus | faecalis | 2.41 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus | faecalis | Enterococcus | faecalis | 2.39 | Enterococcus | faecalis | 2.38 | A | 1 | 1 | 0 | Enterococcus faecalis | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MR11-207 LVU04 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus | faecium | Enterococcus | faecium | 2.42 | Enterococcus | faecium | 2.1 | A | 1 | 1 | 0 | Enterococcus faecium | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus | gallinarum | Enterococcus | gallinarum | 2.42 | Enterococcus | gallinarum | 2.38 | A | 1 | 1 | 0 | Enterococcus gallinarum | Enterococcus gallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus | solitarius | Tetragenococcus | solitarius | 2.14 | Halomonas | elongata | 1.38 | A | 1 | 1 | 0 | Tetragenococcus solitarius | Halomonas elongata | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1549,2 | 3; sequenced | DT | 2016-09-27 | Vagococcus | salmoninarum | Vagococcus | salmoninarum | 2.09 | Limosilactobacillus | oris | 1.43 | A | 1 | 1 | 0 | Vagococcus salmoninarum | Limosilactobacillus oris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus | acidophilus | Lactobacillus | acidophilus | 2.32 | Lactobacillus | acidophilus | 2.24 | A | 1 | 1 | 0 | Lactobacillus acidophilus | Lactobacillus acidophilus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus | johnsonii | Lactobacillus | johnsonii | 2.42 | Lactobacillus | johnsonii | 2.35 | A | 1 | 1 | 0 | Lactobacillus johnsonii | Lactobacillus johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32277,2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus | sp-CVUAS-32277,2 | Limosilactobacillus | gorillae | 1.25 | Secundilactobacillus | malefermentans | 1.24 | B | 0 | 0 | 1 | Limosilactobacillus gorillae | Secundilactobacillus malefermentans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus | claussenii | Pediococcus | claussenii | 2.25 | Pediococcus | claussenii | 2.04 | A | 1 | 1 | 0 | Pediococcus claussenii | Pediococcus claussenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1213 | 2; 0 | DT | 2013-07-09 | Leuconostoc | mesenteroides | Leuconostoc | mesenteroides | 2.28 | Leuconostoc | mesenteroides | 2.09 | A | 1 | 1 | 0 | Leuconostoc mesenteroides | Leuconostoc mesenteroides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LTH 1146 | 2; 0 | DT | 2013-07-09 | Lactococcus | lactis | Lactococcus | lactis | 2.42 | Lactococcus | lactis | 2.4 | A | 1 | 1 | 0 | Lactococcus lactis | Lactococcus lactis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.45 | Streptococcus | agalactiae | 2.44 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 6784 CVUAS | 3; public strain collection | eDT | 2019-07-09 | Streptococcus | agalactiae | Streptococcus | agalactiae | 2.49 | Streptococcus | agalactiae | 2.48 | A | 1 | 1 | 0 | Streptococcus agalactiae | Streptococcus agalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus | canis | Streptococcus | canis | 2.53 | Streptococcus | canis | 2.35 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL Mi-845_17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus | canis | Streptococcus | canis | 2.47 | Streptococcus | canis | 2.45 | A | 1 | 1 | 0 | Streptococcus canis | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.19 | Streptococcus | dysgalactiae | 2.04 | B | 0 | 0 | 1 | Streptococcus castoreus | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17536 CVUAS | 3; type strain | EFEx | 2018-05-03 | Streptococcus | castoreus | Streptococcus | castoreus | 2.35 | Streptococcus | pyogenes | 1.96 | A | 1 | 1 | 0 | Streptococcus castoreus | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus | catagoni | Streptococcus | uberis | 1.74 | Streptococcus | phocae | 1.69 | B | 0 | 0 | 1 | Streptococcus uberis | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 99-1-2017 | 3; type strain | EFEx | 2018-05-03 | Streptococcus | catagoni | Streptococcus | uberis | 1.81 | Streptococcus | uberis | 1.71 | B | 0 | 0 | 1 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.14 | Streptococcus | canis | 2.07 | B | 0 | 0 | 1 | Streptococcus dysgalactiae | Streptococcus canis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.26 | Streptococcus | dysgalactiae | 1.98 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus | dysgalactiae | Streptococcus | canis | 2.12 | Streptococcus | dysgalactiae | 2.07 | B | 0 | 0 | 1 | Streptococcus canis | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus | dysgalactiae | Streptococcus | dysgalactiae | 2.11 | Streptococcus | dysgalactiae | 2.1 | A | 1 | 1 | 0 | Streptococcus dysgalactiae | Streptococcus dysgalactiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32462,2 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi_ssp_equi | 2.12 | Streptococcus | equi_ssp_equi | 2.06 | A | 0 | 0 | 1 | Streptococcus equi_ssp_equi | Streptococcus equi_ssp_equi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32770 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus | equi | Streptococcus | equi_ssp_equi | 2.29 | Streptococcus | equi_ssp_equi | 2.24 | A | 0 | 0 | 1 | Streptococcus equi_ssp_equi | Streptococcus equi_ssp_equi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus | equi | Streptococcus | equi_ssp_zooepidemicus | 2.34 | Streptococcus | equi_ssp_zooepidemicus | 2.34 | A | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus | equi | Streptococcus | equi_ssp_zooepidemicus | 2.21 | Streptococcus | equi_ssp_zooepidemicus | 2.18 | A | 0 | 0 | 1 | Streptococcus equi_ssp_zooepidemicus | Streptococcus equi_ssp_zooepidemicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus | iniae | Streptococcus | iniae | 1.89 | Streptococcus | iniae | 1.74 | B | 0 | 0 | 1 | Streptococcus iniae | Streptococcus iniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus | orisasini | Streptococcus | ratti | 2.02 | Streptococcus | ratti | 1.97 | A | 0 | 0 | 1 | Streptococcus ratti | Streptococcus ratti | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | parauberis | Streptococcus | parauberis | 2.1 | Streptococcus | parauberis | 1.9 | A | 1 | 1 | 0 | Streptococcus parauberis | Streptococcus parauberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus | phocae | Streptococcus | phocae | 2.18 | Streptococcus | phocae | 1.97 | A | 1 | 1 | 0 | Streptococcus phocae | Streptococcus phocae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 12-13862-001-01 LVU11 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus | pluranimalium | Streptococcus | pluranimalium | 2.2 | Streptococcus | pluranimalium | 2.18 | A | 1 | 1 | 0 | Streptococcus pluranimalium | Streptococcus pluranimalium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.48 | Streptococcus | pneumoniae | 2.44 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.16 | Streptococcus | pneumoniae | 2.16 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.29 | Streptococcus | pneumoniae | 2.23 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus | pneumoniae | Streptococcus | pneumoniae | 2.49 | Streptococcus | pneumoniae | 2.46 | A | 1 | 1 | 0 | Streptococcus pneumoniae | Streptococcus pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus | porcinus | Streptococcus | porcinus | 2.1 | Streptococcus | porcinus | 2.09 | A | 1 | 1 | 0 | Streptococcus porcinus | Streptococcus porcinus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.04.2008, K11, [0034] | Streptococcus | pyogenes | Streptococcus | pyogenes | 2.43 | Streptococcus | pyogenes | 2.4 | A | 1 | 1 | 0 | Streptococcus pyogenes | Streptococcus pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10145,3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus | salivarius | Streptococcus | salivarius | 2.13 | Streptococcus | salivarius | 2.13 | A | 1 | 1 | 0 | Streptococcus salivarius | Streptococcus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus | sp-CVUAS-32349 | Streptococcus | dysgalactiae | 1.52 | Streptococcus | hyovaginalis | 1.48 | B | 0 | 0 | 1 | Streptococcus dysgalactiae | Streptococcus hyovaginalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2260,4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus | troglodytidis | Streptococcus | anginosus | 1.57 | Streptococcus | anginosus | 1.51 | B | 0 | 0 | 1 | Streptococcus anginosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5218,4 | 3; sequenced; 16S; NRZ Strep Uni AA | DT | 2015-06-11 | Streptococcus | troglodytidis | Streptococcus | anginosus | 1.69 | Streptococcus | constellatus | 1.58 | B | 0 | 0 | 1 | Streptococcus anginosus | Streptococcus constellatus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.03 | Streptococcus | uberis | 1.81 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus | uberis | Streptococcus | uberis | 2.23 | Streptococcus | uberis | 2.18 | A | 1 | 1 | 0 | Streptococcus uberis | Streptococcus uberis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 1.94 | Clostridioides | difficile | 1.71 | B | 0 | 0 | 1 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL305725 | 4; ribotype; ex Clostridium | eDT | 2019-07-12 | Clostridioides | difficile | Clostridioides | difficile | 2.44 | Clostridioides | difficile | 2.24 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL402952 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.17 | Clostridioides | difficile | 2 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL402975 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.03 | Clostridioides | difficile | 2.02 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL402979 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.11 | Clostridioides | difficile | 2.09 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL403180 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.01 | Clostridioides | difficile | 1.9 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL403181 | 4; ribotype; ex Clostridium | EFEx | 2019-09-13 | Clostridioides | difficile | Clostridioides | difficile | 2.18 | Clostridioides | difficile | 1.84 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL403210 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.14 | Clostridioides | difficile | 2.06 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL403313 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.04 | Clostridioides | difficile | 1.96 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL403467 | 4; ribotype; ex Clostridium | EFEx | 2019-09-13 | Clostridioides | difficile | Clostridioides | difficile | 2.21 | Clostridioides | difficile | 1.97 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL403488 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.02 | Clostridioides | difficile | 1.94 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14EL403625 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides | difficile | Clostridioides | difficile | 2.1 | Clostridioides | difficile | 1.8 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0014-11 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.4 | Clostridioides | difficile | 2.24 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0021-11 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.34 | Clostridioides | difficile | 2.1 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0022-11 | 4; ribotype; ex Clostridium | DT | 2012-12-19 | Clostridioides | difficile | Clostridioides | difficile | 2.38 | Clostridioides | difficile | 2.14 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0059-10 | 4; ribotype; ex Clostridium | DT | 2013-02-06 | Clostridioides | difficile | Clostridioides | difficile | 2.25 | Clostridioides | difficile | 2.17 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0061-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.22 | Clostridioides | difficile | 2.03 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0062-10 | 4; ribotype; ex Clostridium | DT | 2013-01-09 | Clostridioides | difficile | Clostridioides | difficile | 2.24 | Clostridioides | difficile | 2.14 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0063-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.38 | Clostridioides | difficile | 2.22 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 1 | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides | difficile | Clostridioides | difficile | 2.17 | Clostridioides | difficile | 2.06 | A | 1 | 1 | 0 | Clostridioides difficile | Clostridioides difficile | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 2.05 | Clostridium | sporogenes | 1.9 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 1.88 | Clostridium | sporogenes | 1.78 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 2.01 | Clostridium | sporogenes | 1.93 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot B L677 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 1.95 | Clostridium | sporogenes | 1.85 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 1.59 | Clostridium | sporogenes | 1.51 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Clostridium | sporogenes | 1.67 | Clostridium | sporogenes | 1.59 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium | [botulinum] | Arthrobacter | ramosus | 1.51 | Aeromonas | caviae | 1.36 | B | 0 | 0 | 1 | Arthrobacter ramosus | Aeromonas caviae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium | botulinum | Bacillus | subtilis | 1.49 | Clostridium | cochlearium | 1.42 | B | 0 | 0 | 1 | Bacillus subtilis | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32971 | 3; sequenced | EFEx | 2021-02-25 | Clostridium | botulinum | Clostridium | sporogenes | 2.08 | Clostridium | sporogenes | 1.97 | A | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33885,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium | botulinum | Clostridium | sporogenes | 1.71 | Clostridium | sporogenes | 1.52 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33886,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium | botulinum | Clostridium | sporogenes | 1.88 | Clostridium | sporogenes | 1.84 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33891,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium | botulinum | Clostridium | sporogenes | 1.86 | Clostridium | sporogenes | 1.74 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium | chauvoei | Clostridium | chauvoei | 2.38 | Clostridium | chauvoei | 2.17 | A | 1 | 1 | 0 | Clostridium chauvoei | Clostridium chauvoei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 849,2 | 3; sequenced | DT | 2014-12-17 | Clostridium | isatidis | Clostridium | isatidis | 1.6 | Pseudomonas | boreopolis | 1.34 | B | 0 | 0 | 1 | Clostridium isatidis | Pseudomonas boreopolis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108256 | 3; Public strain collection | EFEx | 0 | Clostridium | leptum | Clostridium | paraputrificum | 1.43 | Staphylococcus | simulans | 1.4 | B | 0 | 0 | 1 | Clostridium paraputrificum | Staphylococcus simulans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_23_22 | 3; sequenced [0207] | EFEx | 0 | Clostridium | nexile | Aeromonas | caviae | 1.24 | Pseudomonas | agarici | 1.24 | B | 0 | 0 | 1 | Aeromonas caviae | Pseudomonas agarici | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium | nigeriense | Clostridium | tertium | 1.75 | Clostridium | tertium | 1.68 | B | 0 | 0 | 1 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM_10421 | 3; Public strain collection | eDT | 0 | Clostridium | scindens | Clostridium | scindens | 2.26 | Clostridium | scindens | 1.98 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM_10419 | 3; Public strain collection | eDT | 2021-10-26 | Clostridium | scindens | Clostridium | scindens | 2.23 | Clostridium | scindens | 2.15 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM_10420 | 3; Public strain collection | EFEx | 2021-10-20 | Clostridium | scindens | Clostridium | scindens | 2.19 | Clostridium | scindens | 2.15 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_1_26 | 3; sequenced [0207] | eDT | 0 | Clostridium | scindens | Clostridium | scindens | 2.2 | Clostridium | scindens | 2.06 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_5_24 | 3; sequenced [0207] | eDT | 0 | Clostridium | scindens | Clostridium | scindens | 2.01 | Clostridium | scindens | 1.78 | A | 1 | 1 | 0 | Clostridium scindens | Clostridium scindens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium | septicum | Clostridium | septicum | 2.27 | Clostridium | septicum | 2.23 | A | 1 | 1 | 0 | Clostridium septicum | Clostridium septicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34091 | 3; sequenced | EFex | 2022-08-05 | Clostridium | sp-CVUAS-34091 | Clostridium | sartagoforme | 1.73 | Clostridium | tertium | 1.62 | B | 0 | 0 | 1 | Clostridium sartagoforme | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium | sp-DSM-1985 | Campylobacter | jejuni | 1.33 | Campylobacter | jejuni | 1.32 | B | 0 | 0 | 1 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium | sporogenes | Clostridium | sporogenes | 1.95 | Clostridium | sporogenes | 1.79 | B | 0 | 0 | 1 | Clostridium sporogenes | Clostridium sporogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_1_9 | 3; sequenced [0207] | eDT | 0 | Clostridium | symbiosum | Clostridium | symbiosum | 2.43 | Clostridium | symbiosum | 1.96 | A | 1 | 1 | 0 | Clostridium symbiosum | Clostridium symbiosum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_7_21 | 3; sequenced [0207] | eDT | 0 | Clostridium | symbiosum | Clostridium | symbiosum | 1.93 | Clostridium | symbiosum | 1.84 | B | 0 | 0 | 1 | Clostridium symbiosum | Clostridium symbiosum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | WAL-14163 | 3; sequenced [0207] | eDT | 0 | Clostridium | symbiosum | Clostridium | symbiosum | 2.45 | Clostridium | symbiosum | 2.07 | A | 1 | 1 | 0 | Clostridium symbiosum | Clostridium symbiosum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 20 | 2; confirmed proficency test | eDT | 2022-02-14 | Clostridium | tertium | Clostridium | tertium | 2.2 | Clostridium | tertium | 2.18 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10003 | 2; 0 | DT | 2013-02-20 | Clostridium | tertium | Clostridium | tertium | 2.34 | Clostridium | tertium | 2.25 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium | tertium | Clostridium | tertium | 2.33 | Clostridium | tertium | 2.28 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7893 | 3; sequenced | DT | 2012-12-19 | Clostridium | tertium | Clostridium | tertium | 2.35 | Clostridium | tertium | 2.32 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8330 | 3; sequenced | DT | 2012-12-19 | Clostridium | tertium | Clostridium | tertium | 2.36 | Clostridium | tertium | 2.34 | A | 1 | 1 | 0 | Clostridium tertium | Clostridium tertium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium | tetani | Clostridium | tetani | 2.55 | Clostridium | tetani | 2.47 | A | 1 | 1 | 0 | Clostridium tetani | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16296 | 3; Public strain collection | EFEx | 0 | Eubacterium | ramulus | Lacticaseibacillus | paracasei | 1.38 | Terrimonas | ferruginea | 1.33 | B | 0 | 0 | 1 | Lacticaseibacillus paracasei | Terrimonas ferruginea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H258 | 3; sequenced [0207] | eDT | 0 | Agathobacter | rectalis | Pseudomonas | putida | 1.45 | Pseudomonas | monteilii | 1.38 | B | 0 | 0 | 1 | Pseudomonas putida | Pseudomonas monteilii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_9_13 | 3; sequenced [0207] | eDT | 0 | Agathobacter | rectalis | Pseudomonas | putida | 1.39 | Lactococcus | plantarum | 1.31 | B | 0 | 0 | 1 | Pseudomonas putida | Lactococcus plantarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Clinical isolate | 3; sequenced [0207] | EFEx | 0 | Anaerobutyricum | hallii | Staphylococcus | simulans | 1.2 | Bacillus | subtilis | 1.16 | B | 0 | 0 | 1 | Staphylococcus simulans | Bacillus subtilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H53 | 3; sequenced [0207] | EFEx | 0 | Anaerostipes | hadrus | Arthrobacter | parietis | 1.22 | Delftia | tsuruhatensis | 1.22 | B | 0 | 0 | 1 | Arthrobacter parietis | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 29492 | 3; type strain | EFEx | 0 | Blautia | caecimuris | Rhizorhabdus | wittichii | 1.27 | Enterococcus | faecium | 1.27 | B | 0 | 0 | 1 | Rhizorhabdus wittichii | Enterococcus faecium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_18_51 | 3; sequenced [0207] | EFEx | 0 | Blautia | faecis | Terrimonas | ferruginea | 1.29 | Limosilactobacillus | antri | 1.21 | B | 0 | 0 | 1 | Terrimonas ferruginea | Limosilactobacillus antri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22028 | 3; type strain | EFEx | 0 | Blautia | glucerasea | Neisseria | oralis | 1.29 | Parvimonas | micra | 1.28 | B | 0 | 0 | 1 | Neisseria oralis | Parvimonas micra | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20583 | 3; type strain | EFEx | 0 | Blautia | hansenii | Aeromonas | molluscorum | 1.4 | Staphylococcus | simiae | 1.34 | B | 0 | 0 | 1 | Aeromonas molluscorum | Staphylococcus simiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_19_60 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Halococcus | morrhuae | 1.29 | Acinetobacter | johnsonii | 1.28 | B | 0 | 0 | 1 | Halococcus morrhuae | Acinetobacter johnsonii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_20_85 | 3; sequenced [0207] | EFEx | 0 | Blautia | massiliensis | Acinetobacter | gandensis | 1.45 | Acinetobacter | gandensis | 1.29 | B | 0 | 0 | 1 | Acinetobacter gandensis | Acinetobacter gandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_21_1 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Blastomonas | ursincola | 1.44 | Lacticaseibacillus | paracasei | 1.43 | B | 0 | 0 | 1 | Blastomonas ursincola | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_22_56 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Sphingomonas | yabuuchiae | 1.36 | Clostridium | merdae | 1.25 | B | 0 | 0 | 1 | Sphingomonas yabuuchiae | Clostridium merdae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_9_19 | 3; sequenced [0207] | eDT | 0 | Blautia | massiliensis | Weissella | viridescens | 1.44 | Aeromonas | eucrenophila | 1.22 | B | 0 | 0 | 1 | Weissella viridescens | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_20_66 | 3; sequenced [0207] | eDT | 0 | Blautia | obeum | Streptomyces | phaeochromogenes | 1.46 | Aeromonas | eucrenophila | 1.43 | B | 0 | 0 | 1 | Streptomyces phaeochromogenes | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9990 | 3; type strain | eDT | 0 | Blautia | producta | Blautia | coccoides | 2.03 | Lactobacillus | paraplantarum | 1.29 | A | 0 | 0 | 1 | Blautia coccoides | Lactobacillus paraplantarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_4_13 | 2; data not listed [0207] | eDT | 0 | Blautia | producta | Clostridium | haemolyticum | 1.43 | Clostridium | baratii | 1.4 | B | 0 | 0 | 1 | Clostridium haemolyticum | Clostridium baratii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_4_17 | 3; sequenced [0207] | eDT | 0 | Blautia | producta | Clostridium | baratii | 1.38 | Sinomonas | atrocyanea | 1.37 | B | 0 | 0 | 1 | Clostridium baratii | Sinomonas atrocyanea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 26115 | 3; Public strain collection | EFEx | 0 | Blautia | pseudococcoides | Blautia | coccoides | 1.61 | Acinetobacter | tjernbergiae | 1.28 | B | 0 | 0 | 1 | Blautia coccoides | Acinetobacter tjernbergiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_6_27 | 3; sequenced [0207] | EFEx | 2021-10-20 | Blautia | schinkii | Flavobacterium | pectinovorum | 1.28 | Staphylococcus | capitis | 1.27 | B | 0 | 0 | 1 | Flavobacterium pectinovorum | Staphylococcus capitis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_6_32 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia | schinkii | Paracoccus | denitrificans | 1.34 | Staphylococcus | epidermidis | 1.31 | B | 0 | 0 | 1 | Paracoccus denitrificans | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_18_16 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia | wexlerae | Staphylococcus | sciuri | 1.29 | Achromobacter | spanius | 1.28 | B | 0 | 0 | 1 | Staphylococcus sciuri | Achromobacter spanius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31264 | 3; Public strain collection | eDT | 2021-10-20 | Coprococcus | comes | Clostridium | baratii | 1.28 | Staphylococcus | epidermidis | 1.25 | B | 0 | 0 | 1 | Clostridium baratii | Staphylococcus epidermidis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_17_21 | 3; sequenced [0207] | eDT | 2021-10-20 | Coprococcus | comes | Clostridium | baratii | 1.38 | Listeria | ivanovii | 1.35 | B | 0 | 0 | 1 | Clostridium baratii | Listeria ivanovii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_10_6 | 3; sequenced [0207] | EFEx | 2021-10-20 | Coprococcus | eutactus | Staphylococcus | schleiferi | 1.4 | Pseudomonas | thermotolerans | 1.33 | B | 0 | 0 | 1 | Staphylococcus schleiferi | Pseudomonas thermotolerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31256 | 3; Public strain collection | eDT | 0 | Dorea | formicigenerans | Dorea | formicigenerans | 1.56 | Veillonella | atypica | 1.25 | B | 0 | 0 | 1 | Dorea formicigenerans | Veillonella atypica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_19_39 | 3; sequenced [0207] | eDT | 0 | Dorea | formicigenerans | Dorea | formicigenerans | 1.52 | Lacticaseibacillus | paracasei | 1.28 | B | 0 | 0 | 1 | Dorea formicigenerans | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_23_29 | 3; sequenced [0207] | eDT | 0 | Dorea | formicigenerans | Dorea | formicigenerans | 1.48 | Pantoea | agglomerans | 1.34 | B | 0 | 0 | 1 | Dorea formicigenerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_18_49 | 3; sequenced [0207] | eDT | 0 | Dorea | longicatena | Lactobacillus | helveticus | 1.37 | Actinomyces | naeslundii | 1.36 | B | 0 | 0 | 1 | Lactobacillus helveticus | Actinomyces naeslundii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108191 | 3; Public strain collection | EFEx | 0 | Enterocloster | clostridioformis | Enterocloster | clostridioformis | 2.34 | Enterocloster | clostridioformis | 2.32 | A | 1 | 1 | 0 | Enterocloster clostridioformis | Enterocloster clostridioformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_20_55 | 3; sequenced [0207] | eDT | 0 | Faecalicatena | fissicatena | Stenotrophomonas | maltophilia | 1.35 | Clostridium | novyi | 1.29 | B | 0 | 0 | 1 | Stenotrophomonas maltophilia | Clostridium novyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_9_3 | 3; sequenced [0207] | eDT | 0 | Faecalicatena | fissicatena | Metabacillus | indicus | 1.25 | Paracoccus | denitrificans | 1.24 | B | 0 | 0 | 1 | Metabacillus indicus | Paracoccus denitrificans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_14_54 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter | saccharivorans | Serratia | rubidaea | 1.35 | Terrimonas | ferruginea | 1.34 | B | 0 | 0 | 1 | Serratia rubidaea | Terrimonas ferruginea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_15_44 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter | saccharivorans | Pseudomonas | antarctica | 1.38 | Terrimonas | ferruginea | 1.37 | B | 0 | 0 | 1 | Pseudomonas antarctica | Terrimonas ferruginea | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H260 | 3; sequenced [0207] | EFEx | 0 | Lachnospira | eligens | Bordetella | pertussis | 1.27 | Staphylococcus | auricularis | 1.21 | B | 0 | 0 | 1 | Bordetella pertussis | Staphylococcus auricularis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 105336 | 3; Public strain collection | EFEx | 0 | Lacrimispora | celerecrescens | Lacrimispora | celerecrescens | 2.33 | Lacrimispora | sphenoides | 2.18 | B | 0 | 0 | 1 | Lacrimispora celerecrescens | Lacrimispora sphenoides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H164 | 3; sequenced [0207] | EFEx | 0 | Roseburia | faecis | Staphylococcus | ureilyticus | 1.42 | Staphylococcus | ureilyticus | 1.33 | B | 0 | 0 | 1 | Staphylococcus ureilyticus | Staphylococcus ureilyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16839 | 3; type strain | EFEx | 0 | Roseburia | hominis | Limosilactobacillus | mucosae | 1.4 | Neisseria | gonorrhoeae | 1.37 | B | 0 | 0 | 1 | Limosilactobacillus mucosae | Neisseria gonorrhoeae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 108070 | 3; Public strain collection | EFEx | 0 | Roseburia | inulinivorans | Bacillus | pumilus | 1.49 | Enterocloster | clostridioformis | 1.4 | B | 0 | 0 | 1 | Bacillus pumilus | Enterocloster clostridioformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_1_25 | 3; sequenced [0207] | eDT | 0 | Sellimonas | intestinalis | Magnusiomyces | capitatus | 1.4 | Hungatella | hathewayi | 1.37 | B | 0 | 0 | 1 | Magnusiomyces capitatus | Hungatella hathewayi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | JCM 31255 | 3; Public strain collection | eDT | 0 | Anaerotruncus | colihominis | Anaerotruncus | colihominis | 2.38 | Anaerotruncus | colihominis | 2.24 | A | 1 | 1 | 0 | Anaerotruncus colihominis | Anaerotruncus colihominis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H281 | 3; sequenced [0207] | EFEx | 0 | Faecalibacterium | prausnitzii | Sphingomonas | aurantiaca | 1.28 | Agromyces | mediolanus | 1.26 | B | 0 | 0 | 1 | Sphingomonas aurantiaca | Agromyces mediolanus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H271 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | bromii | Clostridium | budayi | 1.31 | Ralstonia | sp | 1.29 | B | 0 | 0 | 1 | Clostridium budayi | Ralstonia sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_22_53 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | gnavus | Ruminococcus | gnavus | 2.13 | Ruminococcus | gnavus | 1.97 | A | 1 | 1 | 0 | Ruminococcus gnavus | Ruminococcus gnavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | MSK_23_60 | 3; sequenced [0207] | eDT | 0 | Ruminococcus | gnavus | Ruminococcus | gnavus | 2.13 | Ruminococcus | gnavus | 2.07 | A | 1 | 1 | 0 | Ruminococcus gnavus | Ruminococcus gnavus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H279 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus | lactaris | Clostridioides | difficile | 1.32 | Alteromonas | stellipolaris | 1.32 | B | 0 | 0 | 1 | Clostridioides difficile | Alteromonas stellipolaris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CLA-AA-H261 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus | torques | Staphylococcus | arlettae | 1.46 | Moraxella | catarrhalis | 1.27 | B | 0 | 0 | 1 | Staphylococcus arlettae | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix | anatis | Lactobacillus | amylovorus | 1.43 | Erysipelothrix | tonsillarum | 1.41 | B | 0 | 0 | 1 | Lactobacillus amylovorus | Erysipelothrix tonsillarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix | anatis | Staphylococcus | carnosus | 1.36 | Erysipelothrix | rhusiopathiae | 1.26 | B | 0 | 0 | 1 | Staphylococcus carnosus | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix | aquatica | Erysipelothrix | rhusiopathiae | 1.42 | Erysipelothrix | rhusiopathiae | 1.37 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | VA-92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | aquatica | Neisseria_flavescens_subflava | group | 1.35 | Fusobacterium | necrophorum | 1.31 | B | 0 | 0 | 1 | Neisseria_flavescens_subflava group | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | larvae | Loigolactobacillus | bifermentans | 1.46 | Fusobacterium | necrophorum | 1.44 | B | 0 | 0 | 1 | Loigolactobacillus bifermentans | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 110099 CVUAS | 3; type strain | EFex [BRU03] | 2021-10-18 | Erysipelothrix | piscisicarius | Erysipelothrix | rhusiopathiae | 1.75 | Erysipelothrix | rhusiopathiae | 1.63 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.28 | Erysipelothrix | rhusiopathiae | 2.16 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix | rhusiopathiae | Erysipelothrix | rhusiopathiae | 2.03 | Erysipelothrix | rhusiopathiae | 2.02 | A | 1 | 1 | 0 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix | tonsillarum | Erysipelothrix | rhusiopathiae | 1.92 | Erysipelothrix | rhusiopathiae | 1.83 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-23 | Erysipelothrix | tonsillarum | Erysipelothrix | rhusiopathiae | 1.88 | Erysipelothrix | rhusiopathiae | 1.79 | B | 0 | 0 | 1 | Erysipelothrix rhusiopathiae | Erysipelothrix rhusiopathiae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus | sp-CVUAS-10169 | Pseudomonas | putida | 1.34 | Pseudomonas | putida | 1.21 | B | 0 | 0 | 1 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32948 | 3; confirmed proficency test | EFEx | 2021-02-18 | Finegoldia | magna | Finegoldia | magna | 2.26 | Finegoldia | magna | 2.08 | A | 1 | 1 | 0 | Finegoldia magna | Finegoldia magna | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.25 | Helcococcus | kunzii | 2.25 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus | kunzii | Helcococcus | kunzii | 2.45 | Helcococcus | kunzii | 2.4 | A | 1 | 1 | 0 | Helcococcus kunzii | Helcococcus kunzii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 433,2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus | ovis | Helcococcus | ovis | 2.39 | Helcococcus | ovis | 2.3 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus | ovis | Helcococcus | ovis | 2.28 | Helcococcus | ovis | 2.17 | A | 1 | 1 | 0 | Helcococcus ovis | Helcococcus ovis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus | sp-CVUAS-32321 | Helcococcus | sueciensis | 1.77 | Helcococcus | sueciensis | 1.73 | B | 0 | 0 | 1 | Helcococcus sueciensis | Helcococcus sueciensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 692,2 | 3; sequenced | DT | 2013-02-19 | Parvimonas | sp-CVUAS-692,2 | Actinomyces | bovis | 1.44 | Actinomyces | bovis | 1.31 | B | 0 | 0 | 1 | Actinomyces bovis | Actinomyces bovis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides | finegoldii | Bacteroides | finegoldii | 2.6 | Bacteroides | finegoldii | 2.53 | A | 1 | 1 | 0 | Bacteroides finegoldii | Bacteroides finegoldii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides | fragilis | Bacteroides | fragilis | 2.35 | Bacteroides | fragilis | 2.33 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181001478 Bat | 2; 0 | DT | 2018-05-02 | Bacteroides | pyogenes | Bacteroides | pyogenes | 2.48 | Bacteroides | pyogenes | 2.48 | A | 1 | 1 | 0 | Bacteroides pyogenes | Bacteroides pyogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella | heparinolytica | Prevotella | heparinolytica | 2.52 | Prevotella | heparinolytica | 2.49 | A | 1 | 1 | 0 | Prevotella heparinolytica | Prevotella heparinolytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium | gleum | Chryseobacterium | gleum | 1.94 | Chryseobacterium | gleum | 1.9 | B | 0 | 0 | 1 | Chryseobacterium gleum | Chryseobacterium gleum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 2.1 | Chryseobacterium | tructae | 2.07 | A | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium tructae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium | indologenes | Chryseobacterium | tructae | 1.85 | Chryseobacterium | indologenes | 1.8 | B | 0 | 0 | 1 | Chryseobacterium tructae | Chryseobacterium indologenes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32947 | 3; confirmed proficency test | EFEx | 2021-02-18 | Bacteroides | fragilis | Bacteroides | fragilis | 2.22 | Bacteroides | fragilis | 2.2 | A | 1 | 1 | 0 | Bacteroides fragilis | Bacteroides fragilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium | hibernum | Flavobacterium | hibernum | 2 | Flavobacterium | saccharophilum | 1.97 | A | 1 | 1 | 0 | Flavobacterium hibernum | Flavobacterium saccharophilum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium | sp | Flavobacterium | saccharophilum | 2.04 | Flavobacterium | hibernum | 1.97 | A | 0 | 0 | 1 | Flavobacterium saccharophilum | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium | chaponense | Chryseobacterium | chaponense | 1.81 | Chryseobacterium | chaponense | 1.65 | B | 0 | 0 | 1 | Chryseobacterium chaponense | Chryseobacterium chaponense | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium | hominis | Chryseobacterium | hominis | 2.08 | Chryseobacterium | hominis | 1.82 | A | 1 | 1 | 0 | Chryseobacterium hominis | Chryseobacterium hominis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium | manosquense | Deinococcus | wulumuqiensis | 1.45 | Streptomyces | hygroscopicus | 1.23 | B | 0 | 0 | 1 | Deinococcus wulumuqiensis | Streptomyces hygroscopicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 185-U1-6 EFex | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium | manosquense | Staphylococcus | simiae | 1.28 | Chryseobacterium | joostei | 1.25 | B | 0 | 0 | 1 | Staphylococcus simiae | Chryseobacterium joostei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11450,4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium | sp-I-7 | Chryseobacterium | geocarposphaerae | 1.42 | Chryseobacterium | indologenes | 1.3 | B | 0 | 0 | 1 | Chryseobacterium geocarposphaerae | Chryseobacterium indologenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | aquidurense | Flavobacterium | araucananum | 1.89 | Flavobacterium | piscis | 1.88 | B | 0 | 0 | 1 | Flavobacterium araucananum | Flavobacterium piscis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5322,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | branchiarum | Flavobacterium | pectinovorum | 1.94 | Flavobacterium | hibernum | 1.93 | B | 0 | 0 | 1 | Flavobacterium pectinovorum | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | chungangense | Flavobacterium | plurextorum | 1.76 | Flavobacterium | hercynium | 1.72 | B | 0 | 0 | 1 | Flavobacterium plurextorum | Flavobacterium hercynium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium | columnare | Flavobacterium | columnare | 2.3 | Flavobacterium | columnare | 2.16 | A | 1 | 1 | 0 | Flavobacterium columnare | Flavobacterium columnare | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 2.04 | Flavobacterium | hibernum | 1.94 | A | 1 | 1 | 0 | Flavobacterium frigidimaris | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4217,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | frigidimaris | Flavobacterium | frigidimaris | 1.94 | Flavobacterium | hibernum | 1.88 | B | 0 | 0 | 1 | Flavobacterium frigidimaris | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium | piscis | Flavobacterium | piscis | 2.29 | Flavobacterium | piscis | 2.26 | A | 1 | 1 | 0 | Flavobacterium piscis | Flavobacterium piscis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium | plurextorum | Flavobacterium | plurextorum | 2.35 | Flavobacterium | oncorhynchi | 2.25 | B | 0 | 0 | 1 | Flavobacterium plurextorum | Flavobacterium oncorhynchi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5323,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium | psychoterrae | Flavobacterium | hydatis | 1.81 | Flavobacterium | hibernum | 1.68 | B | 0 | 0 | 1 | Flavobacterium hydatis | Flavobacterium hibernum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11550,4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.52 | Flavobacterium | psychrophilum | 2.36 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium | psychrophilum | Flavobacterium | psychrophilum | 2.3 | Flavobacterium | psychrophilum | 2.11 | A | 1 | 1 | 0 | Flavobacterium psychrophilum | Flavobacterium psychrophilum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | reichenbachii | Flavobacterium | reichenbachii | 1.9 | Flavobacterium | cutihirudinis | 1.58 | B | 0 | 0 | 1 | Flavobacterium reichenbachii | Flavobacterium cutihirudinis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | sp-C2 | Paraburkholderia | phenazinium | 1.57 | Myroides | odoratimimus | 1.45 | B | 0 | 0 | 1 | Paraburkholderia phenazinium | Myroides odoratimimus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-10323 | Flavobacterium | chungbukense | 1.65 | Flavobacterium | anhuiense | 1.6 | B | 0 | 0 | 1 | Flavobacterium chungbukense | Flavobacterium anhuiense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium | sp-CVUAS-4146 | Flavobacterium | pectinovorum | 1.96 | Flavobacterium | hercynium | 1.95 | B | 0 | 0 | 1 | Flavobacterium pectinovorum | Flavobacterium hercynium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4216,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-4216,2 | Flavobacterium | degerlachei | 1.46 | Paraburkholderia | xenovorans | 1.4 | B | 0 | 0 | 1 | Flavobacterium degerlachei | Paraburkholderia xenovorans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium | sp-CVUAS-5506 | Flavobacterium | johnsoniae | 1.82 | Flavobacterium | anhuiense | 1.79 | B | 0 | 0 | 1 | Flavobacterium johnsoniae | Flavobacterium anhuiense | 0 | 0 | 0 | 0
|  |  | 1 | 0 | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium | sp-HMD2009 | Limosilactobacillus | oris | 1.4 | Lactobacillus | curvatus | 1.38 | B | 0 | 0 | 1 | Limosilactobacillus oris | Lactobacillus curvatus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium | sp-JQMS | Flavobacterium | limicola | 1.45 | Malikia | spinosa | 1.33 | B | 0 | 0 | 1 | Flavobacterium limicola | Malikia spinosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium | succinicans | Flavobacterium | tiangeerense | 1.78 | Flavobacterium | sinopsychrotolerans | 1.6 | B | 0 | 0 | 1 | Flavobacterium tiangeerense | Flavobacterium sinopsychrotolerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium | tiangeerense | Flavobacterium | tiangeerense | 2.02 | Flavobacterium | glaciei | 1.57 | A | 1 | 1 | 0 | Flavobacterium tiangeerense | Flavobacterium glaciei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella | columbina | Lactobacillus | plantarum | 1.25 | Lactobacillus | plantarum | 1.21 | B | 0 | 0 | 1 | Lactobacillus plantarum | Lactobacillus plantarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2589,2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas | tenax | Chryseobacterium | aquaticum | 1.4 | Aromatoleum | aromaticum | 1.36 | B | 0 | 0 | 1 | Chryseobacterium aquaticum | Aromatoleum aromaticum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.27 | Ornithobacterium | rhinotracheale | 2.24 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium | rhinotracheale | Ornithobacterium | rhinotracheale | 2.21 | Ornithobacterium | rhinotracheale | 2.04 | A | 1 | 1 | 0 | Ornithobacterium rhinotracheale | Ornithobacterium rhinotracheale | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9288,2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca | zopfii | Gordonia | sputi | 1.28 | Tissierella | praeacuta | 1.23 | B | 0 | 0 | 1 | Gordonia sputi | Tissierella praeacuta | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9876,2 | 2; 0 | Pextr | 2016-09-01 | Prototheca | zopfii | Herbaspirillum | huttiense | 1.43 | Bergeyella | zoohelcum | 1.37 | B | 0 | 0 | 1 | Herbaspirillum huttiense | Bergeyella zoohelcum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus | sp-CVUAS-11266 | Acidovorax | facilis | 1.23 | Lacticaseibacillus | paracasei | 1.17 | B | 0 | 0 | 1 | Acidovorax facilis | Lacticaseibacillus paracasei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium | necrophorum | Fusobacterium | necrophorum | 2.48 | Fusobacterium | necrophorum | 2.41 | A | 1 | 1 | 0 | Fusobacterium necrophorum | Fusobacterium necrophorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium | varium | Fusobacterium | varium | 2.09 | Fusobacterium | varium | 2.04 | A | 1 | 1 | 0 | Fusobacterium varium | Fusobacterium varium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.47 | Streptobacillus | moniliformis | 2.43 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter | abscessus | Pseudoclavibacter | caeni | 1.3 | Anaerobiospirillum | succiniciproducens | 1.22 | B | 0 | 0 | 1 | Pseudoclavibacter caeni | Anaerobiospirillum succiniciproducens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga | salmonicida | Candida | parapsilosis | 1.66 | Candida | parapsilosis | 1.59 | B | 0 | 0 | 1 | Candida parapsilosis | Candida parapsilosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus | hongkongensis | Pseudomonas | trivialis | 1.46 | Pseudomonas | oleovorans | 1.44 | B | 0 | 0 | 1 | Pseudomonas trivialis | Pseudomonas oleovorans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella | termitidis | Pichia | kudriavzevii | 1.2 | Acinetobacter | radioresistens | 1.14 | B | 0 | 0 | 1 | Pichia kudriavzevii | Acinetobacter radioresistens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella | termitidis | Clostridium | chauvoei | 1.24 | Novosphingobium | resinovorum | 1.22 | B | 0 | 0 | 1 | Clostridium chauvoei | Novosphingobium resinovorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCUG 38322 CVUAS | 3; public strain collection | EFEx | 16.12.2015 | Sneathia | sanguinegens | Nocardia | asteroides | 1.24 | Pseudomonas | umsongensis | 1.23 | B | 0 | 0 | 1 | Nocardia asteroides | Pseudomonas umsongensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCUG 41628 CVUAS | 3; type strain | EFEx | 2020-10-08 | Sneathia | sanguinegens | Staphylococcus | aureus | 1.25 | Kandleria | vitulina | 1.22 | B | 0 | 0 | 1 | Staphylococcus aureus | Kandleria vitulina | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Arthrobacter | monumenti | 1.52 | Latilactobacillus | fuchuensis | 1.31 | B | 0 | 0 | 1 | Arthrobacter monumenti | Latilactobacillus fuchuensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia | vaginalis | Arthrobacter | monumenti | 1.42 | Paecilomyces | variotii | 1.32 | B | 0 | 0 | 1 | Arthrobacter monumenti | Paecilomyces variotii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus | felis | Streptobacillus | moniliformis | 1.33 | Paenibacillus | cellulosilyticus | 1.31 | B | 0 | 0 | 1 | Streptobacillus moniliformis | Paenibacillus cellulosilyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus | moniliformis | Streptobacillus | moniliformis | 2.58 | Streptobacillus | moniliformis | 2.52 | A | 1 | 1 | 0 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus | notomytis | Streptobacillus | moniliformis | 1.4 | Streptobacillus | moniliformis | 1.34 | B | 0 | 0 | 1 | Streptobacillus moniliformis | Streptobacillus moniliformis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus | ratti | Pseudomonas | abietaniphila | 1.33 | Malikia | spinosa | 1.32 | B | 0 | 0 | 1 | Pseudomonas abietaniphila | Malikia spinosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012806 | 2; 0 | DT | 2017-10-04 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.18 | Ochrobactrum | intermedium | 2.01 | A | 1 | 1 | 0 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum | intermedium | Ochrobactrum | intermedium | 2.55 | Ochrobactrum | intermedium | 2.32 | A | 1 | 1 | 0 | Ochrobactrum intermedium | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum | sp | Ochrobactrum | anthropi | 2.46 | Ochrobactrum | anthropi | 2.41 | A | 0 | 0 | 1 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum | sp | Ochrobactrum | intermedium | 1.78 | Ochrobactrum | anthropi | 1.68 | B | 0 | 0 | 1 | Ochrobactrum intermedium | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012040 Och | 2; 0 | DT | 2017-10-04 | Ochrobactrum | sp | Ochrobactrum | tritici | 1.78 | Ochrobactrum | intermedium | 1.65 | B | 0 | 0 | 1 | Ochrobactrum tritici | Ochrobactrum intermedium | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia | coli | Escherichia | coli | 2.32 | Escherichia | coli | 2.29 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.4 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella | variicola | Klebsiella | variicola | 2.08 | Klebsiella | variicola | 2.05 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1816,2 | 3; sequenced | DT | 2013-07-24 | Basfia | succiniciproducens | Pasteurella | multocida | 1.6 | Pasteurella | stomatis | 1.54 | B | 0 | 0 | 1 | Pasteurella multocida | Pasteurella stomatis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas | alcaligenes | Pseudomonas | alcaligenes | 2.44 | Pseudomonas | alcaligenes | 2.35 | A | 1 | 1 | 0 | Pseudomonas alcaligenes | Pseudomonas alcaligenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas | anguilliseptica | Pseudomonas | anguilliseptica | 2.02 | Pseudomonas | borbori | 1.95 | A | 1 | 1 | 0 | Pseudomonas anguilliseptica | Pseudomonas borbori | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas | brenneri | Pseudomonas | brenneri | 2.14 | Pseudomonas | veronii | 1.84 | A | 1 | 1 | 0 | Pseudomonas brenneri | Pseudomonas veronii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas | fulva | Pseudomonas | fulva | 2.29 | Pseudomonas | fulva | 1.9 | A | 1 | 1 | 0 | Pseudomonas fulva | Pseudomonas fulva | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas | koreensis | Pseudomonas | koreensis | 2.26 | Pseudomonas | corrugata | 1.89 | A | 1 | 1 | 0 | Pseudomonas koreensis | Pseudomonas corrugata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas | monteilii | Pseudomonas | monteilii | 2.01 | Pseudomonas | putida | 1.92 | A | 1 | 1 | 0 | Pseudomonas monteilii | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1_1 nit | 2; 0 | DT | 2017-08-31 | Pseudomonas | nitroreducens | Pseudomonas | nitroreducens | 2.07 | Pseudomonas | nitroreducens | 1.97 | A | 1 | 1 | 0 | Pseudomonas nitroreducens | Pseudomonas nitroreducens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas | oryzihabitans | Pseudomonas | oryzihabitans | 2.12 | Pseudomonas | oryzihabitans | 2.09 | A | 1 | 1 | 0 | Pseudomonas oryzihabitans | Pseudomonas oryzihabitans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171011722-1_2 put | 2; 0 | DT | 2017-08-31 | Pseudomonas | putida | Pseudomonas | putida | 2.49 | Pseudomonas | putida | 2.31 | A | 1 | 1 | 0 | Pseudomonas putida | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas | taetrolens | Pseudomonas | taetrolens | 2 | Pseudomonas | lundensis | 1.95 | A | 1 | 1 | 0 | Pseudomonas taetrolens | Pseudomonas lundensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2280,3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas | glaciistagni | Limosilactobacillus | gastricus | 1.28 | Agromyces | rhizospherae | 1.26 | B | 0 | 0 | 1 | Limosilactobacillus gastricus | Agromyces rhizospherae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_08 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 02.04.2012 K8 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.42 | Ochrobactrum | anthropi | 2.39 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM373 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; D02 | Brucella | abortus | Ligilactobacillus | ruminis | 1.43 | Ligilactobacillus | ruminis | 1.34 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Ligilactobacillus ruminis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM369 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; E01 | Brucella | abortus | Staphylococcus | schleiferi | 1.39 | Ligilactobacillus | acidipiscis | 1.37 | B | 0 | 0 | 1 | Staphylococcus schleiferi | Ligilactobacillus acidipiscis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; I16 | Brucella | canis | Agromyces | fucosus | 1.36 | Ochrobactrum | daejeonense | 1.35 | B | 0 | 0 | 1 | Agromyces fucosus | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.03.2014; E17 | Brucella | canis | Limosilactobacillus | reuteri | 1.46 | Ochrobactrum | daejeonense | 1.41 | B | 0 | 0 | 1 | Limosilactobacillus reuteri | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM389 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; F04 | Brucella | ceti | Ochrobactrum | daejeonense | 1.43 | Ochrobactrum | gallinifaecis | 1.42 | B | 0 | 0 | 1 | Ochrobactrum daejeonense | Ochrobactrum gallinifaecis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM403 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; K04 | Brucella | ceti | Lactobacillus | graminis | 1.44 | Pseudomonas | jinjuensis | 1.41 | B | 0 | 0 | 1 | Lactobacillus graminis | Pseudomonas jinjuensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM374 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I02 | Brucella | melitensis | Staphylococcus | sciuri | 1.34 | Agrobacterium | radiobacter | 1.32 | B | 0 | 0 | 1 | Staphylococcus sciuri | Agrobacterium radiobacter | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM407 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; J05 | Brucella | melitensis | Ochrobactrum | anthropi | 1.42 | Bacteroides | thetaiotaomicron | 1.39 | B | 0 | 0 | 1 | Ochrobactrum anthropi | Bacteroides thetaiotaomicron | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM408 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; H06 | Brucella | melitensis | Schleiferilactobacillus | harbinensis | 1.43 | Ochrobactrum | daejeonense | 1.42 | B | 0 | 0 | 1 | Schleiferilactobacillus harbinensis | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM410 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I06 | Brucella | melitensis | Bacteroides | fragilis | 1.37 | Ochrobactrum | tritici | 1.29 | B | 0 | 0 | 1 | Bacteroides fragilis | Ochrobactrum tritici | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CCM 4915 CVUAS | 3; type strain | DT | 2013-11-26 | Brucella | microti | Lacticaseibacillus | pantheris | 1.5 | Ochrobactrum | daejeonense | 1.47 | B | 0 | 0 | 1 | Lacticaseibacillus pantheris | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCM 4915 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.03.2015 C7 | Brucella | microti | Ligilactobacillus | aviarius | 1.35 | Ochrobactrum | daejeonense | 1.34 | B | 0 | 0 | 1 | Ligilactobacillus aviarius | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM421 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; F07 | Brucella | neotomae | Ligilactobacillus | ruminis | 1.49 | Stenotrophomonas | acidaminiphila | 1.48 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Stenotrophomonas acidaminiphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.03.2015; C15 | Brucella | neotomae | Klebsiella | pneumoniae | 1.52 | Ochrobactrum | daejeonense | 1.52 | B | 0 | 0 | 1 | Klebsiella pneumoniae | Ochrobactrum daejeonense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM423 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; L07 | Brucella | ovis | Ochrobactrum | daejeonense | 1.43 | Pseudomonas | jinjuensis | 1.3 | B | 0 | 0 | 1 | Ochrobactrum daejeonense | Pseudomonas jinjuensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; F16 | Brucella | ovis | Ochrobactrum | tritici | 1.35 | Lactobacillus | pentosus | 1.31 | B | 0 | 0 | 1 | Ochrobactrum tritici | Lactobacillus pentosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM405 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; E05 | Brucella | pinnipedialis | Staphylococcus | vitulinus | 1.5 | Corynebacterium | flavescens | 1.45 | B | 0 | 0 | 1 | Staphylococcus vitulinus | Corynebacterium flavescens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; C03 | Brucella | pinnipedialis | Pandoraea | norimbergensis | 1.4 | Pannonibacter | phragmitetus | 1.33 | B | 0 | 0 | 1 | Pandoraea norimbergensis | Pannonibacter phragmitetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM431 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; B10 | Brucella | suis | Ligilactobacillus | ruminis | 1.39 | Actinocorallia | libanotica | 1.38 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Actinocorallia libanotica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM427 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; G08 | Brucella | suis | Proteus | vulgaris | 1.36 | Pseudomonas | putida | 1.34 | B | 0 | 0 | 1 | Proteus vulgaris | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella | suis | Ochrobactrum | anthropi | 1.23 | Ochrobactrum | anthropi | 1.19 | B | 0 | 0 | 1 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM428 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I08 | Brucella | suis | Lactobacillus | delbrueckii | 1.37 | Lactobacillus | delbrueckii | 1.34 | B | 0 | 0 | 1 | Lactobacillus delbrueckii | Lactobacillus delbrueckii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BM430 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; I09 | Brucella | suis | Ligilactobacillus | ruminis | 1.34 | Acinetobacter | calcoaceticus | 1.29 | B | 0 | 0 | 1 | Ligilactobacillus ruminis | Acinetobacter calcoaceticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 15.15.2016 I7 | Brucella | vulpis | Pannonibacter | phragmitetus | 1.36 | Paenibacillus | turicensis | 1.33 | B | 0 | 0 | 1 | Pannonibacter phragmitetus | Paenibacillus turicensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum | ovis | Falsochrobactrum | ovis | 2.3 | Shewanella | profunda | 1.21 | A | 1 | 1 | 0 | Falsochrobactrum ovis | Shewanella profunda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 03.06.2016 H12 | Ochrobactrum | anthropi | Ochrobactrum | anthropi | 2.54 | Ochrobactrum | anthropi | 2.42 | A | 1 | 1 | 0 | Ochrobactrum anthropi | Ochrobactrum anthropi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter | aganoensis | Pseudomonas | mucidolens | 1.45 | Pseudomonas | orientalis | 1.44 | B | 0 | 0 | 1 | Pseudomonas mucidolens | Pseudomonas orientalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter | aminovorans | Lactobacillus | curvatus | 1.5 | Aromatoleum | toluolicum | 1.42 | B | 0 | 0 | 1 | Lactobacillus curvatus | Aromatoleum toluolicum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2015 H22 | Paracoccus | yeei | Paracoccus | yeei | 2.45 | Paracoccus | yeei | 2.24 | A | 1 | 1 | 0 | Paracoccus yeei | Paracoccus yeei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas | mucosa | Roseomonas | mucosa | 2.53 | Roseomonas | mucosa | 2.49 | A | 1 | 1 | 0 | Roseomonas mucosa | Roseomonas mucosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 24.01.2007, H7, [0034] | Inquilinus | limosus | Inquilinus | limosus | 2.63 | Inquilinus | limosus | 2.62 | A | 1 | 1 | 0 | Inquilinus limosus | Inquilinus limosus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.32 | Taylorella | equigenitalis | 1.32 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter | denitrificans | Achromobacter | denitrificans | 2.22 | Achromobacter | denitrificans | 2.04 | A | 1 | 1 | 0 | Achromobacter denitrificans | Achromobacter denitrificans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter | sp-DSM-30031 | Achromobacter | piechaudii | 1.95 | Achromobacter | spanius | 1.9 | B | 0 | 0 | 1 | Achromobacter piechaudii | Achromobacter spanius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2008 E5 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.33 | Achromobacter | xylosoxidans | 2.16 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.02.2016 N7 | Achromobacter | xylosoxidans | Achromobacter | xylosoxidans | 2.26 | Achromobacter | xylosoxidans | 2.23 | A | 1 | 1 | 0 | Achromobacter xylosoxidans | Achromobacter xylosoxidans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea | sp-CVUAS-31213 | Streptomyces | badius | 1.35 | Staphylococcus | aureus | 1.32 | B | 0 | 0 | 1 | Streptomyces badius | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella | avium | Bordetella | avium | 2.39 | Bordetella | avium | 2.29 | A | 1 | 1 | 0 | Bordetella avium | Bordetella avium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32946 | 3; confirmed proficency test | EFEx | 2021-02-18 | Kerstersia | gyiorum | Kerstersia | gyiorum | 2.48 | Kerstersia | gyiorum | 2.48 | A | 1 | 1 | 0 | Kerstersia gyiorum | Kerstersia gyiorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 1.67 | Oligella | urethralis | 1.66 | B | 0 | 0 | 1 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella | urethralis | Oligella | urethralis | 1.6 | Oligella | urethralis | 1.56 | B | 0 | 0 | 1 | Oligella urethralis | Oligella urethralis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga | kullae | Pigmentiphaga | daeguensis | 1.57 | Clostridium | spiroforme | 1.37 | B | 0 | 0 | 1 | Pigmentiphaga daeguensis | Clostridium spiroforme | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella | asinigenitalis | Taylorella | asinigenitalis | 2.09 | Taylorella | equigenitalis | 1.28 | A | 1 | 1 | 0 | Taylorella asinigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.15 | Taylorella | equigenitalis | 2.1 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella | equigenitalis | Taylorella | equigenitalis | 2.35 | Taylorella | equigenitalis | 2.24 | A | 1 | 1 | 0 | Taylorella equigenitalis | Taylorella equigenitalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012; K10 | Burkholderia | pseudomallei | Burkholderia | thailandensis | 2.03 | Staphylococcus | aureus | 1.49 | A | 0 | 0 | 1 | Burkholderia thailandensis | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; H21 | Burkholderia | pseudomallei | Burkholderia | thailandensis | 2.03 | Paeniclostridium | sordellii | 1.49 | A | 0 | 0 | 1 | Burkholderia thailandensis | Paeniclostridium sordellii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.04.2012, K6 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.38 | Burkholderia | ambifaria | 1.76 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007, K7 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.27 | Burkholderia | ambifaria | 2.24 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16087 LMG 19182 | 3; type starin | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K19 | Burkholderia | ambifaria | Burkholderia | ambifaria | 2.36 | Burkholderia | ambifaria | 2.15 | A | 1 | 1 | 0 | Burkholderia ambifaria | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K21 | Burkholderia | anthina | Burkholderia | anthina | 2.35 | Burkholderia | anthina | 2.27 | A | 1 | 1 | 0 | Burkholderia anthina | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.08.2007, N10 | Burkholderia | anthina | Burkholderia | pyrrocinia | 2.17 | Burkholderia | anthina | 2.16 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia anthina | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2010, K12 | Burkholderia | arboris | Burkholderia | pyrrocinia | 2.34 | Burkholderia | cepacia | 2.23 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 3.4.2017; L6 | Burkholderia | arboris | Burkholderia | cenocepacia | 2.29 | Burkholderia | vietnamiensis | 2.22 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K11 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.48 | Burkholderia | cenocepacia | 2.36 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.07.2009; K21 | Burkholderia | cenocepacia | Burkholderia | cenocepacia | 2.42 | Burkholderia | cenocepacia | 2.17 | A | 1 | 1 | 0 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 4.5.2006; H15 | Burkholderia | cepacia | Burkholderia | cenocepacia | 2.46 | Burkholderia | cenocepacia | 2.15 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 9.5.2007; N9 | Burkholderia | cepacia | Burkholderia | cepacia | 2.4 | Burkholderia | pyrrocinia | 2.34 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.11.2006; H10 | Burkholderia | contaminans | Burkholderia | cepacia | 2.28 | Burkholderia | cepacia | 2.16 | A | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 01 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E10 | Burkholderia | contaminans | Burkholderia | cepacia | 2.42 | Burkholderia | cenocepacia | 2.17 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016 C13 | Burkholderia | diffusa | Burkholderia | ambifaria | 2.35 | Burkholderia | vietnamiensis | 2.32 | B | 0 | 0 | 1 | Burkholderia ambifaria | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L17 | Burkholderia | diffusa | Burkholderia | cenocepacia | 2.31 | Burkholderia | ambifaria | 2.21 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K16 | Burkholderia | dolosa | Burkholderia | dolosa | 2.29 | Burkholderia | multivorans | 2.15 | B | 0 | 0 | 1 | Burkholderia dolosa | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.11.2006; E09 | Burkholderia | dolosa | Burkholderia | dolosa | 2.08 | Burkholderia | cenocepacia | 2.07 | B | 0 | 0 | 1 | Burkholderia dolosa | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.01.2007; H22 | Burkholderia | gladioli | Burkholderia | gladioli | 2.45 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H15 | Burkholderia | gladioli | Burkholderia | gladioli | 2.42 | Burkholderia | gladioli | 2.27 | A | 1 | 1 | 0 | Burkholderia gladioli | Burkholderia gladioli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K03 | Burkholderia | glumae | Burkholderia | glumae | 2.33 | Burkholderia | plantarii | 2.06 | B | 0 | 0 | 1 | Burkholderia glumae | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.05.2012; K22 | Burkholderia | glumae | Burkholderia | glumae | 2.33 | Burkholderia | plantarii | 2.13 | B | 0 | 0 | 1 | Burkholderia glumae | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016; C10 | Burkholderia | lata | Burkholderia | cepacia | 2.49 | Burkholderia | lata | 2.41 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia lata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.08.2016; K19 | Burkholderia | lata | Burkholderia | cepacia | 2.47 | Burkholderia | lata | 2.38 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia lata | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.05.2016; L06 | Burkholderia | latens | Burkholderia | latens | 2.32 | Burkholderia | vietnamiensis | 2.29 | B | 0 | 0 | 1 | Burkholderia latens | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.09.2009; K20 | Burkholderia | mallei | Burkholderia | thailandensis | 1.65 | Staphylococcus | aureus | 1.38 | B | 0 | 0 | 1 | Burkholderia thailandensis | Staphylococcus aureus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.08.2007; H13 | Burkholderia | mallei | Burkholderia | thailandensis | 1.73 | Lactobacillus | paralimentarius | 1.38 | B | 0 | 0 | 1 | Burkholderia thailandensis | Lactobacillus paralimentarius | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2010; H14 | Burkholderia | metallica | Burkholderia | metallica | 2.4 | Burkholderia | cenocepacia | 2.27 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L21 | Burkholderia | metallica | Burkholderia | metallica | 2.31 | Burkholderia | cenocepacia | 2.23 | B | 0 | 0 | 1 | Burkholderia metallica | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.12.2006; K19 | Burkholderia | multivorans | Burkholderia | multivorans | 2.22 | Burkholderia | multivorans | 2.21 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.11.2006; K04 | Burkholderia | multivorans | Burkholderia | multivorans | 2.13 | Burkholderia | multivorans | 2 | A | 1 | 1 | 0 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K17 | Burkholderia | oklahomensis | Burkholderia | multivorans | 1.77 | Burkholderia | ambifaria | 1.76 | B | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K18 | Burkholderia | oklahomensis | Burkholderia | thailandensis | 1.72 | Burkholderia | ambifaria | 1.65 | B | 0 | 0 | 1 | Burkholderia thailandensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K06 | Burkholderia | plantarii | Burkholderia | plantarii | 2.49 | Burkholderia | pyrrocinia | 1.89 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K05 | Burkholderia | plantarii | Burkholderia | plantarii | 2.5 | Burkholderia | pyrrocinia | 1.94 | A | 1 | 1 | 0 | Burkholderia plantarii | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K07 | Burkholderia | pseudomultivorans | Burkholderia | multivorans | 2.1 | Burkholderia | multivorans | 2.07 | A | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia multivorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K08 | Burkholderia | pseudomultivorans | Burkholderia | multivorans | 2.03 | Burkholderia | cenocepacia | 1.93 | A | 0 | 0 | 1 | Burkholderia multivorans | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CAMPA 1043 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; K04 | Burkholderia | puraquae | Burkholderia | pyrrocinia | 2.27 | Burkholderia | cepacia_Group | 2.2 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | R-50214 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.03.2016; C04 | Burkholderia | puraquae | Burkholderia | pyrrocinia | 2.41 | Burkholderia | cepacia_Group | 2.32 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; N06 | Burkholderia | pyrrocinia | Burkholderia | pyrrocinia | 2.57 | Burkholderia | cepacia_Group | 2.33 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cepacia_Group | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; I21 | Burkholderia | pyrrocinia | Burkholderia | cenocepacia | 2.13 | Burkholderia | pyrrocinia | 2.12 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H08 | Burkholderia | seminalis | Burkholderia | seminalis | 2.41 | Burkholderia | cepacia | 2.32 | B | 0 | 0 | 1 | Burkholderia seminalis | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.01.2007; K12 | Burkholderia | seminalis | Burkholderia | cepacia | 2.29 | Burkholderia | seminalis | 2.11 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia seminalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K13 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.56 | Burkholderia | stabilis | 2.44 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E08 | Burkholderia | stabilis | Burkholderia | cepacia_Group | 2.51 | Burkholderia | stabilis | 2.28 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia stabilis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; C08 | Burkholderia | stagnalis | Burkholderia | anthina | 2.14 | Burkholderia | ambifaria | 2.11 | B | 0 | 0 | 1 | Burkholderia anthina | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F14 | Burkholderia | stagnalis | Burkholderia | vietnamiensis | 2.16 | Burkholderia | ambifaria | 2.01 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia ambifaria | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007; H19 | Burkholderia | taxonK | Burkholderia | cepacia | 2.45 | Burkholderia | cenocepacia | 2.16 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.06.2009; K13 | Burkholderia | taxonK | Burkholderia | cepacia | 2.23 | Burkholderia | cenocepacia | 2.13 | B | 0 | 0 | 1 | Burkholderia cepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F16 | Burkholderia | territorii | Burkholderia | cenocepacia | 2.3 | Burkholderia | cenocepacia | 2.05 | A | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; E14 | Burkholderia | territorii | Burkholderia | cenocepacia | 2.28 | Burkholderia | vietnamiensis | 2.21 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia vietnamiensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13276 n RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 24.11.2010, K19 | Burkholderia | thailandensis | Burkholderia | thailandensis | 2.34 | Burkholderia | glumae | 1.59 | A | 1 | 1 | 0 | Burkholderia thailandensis | Burkholderia glumae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H11 | Burkholderia | ubonensis | Burkholderia | cenocepacia | 2.19 | Burkholderia | cepacia | 2.07 | B | 0 | 0 | 1 | Burkholderia cenocepacia | Burkholderia cepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; I04 | Burkholderia | ubonensis | Burkholderia | cepacia_Group | 2.1 | Burkholderia | plantarii | 2.09 | B | 0 | 0 | 1 | Burkholderia cepacia_Group | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CCUG 9631=LMG 6999 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.05.2007 N15 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.33 | Burkholderia | cenocepacia | 2.2 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 10929 V | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.12.2006 N9 | Burkholderia | vietnamiensis | Burkholderia | vietnamiensis | 2.43 | Burkholderia | pyrrocinia | 2.14 | B | 0 | 0 | 1 | Burkholderia vietnamiensis | Burkholderia pyrrocinia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.04.2012, N07 | Caballeronia | glathei | Caballeronia | glathei | 2.55 | Caballeronia | turbans | 1.53 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H12 | Caballeronia | glathei | Caballeronia | glathei | 2.47 | Caballeronia | turbans | 1.56 | A | 1 | 1 | 0 | Caballeronia glathei | Caballeronia turbans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H15 | Pandoraea | apista | Pandoraea | apista | 2.36 | Pandoraea | apista | 2.28 | A | 1 | 1 | 0 | Pandoraea apista | Pandoraea apista | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H21 | Pandoraea | pnomenusa | Pandoraea | pnomenusa | 2.4 | Pandoraea | pnomenusa | 2.29 | A | 1 | 1 | 0 | Pandoraea pnomenusa | Pandoraea pnomenusa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H16 | Pandoraea | pulmonicola | Pandoraea | pulmonicola | 2.49 | Pandoraea | pulmonicola | 2.25 | A | 1 | 1 | 0 | Pandoraea pulmonicola | Pandoraea pulmonicola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K7 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.52 | Paraburkholderia | caledonica | 2.45 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K8 | Paraburkholderia | caledonica | Paraburkholderia | phymatum | 2.46 | Paraburkholderia | caledonica | 2.37 | B | 0 | 0 | 1 | Paraburkholderia phymatum | Paraburkholderia caledonica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N9 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.25 | Burkholderia | thailandensis | 1.6 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N10 | Paraburkholderia | caribensis | Paraburkholderia | caribensis | 2.23 | Burkholderia | thailandensis | 1.42 | A | 1 | 1 | 0 | Paraburkholderia caribensis | Burkholderia thailandensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H03 | Paraburkholderia | hospita | Lactobacillus | fructivorans | 1.71 | Herbaspirillum | frisingense | 1.44 | B | 0 | 0 | 1 | Lactobacillus fructivorans | Herbaspirillum frisingense | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H04 | Paraburkholderia | hospita | Paraburkholderia | phenazinium | 1.6 | Paraburkholderia | caribensis | 1.58 | B | 0 | 0 | 1 | Paraburkholderia phenazinium | Paraburkholderia caribensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 1.12.2010; H15 | Paraburkholderia | kururiensis | Burkholderia | pyrrocinia | 1.64 | Burkholderia | cenocepacia | 1.54 | B | 0 | 0 | 1 | Burkholderia pyrrocinia | Burkholderia cenocepacia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.11.2010; N19 | Paraburkholderia | phymatum | Paraburkholderia | caribensis | 1.7 | Paraburkholderia | fungorum | 1.51 | B | 0 | 0 | 1 | Paraburkholderia caribensis | Paraburkholderia fungorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.11.2010; K17 | Paraburkholderia | phytofirmans | Paraburkholderia | xenovorans | 2.25 | Paraburkholderia | phymatum | 2 | B | 0 | 0 | 1 | Paraburkholderia xenovorans | Paraburkholderia phymatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H16 | Paraburkholderia | tuberum | Paraburkholderia | tuberum | 2.36 | Paraburkholderia | xenovorans | 1.59 | A | 1 | 1 | 0 | Paraburkholderia tuberum | Paraburkholderia xenovorans | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.11.2010, N21 | Paraburkholderia | xenovorans | Paraburkholderia | xenovorans | 2.49 | Paraburkholderia | phymatum | 2.06 | B | 0 | 0 | 1 | Paraburkholderia xenovorans | Paraburkholderia phymatum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 E07 | Ralstonia | insidiosa | Ralstonia | insidiosa | 2.38 | Ralstonia | insidiosa | 2.35 | A | 1 | 1 | 0 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 C06 | Ralstonia | mannitolilytica | Ralstonia | mannitolilytica | 2.46 | Ralstonia | mannitolilytica | 2.22 | A | 1 | 1 | 0 | Ralstonia mannitolilytica | Ralstonia mannitolilytica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27511 DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 22.01.2007 H03 | Ralstonia | pickettii | Ralstonia | pickettii | 2.15 | Ralstonia | pickettii | 2.12 | A | 1 | 1 | 0 | Ralstonia pickettii | Ralstonia pickettii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 07.05.2009 N15 | Ralstonia | pickettii | Ralstonia | insidiosa | 2.07 | Ralstonia | insidiosa | 2.07 | A | 0 | 0 | 1 | Ralstonia insidiosa | Ralstonia insidiosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia | acidovorans | Delftia | acidovorans | 2.54 | Delftia | tsuruhatensis | 2.44 | B | 0 | 0 | 1 | Delftia acidovorans | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia | tsuruhatensis | Delftia | tsuruhatensis | 2.31 | Delftia | tsuruhatensis | 2.23 | A | 1 | 1 | 0 | Delftia tsuruhatensis | Delftia tsuruhatensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga | intermedia | Hydrogenophaga | intermedia | 2.4 | Novosphingobium | resinovorum | 1.34 | A | 1 | 1 | 0 | Hydrogenophaga intermedia | Novosphingobium resinovorum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.07.2013 L13 | Xenophilus | sp | Aromatoleum | evansii | 1.43 | Ligilactobacillus | salivarius | 1.32 | B | 0 | 0 | 1 | Aromatoleum evansii | Ligilactobacillus salivarius | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 24713 EFex | 3; type strain | Extr | 2016-02-02 | Deefgea | chitinilytica | Mannheimia | haemolytica | 1.42 | Clostridium | cochlearium | 1.32 | B | 0 | 0 | 1 | Mannheimia haemolytica | Clostridium cochlearium | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 24713 eDT | 3; type strain | eDT | 2015-09-22 | Deefgea | chitinilytica | Pichia | cactophila | 1.24 | Prevotella | oralis | 1.14 | B | 0 | 0 | 1 | Pichia cactophila | Prevotella oralis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea | rivuli | Arthrobacter | woluwensis | 1.45 | Methylobacterium | mesophilicum | 1.41 | B | 0 | 0 | 1 | Arthrobacter woluwensis | Methylobacterium mesophilicum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 08-F17,2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Lactobacillus | crispatus | 1.31 | Agromyces | salentinus | 1.27 | B | 0 | 0 | 1 | Lactobacillus crispatus | Agromyces salentinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea | sp-CVUAS-2703,2 | Lactobacillus | crispatus | 1.38 | Pseudomonas | putida | 1.33 | B | 0 | 0 | 1 | Lactobacillus crispatus | Pseudomonas putida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6995,3 | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea | sp-CVUAS-6995,3 | Pseudomonas | chlororaphis | 1.5 | Pseudomonas | orientalis | 1.45 | B | 0 | 0 | 1 | Pseudomonas chlororaphis | Pseudomonas orientalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-6995,3 | Pseudomonas | orientalis | 1.48 | Corynebacterium | ulcerans | 1.36 | B | 0 | 0 | 1 | Pseudomonas orientalis | Corynebacterium ulcerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea | sp-CVUAS-9920 | Flavobacterium | flevense | 1.27 | Bacteroides | pyogenes | 1.27 | B | 0 | 0 | 1 | Flavobacterium flevense | Bacteroides pyogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea | sp-CVUAS-9920 | Raoultella | ornithinolytica | 1.22 | Pediococcus | acidilactici | 1.21 | B | 0 | 0 | 1 | Raoultella ornithinolytica | Pediococcus acidilactici | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter | limnosediminis | Iodobacter | fluviatilis | 1.45 | Aromatoleum | anaerobicus | 1.36 | B | 0 | 0 | 1 | Iodobacter fluviatilis | Aromatoleum anaerobicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella | kingae | Kingella | kingae | 2.49 | Kingella | kingae | 2.41 | A | 1 | 1 | 0 | Kingella kingae | Kingella kingae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10067,2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria | animaloris | Neisseria | animaloris | 2.17 | Neisseria | animaloris | 2.12 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria | animaloris | Neisseria | animaloris | 2.42 | Neisseria | animaloris | 2.4 | A | 1 | 1 | 0 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-4931 | Neisseria | animaloris | 1.42 | Neisseria | animaloris | 1.4 | B | 0 | 0 | 1 | Neisseria animaloris | Neisseria animaloris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-CVUAS-8997 | Neisseria_flavescens_subflava | group | 1.48 | Providencia | rettgeri | 1.45 | B | 0 | 0 | 1 | Neisseria_flavescens_subflava group | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria | sp-LHL-1296 | Neisseria | animaloris | 1.61 | Neisseria | cinerea | 1.37 | B | 0 | 0 | 1 | Neisseria animaloris | Neisseria cinerea | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5699,3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria | sp-N95-16 | Bergeriella | denitrificans | 1.99 | Neisseria_flavescens_subflava | group | 1.73 | C | 0 | 0 | 1 | Bergeriella denitrificans | Neisseria_flavescens_subflava group | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae | sp-CVUAS-8636 | Pseudomonas | taetrolens | 1.43 | Pseudomonas | fluorescens | 1.37 | B | 0 | 0 | 1 | Pseudomonas taetrolens | Pseudomonas fluorescens | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter | fetus | Campylobacter | fetus | 2.34 | Campylobacter | fetus | 2.28 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter | fetus | Campylobacter | fetus | 2.39 | Campylobacter | fetus | 2.37 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.19 | Campylobacter | hyointestinalis | 2.12 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-07048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Burkholderia | plantarii | 1.42 | A | 1 | 1 | 0 | Campylobacter peloridis | Burkholderia plantarii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | peloridis | Campylobacter | peloridis | 2.33 | Campylobacter | lari | 1.43 | A | 1 | 1 | 0 | Campylobacter peloridis | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter | cryaerophilus | Arcobacter | cryaerophilus | 2.03 | Arcobacter | cryaerophilus | 2.01 | A | 0 | 0 | 1 | Arcobacter cryaerophilus | Arcobacter cryaerophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.06 | Aliarcobacter | butzleri | 1.94 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter | butzleri | Aliarcobacter | butzleri | 2.24 | Aliarcobacter | butzleri | 2.21 | A | 0 | 0 | 1 | Aliarcobacter butzleri | Aliarcobacter butzleri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter | coli | Campylobacter | coli | 2.18 | Campylobacter | coli | 1.97 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4689 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter | coli | Campylobacter | coli | 2.24 | Campylobacter | coli | 2.21 | A | 1 | 1 | 0 | Campylobacter coli | Campylobacter coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter | jejuni | Campylobacter | jejuni | 2.19 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter | jejuni | Campylobacter | jejuni | 2.39 | Campylobacter | jejuni | 2.32 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter | jejuni | Campylobacter | jejuni | 2.16 | Campylobacter | jejuni | 2.14 | A | 1 | 1 | 0 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter | lari | Campylobacter | lari | 2.19 | Campylobacter | lari | 2.04 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-06556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.04 | Campylobacter | lari | 1.82 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter | lari | Campylobacter | lari | 2.2 | Campylobacter | lari | 2.11 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 11375 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter | lari | Campylobacter | lari | 2.42 | Campylobacter | lari | 2.19 | A | 1 | 1 | 0 | Campylobacter lari | Campylobacter lari | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 5361 CVUAS | 3; type strain | DT | 2013-09-24 | Campylobacter | fetus | Campylobacter | fetus | 2.51 | Campylobacter | fetus | 2.38 | A | 1 | 1 | 0 | Campylobacter fetus | Campylobacter fetus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6673,2 | 3; sequenced | DT | 2013-08-28 | Campylobacter | hyointestinalis | Campylobacter | hyointestinalis | 2.15 | Campylobacter | hyointestinalis | 2.06 | A | 1 | 1 | 0 | Campylobacter hyointestinalis | Campylobacter hyointestinalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter | sputorum | Campylobacter | sputorum | 2.01 | Campylobacter | sputorum | 1.9 | A | 1 | 1 | 0 | Campylobacter sputorum | Campylobacter sputorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter | sp-W1 | Campylobacter | jejuni | 1.49 | Campylobacter | jejuni | 1.48 | B | 0 | 0 | 1 | Campylobacter jejuni | Campylobacter jejuni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter | upsaliensis | Campylobacter | upsaliensis | 2.05 | Campylobacter | upsaliensis | 1.96 | A | 1 | 1 | 0 | Campylobacter upsaliensis | Campylobacter upsaliensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11185,3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter | canadensis | Clostridium | cochlearium | 1.42 | Helicobacter | pullorum | 1.4 | B | 0 | 0 | 1 | Clostridium cochlearium | Helicobacter pullorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas | bestiarum | Aeromonas | bestiarum | 2.34 | Aeromonas | popoffii | 2.15 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas popoffii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas | bivalvium | Aeromonas | veronii | 1.93 | Aeromonas | ichthiosmia | 1.9 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas ichthiosmia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas | caviae | Aeromonas | caviae | 2.29 | Aeromonas | caviae | 2.25 | A | 1 | 1 | 0 | Aeromonas caviae | Aeromonas caviae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | hydrophila | 2.03 | Aeromonas | jandaei | 1.96 | A | 0 | 0 | 1 | Aeromonas hydrophila | Aeromonas jandaei | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas | dhakensis | Aeromonas | hydrophila | 2.13 | Aeromonas | hydrophila | 2.11 | A | 0 | 0 | 1 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30187 RKI | 4; type strain, WGS sequenced (BfR) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E17 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.57 | Aeromonas | hydrophila | 2.45 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas | hydrophila | Aeromonas | hydrophila | 2.25 | Aeromonas | hydrophila | 2.09 | A | 1 | 1 | 0 | Aeromonas hydrophila | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas | jandaei | Aeromonas | jandaei | 2.05 | Aeromonas | hydrophila | 1.97 | A | 1 | 1 | 0 | Aeromonas jandaei | Aeromonas hydrophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | media | Aeromonas | media | 2.2 | Aeromonas | veronii | 2.13 | B | 0 | 0 | 1 | Aeromonas media | Aeromonas veronii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas | molluscorum | Aeromonas | molluscorum | 2.15 | Aeromonas | molluscorum | 1.89 | A | 1 | 1 | 0 | Aeromonas molluscorum | Aeromonas molluscorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas | popoffii | Aeromonas | popoffii | 2.11 | Aeromonas | bestiarum | 2.05 | B | 0 | 0 | 1 | Aeromonas popoffii | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 1.89 | Aeromonas | salmonicida | 1.86 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas salmonicida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.1 | Aeromonas | bestiarum | 2.02 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | bestiarum | 2.1 | Aeromonas | popoffii | 2.1 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas popoffii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas | salmonicida | Aeromonas | bestiarum | 2.04 | Aeromonas | eucrenophila | 2.04 | B | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas eucrenophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8400,2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.19 | Aeromonas | bestiarum | 2.05 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.14 | Aeromonas | bestiarum | 2.12 | B | 0 | 0 | 1 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas | salmonicida | Aeromonas | bestiarum | 2.11 | Aeromonas | bestiarum | 2.02 | A | 0 | 0 | 1 | Aeromonas bestiarum | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas | salmonicida | Aeromonas | salmonicida | 2.01 | Aeromonas | bestiarum | 1.96 | A | 1 | 1 | 0 | Aeromonas salmonicida | Aeromonas bestiarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas | simiae | Aeromonas | simiae | 2.07 | Rhizorhabdus | wittichii | 1.68 | A | 1 | 1 | 0 | Aeromonas simiae | Rhizorhabdus wittichii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas | veronii | Aeromonas | ichthiosmia | 2.18 | Aeromonas | veronii | 2.17 | B | 0 | 0 | 1 | Aeromonas ichthiosmia | Aeromonas veronii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7386 CVUAS | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas | veronii | Aeromonas | veronii | 1.94 | Aeromonas | enteropelogenes | 1.91 | B | 0 | 0 | 1 | Aeromonas veronii | Aeromonas enteropelogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas | pectinilytica | Aeromonas | eucrenophila | 1.4 | Pseudaeromonas | sharmana | 1.38 | B | 0 | 0 | 1 | Aeromonas eucrenophila | Pseudaeromonas sharmana | 0 | 0 | 0 | 0
|  |  | 1 | 0 | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella | viscosa | Moritella | japonica | 1.97 | Moritella | dasanensis | 1.51 | B | 0 | 0 | 1 | Moritella japonica | Moritella dasanensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella | viscosa | Moritella | japonica | 1.62 | Pseudomonas | tolaasii | 1.27 | B | 0 | 0 | 1 | Moritella japonica | Pseudomonas tolaasii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | profunda | 1.46 | Lactobacillus | paraplantarum | 1.45 | B | 0 | 0 | 1 | Shewanella profunda | Lactobacillus paraplantarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella | sp-CVUAS-9366 | Shewanella | baltica | 1.42 | Shewanella | profunda | 1.41 | B | 0 | 0 | 1 | Shewanella baltica | Shewanella profunda | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 81,2 | 3; sequenced | EFEx | 28:11.2013 | Shewanella | sp-RKVR22 | Shewanella | corallii | 1.59 | Shewanella | corallii | 1.55 | B | 0 | 0 | 1 | Shewanella corallii | Shewanella corallii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter | condimenti | Cronobacter | sp | 2.14 | Cronobacter | sp | 2.13 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter | dublinensis | Cronobacter | sp | 2.2 | Cronobacter | sp | 2.19 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18702 LALLF | 3; type strain | DT | 2019-03-05 | Cronobacter | malonaticus | Cronobacter | sp | 2.46 | Cronobacter | sp | 2.42 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter | turicensis | Cronobacter | sp | 2.36 | Cronobacter | sp | 2.29 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.11.2011 E6 | Enterobacter | cloacae | Enterobacter | hormaechei | 2.06 | Enterobacter | hormaechei | 2.04 | A | 0 | 0 | 1 | Enterobacter hormaechei | Enterobacter hormaechei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, H5 | Citrobacter | amalonaticus | Citrobacter | amalonaticus | 2.41 | Citrobacter | amalonaticus | 2.3 | A | 1 | 1 | 0 | Citrobacter amalonaticus | Citrobacter amalonaticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | freundii | 2.39 | Citrobacter | freundii | 2.37 | A | 1 | 1 | 0 | Citrobacter freundii | Citrobacter freundii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter | freundii | Citrobacter | braakii | 2.29 | Citrobacter | braakii | 2.23 | A | 0 | 0 | 1 | Citrobacter braakii | Citrobacter braakii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25408 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 11.04.2008, N11, [0034]: Citrobacter diversus ATCC 25408\LMTG70N11_0\0_N11\1\1SLin\fid | Citrobacter | koseri | Citrobacter | koseri | 2.52 | Citrobacter | koseri | 2.47 | A | 1 | 1 | 0 | Citrobacter koseri | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-14 | Cronobacter | malonaticus | Cronobacter | sp | 2.29 | Cronobacter | sp | 2.19 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32945 | 3; confirmed proficency test | EFEx | 2021-02-18 | Cronobacter | sakazakii | Cronobacter | sp | 2.14 | Cronobacter | sp | 1.99 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter | sakazakii | Cronobacter | sp | 2.44 | Cronobacter | sp | 2.33 | A | 0 | 0 | 1 | Cronobacter sp | Cronobacter sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30054 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E18, [0034] | Enterobacter | cloacae | Enterobacter | cloacae | 2.34 | Enterobacter | cloacae | 2.08 | A | 1 | 1 | 0 | Enterobacter cloacae | Enterobacter cloacae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K4, [0034] | Enterobacter | sp-ATCC-27988 | Kosakonia | radicincitans | 1.94 | Raoultella | ornithinolytica | 1.68 | B | 0 | 0 | 1 | Kosakonia radicincitans | Raoultella ornithinolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae | sp-CVUAS-30072 | Kosakonia | cowanii | 1.45 | Escherichia | coli | 1.42 | B | 0 | 0 | 1 | Kosakonia cowanii | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0500,2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia | coli | Escherichia | coli | 2.41 | Escherichia | coli | 2.34 | A | 1 | 1 | 0 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella | grimontii | Klebsiella | oxytoca | 2.42 | Klebsiella | oxytoca | 2.24 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2486,3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella | grimontii | Klebsiella | oxytoca | 2.3 | Klebsiella | oxytoca | 2.27 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11699 | 3; sequenced [BfR] | eDT | 2017-04-20 06:56 | Klebsiella | michiganensis | Klebsiella | oxytoca | 2.38 | Klebsiella | oxytoca | 2.34 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2404,3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella | michiganensis | Klebsiella | oxytoca | 2.35 | Klebsiella | oxytoca | 2.27 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10038,2 | 3; sequenced [BfR] | EFex | 2019-02-12 14:02 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.41 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11572 | 3; sequenced [BfR] | eDT | 2017-04-20 06:51 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.39 | Klebsiella | oxytoca | 2.37 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1482,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:35 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.24 | Klebsiella | pneumoniae | 2.24 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1616,4 | 3; sequenced [BfR] | eDT | 2017-11-22 09:55 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.42 | Klebsiella | pneumoniae | 2.34 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1678 | 3; sequenced [BfR] | eDT | 2016-10-18 06:38 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.44 | Klebsiella | pneumoniae | 2.43 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1701 | 3; sequenced [BfR] | eDT | 2016-10-18 06:41 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.5 | Klebsiella | pneumoniae | 2.44 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30499 | 3; sequenced [BfR] | eDT | 2017-11-22 09:59 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.24 | Klebsiella | pneumoniae | 2.2 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 753,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:34 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.32 | Klebsiella | pneumoniae | 2.31 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10109 | 3; sequenced [BfR] | eDT | 2016-10-18 07:00 | Klebsiella | quasipneumoniae | Klebsiella | pneumoniae | 2.34 | Klebsiella | pneumoniae | 2.24 | A | 0 | 0 | 1 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11677 | 3; sequenced [BfR] | eDT | 2017-04-20 06:54 | Klebsiella | variicola | Klebsiella | variicola | 2.42 | Klebsiella | variicola | 2.4 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4888,2 | 3; sequenced [BfR] | EFex | 2017-07-19 11:36 | Klebsiella | variicola | Klebsiella | variicola | 2.39 | Klebsiella | variicola | 2.35 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30053 CVUAS | 3; type strain | EFEx | 2019-02-12 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.53 | Klebsiella | aerogenes | 2.46 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NM 20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E20 | Klebsiella | aerogenes | Klebsiella | aerogenes | 2.48 | Klebsiella | aerogenes | 2.29 | A | 1 | 1 | 0 | Klebsiella aerogenes | Klebsiella aerogenes | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella | michiganensis | Klebsiella | oxytoca | 2.28 | Klebsiella | oxytoca | 2.22 | A | 0 | 0 | 1 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.11.2013 E18 | Klebsiella | oxytoca | Klebsiella | oxytoca | 2.45 | Klebsiella | oxytoca | 2.42 | A | 1 | 1 | 0 | Klebsiella oxytoca | Klebsiella oxytoca | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 11.04.2008 N21 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.33 | Klebsiella | pneumoniae | 2.33 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 30104 | 3; type strain | EFEx | 2013-09-17 | Klebsiella | pneumoniae | Klebsiella | pneumoniae | 2.52 | Klebsiella | pneumoniae | 2.39 | A | 1 | 1 | 0 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | pneumoniae | 2.22 | Klebsiella | pneumoniae | 2.14 | A | 0 | 0 | 1 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | quasipneumoniae | Klebsiella | pneumoniae | 1.86 | Klebsiella | pneumoniae | 1.85 | B | 0 | 0 | 1 | Klebsiella pneumoniae | Klebsiella pneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella | variicola | Klebsiella | variicola | 2.42 | Klebsiella | variicola | 2.32 | A | 1 | 1 | 0 | Klebsiella variicola | Klebsiella variicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia | sacchari | Salmonella | sp | 1.72 | Salmonella | sp | 1.7 | B | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.36 | Plesiomonas | shigelloides | 2.34 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas | shigelloides | Plesiomonas | shigelloides | 2.42 | Plesiomonas | shigelloides | 2.41 | A | 1 | 1 | 0 | Plesiomonas shigelloides | Plesiomonas shigelloides | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N17, [0034] (Enterobacter gergoviae) | Pluralibacter | gergoviae | Pluralibacter | gergoviae | 2.32 | Pluralibacter | gergoviae | 2.24 | A | 1 | 1 | 0 | Pluralibacter gergoviae | Pluralibacter gergoviae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | LT21 Ib 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H9 | Salmonella | enterica | Salmonella | sp | 2.36 | Salmonella | sp | 2.36 | A | 0 | 0 | 1 | Salmonella sp | Salmonella sp | 0 | 0 | 0 | 0
|  |  | 0 | 0 | II 06 06820 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H6 | Shigella | boydii | Escherichia | coli | 2.43 | Escherichia | coli | 2.39 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 3A 07 02503 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H7 | Shigella | flexneri | Escherichia | coli | 2.46 | Escherichia | coli | 2.37 | A | 0 | 0 | 1 | Escherichia coli | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea | agglomerans | Pantoea | agglomerans | 2.52 | Pantoea | agglomerans | 2.47 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 3493 CVUAS | 3; type strain | eDT | 2015-07-09 | Pantoea | agglomerans | Pantoea | agglomerans | 2.35 | Pantoea | agglomerans | 2.31 | A | 1 | 1 | 0 | Pantoea agglomerans | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea | sp | Pantoea | sp | 1.99 | Pantoea | agglomerans | 1.93 | B | 0 | 0 | 1 | Pantoea sp | Pantoea agglomerans | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4263,2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea | sp-FF5 | Pantoea | septica | 1.84 | Pantoea | septica | 1.45 | B | 0 | 0 | 1 | Pantoea septica | Pantoea septica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella | tarda | Edwardsiella | tarda | 2.47 | Edwardsiella | tarda | 2.47 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N13, [0034] | Edwardsiella | tarda | Edwardsiella | tarda | 2.49 | Edwardsiella | tarda | 2.37 | A | 1 | 1 | 0 | Edwardsiella tarda | Edwardsiella tarda | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 18.04.2008, K9, [0034] | Hafnia | alvei | Hafnia | alvei | 2.42 | Hafnia | alvei | 2.32 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia | alvei | Hafnia | alvei | 2.36 | Hafnia | alvei | 2.29 | A | 1 | 1 | 0 | Hafnia alvei | Hafnia alvei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea | myxofaciens | Cosenzaea | myxofaciens | 2.54 | Cosenzaea | myxofaciens | 2.37 | A | 1 | 1 | 0 | Cosenzaea myxofaciens | Cosenzaea myxofaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella | morganii | Morganella | morganii | 2.56 | Morganella | morganii | 2.55 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella | morganii | Morganella | morganii | 2.68 | Morganella | morganii | 2.61 | A | 1 | 1 | 0 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus | cibarius | Proteus | hauseri | 2.26 | Proteus | hauseri | 2.23 | A | 0 | 0 | 1 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | hauseri | Proteus | hauseri | 2.55 | Proteus | hauseri | 2.42 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus | hauseri | Proteus | hauseri | 2.45 | Proteus | hauseri | 2.36 | A | 1 | 1 | 0 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus | mirabilis | Proteus | mirabilis | 2.56 | Proteus | mirabilis | 2.36 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4479 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | mirabilis | Proteus | mirabilis | 2.46 | Proteus | mirabilis | 2.41 | A | 1 | 1 | 0 | Proteus mirabilis | Proteus mirabilis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | penneri | Proteus | penneri | 2.43 | Proteus | penneri | 2.4 | A | 1 | 1 | 0 | Proteus penneri | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | vulgaris | 2.24 | Proteus | penneri | 2.13 | B | 0 | 0 | 1 | Proteus vulgaris | Proteus penneri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | vulgaris | 2.45 | Proteus | hauseri | 2.33 | B | 0 | 0 | 1 | Proteus vulgaris | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus | sp | Proteus | hauseri | 2.39 | Proteus | hauseri | 2.36 | A | 0 | 0 | 1 | Proteus hauseri | Proteus hauseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus | terrae | Proteus | hauseri | 2.31 | Proteus | vulgaris | 2.28 | B | 0 | 0 | 1 | Proteus hauseri | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, E5 K16, [0034]: Proteus vulgaris ATCC 33420\LMTG72K16_0\0_K16\1\1SLin\fid | Proteus | vulgaris | Proteus | vulgaris | 2.28 | Proteus | vulgaris | 2.24 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 2140 | 3; strain from public collection; sequenced (rpoB) | eDT | 30.11.2017 | Proteus | vulgaris | Proteus | vulgaris | 2.38 | Proteus | vulgaris | 2.33 | A | 1 | 1 | 0 | Proteus vulgaris | Proteus vulgaris | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia | rettgeri | Providencia | rettgeri | 2.57 | Providencia | rettgeri | 2.52 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia | rettgeri | Providencia | rettgeri | 2.54 | Providencia | rettgeri | 2.31 | A | 1 | 1 | 0 | Providencia rettgeri | Providencia rettgeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 25827 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K15, [0034]: Proteus stuartii ATCC 25827\LMTG72K15_0\0_K15\1\1SLin\fid in vers 2016 | Providencia | stuartii | Providencia | stuartii | 2.32 | Providencia | stuartii | 2.31 | A | 1 | 1 | 0 | Providencia stuartii | Providencia stuartii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.24 | Yersinia | enterocolitica | 2.08 | B | 0 | 0 | 1 | Yersinia frederiksenii | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia | mollaretii | Yersinia | mollaretii | 2.32 | Yersinia | mollaretii | 2.25 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 2028 | 3; public strain collection | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H13 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.48 | Yersinia | pseudotuberculosis | 2.48 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H14 | Yersinia | pestis | Yersinia | pseudotuberculosis | 2.57 | Yersinia | pseudotuberculosis | 2.43 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia | wautersii | Yersinia | pseudotuberculosis | 2.23 | Yersinia | pseudotuberculosis | 2.22 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia | fonticola | Serratia | fonticola | 2.56 | Serratia | fonticola | 2.55 | A | 1 | 1 | 0 | Serratia fonticola | Serratia fonticola | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.11.2008, E11, [0034] | Serratia | grimesii | Serratia | grimesii | 2.42 | Serratia | liquefaciens | 2.03 | B | 0 | 0 | 1 | Serratia grimesii | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.19 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia | marcescens | Serratia | marcescens | 2.21 | Serratia | marcescens | 2.13 | A | 1 | 1 | 0 | Serratia marcescens | Serratia marcescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia | quinivorans | Serratia | proteamaculans | 2.16 | Serratia | liquefaciens | 2.12 | B | 0 | 0 | 1 | Serratia proteamaculans | Serratia liquefaciens | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6221 | 4; sequenced | DT | 2013-01-08 | Yersinia | aldovae | Yersinia | aldovae | 2.33 | Yersinia | aldovae | 2.3 | A | 1 | 1 | 0 | Yersinia aldovae | Yersinia aldovae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 09.03.2007, H12 | Yersinia | aleksiciae | Yersinia | aleksiciae | 2.28 | Yersinia | aleksiciae | 2.28 | A | 1 | 1 | 0 | Yersinia aleksiciae | Yersinia aleksiciae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia | bercovieri | Yersinia | bercovieri | 2.33 | Yersinia | bercovieri | 2.25 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.56 | Yersinia | bercovieri | 2.45 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SL2012 001153 | 3; sequenced | DT | 2013-01-30 | Yersinia | bercovieri | Yersinia | bercovieri | 2.49 | Yersinia | bercovieri | 2.44 | A | 1 | 1 | 0 | Yersinia bercovieri | Yersinia bercovieri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 4780 CVUAS | 3; type strain | DT | 2013-05-08 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.5 | Yersinia | enterocolitica | 2.49 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.48 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.49 | Yersinia | enterocolitica | 2.38 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.34 | Yersinia | enterocolitica | 2.31 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.45 | Yersinia | enterocolitica | 2.42 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia | enterocolitica | Yersinia | enterocolitica | 2.57 | Yersinia | enterocolitica | 2.46 | A | 1 | 1 | 0 | Yersinia enterocolitica | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.ma_01\0_C13, [0034] | Yersinia | entomophaga | Yersinia | ruckeri | 2.06 | Yersinia | ruckeri | 2.06 | A | 0 | 0 | 1 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 21 RKI | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 22.01.2007 H06 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.41 | Yersinia | frederiksenii | 2.36 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 21 CVUAS | 3; well described strain | DT | 2013-01-30 | Yersinia | frederiksenii | Yersinia | frederiksenii | 2.57 | Yersinia | frederiksenii | 2.31 | A | 1 | 1 | 0 | Yersinia frederiksenii | Yersinia frederiksenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1824 | 2; sequenced | DT | 2012-12-05 | Yersinia | intermedia | Yersinia | intermedia | 2.46 | Yersinia | intermedia | 2.35 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia | intermedia | Yersinia | intermedia | 2.51 | Yersinia | intermedia | 2.45 | A | 1 | 1 | 0 | Yersinia intermedia | Yersinia intermedia | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.57 | Yersinia | kristensenii | 2.53 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 302 | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia | kristensenii | Yersinia | kristensenii | 2.46 | Yersinia | kristensenii | 2.39 | A | 1 | 1 | 0 | Yersinia kristensenii | Yersinia kristensenii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 25.11.2015, Y.ma_01\0_C11, [0034] | Yersinia | massiliensis | Yersinia | massiliensis | 2.33 | Yersinia | massiliensis | 2.06 | A | 1 | 1 | 0 | Yersinia massiliensis | Yersinia massiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | VB 925699 | 3; sequenced | DT | 2013-03-19 | Yersinia | massiliensis | Yersinia | massiliensis | 1.74 | Yersinia | pseudotuberculosis | 1.63 | B | 0 | 0 | 1 | Yersinia massiliensis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 18520 CVUAS | 3; type strain | DT | 16.01.2013 | Yersinia | mollaretii | Yersinia | mollaretii | 2.13 | Yersinia | mollaretii | 2.11 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | RKI 51 | 2; 0 | DT | 2013-01-16 | Yersinia | mollaretii | Yersinia | mollaretii | 2.42 | Yersinia | mollaretii | 2.4 | A | 1 | 1 | 0 | Yersinia mollaretii | Yersinia mollaretii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.nu_04\0_E10, [0034] | Yersinia | nurmii | Yersinia | frederiksenii | 1.89 | Yersinia | ruckeri | 1.87 | B | 0 | 0 | 1 | Yersinia frederiksenii | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.pe_04\0_E8, [0034] | Yersinia | pekkanenii | Yersinia | intermedia | 1.66 | Yersinia | enterocolitica | 1.62 | B | 0 | 0 | 1 | Yersinia intermedia | Yersinia enterocolitica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 0304,2 DT | 2; biochemistry | DT | 2013-01-31 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.45 | Yersinia | pseudotuberculosis | 2.44 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 8992 RKI | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.04.2009, E12 | Yersinia | pseudotuberculosis | Yersinia | pseudotuberculosis | 2.69 | Yersinia | pseudotuberculosis | 2.67 | A | 1 | 1 | 0 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia | rohdei | Yersinia | rohdei | 2.46 | Yersinia | rohdei | 2.41 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 02.03.2007, E07 | Yersinia | rohdei | Yersinia | rohdei | 2.83 | Yersinia | rohdei | 2.31 | A | 1 | 1 | 0 | Yersinia rohdei | Yersinia rohdei | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2772 CVUAS | 3; sequenced | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.52 | Yersinia | ruckeri | 2.35 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 1 | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia | ruckeri | Yersinia | ruckeri | 2.45 | Yersinia | ruckeri | 2.31 | A | 1 | 1 | 0 | Yersinia ruckeri | Yersinia ruckeri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia | similis | Yersinia | pseudotuberculosis | 2.36 | Yersinia | pseudotuberculosis | 2.36 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | IMB 4354 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 13.07.2007, E22 | Yersinia | similis | Yersinia | pseudotuberculosis | 2.52 | Yersinia | pseudotuberculosis | 2.49 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.wa_04\0_E5, [0034] | Yersinia | wautersii | Yersinia | pseudotuberculosis | 2.39 | Yersinia | pseudotuberculosis | 2.39 | A | 0 | 0 | 1 | Yersinia pseudotuberculosis | Yersinia pseudotuberculosis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella | londiniensis | Lactobacillus | pentosus | 1.33 | Streptococcus | anginosus | 1.23 | B | 0 | 0 | 1 | Lactobacillus pentosus | Streptococcus anginosus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella | pneumophila | Legionella | pneumophila | 1.73 | Legionella | pneumophila | 1.61 | B | 0 | 0 | 1 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 4122 | 3; 0 | eDT | 20.01.2016 | Legionella | pneumophila | Legionella | pneumophila | 2.33 | Legionella | pneumophila | 2.29 | A | 1 | 1 | 0 | Legionella pneumophila | Legionella pneumophila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella | taurinensis | Legionella | sp | 1.86 | Legionella | erythra | 1.43 | B | 0 | 0 | 1 | Legionella sp | Legionella erythra | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10873,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Halomonas | aquamarina | 1.54 | Halomonas | aquamarina | 1.41 | B | 0 | 0 | 1 | Halomonas aquamarina | Halomonas aquamarina | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10874,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas | stevensii | Dermacoccus | sp[3] | 1.35 | Halomonas | aquamarina | 1.24 | B | 0 | 0 | 1 | Dermacoccus sp[3] | Halomonas aquamarina | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) | caballi | Actinobacillus | rossii | 1.63 | Actinobacillus | rossii | 1.59 | B | 0 | 0 | 1 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) | caballi | Actinobacillus | rossii | 2.13 | Actinobacillus | rossii | 2.06 | A | 0 | 0 | 1 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) | testudinis | Chelonobacter | oris | 1.52 | Acinetobacter | junii | 1.49 | B | 0 | 0 | 1 | Chelonobacter oris | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 957 | 3; sequenced | eDT | 02.12.2015 | (Pasteurella) | testudinis | Lentilactobacillus | parabuchneri | 1.33 | Escherichia | coli | 1.33 | B | 0 | 0 | 1 | Lentilactobacillus parabuchneri | Escherichia coli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | anseriformium | Actinobacillus | suis | 2.01 | Actinobacillus | suis | 1.99 | A | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UGF 198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus | anseriformium | Actinobacillus | suis | 1.96 | Actinobacillus | suis | 1.93 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus | arthritidis | Actinobacillus | arthritidis | 1.93 | Actinobacillus | suis | 1.69 | B | 0 | 0 | 1 | Actinobacillus arthritidis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.14 | Actinobacillus | equuli | 1.88 | A | 1 | 1 | 0 | Actinobacillus capsulatus | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus | capsulatus | Actinobacillus | capsulatus | 2.41 | Actinobacillus | suis | 2.21 | B | 0 | 0 | 1 | Actinobacillus capsulatus | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.34 | Actinobacillus | suis | 2.23 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.28 | Actinobacillus | suis | 2.22 | B | 0 | 0 | 1 | Actinobacillus equuli | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | equuli | Actinobacillus | equuli | 2.36 | Actinobacillus | equuli | 2.3 | A | 1 | 1 | 0 | Actinobacillus equuli | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus | equuli | Actinobacillus | suis | 2.52 | Actinobacillus | suis | 2.46 | A | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus | lignieresii | Actinobacillus | lignieresii | 2.33 | Actinobacillus | pleuropneumoniae | 2.27 | B | 0 | 0 | 1 | Actinobacillus lignieresii | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9945,2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | lignieresii | Actinobacillus | lignieresii | 2.37 | Actinobacillus | pleuropneumoniae | 2.36 | B | 0 | 0 | 1 | Actinobacillus lignieresii | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus | rossii | Actinobacillus | rossii | 2 | Actinobacillus | rossii | 1.94 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus | rossii | Actinobacillus | rossii | 2.64 | Actinobacillus | rossii | 2.56 | A | 1 | 1 | 0 | Actinobacillus rossii | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32505,2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus | sp-CVUAS-32505,2 | Actinobacillus | lignieresii | 2.08 | Actinobacillus | pleuropneumoniae | 2.07 | B | 0 | 0 | 1 | Actinobacillus lignieresii | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | pleuropneumoniae | 2.1 | Actinobacillus | pleuropneumoniae | 2.05 | A | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus pleuropneumoniae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-genomospecies-1 | Actinobacillus | pleuropneumoniae | 2.08 | Actinobacillus | lignieresii | 2.01 | B | 0 | 0 | 1 | Actinobacillus pleuropneumoniae | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus | sp-MK-2012 | Actinobacillus | suis | 2.22 | Actinobacillus | equuli | 2.15 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus | sp-MK-2012 | Actinobacillus | suis | 2.27 | Actinobacillus | equuli | 2.18 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus equuli | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus | suis | Actinobacillus | suis | 2.6 | Actinobacillus | suis | 2.55 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus | suis | Actinobacillus | suis | 2.5 | Actinobacillus | suis | 2.47 | A | 1 | 1 | 0 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium | gallinarum | Avibacterium | volantium | 1.74 | Avibacterium | endocarditidis | 1.68 | B | 0 | 0 | 1 | Avibacterium volantium | Avibacterium endocarditidis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium | gallinarum | Avibacterium | paragallinarum | 1.74 | Avibacterium | paragallinarum | 1.61 | B | 0 | 0 | 1 | Avibacterium paragallinarum | Avibacterium paragallinarum | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.52 | Bibersteinia | trehalosi | 2.48 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia | trehalosi | Bibersteinia | trehalosi | 2.47 | Bibersteinia | trehalosi | 2.39 | A | 1 | 1 | 0 | Bibersteinia trehalosi | Bibersteinia trehalosi | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter | oris | Chelonobacter | oris | 2.41 | Chelonobacter | oris | 2.4 | A | 1 | 1 | 0 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter | oris | Chelonobacter | oris | 1.96 | Chelonobacter | oris | 1.89 | B | 0 | 0 | 1 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 3802,2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter | osteomyelitidis | Haemophilus | haemolyticus | 1.43 | Aggregatibacter | segnis | 1.41 | B | 0 | 0 | 1 | Haemophilus haemolyticus | Aggregatibacter segnis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia | canicola | Frederiksenia | canicola | 1.98 | Frederiksenia | canicola | 1.92 | B | 0 | 0 | 1 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia | canicola | Frederiksenia | canicola | 2.32 | Frederiksenia | canicola | 2.18 | A | 1 | 1 | 0 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium | anatis | Gallibacterium | anatis | 2.07 | Gallibacterium | anatis | 2.04 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium | anatis | Gallibacterium | anatis | 2.45 | Gallibacterium | anatis | 2.36 | A | 1 | 1 | 0 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium | sp-F151 | Gallibacterium | anatis | 1.44 | Ligilactobacillus | murinus | 1.33 | B | 0 | 0 | 1 | Gallibacterium anatis | Ligilactobacillus murinus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium | trehalosifermentans | Gallibacterium | anatis | 1.58 | Gallibacterium | anatis | 1.58 | B | 0 | 0 | 1 | Gallibacterium anatis | Gallibacterium anatis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus | haemoglobinophilus | Haemophilus | haemoglobinophilus | 1.74 | Haemophilus | haemoglobinophilus | 1.55 | B | 0 | 0 | 1 | Haemophilus haemoglobinophilus | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | haemoglobinophilus | 2.28 | Haemophilus | haemoglobinophilus | 1.99 | A | 0 | 0 | 1 | Haemophilus haemoglobinophilus | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus | sp-2269 | Haemophilus | haemoglobinophilus | 2.38 | Haemophilus | haemoglobinophilus | 2.34 | A | 0 | 0 | 1 | Haemophilus haemoglobinophilus | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | P1066-2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus | somni | Histophilus | somni | 1.82 | Histophilus | somni | 1.77 | B | 0 | 0 | 1 | Histophilus somni | Histophilus somni | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia | caviae | Mannheimia | caviae | 2.22 | Mannheimia | haemolytica | 1.76 | A | 1 | 1 | 0 | Mannheimia caviae | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1064,4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia | glucosida | Mannheimia | haemolytica | 2.09 | Mannheimia | haemolytica | 2.08 | A | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia | glucosida | Mannheimia | haemolytica | 2.36 | Mannheimia | haemolytica | 2.36 | A | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia | granulomatis | Mannheimia | granulomatis | 2.05 | Mannheimia | haemolytica | 1.77 | A | 1 | 1 | 0 | Mannheimia granulomatis | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia | granulomatis | Mannheimia | granulomatis | 2.03 | Actinobacillus | lignieresii | 1.87 | C | 0 | 0 | 1 | Mannheimia granulomatis | Actinobacillus lignieresii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.47 | Mannheimia | haemolytica | 2.44 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 6197,2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia | haemolytica | Mannheimia | haemolytica | 2.52 | Mannheimia | haemolytica | 2.51 | A | 1 | 1 | 0 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia | pernigra | Mannheimia | haemolytica | 1.5 | Mannheimia | haemolytica | 1.39 | B | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia | pernigra | Mannheimia | caviae | 1.33 | Glaesserella | parasuis | 1.3 | B | 0 | 0 | 1 | Mannheimia caviae | Glaesserella parasuis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LLBB 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.3 | Mannheimia | ruminalis | 2.22 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia | ruminalis | Mannheimia | ruminalis | 2.19 | Mannheimia | ruminalis | 2.13 | A | 1 | 1 | 0 | Mannheimia ruminalis | Mannheimia ruminalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10396,2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia | sp-10-12559-1 | Mannheimia | caviae | 2.03 | Mannheimia | haemolytica | 1.73 | A | 0 | 0 | 1 | Mannheimia caviae | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10183,3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia | sp-14UCF195 | Mannheimia | granulomatis | 2.24 | Mannheimia | haemolytica | 1.73 | A | 0 | 0 | 1 | Mannheimia granulomatis | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia | sp-14UCF195 | Mannheimia | granulomatis | 2.02 | Mannheimia | haemolytica | 1.85 | A | 0 | 0 | 1 | Mannheimia granulomatis | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia | sp-15USF15 | Mannheimia | caviae | 2.07 | Mannheimia | varigena | 2 | B | 0 | 0 | 1 | Mannheimia caviae | Mannheimia varigena | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia | sp-15USF15 | Mannheimia | caviae | 2.13 | Mannheimia | haemolytica | 1.98 | A | 0 | 0 | 1 | Mannheimia caviae | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia | sp-CVUAS-11240 | Mannheimia | varigena | 1.87 | Mannheimia | haemolytica | 1.74 | B | 0 | 0 | 1 | Mannheimia varigena | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia | sp-CVUAS-1496 | Mannheimia | haemolytica | 2.31 | Mannheimia | haemolytica | 2.24 | A | 0 | 0 | 1 | Mannheimia haemolytica | Mannheimia haemolytica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.43 | Haemophilus | haemoglobinophilus | 1.74 | C | 0 | 0 | 1 | Necropsobacter rosorum | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter | rosorum | Necropsobacter | rosorum | 2.31 | Haemophilus | haemoglobinophilus | 1.75 | C | 0 | 0 | 1 | Necropsobacter rosorum | Haemophilus haemoglobinophilus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella | semolina | Actinobacillus | equuli | 1.57 | Bisgaardia | hudsonensis | 1.44 | B | 0 | 0 | 1 | Actinobacillus equuli | Bisgaardia hudsonensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella | semolina | Mannheimia | caviae | 1.49 | Glaesserella | parasuis | 1.4 | B | 0 | 0 | 1 | Mannheimia caviae | Glaesserella parasuis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4060,3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella | aerogenes | Pasteurella | aerogenes | 2.01 | Actinobacillus | seminis | 2 | C | 0 | 0 | 1 | Pasteurella aerogenes | Actinobacillus seminis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | canis | Pasteurella | canis | 2.07 | Pasteurella | canis | 2.04 | A | 1 | 1 | 0 | Pasteurella canis | Pasteurella canis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | Pasteurella | mairii | Pasteurella | mairii | 2.11 | Actinobacillus | seminis | 2.06 | C | 0 | 0 | 1 | Pasteurella mairii | Actinobacillus seminis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella | multocida | Pasteurella | multocida | 2.32 | Pasteurella | multocida | 2.3 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 32055 | 3; specific PCR | eDT | 2019-11-14 | Pasteurella | multocida | Pasteurella | multocida | 2.35 | Pasteurella | multocida | 2.35 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella | multocida | Pasteurella | multocida | 2.45 | Pasteurella | multocida | 2.4 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.32 | Pasteurella | multocida | 2.28 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella | multocida | Pasteurella | multocida | 2.42 | Pasteurella | multocida | 2.32 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 1 | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella | multocida | Pasteurella | multocida | 2.35 | Pasteurella | multocida | 2.31 | A | 1 | 1 | 0 | Pasteurella multocida | Pasteurella multocida | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32506,2 | 3; sequenced | EFEx | 2020-11-19 | Pasteurellaceae | Bisgaard-Taxon-10 | Actinobacillus | succinogenes | 1.72 | Actinobacillus | rossii | 1.66 | B | 0 | 0 | 1 | Actinobacillus succinogenes | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Pasteurellaceae | Bisgaard-Taxon-10 | Actinobacillus | succinogenes | 1.59 | Actinobacillus | rossii | 1.56 | B | 0 | 0 | 1 | Actinobacillus succinogenes | Actinobacillus rossii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | Bisgaard-Taxon-40 | Haemophilus | parainfluenzae | 1.36 | Pseudomonas | marginalis | 1.31 | B | 0 | 0 | 1 | Haemophilus parainfluenzae | Pseudomonas marginalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9987,2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae | canine-oral-taxon-080 | Frederiksenia | canicola | 1.83 | Frederiksenia | canicola | 1.8 | B | 0 | 0 | 1 | Frederiksenia canicola | Frederiksenia canicola | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 261,4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae | sp-CVUAS-261,4 | Klebsiella | oxytoca | 1.47 | Pseudomonas | savastanoi | 1.42 | B | 0 | 0 | 1 | Klebsiella oxytoca | Pseudomonas savastanoi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae | sp-CVUAS-31988 | Haemophilus | haemolyticus | 1.65 | Haemophilus | haemolyticus | 1.61 | B | 0 | 0 | 1 | Haemophilus haemolyticus | Haemophilus haemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae | sp-CVUAS-4868 | Enterococcus | faecalis | 1.34 | Paucilactobacillus | suebicus | 1.3 | B | 0 | 0 | 1 | Enterococcus faecalis | Paucilactobacillus suebicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 724,3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae | sp-CVUAS-724,3 | Rodentibacter | heylii | 1.47 | Rodentibacter | heylii | 1.44 | B | 0 | 0 | 1 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae | sp-CVUAS-9305 | Actinobacillus | suis | 1.33 | Actinobacillus | suis | 1.3 | B | 0 | 0 | 1 | Actinobacillus suis | Actinobacillus suis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9967,2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae | sp-CVUAS-9967,2 | Haemophilus | haemolyticus | 1.74 | Haemophilus | parainfluenzae | 1.62 | B | 0 | 0 | 1 | Haemophilus haemolyticus | Haemophilus parainfluenzae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae | sp-LHL-U-309 | Lactobacillus | pentosus | 1.45 | Chelonobacter | oris | 1.39 | B | 0 | 0 | 1 | Lactobacillus pentosus | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL U-309 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-309 | Chelonobacter | oris | 1.66 | Chelonobacter | oris | 1.58 | B | 0 | 0 | 1 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | LHL U-3444 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae | sp-LHL-U-3444 | Chelonobacter | oris | 1.54 | Chelonobacter | oris | 1.51 | B | 0 | 0 | 1 | Chelonobacter oris | Chelonobacter oris | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae | sp-MCCM-2539 | Haemophilus | pittmaniae | 1.53 | Citrobacter | koseri | 1.51 | B | 0 | 0 | 1 | Haemophilus pittmaniae | Citrobacter koseri | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 308,2 | 3; sequenced | DT | 2013-11-06 | Psittacicella | hinzii | Agromyces | cerinus | 1.43 | Clostridium | tetani | 1.29 | B | 0 | 0 | 1 | Agromyces cerinus | Clostridium tetani | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.4 | Rodentibacter | heylii | 2.25 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter | heylii | Rodentibacter | heylii | 2.5 | Rodentibacter | heylii | 2.48 | A | 1 | 1 | 0 | Rodentibacter heylii | Rodentibacter heylii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium | mucosae | Staphylococcus | sciuri | 1.42 | Photorhabdus | akhurstii | 1.36 | B | 0 | 0 | 1 | Staphylococcus sciuri | Photorhabdus akhurstii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium | mucosae | Sphingomonas | aurantiaca | 1.28 | Xanthomonas | hortorum | 1.26 | B | 0 | 0 | 1 | Sphingomonas aurantiaca | Xanthomonas hortorum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 8966 | 3; sequenced | DT | 23.07.2013 | Volucribacter | amazonae | Haemophilus | influenzae | 1.53 | Haemophilus | influenzae | 1.49 | B | 0 | 0 | 1 | Haemophilus influenzae | Haemophilus influenzae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E22 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.51 | Acinetobacter | baumannii | 2.48 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2014 C05 | Acinetobacter | baumannii | Acinetobacter | baumannii | 2.46 | Acinetobacter | baumannii | 2.41 | A | 1 | 1 | 0 | Acinetobacter baumannii | Acinetobacter baumannii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter | defluvii | Acinetobacter | defluvii | 2.16 | Acinetobacter | defluvii | 2.05 | A | 1 | 1 | 0 | Acinetobacter defluvii | Acinetobacter defluvii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | guillouiae | Acinetobacter | guillouiae | 2.34 | Acinetobacter | guillouiae | 2.24 | A | 1 | 1 | 0 | Acinetobacter guillouiae | Acinetobacter guillouiae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | johnsonii | Acinetobacter | johnsonii | 2.3 | Acinetobacter | johnsonii | 2.26 | A | 1 | 1 | 0 | Acinetobacter johnsonii | Acinetobacter johnsonii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | junii | Acinetobacter | junii | 2.33 | Acinetobacter | junii | 2.31 | A | 1 | 1 | 0 | Acinetobacter junii | Acinetobacter junii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | lwoffii | Acinetobacter | lwoffii | 2.35 | Acinetobacter | lwoffii | 2.27 | A | 1 | 1 | 0 | Acinetobacter lwoffii | Acinetobacter lwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7379,3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.29 | Acinetobacter | pseudolwoffii | 2.24 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | pseudolwoffii | Acinetobacter | pseudolwoffii | 2.08 | Acinetobacter | pseudolwoffii | 2.07 | A | 1 | 1 | 0 | Acinetobacter pseudolwoffii | Acinetobacter pseudolwoffii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter | variabilis | Acinetobacter | variabilis | 2.17 | Acinetobacter | variabilis | 2.15 | A | 1 | 1 | 0 | Acinetobacter variabilis | Acinetobacter variabilis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | bovoculi | 2.16 | Moraxella | ovis | 2.1 | B | 0 | 0 | 1 | Moraxella bovoculi | Moraxella ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella | bovoculi | Moraxella | bovoculi | 2.06 | Moraxella | ovis | 1.96 | A | 1 | 1 | 0 | Moraxella bovoculi | Moraxella ovis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella | catarrhalis | Moraxella | catarrhalis | 2.53 | Moraxella | catarrhalis | 2.51 | A | 1 | 1 | 0 | Moraxella catarrhalis | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 11674,3 | 3; sequenced | eDT | 2022-03-22 | Moraxella | cuniculi | Moraxella | catarrhalis | 1.63 | Moraxella | catarrhalis | 1.51 | B | 0 | 0 | 1 | Moraxella catarrhalis | Moraxella catarrhalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella | osloensis | Moraxella | osloensis | 1.94 | Moraxella | osloensis | 1.87 | B | 0 | 0 | 1 | Moraxella osloensis | Moraxella osloensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella | ovis | Moraxella | bovoculi | 2.32 | Moraxella | ovis | 1.97 | A | 0 | 0 | 1 | Moraxella bovoculi | Moraxella ovis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | frigidicola | Pseudomonas | rhodesiae | 1.33 | Acinetobacter | parvus | 1.24 | B | 0 | 0 | 1 | Pseudomonas rhodesiae | Acinetobacter parvus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter | pasteurii | Psychrobacter | pasteurii | 2.04 | Psychrobacter | phenylpyruvicus | 1.92 | A | 1 | 1 | 0 | Psychrobacter pasteurii | Psychrobacter phenylpyruvicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | pulmonis | 2.35 | Psychrobacter | faecalis | 2.19 | B | 0 | 0 | 1 | Psychrobacter pulmonis | Psychrobacter faecalis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter | pulmonis | Psychrobacter | maritimus | 1.78 | Psychrobacter | pulmonis | 1.71 | B | 0 | 0 | 1 | Psychrobacter maritimus | Psychrobacter pulmonis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas | (alkylphenolica) | Pseudomonas | japonica | 1.59 | Pseudomonas | graminis | 1.57 | B | 0 | 0 | 1 | Pseudomonas japonica | Pseudomonas graminis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.33 | Pseudomonas | aeruginosa | 2.28 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007 H8 | Pseudomonas | aeruginosa | Pseudomonas | aeruginosa | 2.46 | Pseudomonas | aeruginosa | 2.38 | A | 1 | 1 | 0 | Pseudomonas aeruginosa | Pseudomonas aeruginosa | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas | baetica | Pseudomonas | koreensis | 2.06 | Pseudomonas | rhodesiae | 1.94 | A | 0 | 0 | 1 | Pseudomonas koreensis | Pseudomonas rhodesiae | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.02.2008 N11 | Pseudomonas | chlororaphis | Pseudomonas | chlororaphis | 2.44 | Pseudomonas | chlororaphis | 2.3 | A | 1 | 1 | 0 | Pseudomonas chlororaphis | Pseudomonas chlororaphis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 02.04.2012, H21 | Francisella | tularensis | Arthrobacter | pyridinolis | 1.43 | Lactococcus | lactis | 1.33 | B | 0 | 0 | 1 | Arthrobacter pyridinolis | Lactococcus lactis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 12.12.2016, F19, [0096] | Allofrancisella | guangzhouensis | Paracoccus | denitrificans | 1.33 | Enterococcus | faecalis | 1.29 | B | 0 | 0 | 1 | Paracoccus denitrificans | Enterococcus faecalis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 18.11.2015, L16, [0096] | Francisella | hispaniensis | Lactobacillus | plantarum | 1.42 | Francisella | philomiragia | 1.38 | B | 0 | 0 | 1 | Lactobacillus plantarum | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 25.11.2015, I3, [0096] | Francisella | naotunensis | Francisella | philomiragia | 1.57 | Francisella | philomiragia | 1.53 | B | 0 | 0 | 1 | Francisella philomiragia | Francisella philomiragia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4680,2 | 3; FLI: specific PCR | DT | 2013-04-02 | Francisella | tularensis | Morganella | morganii | 1.27 | Morganella | morganii | 1.26 | B | 0 | 0 | 1 | Morganella morganii | Morganella morganii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 23.04.2013, F20, [0096] | Francisella | tularensis | Clostridium | beijerinckii | 1.49 | Loigolactobacillus | coryniformis | 1.46 | B | 0 | 0 | 1 | Clostridium beijerinckii | Loigolactobacillus coryniformis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 26.04.2013, I20, [0096] | Francisella | tularensis | Agrobacterium | radiobacter | 1.46 | Latilactobacillus | fuchuensis | 1.43 | B | 0 | 0 | 1 | Agrobacterium radiobacter | Latilactobacillus fuchuensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 01.06.2017, F3, [0096] | Francisella | tularensis | Lactobacillus | plantarum | 1.29 | Pandoraea | nosoerga | 1.27 | B | 0 | 0 | 1 | Lactobacillus plantarum | Pandoraea nosoerga | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio | aestuarianus | Vibrio | aestuarianus | 2.39 | Vibrio | aestuarianus | 2.39 | A | 1 | 1 | 0 | Vibrio aestuarianus | Vibrio aestuarianus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio | anguillarum | Vibrio | anguillarum | 2.34 | Vibrio | anguillarum | 2.3 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio | campbelli | Vibrio | harveyi | 2.03 | Vibrio | harveyi | 1.97 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio | campbelli | Vibrio | harveyi | 2.17 | Vibrio | harveyi | 2.13 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 16-VB00067-0_BfR 1_eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio | cholerae | Vibrio | mimicus | 1.93 | Vibrio | albensis | 1.87 | B | 0 | 0 | 1 | Vibrio mimicus | Vibrio albensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 10-VBH0202 DT-48h | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 1.52 | Vibrio | alginolyticus | 1.4 | B | 0 | 0 | 1 | Vibrio cincinnatiensis | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00004 LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio | cincinnatiensis | Vibrio | cincinnatiensis | 2.14 | Vibrio | europaeus | 1.44 | A | 1 | 1 | 0 | Vibrio cincinnatiensis | Vibrio europaeus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.08 | Vibrio | diazotrophicus | 1.98 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio | fluvialis | Vibrio | fluvialis | 2.42 | Vibrio | fluvialis | 2.16 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | BfR-VB-00036 LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio | fluvialis | Vibrio | fluvialis | 2.22 | Vibrio | fluvialis | 2.18 | A | 1 | 1 | 0 | Vibrio fluvialis | Vibrio fluvialis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 16.08.2012, K12, [0034] | Vibrio | harveyi | Vibrio | harveyi | 2.28 | Vibrio | harveyi | 2.18 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio | harveyi | Vibrio | harveyi | 2.32 | Vibrio | harveyi | 2.04 | A | 1 | 1 | 0 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00047 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio | jasicida | Vibrio | harveyi | 2.49 | Vibrio | harveyi | 2.43 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00053 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio | jasicida | Vibrio | harveyi | 2.29 | Vibrio | harveyi | 2.25 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 12-VBH0498 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metoecus | Vibrio | mimicus | 2.34 | Vibrio | diazotrophicus | 1.58 | A | 0 | 0 | 1 | Vibrio mimicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00016 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | metoecus | Vibrio | mimicus | 2.06 | Pseudomonas | cichorii | 1.46 | A | 0 | 0 | 1 | Vibrio mimicus | Pseudomonas cichorii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio | metschnikovii | Vibrio | metschnikovii | 1.67 | Vibrio | metschnikovii | 1.58 | B | 0 | 0 | 1 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio | navarrensis | Vibrio | navarrensis | 2.16 | Vibrio | navarrensis | 1.99 | A | 1 | 1 | 0 | Vibrio navarrensis | Vibrio navarrensis | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio | owensii | Vibrio | harveyi | 2.19 | Vibrio | harveyi | 2.17 | A | 0 | 0 | 1 | Vibrio harveyi | Vibrio harveyi | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 17-VB00254-0_BfR 18_DT | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.28 | Vibrio | parahaemolyticus | 2.26 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio | vulnificus | Vibrio | vulnificus | 2.32 | Vibrio | vulnificus | 2.3 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 33180 25C | 3; sequenced | DT | 2021-08-04 | Photobacterium | damselae | Photobacterium | damselae | 2.22 | Photobacterium | damselae | 2.02 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium | damselae | Photobacterium | damselae | 2.14 | Photobacterium | damselae | 2.07 | A | 1 | 1 | 0 | Photobacterium damselae | Photobacterium damselae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 1266,2 | 3; sequenced | EFEx | 25.09.1019 | Photobacterium | piscicola | Photobacterium | iliopiscarium | 1.57 | Photobacterium | leiognathi | 1.56 | B | 0 | 0 | 1 | Photobacterium iliopiscarium | Photobacterium leiognathi | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 842,2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium | swingsii | Providencia | alcalifaciens | 1.37 | Arthrobacter | luteolus | 1.37 | B | 0 | 0 | 1 | Providencia alcalifaciens | Arthrobacter luteolus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7079 sir | 3; spezific PCR (CVFR) | eDT | 2022-06-23 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.06 | Vibrio | alginolyticus | 1.91 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 2171 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 12.07.2012 E12 | Vibrio | alginolyticus | Vibrio | alginolyticus | 2.3 | Vibrio | alginolyticus | 2.29 | A | 1 | 1 | 0 | Vibrio alginolyticus | Vibrio alginolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | A8-12 sir | 3; gyrB sequence | eDT | 2022-06-23 | Vibrio | anguillarum | Vibrio | anguillarum | 2.22 | Vibrio | anguillarum | 2.2 | A | 1 | 1 | 0 | Vibrio anguillarum | Vibrio anguillarum | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 10227 | 3; sequenced | eDT | 2013-05-22 | Vibrio | cholerae | Vibrio | mimicus | 2.04 | Vibrio | albensis | 1.95 | A | 0 | 0 | 1 | Vibrio mimicus | Vibrio albensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 8995 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio | diazotrophicus | Vibrio | diazotrophicus | 2.14 | Vibrio | diazotrophicus | 1.9 | A | 1 | 1 | 0 | Vibrio diazotrophicus | Vibrio diazotrophicus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | DSM 14383 sir | 3; public strain collection, sequenced | eDT | 2022-06-23 | Vibrio | furnissi | Vibrio | furnissii | 2.21 | Vibrio | furnissii | 2.06 | A | 0 | 0 | 1 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 1754 eDT | 3; sequenced | eDT | 2013-04-16 | Vibrio | furnissii | Vibrio | furnissii | 2.2 | Vibrio | furnissii | 2.05 | A | 1 | 1 | 0 | Vibrio furnissii | Vibrio furnissii | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 619 sir | 2; 0 | eDT | 2022-06-23 | Vibrio | metschnikovii | Vibrio | metschnikovii | 1.82 | Vibrio | metschnikovii | 1.64 | B | 0 | 0 | 1 | Vibrio metschnikovii | Vibrio metschnikovii | 0 | 0 | 0 | 0
|  |  | 1 | 1 | CVUAS 7872 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio | mimicus | Vibrio | mimicus | 2.22 | Vibrio | mimicus | 1.34 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 0 | 0 | DSM 19130 RKI | 4; type strain; WGS sequenced | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 24.01.2011, N4, [0034] | Vibrio | mimicus | Vibrio | mimicus | 2.08 | Vibrio | mimicus | 1.98 | A | 1 | 1 | 0 | Vibrio mimicus | Vibrio mimicus | 0 | 0 | 0 | 0
|  |  | 1 | 1 | DSM 10027 eDT | 3; type strain | eDT | 2022-06-23 | Vibrio | parahaemolyticus | Vibrio | parahaemolyticus | 2.21 | Vibrio | parahaemolyticus | 2.16 | A | 1 | 1 | 0 | Vibrio parahaemolyticus | Vibrio parahaemolyticus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 871,2 | 3; sequenced | eDT | 2014-08-21 | Vibrio | sp-CVUAS-871,2 | Vibrio | mimicus | 1.45 | Weissella | viridescens | 1.33 | B | 0 | 0 | 1 | Vibrio mimicus | Weissella viridescens | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778,2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio | tubiashii | Vibrio | europaeus | 1.65 | Vibrio | brasiliensis | 1.56 | B | 0 | 0 | 1 | Vibrio europaeus | Vibrio brasiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 2778,2 eDT | 3; sequenced | eDT | 2019-10-24 | Vibrio | tubiashii | Vibrio | europaeus | 1.77 | Vibrio | brasiliensis | 1.52 | B | 0 | 0 | 1 | Vibrio europaeus | Vibrio brasiliensis | 0 | 0 | 0 | 0
|  |  | 1 | 1 | LVM05 LVU05 | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-14 | Vibrio | vulnificus | Vibrio | vulnificus | 2.46 | Vibrio | vulnificus | 2.31 | A | 1 | 1 | 0 | Vibrio vulnificus | Vibrio vulnificus | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas | sp-G9077 | Wohlfahrtiimonas | chitiniclastica | 2.35 | Wohlfahrtiimonas | chitiniclastica | 2.31 | A | 0 | 0 | 1 | Wohlfahrtiimonas chitiniclastica | Wohlfahrtiimonas chitiniclastica | 0 | 0 | 0 | 0
|  |  | 1 | 1 | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria | indica | Ignatzschineria | indica | 2.18 | Ignatzschineria | indica | 1.92 | A | 1 | 1 | 0 | Ignatzschineria indica | Ignatzschineria indica | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria | ureiclastica | Ignatzschineria | ureiclastica | 2.42 | Ignatzschineria | larvae | 2.15 | B | 0 | 0 | 1 | Ignatzschineria ureiclastica | Ignatzschineria larvae | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas | chelatiphaga | Stenotrophomonas | maltophilia | 1.61 | Stenotrophomonas | acidaminiphila | 1.54 | B | 0 | 0 | 1 | Stenotrophomonas maltophilia | Stenotrophomonas acidaminiphila | 0 | 0 | 0 | 0
|  |  | 0 | 0 | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2009 K17 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.26 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.05.2009 K22 | Stenotrophomonas | maltophilia | Stenotrophomonas | maltophilia | 2.44 | Stenotrophomonas | maltophilia | 2.22 | A | 1 | 1 | 0 | Stenotrophomonas maltophilia | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 1 | 0 | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 1.79 | Stenotrophomonas | maltophilia | 1.63 | B | 0 | 0 | 1 | Stenotrophomonas rhizophila | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0
|  |  | 0 | 0 | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.05.2015 C17 | Stenotrophomonas | rhizophila | Stenotrophomonas | rhizophila | 2.4 | Stenotrophomonas | acidaminiphila | 1.51 | A | 1 | 1 | 0 | Stenotrophomonas rhizophila | Stenotrophomonas acidaminiphila | 0 | 0 | 0 | 0
|  |  | 1 | 0 | CVUAS 4239,2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas | tumulicola | Stenotrophomonas | acidaminiphila | 1.49 | Stenotrophomonas | maltophilia | 1.43 | B | 0 | 0 | 1 | Stenotrophomonas acidaminiphila | Stenotrophomonas maltophilia | 0 | 0 | 0 | 0

## #Parameter (Spezies) ggPM
Parameter: | Clostridium | perfringens | Codierung | Validität | Methode | Datum | verwendeter Organismus | Score | richtig-negativ | falsch-positiv
Test auf spektrale Nähe zum Parameter |  |  | DSM 43014 | 3; type strain | EFEx | 2018-01-04 | Actinomyces bovis | 0.48 | 1 | 0
Anzahl Isolate: | 1144 |  | CVUAS 31891 | 3; sequenced | EFEx | 2019-09-04 | Actinomyces denticolens | 0.57 | 1 | 0
|  |  | CVUAS 31303 | 3; sequenced | EFEx | 2019-07-05 | Actinomyces sp-CVUAS-31303 | 0.19 | 1 | 0
richtig-negativ: | 1144 |  | Wü6 | 3; sequenced | EFEx | 2018-01-10 | Actinomyces viscosus | 0.27 | 1 | 0
|  |  | DSM 24894 | 3; type strain | EFEx | 2018-01-04 | Actinomyces weissii | 0.69 | 1 | 0
falsch-positiv: | 0 |  | C605018-01-1T | 3; type strain | EFEx | 2017-08-16 | Arcanobacterium bovis | 0.64 | 1 | 0
|  |  | CIP 110339 | 3; type strain | DT | 2013-04-09 | Arcanobacterium canis | 0 | 1 | 0
| 0% |  | DSM 20595 | 3; type strain | EFEx | 2017-10-06 | Arcanobacterium haemolyticum | 0.59 | 1 | 0
|  |  | 17-6203 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium hippocoleae | 0.82 | 1 | 0
|  |  | DSM 15539 | 3; type strain | DT | 2013-08-07 | Arcanobacterium hippocoleae | 0.74 | 1 | 0
Scorewertgrenzen: | 2 | Spezies | DSM 10002 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium phocae | 0.9 | 1 | 0
| 1.7 | Genus | LALLF 21TRD0572 | 3; sequenced | EFEx | 2022-05-11 | Arcanobacterium phocae | 0.63 | 1 | 0
Es wird gegen eine nur aus dem Parameter bestehende DB geprüft. |  |  | 2698T | 3; type strain | DT | 2013-08-07 | Arcanobacterium phocisimile | 0.56 | 1 | 0
|  |  | 014418 | 3; sequenced | EFEx | 2020-01-15 | Arcanobacterium pinnipediorum | 0.61 | 1 | 0
|  |  | DSM 28752 | 3; type strain | EFex | 2017-09-06 | Arcanobacterium pinnipediorum | 0.4 | 1 | 0
|  |  | DSM 13483 | 3; type strain | EFex | 04.10.2017 | Arcanobacterium pluranimalium | 0.61 | 1 | 0
|  |  | 17-2065.01 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 0.81 | 1 | 0
|  |  | 17-3515 | 2; 0 | EFex [BRU03] | 2019-02-13 | Arcanobacterium sp-17-2065.02 | 0.58 | 1 | 0
|  |  | 21044 | 2; description in progress | EFex | 2019-11-25 | Arcanobacterium sp-2701 | 0.57 | 1 | 0
|  |  | 2701 | 2; description in progress | EFex | 2017-08-16 | Arcanobacterium sp-2701 | 0.46 | 1 | 0
|  |  | CVUAS 1060,2 | 3; sequenced | DT | 2013-03-15 | Arcanobacterium wilhelmae | 0.88 | 1 | 0
|  |  | CVUAS 647 | 3; type strain | DT | 2014-09-04 | Arcanobacterium wilhelmae | 0.52 | 1 | 0
|  |  | CVUAS 30561 | 3; sequenced | EFEx | 2017-11-15 | Boudabousia sp-CVUAS-30561 | 0.45 | 1 | 0
|  |  | DSM 15536 | 3; type strain | EFEx | 2018-01-04 | Schaalia canis | 0.5 | 1 | 0
|  |  | DSM 15803 | 3; type strain | EFEx | 2020-03-27 | Schaalia cardiffensis | 0.19 | 1 | 0
|  |  | CVUAS 32021 | 3; sequenced | EFEx | 2019-10-23 | Schaalia hyovaginalis | 0.89 | 1 | 0
|  |  | DSM 10695 | 3; type strain | EFEx | 2018-08-31 | Schaalia hyovaginalis | 0.56 | 1 | 0
|  |  | CVUAS 32183 | 3; sequenced | EFEx | 2020-02-13 | Schaalia sp-CVUAS-32183 | 0.61 | 1 | 0
|  |  | CVUAS 31838 | 3; sequenced | EFEx | 2019-08-23 | Schaalia sp-CVUAS-8688 | 0.01 | 1 | 0
|  |  | CVUAS 31845,2 | 3; sequenced | EFEx | 2019-09-25 | Schaalia sp-CVUAS-8688 | 0.57 | 1 | 0
|  |  | 874-14 | 3; sequenced [0070] | EFEx | 2017-05-04 | Trueperella abortisuis | 0.69 | 1 | 0
|  |  | LALLF 22TRD0373 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella abortisuis | 0.27 | 1 | 0
|  |  | D12_0613-1-4-3 | 3; 0 | EFEx | 2019-07-10 | Trueperella bernardiae | 0.82 | 1 | 0
|  |  | DSM 9152 | 3; type strain | EFEx | 2017-05-09 | Trueperella bernardiae | 0.41 | 1 | 0
|  |  | DSM 17162 | 3; type strain | EFEx | 2017-05-04 | Trueperella bialowiezensis | 0.49 | 1 | 0
|  |  | DSM 17163 | 3; type strain | EFEx | 2018-02-26 | Trueperella bonasi | 0.37 | 1 | 0
|  |  | LALLF 20TRD0366 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.57 | 1 | 0
|  |  | LALLF 20TRD0808 | 3; sequenced | EFex [BRU03] | 2022-05-11 | Trueperella pecoris | 0.53 | 1 | 0
|  |  | 11-7-D-03394 | 3; 0 | EFEx | 2017-05-09 | Trueperella pyogenes | 0.75 | 1 | 0
|  |  | DSM 20630 | 3; type strain | EFEx | 2017-05-09 | Trueperella pyogenes | 1.18 | 1 | 0
|  |  | C19 | 3; sequenced (rpoB) UGent | eDT | 2020-06-18 | Corynebacterium amycolatum | 0.24 | 1 | 0
|  |  | RRLK02_02 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2014 E8 | Corynebacterium amycolatum | 0.87 | 1 | 0
|  |  | 273U3-1 | 3; sequenced | eDT | 2016-12-28 | Corynebacterium aurimucosum | 0.55 | 1 | 0
|  |  | DSM 105776 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium belfantii | 1.04 | 1 | 0
|  |  | KL0236 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium belfantii | 0.42 | 1 | 0
|  |  | DSM 20582 | 3; type strain | DT | 2013-12-20 | Corynebacterium bovis | 0.58 | 1 | 0
|  |  | 234-U2-11 | 3; sequenced | EFEx | 2017-03-15 | Corynebacterium camporealensis | 0.71 | 1 | 0
|  |  | CVUAS 1764 | 3; sequenced | DT | 2015-01-14 | Corynebacterium casei | 0.35 | 1 | 0
|  |  | 253-y-4 | 3; sequenced | EFEx | 2020-10-15 | Corynebacterium confusum | 0.52 | 1 | 0
|  |  | 54y-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium confusum | 0.52 | 1 | 0
|  |  | CVUAS 32184 | 3; sequenced | eDT | 2020-01-30 | Corynebacterium cystitidis | 0.64 | 1 | 0
|  |  | KL0163 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.54 | 1 | 0
|  |  | KL0167 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.23 | 1 | 0
|  |  | KL0173 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-04 | Corynebacterium diphtheriae | 0.62 | 1 | 0
|  |  | KL0956 | 3; MLST | 0 | 0 | Corynebacterium diphtheriae | 0.87 | 1 | 0
|  |  | 202y-3 | 3; sequenced | eDT | 2016-03-02 | Corynebacterium frankenforstense | 0.42 | 1 | 0
|  |  | ST18T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium frankenforstense | 0.33 | 1 | 0
|  |  | DSM 45254 | 3; type strain | DT | 2012-12-18 | Corynebacterium freiburgense | 0.66 | 1 | 0
|  |  | LALLF 19TRB1509 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium hansenii | 0.6 | 1 | 0
|  |  | LHL 1380 | 3; sequenced | DT | 2013-02-13 | Corynebacterium hansenii | 0.69 | 1 | 0
|  |  | 216x-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium humireducens | 0.54 | 1 | 0
|  |  | CVUAS 4986 | 3; sequenced | DT | 2012-12-12 | Corynebacterium kutscheri | 0.56 | 1 | 0
|  |  | 234z-1 | 3; sequenced | eDT | 2016-08-31 | Corynebacterium lactis | 0.36 | 1 | 0
|  |  | RW2-5T | 3; type strain | EFEx | 2016-12-28 | Corynebacterium lactis | 0.57 | 1 | 0
|  |  | 277U1-1 | 3; sequenced, rpoB; 16S | EFEx | 2016-12-28 | Corynebacterium phoceense | 1.12 | 1 | 0
|  |  | 1403-y-2 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium pilosum | 0.49 | 1 | 0
|  |  | CVUAS 5197 | 3; sequenced | DT | 2012-10-18 | Corynebacterium pseudotuberculosis | 0.44 | 1 | 0
|  |  | DSM 7177 | 3; public strain collection | DT | 2012-10-08 | Corynebacterium pseudotuberculosis | 0.74 | 1 | 0
|  |  | CVUAS 32931 | 2; 0 | eDT | 2021-06-16 | Corynebacterium pseudotuberculosis | 0.8 | 1 | 0
|  |  | DSM 20689 CVUAS | 3; type strain | DT | 2012-12-18 | Corynebacterium pseudotuberculosis | 0.61 | 1 | 0
|  |  | C138 | 3; sequenced (rpoB) UGent | EFEx | 2020-07-30 | Corynebacterium rouxii | 0.46 | 1 | 0
|  |  | 191012535 | 3; KL Diphtheriae (LGL Bayern) | EFEx | 2020-07-30 | Corynebacterium rouxii | 0.37 | 1 | 0
|  |  | CVUAS 3559,2 | 3; KL Diphtheriae, LGL | EFEx | 2020-08-12 | Corynebacterium rouxii | 0.7 | 1 | 0
|  |  | DSM 110354 | 3; type strain | EFEx | 2020-10-20 | Corynebacterium rouxii | 0.49 | 1 | 0
|  |  | 171007405 | 3; strain of the species description [0127] | EFEX | 2019-05-07 | Corynebacterium silvaticum | 0.6 | 1 | 0
|  |  | AGES WS 007 | 3; FT-IR [see 0015]; sequenced | EFEx | 2019-01-30 | Corynebacterium silvaticum | 0.73 | 1 | 0
|  |  | 208-U2-6 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Corynebacterium sp-208-U2-6 | 0.5 | 1 | 0
|  |  | CVUAS 32165,2 | 3; sequenced | eDT | 2020-05-29 | Corynebacterium sp-2300500 | 0 | 1 | 0
|  |  | CVUAS 5775,2 | 3; sequenced | DT | 2013-07-24 | Corynebacterium sp-2300500 | 0 | 1 | 0
|  |  | CVUAS 2554,2 | 3; sequenced | eDT | 2014-11-18 | Corynebacterium sp-2396 | 0.4 | 1 | 0
|  |  | 243-z-3 | 3; sequenced, rpoB; 16S | EFEx | 2017-03-15 | Corynebacterium sp-243-z-3 | 0.43 | 1 | 0
|  |  | 257-y-4 | 3; sequenced (CVUAS) | eDT | 2020-10-20 | Corynebacterium sp-257-y-4 | 0.89 | 1 | 0
|  |  | 284z-2 | 2; sequenced | eDT | 2016-12-28 | Corynebacterium sp-284z-2 | 0.62 | 1 | 0
|  |  | CVUAS 11664 | 3; sequenced | EFEx | 2017-01-18 | Corynebacterium sp-CVUAS-11664 | 0.72 | 1 | 0
|  |  | CVUAS 1888 | 3; sequenced | DT | 2012-10-18 | Corynebacterium sp-CVUAS-1888 | 0.76 | 1 | 0
|  |  | CVUAS 1721,3 | 3; sequenced rpoB | eDT | 2021-06-08 | Corynebacterium sp-CVUAS-2484.4 | 0.36 | 1 | 0
|  |  | CVUAS 2484,4 | 3; 16S rDNA: 1393bp (96.7% of ref-sequence) > 33/1388; 97,62% similarity with Corynebacterium lactis RW2-5(T) 100% completeness (15.04.2016; EzBioCloud); Genbank accession number: rpoB: | EFEx | 2016-04-26 | Corynebacterium sp-CVUAS-2484.4 | 0.93 | 1 | 0
|  |  | CVUAS 30404 | 3; sequenced | EFEx | 2017-09-05 | Corynebacterium sp-CVUAS-30404 | 0.13 | 1 | 0
|  |  | LALLF 22TRD0272 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-CVUAS-30404 | 0.75 | 1 | 0
|  |  | CVUAS 31533 | 3; sequenced; 16S, rpoB | EFex | 2019-02-01 | Corynebacterium sp-CVUAS-31533 | 0.93 | 1 | 0
|  |  | CVUAS 32278,2 | 3; sequenced | EFEx | 2020-07-21 | Corynebacterium sp-CVUAS-32278.2 | 0.24 | 1 | 0
|  |  | 285-y-1 | 3; sequenced | eDT | 2017-06-23 | Corynebacterium sp-CVUAS-874,4 | 0.1 | 1 | 0
|  |  | CVUAS 874,4 | 3; sequenced, rpoB | EFEx | 2017-06-22 | Corynebacterium sp-CVUAS-874,4 | 0.48 | 1 | 0
|  |  | LALLF 21TRB0553 | 3; sequenced | eDT | 2022-05-09 | Corynebacterium sp-LALLF-21TRB0553 | 0.75 | 1 | 0
|  |  | LHL 0497 | 3; sequenced | DT | 2013-07-16 | Corynebacterium sp-LHL-0497 | 0.74 | 1 | 0
|  |  | LHL 0558 | 3; sequenced | DT | 2012-12-18 | Corynebacterium sp-LHL-0558 | 0.94 | 1 | 0
|  |  | LHL 1322 | 3; sequenced | DT | 2013-03-14 | Corynebacterium sp-LHL-1322 | 0.87 | 1 | 0
|  |  | CVUAS 4315 | 3; sequenced | DT | 2013-02-13 | Corynebacterium stationis | 0.57 | 1 | 0
|  |  | LHL 7774 | 3; sequenced, rpoB | DT | 2013-06-26 | Corynebacterium testudinoris | 1.05 | 1 | 0
|  |  | CCM 2823 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 0.44 | 1 | 0
|  |  | CVUAS 32709 | 3; sequenced | eDT | 2020-10-21 | Corynebacterium ulcerans | 0.23 | 1 | 0
|  |  | CVUAS 11452 | 3; FT-IR no NTTB-D, MLST | eDT | 2016-11-23 | Corynebacterium ulcerans | 0.59 | 1 | 0
|  |  | JB362081 | 3; 0 | eDT | 2013-01-31 | Corynebacterium ulcerans | 0.06 | 1 | 0
|  |  | CIP 102462 | 3; reference strain collection | DT | 2012-10-18 | Corynebacterium ulcerans | 0.65 | 1 | 0
|  |  | CVUAS 6522,4 | 3; FT-IR (NTTB-D), tox sequence | eDT | 2018-07-04 | Corynebacterium ulcerans | 0.49 | 1 | 0
|  |  | CVUAS 5437,3 | 3; sequenced | EFEx | 2017-07-19 | Corynebacterium variabile | 0.21 | 1 | 0
|  |  | 161z-3 | 3; sequenced | EFEx | 2016-08-16 | Corynebacterium xerosis | 0.41 | 1 | 0
|  |  | 272-U3-7 | 2; sequenced | EFEx | 2017-03-15 | Corynebacterium xerosis-group | 0.62 | 1 | 0
|  |  | CVUAS 2379,2 | 3; sequenced | EFex | 2021-04-08 | Gordonia paraffinivorans | 0 | 1 | 0
|  |  | CVUAS 32967 | 3; sequenced | eDT | 2021-02-26 | Gordonia paraffinivorans | 0.41 | 1 | 0
|  |  | DSM 43248 | 3; public strain collection | EFex | 2021-03-25 | Gordonia rubripertincta | 0.62 | 1 | 0
|  |  | CVUAS 10185.3 | 3; sequenced | EFEx | 2020-03-18 | Mycolicibacterium smegmatis | 0.62 | 1 | 0
|  |  | CVUAS 31296,3 | 3; sequenced | EFEx | 2019-01-03 | Nocardia cyriacigeorgica | 0.76 | 1 | 0
|  |  | CVUAS 31377 | 3; sequenced | EFEx | 2018-12-04 | Nocardia farcinica | 0.57 | 1 | 0
|  |  | ATCC 33701 | 3; public strain collection | EFEx | 2021-07-22 | Rhodococcus equi | 0.77 | 1 | 0
|  |  | CVUAS 952 | 2; 0 | EFEx | 2021-07-22 | Rhodococcus equi | 0.81 | 1 | 0
|  |  | CVUAS 1447,3 | 3; sequenced | DT | 2015-04-14 | Tsukamurella sp | 0.64 | 1 | 0
|  |  | CVUAS 2895,3 | 3; sequenced | eDT | 2017-06-23 | Brevibacterium luteolum | 0.41 | 1 | 0
|  |  | CVUAS 31260 | 3; sequenced | EFEx | 2018-10-16 | Cellulomonas flavigena | 0.75 | 1 | 0
|  |  | 234z-2 | 3; sequenced | EFEx | 2016-08-16 | Brachybacterium sp-234z-2 | 1.08 | 1 | 0
|  |  | CVUAS 30560 | 3; sequenced | EFEx | 2017-11-15 | Brachybacterium sp-CVUAS-30560 | 0.68 | 1 | 0
|  |  | 285-U5-2 | 3; 16S rDNA-sequencing | eDT | 2017-03-14 | Kytococcus sedentarius | 0.2 | 1 | 0
|  |  | LALLF 19TRD0226 | 3; sequenced | EFex | 2022-05-09 | Dermatophilus congolensis | 0.58 | 1 | 0
|  |  | LHL 141010120-1 | 3; sequenced | EFex | 2022-05-11 | Dermatophilus congolensis | 0.35 | 1 | 0
|  |  | CVUAS 31268 | 3; sequenced | EFEx | 2018-10-26 | Curtobacterium flaccumfaciens | 0.38 | 1 | 0
|  |  | MUVA-A-508948 | 2; sequenced (Genus) | EFex [BRU03] | 2022-02-23 | Leifsonia sp-MUVA-A-508948 | 0.47 | 1 | 0
|  |  | 284U2-2-2 | 2; sequenced | EFEx | 2016-12-29 | Microbacterium sp | 0.63 | 1 | 0
|  |  | IMET 10353 DSM 20123 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.09.2011 L8 | Glutamicibacter nicotianae | 0.44 | 1 | 0
|  |  | IMET 10432 DSM 20127 | 3; type strain | unknown | 08.09.2011, L13, [0034] | Sinomonas atrocyanea | 0.15 | 1 | 0
|  |  | CVUAS 9434 | 3; sequenced | DT | 2013-09-25 | Glutamicibacter protophormiae | 0.69 | 1 | 0
|  |  | 161-U2-7 | 3; sequenced | EFex | 2017-03-14 | Kocuria atrinae | 0.74 | 1 | 0
|  |  | 56a-5 | 3; sequenced | EFEx | 2020-05-20 | Kocuria gwangalensis | 1.06 | 1 | 0
|  |  | RRLK06_02 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.04.2014 E22 | Kocuria palustris | 0.37 | 1 | 0
|  |  | RRLK07_03 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.04.2015 K12 | Micrococcus luteus | 0.29 | 1 | 0
|  |  | DSM 20032 | 3; type strain | eDT | 2015-07-09 | Rothia kristinae | 0.46 | 1 | 0
|  |  | CVUAS 11689 | 3; sequenced; GLC pos, XYL pos, MOT neg | eDT | 2017-01-31 | Cellulosimicrobium cellulans | 0.22 | 1 | 0
|  |  | CVUAS 749,2 | 3; sequenced | eDT | 2016-08-19 | Sanguibacter inulinus | 0.38 | 1 | 0
|  |  | DSM 1079 | 3; public strain collection | EFEx | 2017-01-20 | Penicillium roqueforti | 0.71 | 1 | 0
|  |  | CVUAS 10354,2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton verrucosum | 0.19 | 1 | 0
|  |  | CVUAS 10357.2 | 3; sequenced (ITS) | F-Extr | 2017-07-24 | Trichophyton verrucosum | 0.89 | 1 | 0
|  |  | DSM 1386 | 3; public strain collection | Fextr | 2016-03-10 | Candida albicans | 0.79 | 1 | 0
|  |  | IGV 011 IGV | 3; sequenced | EFEx | 2016-12-08 | Sarocladium strictum | 1.21 | 1 | 0
|  |  | CVUAS 2520.4 | 3; sequenced (ITS) | F-Extr | 2018-07-26 | Fusarium sambucinum | 0.38 | 1 | 0
|  |  | 141010301-6 | 2; 0 | DT | 2014-08-19 | Paenibacillus amylolyticus | 0.37 | 1 | 0
|  |  | 161006368 | 2; 0 | DT | 2016-05-23 | Paenibacillus amylolyticus | 0.26 | 1 | 0
|  |  | 151006426-14 | 2; 0 | DT | 2015-05-27 | Paenibacillus barengoltzii | 0.48 | 1 | 0
|  |  | 181005658 | 2; 0 | DT | 2018-04-19 | Paenibacillus barengoltzii | 0.59 | 1 | 0
|  |  | 141010244-2 | 2; 0 | DT | 2014-08-15 | Paenibacillus glucanolyticus | 0.51 | 1 | 0
|  |  | 161004455-3 | 2; 0 | DT | 2016-04-12 | Paenibacillus glucanolyticus | 0.42 | 1 | 0
|  |  | 151008473-1 | 3; FLI | DT | 2015-06-30 | Paenibacillus larvae | 0.6 | 1 | 0
|  |  | 151009681-2 | 3; FLI | DT | 2015-07-28 | Paenibacillus larvae | 0.67 | 1 | 0
|  |  | 151005015-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 0.51 | 1 | 0
|  |  | 151005237-1 | 3; FLI | DT | 2015-04-28 | Paenibacillus larvae | 0.14 | 1 | 0
|  |  | 151005943-3 | 2; 0 | DT | 2015-05-12 | Paenibacillus macerans | 0.52 | 1 | 0
|  |  | 141006464-7 | 2; 0 | DT | 2014-05-23 | Paenibacillus pabuli | 0.49 | 1 | 0
|  |  | 141010955-13 | 2; 0 | DT | 2014-08-29 | Paenibacillus pabuli | 0.96 | 1 | 0
|  |  | 141013794-4 | 2; 0 | DT | 2014-10-30 | Paenibacillus pasadenensis | 0.17 | 1 | 0
|  |  | 171012589-4 | 2; 0 | DT | 2017-09-19 | Paenibacillus validus | 0.14 | 1 | 0
|  |  | CDH 260-11 | 3; public strain collection | DT | 2017-12-06 | Clostridioides difficile | 0.64 | 1 | 0
|  |  | 171012939 | 2; 0 | DT | 2017-10-26 | Clostridium barati | 0.54 | 1 | 0
|  |  | DSM 2478 | 3; public strain collection | DT | 2015-06-23 | Clostridium butyricum | 1.12 | 1 | 0
|  |  | 121017785-2 | 3; FLI | DT | 2018-07-02 | Clostridium chauvoei | 0.58 | 1 | 0
|  |  | 181017156-1 | 3; FLI | DT | 2018-11-27 | Clostridium chauvoei | 0.61 | 1 | 0
|  |  | 181017156-2 | 3; FLI | DT | 2018-11-27 | Clostridium chauvoei | 0.52 | 1 | 0
|  |  | 181017156-3 | 3; FLI | DT | 2018-11-27 | Clostridium chauvoei | 0.63 | 1 | 0
|  |  | 181017156-4 | 3; FLI | DT | 2018-11-27 | Clostridium chauvoei | 0.52 | 1 | 0
|  |  | 171002729_1 coch | 2; 0 | DT | 2017-02-28 | Clostridium cochlearium | 0.37 | 1 | 0
|  |  | 161011787 | 2; 0 | DT | 2016-09-27 | Clostridium colicanis | 0.48 | 1 | 0
|  |  | 151010772 | 2; 0 | DT | 2015-08-17 | Clostridium novyi | 0.89 | 1 | 0
|  |  | 151002484-2 | 2; 0 | DT | 2015-03-03 | Clostridium septicum | 0.6 | 1 | 0
|  |  | DSM 767 | 3; public strain collection | DT | 2015-06-23 | Clostridium sporogenes | 0.49 | 1 | 0
|  |  | 141015223 aerob | 2; 0 | DT | 2014-11-27 | Clostridium tertium | 0.6 | 1 | 0
|  |  | 161002486-2 aerob | 2; 0 | DT | 2016-03-01 | Clostridium tertium | 0.17 | 1 | 0
|  |  | 181003180 aerob | 2; 0 | DT | 2018-06-30 | Clostridium tertium | 0.36 | 1 | 0
|  |  | 181005218 | 2; 0 | DT | 2018-11-04 | Clostridium tetani | 0.72 | 1 | 0
|  |  | 151011769 | 2; 0 | DT | 2015-09-04 | Hathewaya limosa | 0.55 | 1 | 0
|  |  | 151006002 | 2; 0 | DT | 2015-05-12 | Paeniclostridium sordellii | 0.75 | 1 | 0
|  |  | DSM 630 | 3; public strain collection | DT | 2015-06-25 | Paraclostridium bifermentans | 0.55 | 1 | 0
|  |  | B 44 | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N11 | Bacillus badius | 0.65 | 1 | 0
|  |  | DSM 28970 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.03.2017 I18 | Bacillus ectoiniformans | 0.94 | 1 | 0
|  |  | ATCC 14575 B 30 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F22 | Cytobacillus firmus | 0 | 1 | 0
|  |  | CCM 37 B 88 | 3; reference strain collection | 0 | 09.02.2015, L16 | Cytobacillus firmus | 0.22 | 1 | 0
|  |  | B 46 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.12.2014 N14 | Niallia circulans | 0.45 | 1 | 0
|  |  | DSM 26 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 N5 | Virgibacillus pantothenticus | 0.34 | 1 | 0
|  |  | CVUAS 31071,2 | 3; sequenced | EFEx | 2021-09-08 | Alkalihalobacillus clausii | 0.4 | 1 | 0
|  |  | DSM 8716 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015; H10 | Alkalihalobacillus clausii | 0.57 | 1 | 0
|  |  | DSM 19092 | 3; type strain | 80% TFA/1:10 verduennnt/1:1 HCCA:TA2 | 13.02.2015, K10, [0034] | Bacillus alveayuensis | 0.79 | 1 | 0
|  |  | DSM 7 RKI 2015 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCA/TA2 (A) | 16.02.2015, K16, [0034] | Bacillus amyloliquefaciens | 0.94 | 1 | 0
|  |  | DSM 7 CVUAS | 3; type strain | DT | 2014-01-29 | Bacillus amyloliquefaciens | 0.87 | 1 | 0
|  |  | ATCC 11966 A117 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 09.11.2015; E06 | Bacillus anthracis | 0.75 | 1 | 0
|  |  | ATCC 4229 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; E04 | Bacillus anthracis | 0.57 | 1 | 0
|  |  | B 004 | 2; 0 | DT | 2014-01-22 | Bacillus atrophaeus | 0.45 | 1 | 0
|  |  | DSM 7264 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2017 L8 | Bacillus atrophaeus | 0.78 | 1 | 0
|  |  | DSM 28241 | 3; public strain collection (not further listed, 22.09.2021 [0003]) | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H20 | Bacillus badius | 0.79 | 1 | 0
|  |  | ATCC 14893 | 3; public strain collection | eDT (8h) | 2016-01-13 | Bacillus cereus | 0.43 | 1 | 0
|  |  | CVUAS 1107 | 2; 0 | eDT | 27.01.2016 | Bacillus cereus | 0.26 | 1 | 0
|  |  | ATCC 4342 B338 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.05.2015; E09 | Bacillus tropicus | 0.63 | 1 | 0
|  |  | CVUAS 2833 CVUAS | 4; 16S, draft genome [0009] | eDT | 2016-01-27 | Bacillus cytotoxicus | 0.74 | 1 | 0
|  |  | NVH 391-98 | 3; type strain | eDT | 2016-11-16 | Bacillus cytotoxicus | 0.44 | 1 | 0
|  |  | CCEB 627 | 3; type strain | DT | 2014-01-22 | Bacillus firmus | 0.67 | 1 | 0
|  |  | CCM 37 CVUAS | 3; reference strain collection | DT | 2014-01-22 | Bacillus firmus | 0.5 | 1 | 0
|  |  | DSM 18911 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.11.2017 F14 | Bacillus halotolerans | 0.91 | 1 | 0
|  |  | DSM 8802 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.12.2017 I8 | Bacillus halotolerans | 0.76 | 1 | 0
|  |  | DSM 22148 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.12.2016 [H3] | Bacillus inaquosorum | 0.87 | 1 | 0
|  |  | DSM 19098 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.02.2016 C20 | Bacillus infantis | 0.53 | 1 | 0
|  |  | B084-b | 3; sequenced, biochemistry | DT | 2015-07-22 | Bacillus licheniformis | 0.5 | 1 | 0
|  |  | DSM 13 CVUAS | 3; type strain | eDT | 2016-01-27 | Bacillus licheniformis | 0.59 | 1 | 0
|  |  | DSM 9205 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K15 | Bacillus mojavensis | 0.91 | 1 | 0
|  |  | DSM 9206 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2016 F4 | Bacillus mojavensis | 0.98 | 1 | 0
|  |  | DSM 2048 CVUAS | 3; type strain | eDT | 09.12.2015 | Bacillus mycoides | 0.28 | 1 | 0
|  |  | DSM 309 | 3; public strain collection | eDT | 2016-01-27 | Bacillus mycoides | 0.68 | 1 | 0
|  |  | ATCC 11778 RKI | 3; public strain collection | in 80% TFA/ 1:1 HCCATA2 (A) | 05.03.2007, E8, [0034] | Bacillus pacificus | 0.38 | 1 | 0
|  |  | CVUAS 7789 | 3; sequenced | eDT | 20.01.2016 | Bacillus pseudomycoides | 0.79 | 1 | 0
|  |  | DSM 12443 CVUAS | 3; public strain collection | eDT | 09.12.2015 | Bacillus pseudomycoides | 0.49 | 1 | 0
|  |  | B017 CVUAS | 3; sequenced (CVUAS) | DT | 2015-07-22 | Bacillus pumilus | 0.95 | 1 | 0
|  |  | CCM 340 | 3; reference strain collection | DT | 2014-01-22 | Bacillus pumilus | 0.86 | 1 | 0
|  |  | DSM 25261 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.05.2018, N12 | Bacillus siamensis | 0.81 | 1 | 0
|  |  | DSM 13779 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2016, K18 | Bacillus sonorensis | 0.71 | 1 | 0
|  |  | DSM 12369 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 2007-04-18 | Bacillus sp-DSM-12369 | 0.85 | 1 | 0
|  |  | ATCC 6633 CVUAS | 3; public strain collection | eDT | 2015-07-22 | Bacillus spizizenii | 0.68 | 1 | 0
|  |  | DSM 618 | 4; public strain collection | eDT | 27.01.2016 | Bacillus spizizenii | 0.83 | 1 | 0
|  |  | CCM 110 | 3; reference strain collection | DT | 2014-01-22 | Bacillus subtilis | 0.64 | 1 | 0
|  |  | DSM 1092 | 3; reference strain collection | DT | 2014-01-22 | Bacillus subtilis | 0.63 | 1 | 0
|  |  | BBA 847 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.39 | 1 | 0
|  |  | BBA 846 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.45 | 1 | 0
|  |  | CVUAS 9675 | 4; biochemistry | eDT | 2017-07-04 | Bacillus thuringiensis | 0.31 | 1 | 0
|  |  | BBA 845 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.24 | 1 | 0
|  |  | BBA 848 | 4; biochemistry; specific gen | eDT | 2016-01-27 | Bacillus thuringiensis | 0.31 | 1 | 0
|  |  | DSM 6890 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.12.2015, L4 | Bacillus thuringiensis | 0.27 | 1 | 0
|  |  | WSBC 28001 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 07.11.2006 K4 | Bacillus thuringiensis | 0.35 | 1 | 0
|  |  | CVUAS 2786,2 | 3; sequenced | DT | 2013-07-31 | Bacillus timorensis | 0.72 | 1 | 0
|  |  | DSM 11031 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.07.2017, H17 | Bacillus vallismortis | 0.47 | 1 | 0
|  |  | DSM 23495 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.02.2016 K19 | Cytobacillus horneckiae | 0.52 | 1 | 0
|  |  | RRLK02_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.04.2014 H5 | Cytobacillus horneckiae | 0.6 | 1 | 0
|  |  | CVUAS 11431 | 3; sequenced | eDT | 2016-01-27 | Cytobacillus kochii | 0.48 | 1 | 0
|  |  | DSM 1314 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 19.02.2015, E18 | Domibacillus aminovorans | 0.56 | 1 | 0
|  |  | DSM 4337 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 15.02.2015, E8 | Domibacillus aminovorans | 0.58 | 1 | 0
|  |  | DSM 8720 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.02.2015 H5 | Evansella clarkii | 0.49 | 1 | 0
|  |  | DSM 9725 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F7 | Evansella clarkii | 0.76 | 1 | 0
|  |  | DSM 6790 | 3; public strain collection | eDT | 2020-06-17 | Geobacillus stearothermophilus | 0.72 | 1 | 0
|  |  | DSM 497 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 I10; (single sp is included in the corresponding msp) | Halalkalihalobacillus halodurans | 0.49 | 1 | 0
|  |  | DSM 6940 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015 E19; (single sp is included in the corresponding msp) | Halalkalihalobacillus halodurans | 0.94 | 1 | 0
|  |  | B041b | 3; sequenced | DT | 2015-07-22 | Heyndrickxia sporothermodurans | 0.43 | 1 | 0
|  |  | DSM 10599 RKI | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.03.2015, H9 | Heyndrickxia sporothermodurans | 0.58 | 1 | 0
|  |  | DSM 5221 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017; N12 | Lederbergia lenta | 0.6 | 1 | 0
|  |  | DSM 9 CVUAS | 3; type strain | DT | 2014-01-22 | Lederbergia lenta | 1.04 | 1 | 0
|  |  | DSM 15601 | 3; type strain | EFex [BRU03] | 2022-01-13 | Neobacillus bataviensis | 0.78 | 1 | 0
|  |  | DSM 101566 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.01.2017 N13 | Neobacillus cucumis | 0.74 | 1 | 0
|  |  | DSM 7257 | 3; public strain collection | 80% TFA/ 1:1 HCCA/TA2 (A) | 05.03.2007, E14, [0034] | Niallia circulans | 0.76 | 1 | 0
|  |  | CVUAS 2582,2 | 3; sequenced | eDT | 2015-09-23 | Niallia nealsonii | 0.31 | 1 | 0
|  |  | DSM 16288 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.03.2015 K20 | Peribacillus muralis | 0.73 | 1 | 0
|  |  | DSM 6 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016; K20 | Peribacillus psychrosaccharolyticus | 1.01 | 1 | 0
|  |  | 233bU3-2 | 3; sequenced | EFEx | 2016-12-28 | Peribacillus simplex | 0.54 | 1 | 0
|  |  | A379 | 2; strain collection of RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2016 H13 | Peribacillus simplex | 0.54 | 1 | 0
|  |  | DSM 1321 | 3; type strain | in 80% TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 15.12.2016, K15, [0096] | Peribacillus simplex | 0.67 | 1 | 0
|  |  | DSM 1320 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.02.2016 K15 | Priestia flexa | 0.55 | 1 | 0
|  |  | DSM 1667 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H18 | Priestia flexa | 0.48 | 1 | 0
|  |  | B 120 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.03.2015 H21 | Priestia megaterium | 0.62 | 1 | 0
|  |  | B090 | 3; sequenced, biochemistry | DT | 2015-07-22 | Priestia megaterium | 0.48 | 1 | 0
|  |  | DSM 16461 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 30.03.2015, E3 | Salibacillus salarius | 0.6 | 1 | 0
|  |  | DSM 6350 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.02.2015 F13 | Sutcliffiella cohnii | 1.13 | 1 | 0
|  |  | DSM 6351 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.02.2015 K8 | Sutcliffiella cohnii | 0.7 | 1 | 0
|  |  | B 043 | 3; sequenced | eDT | 2014-01-29 | Terribacillus saccharophilus | 0.72 | 1 | 0
|  |  | B 105 | 3; sequenced | DT | 2014-01-29 | Virgibacillus proomii | 0.46 | 1 | 0
|  |  | CVUAS 3616,2 | 3; sequenced | eDT | 2014-04-22 | Virgibacillus proomii | 0.42 | 1 | 0
|  |  | DSM 1 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.01.2017 H15 | Weizmannia coagulans | 0.42 | 1 | 0
|  |  | CVUAS 71,2 | 3; sequenced | DT | 2013-12-13 | Weizmannia ginsengihumi | 0.61 | 1 | 0
|  |  | CVUAS 3637 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.53 | 1 | 0
|  |  | CVUAS 9535 | 3; sequenced | DT | 2012-11-20 | Kurthia gibsonii | 0.38 | 1 | 0
|  |  | CVUAS 9909 | 3; sequenced | eDT | 2012-10-24 | Kurthia zopfii | 0.94 | 1 | 0
|  |  | B092 | 3; sequenced | DT | 2015-07-22 | Lysinibacillus fusiformis | 0.28 | 1 | 0
|  |  | LAZBW B1 | 3; sequenced | EFEx | 2019-04-10 | Lysinibacillus fusiformis | 0.61 | 1 | 0
|  |  | CVUAS 7192,2 | 3; sequenced | EFEx | 2020-02-21 | Lysinibacillus macroides | 0.73 | 1 | 0
|  |  | B066 | 3; sequenced | eDT | 2015-07-22 | Lysinibacillus sphaericus | 0.65 | 1 | 0
|  |  | CCM 1087 | 3; public strain collection | DT | 2014-01-22 | Lysinibacillus sphaericus | 0.8 | 1 | 0
|  |  | CVUAS 32374 | 3; sequenced | eDT | 2020-07-09 | Metalysinibacillus saudimassiliensis | 0.52 | 1 | 0
|  |  | CVUAS 7684 | 3; sequenced | eDT | 2014-03-06 | Solibacillus odysseyi | 0.31 | 1 | 0
|  |  | P-1574267 | 3; sequenced | EFEx | 2020-09-01 | Exiguobacterium acetylicum | 0.48 | 1 | 0
|  |  | CVUAS 10100 | 3; sequenced | EFEx | 2017-09-26 | Exiguobacterium enclense | 0.89 | 1 | 0
|  |  | CVUAS 10045 | 3; sequenced | DT | 2013-05-14 | Exiguobacterium profundum | 0.65 | 1 | 0
|  |  | DSM 20171 | 3; type strain | eDT | 2016-08-09 | Brochothrix thermosphacta | 0.53 | 1 | 0
|  |  | HS33 LVU07 | 3; sequenced (LGL) | EFEx | 2015-04-15 | Brochothrix thermosphacta | 0.26 | 1 | 0
|  |  | CVUAS 2686 | 3; sequenced | DT | 2013-01-09 | Listeria aquatica | 1.21 | 1 | 0
|  |  | LAZBW-A2021-4759-1 | 3; sequenced | EFEx | 2021-11-05 | Listeria aquatica | 0.65 | 1 | 0
|  |  | DSM 24998 | 3; type strain [0002] | EFEx | 2017-02-15 | Listeria fleischmannii | 0.63 | 1 | 0
|  |  | LU2006-3 EFex | 3; isolate of species description [0002]; sequenced | EFEx | 2017-02-15 | Listeria fleischmannii | 0.7 | 1 | 0
|  |  | CIP 105447 EFEx | 3; type strain | EFEx | 2017-02-15 | Listeria grayi | 0.75 | 1 | 0
|  |  | DSM 20596 EFex | 3; offizielle Stammsammlung | EFEx | 2017-02-15 | Listeria grayi | 0.6 | 1 | 0
|  |  | CVUAS 10132,2 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria innocua | 0.76 | 1 | 0
|  |  | CVUAS 8569 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria innocua | 0.7 | 1 | 0
|  |  | CVUAS 30423 | 3; sequenced | EFex | 2017-10-12 | Listeria ivanovii | 0.92 | 1 | 0
|  |  | CVUAS 32100 | 2; sequenced | EFex | 2019-12-13 | Listeria ivanovii | 0.52 | 1 | 0
|  |  | DSM 23813 EFex | 3; type strain | EFEx | 2017-02-15 | Listeria marthii | 0.92 | 1 | 0
|  |  | CVUAS 8071 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.97 | 1 | 0
|  |  | CVUAS 873,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.21 | 1 | 0
|  |  | CVUAS 240,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.43 | 1 | 0
|  |  | CVUAS 5798 | 4; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.56 | 1 | 0
|  |  | 10BfR L0206 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.87 | 1 | 0
|  |  | 10BfR L0066 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.57 | 1 | 0
|  |  | CVUAS 6122 | 3; NRL BfR | EFex | 2017-02-23 | Listeria monocytogenes | 0.68 | 1 | 0
|  |  | 10BfR L0211 | 4; public strain collection | EFex | 2017-02-23 | Listeria monocytogenes | 0.53 | 1 | 0
|  |  | CVUAS 7366,2 | 4; NRL BfR | EFex | 2017-01-25 | Listeria monocytogenes | 0.68 | 1 | 0
|  |  | CVUAS 33416 | 3; sequenced | EFEx | 2021-11-22 | Listeria newyorkensis | 0.9 | 1 | 0
|  |  | DSM 22097 EFEx | 3; type strain | EFEx | 2017-02-15 | Listeria rocourtiae | 0.64 | 1 | 0
|  |  | CVUAS 2763 EFex | 3; 16S Sequenz | EFEx | 2017-02-02 | Listeria seeligeri | 0.9 | 1 | 0
|  |  | CVUAS 9353 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria seeligeri | 0.58 | 1 | 0
|  |  | CVUAS 31185 | 3; sequenced | EFEx | 2018-09-18 | Listeria sp | 0.61 | 1 | 0
|  |  | CVUAS 2412,3 | 3; 16S Sequenz | EFEx | 2017-02-15 | Listeria welshimeri | 0.85 | 1 | 0
|  |  | CVUAS 1561 | 3; sequenced | EFEx | 2017-02-02 | Listeria welshimeri | 1.05 | 1 | 0
|  |  | CCM 1612 | 3; copy of isolate from public strain collection | eDT | 2015-09-23 | Brevibacillus laterosporus | 0.77 | 1 | 0
|  |  | CCM 2116 | 3; type strain | DT | 2014-01-22 | Brevibacillus laterosporus | 0.37 | 1 | 0
|  |  | B 210 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.03.2016 K19 | Paenibacillus alvei | 0.71 | 1 | 0
|  |  | NCTC 7578 | 3; public strain collection | DT | 2014-01-22 | Paenibacillus chibensis | 0.6 | 1 | 0
|  |  | CVUAS 32347 | 3; sequenced | eDT | 2020-06-05 | Paenibacillus nuruki | 0.49 | 1 | 0
|  |  | A418-1 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.09.2013 C8 | Paenibacillus polymyxa | 0.73 | 1 | 0
|  |  | DSM 365 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.11.2017 C18 | Paenibacillus polymyxa | 0.69 | 1 | 0
|  |  | CVUAS 6369 | 3; sequenced | eDT | 2013-09-29 | Paenibacillus sp-CVUAS-6369 | 0.62 | 1 | 0
|  |  | HAMBI 3232 | 3; public strain collection; species position uncertain | eDT | 2013-07-18 | Paenibacillus sp-E8a | 0.31 | 1 | 0
|  |  | B 128 | 3; sequenced | DT | 2014-01-22 | Paenibacillus taichungensis | 0.59 | 1 | 0
|  |  | FLI003 | 3; sequenced | EFEx | 2019-04-16 | Paenibacillus taichungensis | 0.28 | 1 | 0
|  |  | HAMBI 3233 | 3; type strain | EFEx | 2016-08-23 | Paenibacillus tundrae | 0.69 | 1 | 0
|  |  | 190-z-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus halophilus | 0.36 | 1 | 0
|  |  | 257-y-10 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus huakuii | 0.41 | 1 | 0
|  |  | 1403-y-3 | 3; sequenced | EFEx | 2017-10-26 | Jeotgalicoccus psychrophilus | 0.3 | 1 | 0
|  |  | CVUAS 31604 | 3; sequenced | EFEx | 2019-03-13 | Macrococcus carouselicus | 0.85 | 1 | 0
|  |  | CVUAS 121 | 3; 16S rDNA-sequencing | DT | 14.03.2013 | Macrococcus caseolyticus | 0.77 | 1 | 0
|  |  | CVUAS 3051 | 3; 16S rDNA-sequencing | DT | 2013-03-13 | Macrococcus caseolyticus | 0.67 | 1 | 0
|  |  | CVUAS 2968 | 3; 16S rDNA-sequencing | DT | 2013-06-18 | Mammaliicoccus lentus | 0.83 | 1 | 0
|  |  | CVUAS 3565 | 3; sequenced | EFEx | 2017-12-13 | Mammaliicoccus sciuri | 0.69 | 1 | 0
|  |  | CVUAS 4650 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Mammaliicoccus sciuri | 0.71 | 1 | 0
|  |  | CVUAS 9937,2 EFex | 3; 16S rDNA-sequencing | EFex | 2016-04-06 | Mammaliicoccus vitulinus | 0.56 | 1 | 0
|  |  | CVUAS 343,2 | 3; 16S rDNA-sequencing | EFex | 2014-06-06 | Mammaliicoccus vitulinus | 0.24 | 1 | 0
|  |  | CVUAS 5687 | 3; sequenced | DT | 2013-04-16 | Staphylococcus agnetis | 0.54 | 1 | 0
|  |  | TUM 7385 | 3; sequenced | DT | 02.10.2013 | Staphylococcus agnetis | 0.97 | 1 | 0
|  |  | RK308 | 4; WGS | DT | 2018-04-12 | Staphylococcus argenteus | 1 | 1 | 0
|  |  | DSM 28299 | 3; type strain [0003] | DT | 2015-04-21 | Staphylococcus argenteus | 0.91 | 1 | 0
|  |  | 254U4-2 | 3; sequenced | eDT | 2016-12-28 | Staphylococcus arlettae | 0.78 | 1 | 0
|  |  | CVUAS 3762 | 3; sequenced | DT | 2013-02-13 | Staphylococcus arlettae | 0.53 | 1 | 0
|  |  | ATCC 29213 CVUAS | 3; public strain collection | eDT | 2015-05-27 | Staphylococcus aureus | 0.85 | 1 | 0
|  |  | CVUAS 9914,2 sir | 4; sequenztype (NRL BfR) | EFEx | 2022-04-27 | Staphylococcus aureus | 0.86 | 1 | 0
|  |  | 277U2-3 | 3; sequenced | EFEx | 2017-01-03 | Staphylococcus auricularis | 0.24 | 1 | 0
|  |  | DSM 20326 (T) | 3; type strain [0003] | DT | 2013-10-02 | Staphylococcus capitis | 0.69 | 1 | 0
|  |  | TUM 5814 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus capitis | 0.36 | 1 | 0
|  |  | DSM 20608 (T) | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus caprae | 0.29 | 1 | 0
|  |  | TUM 4073 | 3; rpoB-sequencing | DT | 2013-10-04 | Staphylococcus caprae | 0.5 | 1 | 0
|  |  | CVUAS 1426 | 2; API Staph ID32 | eDT | 2015-12-01 | Staphylococcus carnosus | 0.26 | 1 | 0
|  |  | CVUAS 4590 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus carnosus | 0.73 | 1 | 0
|  |  | 5380 | 2; biochemical tube tests | DT | 2013-04-03 | Staphylococcus chromogenes | 0.6 | 1 | 0
|  |  | TS 176 | 2; API Staph ID32 | DT | 2013-04-03 | Staphylococcus chromogenes | 0.49 | 1 | 0
|  |  | DSM 20260 | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus cohnii | 0.72 | 1 | 0
|  |  | 208x-1 | 3; sequenced | EFEx | 2021-04-15 | Staphylococcus cohnii | 0.49 | 1 | 0
|  |  | x Sta | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012 K10 | Staphylococcus delphini | 0.4 | 1 | 0
|  |  | CVUAS 9351 | 3; sequenced | eDT | 2015-04-10 | Staphylococcus delphini | 0.32 | 1 | 0
|  |  | 215-U2-2 | 3; sequenced (CVUAS) | EFEx | 2017-06-28 | Staphylococcus devriesei | 0.13 | 1 | 0
|  |  | CVUAS 3786 | 3; 16S rDNA-sequencing | DT | 2012-12-04 | Staphylococcus epidermidis | 0.09 | 1 | 0
|  |  | DSM 20044 | 3; type strain | DT | 2013-05-15 | Staphylococcus epidermidis | 0.52 | 1 | 0
|  |  | 4357 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus equorum | 0.64 | 1 | 0
|  |  | DSM 20674 T | 3; type strain [0003] | DT | 2013-10-04 | Staphylococcus equorum | 0.62 | 1 | 0
|  |  | CVUAS 8203 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus felis | 0.86 | 1 | 0
|  |  | CVUAS 8599 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus felis | 1.09 | 1 | 0
|  |  | CCM 1798 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.08.2011 H9 | Staphylococcus haemolyticus | 0.02 | 1 | 0
|  |  | CVUAS 5860 | 3; 16S rDNA-sequencing | DT | 2013-02-13 | Staphylococcus haemolyticus | 0.09 | 1 | 0
|  |  | DSM 20328 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N6 [0034] | Staphylococcus hominis | 0.46 | 1 | 0
|  |  | CVUAS 789,2 | 3; 16S rDNA-sequencing | DT | 2014-01-22 | Staphylococcus hominis | 0.85 | 1 | 0
|  |  | CVUAS 2838 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus hyicus | 0.11 | 1 | 0
|  |  | CVUAS 3767 | 3; 16S rDNA-sequencing | DT | 2013-04-03 | Staphylococcus hyicus | 0.33 | 1 | 0
|  |  | CCM 5739 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, N8, [0034] | Staphylococcus intermedius | 0.26 | 1 | 0
|  |  | DSM 20373 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 08.09.2011, N7, [0034] | Staphylococcus intermedius | 0.4 | 1 | 0
|  |  | ATCC 43809 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 05.04.2012, N5, [0034] | Staphylococcus lugdunensis | 0.23 | 1 | 0
|  |  | 09-00266 | 3; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 04.08.2011, H16 | Staphylococcus lugdunensis | 0.33 | 1 | 0
|  |  | CVUAS 6077 | 3; 16S rDNA-sequencing | DT | 2012-10-31 | Staphylococcus microti | 0.94 | 1 | 0
|  |  | CVUAS 3071 | 3; 16S rDNA-sequencing | DT | 2012-11-14 | Staphylococcus pasteuri | 1.03 | 1 | 0
|  |  | DSM 10656 | 3; public strain collection | DT | 2013-10-04 | Staphylococcus pasteuri | 0.92 | 1 | 0
|  |  | DSM 21284 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 02.03.2016, H05, [0034] | Staphylococcus pseudintermedius | 0.64 | 1 | 0
|  |  | CVUAS 8759 | 2; biochemical tube tests | DT | 2013-07-02 | Staphylococcus pseudintermedius | 0.9 | 1 | 0
|  |  | 4824 | 2; biochemical tube tests | DT | 2013-02-13 | Staphylococcus saprophyticus | 0.04 | 1 | 0
|  |  | CVUAS 503,2 | 3; rpoB-sequencing | DT | 2013-09-17 | Staphylococcus saprophyticus | 0.44 | 1 | 0
|  |  | CVUAS 9352 | 2; 16S rDNA-sequencing | DT | 2013-07-02 | Staphylococcus schleiferi | 0.5 | 1 | 0
|  |  | DSM 28300 | 3; type strain | extr | 2015-04-21 | Staphylococcus schweitzeri | 0.24 | 1 | 0
|  |  | 202-42 | 2; API Staph ID32 | DT | 2013-02-05 | Staphylococcus simulans | 0.44 | 1 | 0
|  |  | TUM 7391 | 3; rpoB-sequencing | DT | 2013-10-02 | Staphylococcus simulans | 0.47 | 1 | 0
|  |  | CVUAS 2439,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 0.45 | 1 | 0
|  |  | CVUAS 761,2 | 3; 16S rDNA-sequencing | EFex | 2015-09-23 | Staphylococcus succinus | 0.66 | 1 | 0
|  |  | CVUAS 4714 | 3; 16S rDNA-sequencing | eDT | 2015-12-01 | Staphylococcus warneri | 1.14 | 1 | 0
|  |  | DSM 20316 CVUAS | 3; type strain | DT | 2013-10-02 | Staphylococcus warneri | 0.72 | 1 | 0
|  |  | 5741 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.06 | 1 | 0
|  |  | TS 251 | 2; biochemical tube tests | DT | 2013-04-30 | Staphylococcus xylosus | 0.21 | 1 | 0
|  |  | 185-U1-2 | 3; sequenced | eDT | 2017-03-13 | Aerococcus sp-185-U1-2 | 0.53 | 1 | 0
|  |  | 237z-2 | 3; sequenced | eDT | 2016-08-04 | Aerococcus suis | 0.97 | 1 | 0
|  |  | 257-y-1 | 3; sequenced | EFEx | 2017-10-26 | Aerococcus urinaeequi | 0.41 | 1 | 0
|  |  | CVUAS 2066 | 3; sequenced, biochemical profile | eDT | 27.09.2016 | Aerococcus viridans | 0.52 | 1 | 0
|  |  | JB 379054.2 | 3; sequenced | DT | 2013-09-04 | Atopobacter phocae | 0.67 | 1 | 0
|  |  | CVUAS 2442,2 | 3; sequenced | EFEx | 2017-02-09 | Carnobacterium divergens | 0.47 | 1 | 0
|  |  | CVUAS 30407 | 3; sequenced | EFEx | 2017-09-05 | Carnobacterium divergens | 0.16 | 1 | 0
|  |  | CVUAS 6273 | 3; sequenced | eDT | 13.01.2016 | Carnobacterium maltaromaticum | 0.49 | 1 | 0
|  |  | CVUAS 1519,3 | 3; sequenced | EFEx | 2016-12-07 | Granulicatella adiacens | 0.56 | 1 | 0
|  |  | CVUAS 2616 | 3; sequenced | DT | 25.09.2013 | Marinilactibacillus sp-CVUAS-2616 | 0.46 | 1 | 0
|  |  | CVUAS 8980,3 | 3; sequenced | EFEx | 2020-01-17 | Marinilactibacillus sp-CVUAS-2616 | 0.41 | 1 | 0
|  |  | 16 Ec RKI | 2; strain collection of RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011 H20 | Enterococcus casseliflavus | 0.44 | 1 | 0
|  |  | 6 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H12 | Enterococcus durans | 0.47 | 1 | 0
|  |  | 8 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 16.09.2011; H13 | Enterococcus durans | 0.48 | 1 | 0
|  |  | UW 6847 | 3; confirmed by Sequence typing [0167] | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.07.2011 N8 | Enterococcus faecium | 0.54 | 1 | 0
|  |  | 14 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N13 | Enterococcus gallinaru | 0.54 | 1 | 0
|  |  | 2 | 2; strain collection of the RKI | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCC in TA2 | 23.09.2011; N3 | Enterococcus hirae | 0.92 | 1 | 0
|  |  | MB 2022000094 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Enterococcus casseliflavus | 0.23 | 1 | 0
|  |  | CVUAS 5030 | 3; sequenced | eDT | 23.07.2014 | Enterococcus cecorum | 0.62 | 1 | 0
|  |  | CVUAS 8897 | 3; sequenced | eDT | 23.07.2014 | Enterococcus cecorum | 0.58 | 1 | 0
|  |  | CVUAS 31454 | 3; sequenced | eDT | 2020-08-20 | Enterococcus faecalis | 0.69 | 1 | 0
|  |  | DSM 2570 | 3; public strain collection | eDT | 2016-06-08 | Enterococcus faecalis | 1.03 | 1 | 0
|  |  | MR11-207 LVU04 | 3; sequenced (LGL) | eDT | 2015-04-15 | Enterococcus faecium | 0.91 | 1 | 0
|  |  | BM 4174 | 2; 0 | eDT | 2016-06-08 | Enterococcus gallinarum | 0.9 | 1 | 0
|  |  | 265z-7 | 3; sequenced | EFEx | 2016-12-29 | Tetragenococcus solitarius | 0.93 | 1 | 0
|  |  | CVUAS 1549,2 | 3; sequenced | DT | 2016-09-27 | Vagococcus salmoninarum | 0.69 | 1 | 0
|  |  | ATCC 4356 | 3; type strain | DT | 2016-02-26 | Lactobacillus acidophilus | 0.98 | 1 | 0
|  |  | CVUAS 9322 | 3; sequenced | DT | 2016-02-26 | Lactobacillus johnsonii | 0.52 | 1 | 0
|  |  | CVUAS 32277,2 | 3; sequenced | eDT | 2020-07-21 | Lactobacillus sp-CVUAS-32277,2 | 0.76 | 1 | 0
|  |  | DSM 14800 LVU14 | 3; type strain | eDT | 2015-04-15 | Pediococcus claussenii | 0.78 | 1 | 0
|  |  | LTH 1213 | 2; 0 | DT | 2013-07-09 | Leuconostoc mesenteroides | 0.41 | 1 | 0
|  |  | LTH 1146 | 2; 0 | DT | 2013-07-09 | Lactococcus lactis | 0.1 | 1 | 0
|  |  | 08-7-D-06424 | 3; well defined isolate | EFEx | 2015-03-11 | Streptococcus agalactiae | 0.61 | 1 | 0
|  |  | DSM 6784 CVUAS | 3; public strain collection | eDT | 2019-07-09 | Streptococcus agalactiae | 0.33 | 1 | 0
|  |  | CVUAS 1924 | 3; 16S, confirmed by NRZ Aachen | DT | 2013-04-16 | Streptococcus canis | 0.27 | 1 | 0
|  |  | LHL Mi-845_17 | 3; sequenced | eDT | 2017-10-05 | Streptococcus canis | 0.83 | 1 | 0
|  |  | 10UCF103 | 3; sequenced (LHL), [0114] | EFEx | 2018-05-03 | Streptococcus castoreus | 0.3 | 1 | 0
|  |  | DSM 17536 CVUAS | 3; type strain | EFEx | 2018-05-03 | Streptococcus castoreus | 0.59 | 1 | 0
|  |  | 355-1-2018 | 3; stain of species description [0130] | EFEx | 2019-06-07 | Streptococcus catagoni | 0.87 | 1 | 0
|  |  | 99-1-2017 | 3; type strain | EFEx | 2018-05-03 | Streptococcus catagoni | 0.65 | 1 | 0
|  |  | CVUAS 1514 | 2; API: 99,9% | EFEx | 2015-04-29 | Streptococcus dysgalactiae | 0.72 | 1 | 0
|  |  | CVUAS 4270 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0 | 1 | 0
|  |  | CVUAS 1494 | 3; API: 99,5%; sequenced | DT | 2013-02-19 | Streptococcus dysgalactiae | 0.68 | 1 | 0
|  |  | CVUAS 2151 | 3; sequenced | DT | 2013-02-20 | Streptococcus dysgalactiae | 0.42 | 1 | 0
|  |  | CVUAS 32462,2 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus equi | 0.14 | 1 | 0
|  |  | CVUAS 32770 | 4; specific PCR; (NRZ Strep: 16S Cluster interpretation) | eDT | 2022-05-30 | Streptococcus equi | 0.46 | 1 | 0
|  |  | 10HE 00580 | 3; sequenced, specific PCR | eDT | 2018-03-20 | Streptococcus equi | 0.29 | 1 | 0
|  |  | CVUAS 1004 | 3; specific PCR | eDT | 2021-05-11 | Streptococcus equi | 0.53 | 1 | 0
|  |  | CVUAS 5788 | 3; sequenced | Efex | 2015-05-05 | Streptococcus iniae | 0.13 | 1 | 0
|  |  | CVUAS 2474 | 3; sequenced; 16S, rpoB | eDT | 2013-03-07 | Streptococcus orisasini | 0 | 1 | 0
|  |  | CVUAS 4971 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus parauberis | 0.12 | 1 | 0
|  |  | JB 379054.1 | 3; 16S rRNA gene sequence | DT | 2013-02-20 | Streptococcus phocae | 0.27 | 1 | 0
|  |  | 12-13862-001-01 LVU11 | 3; sequenced (LGL) | eDT | 2015-04-15 | Streptococcus pluranimalium | 0.36 | 1 | 0
|  |  | CVUAS 5040 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.15 | 1 | 0
|  |  | 29800 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.71 | 1 | 0
|  |  | CVUAS 4253 | 4; confirmed by NRL Streptococcus (Uni AA) | eDT | 2021-10-22 | Streptococcus pneumoniae | 0.11 | 1 | 0
|  |  | DSM 24048 | 3; public strain collection | EFEx | 2017-12-28 | Streptococcus pneumoniae | 0.63 | 1 | 0
|  |  | CVUAS 905 | 3; sequenced | DT | 2013-02-19 | Streptococcus porcinus | 0.11 | 1 | 0
|  |  | DSM 20565 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.04.2008, K11, [0034] | Streptococcus pyogenes | 0.52 | 1 | 0
|  |  | CVUAS 10145,3 | 3; sequenced; 16S, Biochemistry | EFEx | 2016-11-23 | Streptococcus salivarius | 0.48 | 1 | 0
|  |  | CVUAS 32349 | 3; sequenced | eDT | 2020-06-05 | Streptococcus sp-CVUAS-32349 | 0.09 | 1 | 0
|  |  | CVUAS 2260,4 | 3; sequenced | eDT | 2021-06-08 | Streptococcus troglodytidis | 0.58 | 1 | 0
|  |  | CVUAS 5218,4 | 3; sequenced; 16S; NRZ Strep Uni AA | DT | 2015-06-11 | Streptococcus troglodytidis | 0.02 | 1 | 0
|  |  | CVUAS 2048 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0.16 | 1 | 0
|  |  | CVUAS 8422 | 2; biochemistry | DT | 2013-02-20 | Streptococcus uberis | 0.23 | 1 | 0
|  |  | 14EL305692 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.57 | 1 | 0
|  |  | 14EL305725 | 4; ribotype; ex Clostridium | eDT | 2019-07-12 | Clostridioides difficile | 0.74 | 1 | 0
|  |  | 14EL402952 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.42 | 1 | 0
|  |  | 14EL402975 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.81 | 1 | 0
|  |  | 14EL402979 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.34 | 1 | 0
|  |  | 14EL403180 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.58 | 1 | 0
|  |  | 14EL403181 | 4; ribotype; ex Clostridium | EFEx | 2019-09-13 | Clostridioides difficile | 0.11 | 1 | 0
|  |  | 14EL403210 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.56 | 1 | 0
|  |  | 14EL403313 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.69 | 1 | 0
|  |  | 14EL403467 | 4; ribotype; ex Clostridium | EFEx | 2019-09-13 | Clostridioides difficile | 0 | 1 | 0
|  |  | 14EL403488 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.58 | 1 | 0
|  |  | 14EL403625 | 4; ribotype; ex Clostridium | DT | 2014-08-27 | Clostridioides difficile | 0.62 | 1 | 0
|  |  | S0014-11 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.48 | 1 | 0
|  |  | S0021-11 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.54 | 1 | 0
|  |  | S0022-11 | 4; ribotype; ex Clostridium | DT | 2012-12-19 | Clostridioides difficile | 0.33 | 1 | 0
|  |  | S0059-10 | 4; ribotype; ex Clostridium | DT | 2013-02-06 | Clostridioides difficile | 0.16 | 1 | 0
|  |  | S0061-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.62 | 1 | 0
|  |  | S0062-10 | 4; ribotype; ex Clostridium | DT | 2013-01-09 | Clostridioides difficile | 0.71 | 1 | 0
|  |  | S0063-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.62 | 1 | 0
|  |  | S0072-10 | 4; ribotype; ex Clostridium | DT | 2013-08-13 | Clostridioides difficile | 0.81 | 1 | 0
|  |  | REB 1750 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.64 | 1 | 0
|  |  | C bot B | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.42 | 1 | 0
|  |  | REB 89 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.58 | 1 | 0
|  |  | C bot B L677 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.77 | 1 | 0
|  |  | C bot F1 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.48 | 1 | 0
|  |  | C bot F2 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.96 | 1 | 0
|  |  | REB 1718 | 2; 0 | DT | 2012-12-19 | Clostridium [botulinum] | 0.81 | 1 | 0
|  |  | CVUAS 30006 | 3; sequenced | eDT | 2017-06-22 | Clostridium botulinum | 0.62 | 1 | 0
|  |  | CVUAS 32971 | 3; sequenced | EFEx | 2021-02-25 | Clostridium botulinum | 0.79 | 1 | 0
|  |  | CVUAS 33885,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium botulinum | 0.79 | 1 | 0
|  |  | CVUAS 33886,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium botulinum | 0.71 | 1 | 0
|  |  | CVUAS 33891,2 | 3; specific PCR | eDT | 2024-07-01 | Clostridium botulinum | 0.92 | 1 | 0
|  |  | DSM 7528 | 3; type strain | eDT | 2018-08-20 | Clostridium chauvoei | 0.32 | 1 | 0
|  |  | CVUAS 849,2 | 3; sequenced | DT | 2014-12-17 | Clostridium isatidis | 0.37 | 1 | 0
|  |  | DSM 108256 | 3; Public strain collection | EFEx | 0 | Clostridium leptum | 0.78 | 1 | 0
|  |  | MSK_23_22 | 3; sequenced [0207] | EFEx | 0 | Clostridium nexile | 0.82 | 1 | 0
|  |  | CVUAS 9474 | 3; sequenced | DT | 2013-02-20 | Clostridium nigeriense | 0.66 | 1 | 0
|  |  | JCM_10421 | 3; Public strain collection | eDT | 0 | Clostridium scindens | 0.34 | 1 | 0
|  |  | JCM_10419 | 3; Public strain collection | eDT | 2021-10-26 | Clostridium scindens | 0.51 | 1 | 0
|  |  | JCM_10420 | 3; Public strain collection | EFEx | 2021-10-20 | Clostridium scindens | 0.9 | 1 | 0
|  |  | MSK_1_26 | 3; sequenced [0207] | eDT | 0 | Clostridium scindens | 0.39 | 1 | 0
|  |  | MSK_5_24 | 3; sequenced [0207] | eDT | 0 | Clostridium scindens | 0.87 | 1 | 0
|  |  | DSM 7534 | 3; type strain | eDT | 2018-08-20 | Clostridium septicum | 0.58 | 1 | 0
|  |  | CVUAS 34091 | 3; sequenced | EFex | 2022-08-05 | Clostridium sp-CVUAS-34091 | 0.55 | 1 | 0
|  |  | DSM 1985 | 2; public strain collection | DT | 2012-12-19 | Clostridium sp-DSM-1985 | 0.99 | 1 | 0
|  |  | CVUAS 7998 | 3; sequenced | DT | 2013-02-20 | Clostridium sporogenes | 0.28 | 1 | 0
|  |  | MSK_1_9 | 3; sequenced [0207] | eDT | 0 | Clostridium symbiosum | 0.98 | 1 | 0
|  |  | MSK_7_21 | 3; sequenced [0207] | eDT | 0 | Clostridium symbiosum | 0.92 | 1 | 0
|  |  | WAL-14163 | 3; sequenced [0207] | eDT | 0 | Clostridium symbiosum | 0.79 | 1 | 0
|  |  | 20 | 2; confirmed proficency test | eDT | 2022-02-14 | Clostridium tertium | 0.21 | 1 | 0
|  |  | CVUAS 10003 | 2; 0 | DT | 2013-02-20 | Clostridium tertium | 0.04 | 1 | 0
|  |  | CVUAS 10004 | 2; 0 | DT | 2013-02-20 | Clostridium tertium | 0.39 | 1 | 0
|  |  | CVUAS 7893 | 3; sequenced | DT | 2012-12-19 | Clostridium tertium | 0.27 | 1 | 0
|  |  | CVUAS 8330 | 3; sequenced | DT | 2012-12-19 | Clostridium tertium | 0.38 | 1 | 0
|  |  | CVUAS 1478 | 2; 0 | DT | 2012-12-19 | Clostridium tetani | 1.07 | 1 | 0
|  |  | DSM 16296 | 3; Public strain collection | EFEx | 0 | Eubacterium ramulus | 0.29 | 1 | 0
|  |  | CLA-AA-H258 | 3; sequenced [0207] | eDT | 0 | Agathobacter rectalis | 1.05 | 1 | 0
|  |  | MSK_9_13 | 3; sequenced [0207] | eDT | 0 | Agathobacter rectalis | 1.01 | 1 | 0
|  |  | Clinical isolate | 3; sequenced [0207] | EFEx | 0 | Anaerobutyricum hallii | 0.14 | 1 | 0
|  |  | CLA-AA-H53 | 3; sequenced [0207] | EFEx | 0 | Anaerostipes hadrus | 0.08 | 1 | 0
|  |  | DSM 29492 | 3; type strain | EFEx | 0 | Blautia caecimuris | 0.67 | 1 | 0
|  |  | MSK_18_51 | 3; sequenced [0207] | EFEx | 0 | Blautia faecis | 1.02 | 1 | 0
|  |  | DSM 22028 | 3; type strain | EFEx | 0 | Blautia glucerasea | 0.81 | 1 | 0
|  |  | DSM 20583 | 3; type strain | EFEx | 0 | Blautia hansenii | 1.13 | 1 | 0
|  |  | MSK_19_60 | 3; sequenced [0207] | eDT | 0 | Blautia massiliensis | 0.76 | 1 | 0
|  |  | MSK_20_85 | 3; sequenced [0207] | EFEx | 0 | Blautia massiliensis | 0.77 | 1 | 0
|  |  | MSK_21_1 | 3; sequenced [0207] | eDT | 0 | Blautia massiliensis | 1.26 | 1 | 0
|  |  | MSK_22_56 | 3; sequenced [0207] | eDT | 0 | Blautia massiliensis | 0.76 | 1 | 0
|  |  | MSK_9_19 | 3; sequenced [0207] | eDT | 0 | Blautia massiliensis | 1.12 | 1 | 0
|  |  | MSK_20_66 | 3; sequenced [0207] | eDT | 0 | Blautia obeum | 0.57 | 1 | 0
|  |  | CCUG 9990 | 3; type strain | eDT | 0 | Blautia producta | 0.57 | 1 | 0
|  |  | MSK_4_13 | 2; data not listed [0207] | eDT | 0 | Blautia producta | 0.76 | 1 | 0
|  |  | MSK_4_17 | 3; sequenced [0207] | eDT | 0 | Blautia producta | 0.63 | 1 | 0
|  |  | DSM 26115 | 3; Public strain collection | EFEx | 0 | Blautia pseudococcoides | 0.61 | 1 | 0
|  |  | MSK_6_27 | 3; sequenced [0207] | EFEx | 2021-10-20 | Blautia schinkii | 0.86 | 1 | 0
|  |  | MSK_6_32 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia schinkii | 1.15 | 1 | 0
|  |  | MSK_18_16 | 3; sequenced [0207] | eDT | 2021-10-20 | Blautia wexlerae | 0.7 | 1 | 0
|  |  | JCM 31264 | 3; Public strain collection | eDT | 2021-10-20 | Coprococcus comes | 0.6 | 1 | 0
|  |  | MSK_17_21 | 3; sequenced [0207] | eDT | 2021-10-20 | Coprococcus comes | 0.47 | 1 | 0
|  |  | MSK_10_6 | 3; sequenced [0207] | EFEx | 2021-10-20 | Coprococcus eutactus | 0 | 1 | 0
|  |  | JCM 31256 | 3; Public strain collection | eDT | 0 | Dorea formicigenerans | 0.27 | 1 | 0
|  |  | MSK_19_39 | 3; sequenced [0207] | eDT | 0 | Dorea formicigenerans | 0.35 | 1 | 0
|  |  | MSK_23_29 | 3; sequenced [0207] | eDT | 0 | Dorea formicigenerans | 0.33 | 1 | 0
|  |  | MSK_18_49 | 3; sequenced [0207] | eDT | 0 | Dorea longicatena | 0.13 | 1 | 0
|  |  | DSM 108191 | 3; Public strain collection | EFEx | 0 | Enterocloster clostridioformis | 0.58 | 1 | 0
|  |  | MSK_20_55 | 3; sequenced [0207] | eDT | 0 | Faecalicatena fissicatena | 0.9 | 1 | 0
|  |  | MSK_9_3 | 3; sequenced [0207] | eDT | 0 | Faecalicatena fissicatena | 0.26 | 1 | 0
|  |  | MSK_14_54 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter saccharivorans | 0.56 | 1 | 0
|  |  | MSK_15_44 | 3; sequenced [0207] | eDT | 0 | Fusicatenibacter saccharivorans | 0.62 | 1 | 0
|  |  | CLA-AA-H260 | 3; sequenced [0207] | EFEx | 0 | Lachnospira eligens | 0.32 | 1 | 0
|  |  | DSM 105336 | 3; Public strain collection | EFEx | 0 | Lacrimispora celerecrescens | 0.37 | 1 | 0
|  |  | CLA-AA-H164 | 3; sequenced [0207] | EFEx | 0 | Roseburia faecis | 0.48 | 1 | 0
|  |  | DSM 16839 | 3; type strain | EFEx | 0 | Roseburia hominis | 0.08 | 1 | 0
|  |  | DSM 108070 | 3; Public strain collection | EFEx | 0 | Roseburia inulinivorans | 0.42 | 1 | 0
|  |  | MSK_1_25 | 3; sequenced [0207] | eDT | 0 | Sellimonas intestinalis | 0.48 | 1 | 0
|  |  | JCM 31255 | 3; Public strain collection | eDT | 0 | Anaerotruncus colihominis | 0.39 | 1 | 0
|  |  | CLA-AA-H281 | 3; sequenced [0207] | EFEx | 0 | Faecalibacterium prausnitzii | 0.1 | 1 | 0
|  |  | CLA-AA-H271 | 3; sequenced [0207] | eDT | 0 | Ruminococcus bromii | 1 | 1 | 0
|  |  | MSK_22_53 | 3; sequenced [0207] | eDT | 0 | Ruminococcus gnavus | 0.76 | 1 | 0
|  |  | MSK_23_60 | 3; sequenced [0207] | eDT | 0 | Ruminococcus gnavus | 0.5 | 1 | 0
|  |  | CLA-AA-H279 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus lactaris | 0.65 | 1 | 0
|  |  | CLA-AA-H261 | 3; sequenced [0207] | EFEx | 0 | Ruminococcus torques | 0.09 | 1 | 0
|  |  | D14_1188-2-1-2 | 3; strain of species description | EFEx [BRU03] | 2018-06-07 | Erysipelothrix anatis | 0.71 | 1 | 0
|  |  | D17_0559-3-2-1 | 3; type strain | EFEx [BRU03] | 2018-05-03 | Erysipelothrix anatis | 0.87 | 1 | 0
|  |  | CVUAS 34515 | 3; sequenced | EFex [BRU03] | 2022-11-16 | Erysipelothrix aquatica | 0.47 | 1 | 0
|  |  | VA-92-K48 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix aquatica | 0.73 | 1 | 0
|  |  | DSM 28480 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix larvae | 0.97 | 1 | 0
|  |  | DSM 110099 CVUAS | 3; type strain | EFex [BRU03] | 2021-10-18 | Erysipelothrix piscisicarius | 0.05 | 1 | 0
|  |  | CIP 105957 | 3; type strain | EFex [BRU03] | 2018-05-03 | Erysipelothrix rhusiopathiae | 0.51 | 1 | 0
|  |  | CVUAS 33853 | 2; 0 | EFEx [BRU03] | 2022-05-17 | Erysipelothrix rhusiopathiae | 0.85 | 1 | 0
|  |  | CIP 105960 | 3; type strain | EFEx [BRU03] | 2018-11-27 | Erysipelothrix tonsillarum | 0.56 | 1 | 0
|  |  | IHIT 36981 | 2; 0 | EFex [BRU03] | 2018-11-23 | Erysipelothrix tonsillarum | 0.1 | 1 | 0
|  |  | CVUAS 10169 | 3; sequenced | eDT | 2014-04-09 | Anaerococcus sp-CVUAS-10169 | 0.35 | 1 | 0
|  |  | CVUAS 32948 | 3; confirmed proficency test | EFEx | 2021-02-18 | Finegoldia magna | 0.47 | 1 | 0
|  |  | CVUAS 11259 | 3; sequenced | eDT | 2016-10-31 | Helcococcus kunzii | 0.37 | 1 | 0
|  |  | DSM 10548 | 3; type strain | eDT | 2016-10-31 | Helcococcus kunzii | 0.47 | 1 | 0
|  |  | CVUAS 433,2 | 3; sequenced | eDT | 2017-02-02 | Helcococcus ovis | 0.65 | 1 | 0
|  |  | DSM 21504 | 3; type strain | eDT | 2016-10-31 | Helcococcus ovis | 0.56 | 1 | 0
|  |  | CVUAS 32321 | 3; sequenced | EFEx | 2020-06-04 | Helcococcus sp-CVUAS-32321 | 0.6 | 1 | 0
|  |  | CVUAS 692,2 | 3; sequenced | DT | 2013-02-19 | Parvimonas sp-CVUAS-692,2 | 0.48 | 1 | 0
|  |  | 151003106 | 2; 0 | DT | 2015-03-16 | Bacteroides finegoldii | 0.85 | 1 | 0
|  |  | 141004228 | 2; 0 | DT | 2014-04-08 | Bacteroides fragilis | 0.8 | 1 | 0
|  |  | 181001478 Bat | 2; 0 | DT | 2018-05-02 | Bacteroides pyogenes | 0.14 | 1 | 0
|  |  | 171011776-9 | 2; 0 | DT | 2017-08-31 | Prevotella heparinolytica | 0.79 | 1 | 0
|  |  | 181004443-8 | 2; 0 | DT | 2018-04-03 | Chryseobacterium gleum | 0.09 | 1 | 0
|  |  | 171014248-7 | 2; 0 | DT | 2017-10-20 | Chryseobacterium indologenes | 0.06 | 1 | 0
|  |  | 171016830 | 2; 0 | DT | 2017-12-13 | Chryseobacterium indologenes | 0.63 | 1 | 0
|  |  | CVUAS 32947 | 3; confirmed proficency test | EFEx | 2021-02-18 | Bacteroides fragilis | 0.12 | 1 | 0
|  |  | 131014917-2 | 2; 0 | DT | 2013-11-21 | Flavobacterium hibernum | 0.45 | 1 | 0
|  |  | 181003048-1 | 2; 0 | DT | 2018-05-03 | Flavobacterium sp | 0.53 | 1 | 0
|  |  | CVUAS 32831 | 3; sequenced | EFEx | 2021-04-23 | Chryseobacterium chaponense | 0.22 | 1 | 0
|  |  | CVUAS 5661 | 3; sequenced | EFEx | 2015-01-27 | Chryseobacterium hominis | 0.14 | 1 | 0
|  |  | 185-U1-6 sir | 3; sequenced | EFEx | 2022-04-27 | Chryseobacterium manosquense | 0.71 | 1 | 0
|  |  | 185-U1-6 EFex | 3; sequenced | EFEx | 2021-03-26 | Chryseobacterium manosquense | 0.44 | 1 | 0
|  |  | CVUAS 11450,4 | 3; sequenced | Efex | 2018-06-11 | Chryseobacterium sp-I-7 | 0.77 | 1 | 0
|  |  | CVUAS 8768 | 3; sequenced | DT | 2013-05-13 | Flavobacterium aquidurense | 0.43 | 1 | 0
|  |  | CVUAS 5322,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium branchiarum | 0.35 | 1 | 0
|  |  | I-28 | 3; sequenced | DT | 2018-04-23 | Flavobacterium chungangense | 0.32 | 1 | 0
|  |  | LMG 10406 | 3; type strain | DT | 2013-03-25 | Flavobacterium columnare | 0.68 | 1 | 0
|  |  | CVUAS 32889 | 3; sequenced | eDT | 2021-01-21 | Flavobacterium frigidimaris | 0.44 | 1 | 0
|  |  | CVUAS 4217,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium frigidimaris | 0.21 | 1 | 0
|  |  | CVUAS 31758 | 3; sequenced | EFEx | 2019-07-10 | Flavobacterium piscis | 0.19 | 1 | 0
|  |  | CVUAS 32762 | 3; sequenced (CVUAS) | eDT | 2020-11-13 | Flavobacterium plurextorum | 0.4 | 1 | 0
|  |  | CVUAS 5323,2 | 3; sequenced | eDT | 2017-02-07 | Flavobacterium psychoterrae | 0.81 | 1 | 0
|  |  | CVUAS 11550,4 | 3; sequenced | eDT | 2022-03-22 | Flavobacterium psychrophilum | 0.06 | 1 | 0
|  |  | DSM 3660 | 3; type strain | eDT | 2016-03-22 | Flavobacterium psychrophilum | 0.51 | 1 | 0
|  |  | CVUAS 10324 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium reichenbachii | 0.51 | 1 | 0
|  |  | CVUAS 2962 | 3; sequenced | DT | 2013-05-13 | Flavobacterium sp-C2 | 0.51 | 1 | 0
|  |  | CVUAS 10323 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-10323 | 0.5 | 1 | 0
|  |  | CVUAS 4146 | 3; sequenced | EFEx | 2014-04-07 | Flavobacterium sp-CVUAS-4146 | 0.91 | 1 | 0
|  |  | CVUAS 4216,2 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-4216,2 | 0.66 | 1 | 0
|  |  | CVUAS 5506 | 3; sequenced | eDT | 2014-01-27 | Flavobacterium sp-CVUAS-5506 | 0.46 | 1 | 0
|  |  | I-26 | 3; sequenced | DT | 2018-04-23 | Flavobacterium sp-HMD2009 | 0.44 | 1 | 0
|  |  | CVUAS 9320 | 2; sequenced | eDT | 2014-02-18 | Flavobacterium sp-JQMS | 0.64 | 1 | 0
|  |  | CVUAS 11530 | 3; sequenced | eDT | 2017-01-16 | Flavobacterium succinicans | 0.8 | 1 | 0
|  |  | CVUAS 8437 | 3; sequenced | DT | 2013-05-13 | Flavobacterium tiangeerense | 1.01 | 1 | 0
|  |  | CVUAS 33792 | 3; sequenced | EFex [BRU03] | 2022-10-21 | Riemerella columbina | 0.62 | 1 | 0
|  |  | CVUAS 2589,2 | 3; sequenced (CVUAS) | EFEx | 2021-08-05 | Epilithonimonas tenax | 0.93 | 1 | 0
|  |  | CVUAS 6132 | 2; 0 | EFex | 2021-02-24 15:17 | Ornithobacterium rhinotracheale | 0.21 | 1 | 0
|  |  | CVUAS 5275 | 2; 0 | DT | 2021-03-10 14:38 | Ornithobacterium rhinotracheale | 0.34 | 1 | 0
|  |  | CVUAS 9288,2 | 3; sequenced | Pextr | 2016-06-24 | Prototheca zopfii | 0.46 | 1 | 0
|  |  | CVUAS 9876,2 | 2; 0 | Pextr | 2016-09-01 | Prototheca zopfii | 0.65 | 1 | 0
|  |  | CVUAS 11266 | 3; sequenced | EFEx | 2016-10-05 | Deinococcus sp-CVUAS-11266 | 0.19 | 1 | 0
|  |  | 141011554-1 | 3; VETQAS | DT | 2014-09-05 | Fusobacterium necrophorum | 0.47 | 1 | 0
|  |  | 151003133 | 2; 0 | DT | 2015-03-17 | Fusobacterium varium | 0.48 | 1 | 0
|  |  | CVUAS 33436 | 2; 0 | eDT | 2021-12-01 | Streptobacillus moniliformis | 0.35 | 1 | 0
|  |  | CCUG 39713 | 3; type strain | EFEx | 2016-02-01 | Caviibacter abscessus | 0 | 1 | 0
|  |  | AVG 2115 | 3; type strain | eDT | 10.02.2016 | Oceanivirga salmonicida | 0.73 | 1 | 0
|  |  | DSM 26322 | 3; typestrain, sequenced | Extr | 2015-03-23 | Pseudostreptobacillus hongkongensis | 0.73 | 1 | 0
|  |  | DSM 100379 | 3; public strain collection; sequenced | eDT | 2015-07-08 | Sebaldella termitidis | 0.17 | 1 | 0
|  |  | NCTC 11300 | 3; type strain | DT | 2013-07-16 | Sebaldella termitidis | 0.56 | 1 | 0
|  |  | CCUG 38322 CVUAS | 3; public strain collection | EFEx | 16.12.2015 | Sneathia sanguinegens | 0.46 | 1 | 0
|  |  | CCUG 41628 CVUAS | 3; type strain | EFEx | 2020-10-08 | Sneathia sanguinegens | 0.58 | 1 | 0
|  |  | DSM 16630 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.39 | 1 | 0
|  |  | DSM 16631 | 3; sequenced (LHL) | EFEx | 2020-09-29 | Sneathia vaginalis | 0.63 | 1 | 0
|  |  | DSM 29248 | 3; type strain | DT | 2013-07-16 | Streptobacillus felis | 0.51 | 1 | 0
|  |  | ATCC 27747 | 3; public strain collection; well defined isolate [0008] | DT | 2013-07-16 | Streptobacillus moniliformis | 0.37 | 1 | 0
|  |  | AHL 370-1 | 3; type strain, sequenced | Extr | 2015-03-23 | Streptobacillus notomytis | 0.44 | 1 | 0
|  |  | OGS16T | 3; type strain | eDT | 2016-08-06 | Streptobacillus ratti | 0.45 | 1 | 0
|  |  | 171012806 | 2; 0 | DT | 2017-10-04 | Ochrobactrum intermedium | 0.52 | 1 | 0
|  |  | 181003400 | 2; 0 | DT | 2018-03-19 | Ochrobactrum intermedium | 0.37 | 1 | 0
|  |  | 171002894-4 | 2; 0 | DT | 2017-03-01 | Ochrobactrum sp | 0.9 | 1 | 0
|  |  | 171007549-1 | 3; sequenced, 16S rRNA gene | DT | 2017-06-06 | Ochrobactrum sp | 0.41 | 1 | 0
|  |  | 171012040 Och | 2; 0 | DT | 2017-10-04 | Ochrobactrum sp | 0.53 | 1 | 0
|  |  | ATCC 25922 | 3; public strain collection | DT | 2016-02-10 | Escherichia coli | 0.73 | 1 | 0
|  |  | 171012375 | 2; 0 | DT | 2017-09-14 | Klebsiella oxytoca | 0.48 | 1 | 0
|  |  | 171011727 | 2; 0 | DT | 2017-08-31 | Klebsiella variicola | 0.43 | 1 | 0
|  |  | CVUAS 1816,2 | 3; sequenced | DT | 2013-07-24 | Basfia succiniciproducens | 0.38 | 1 | 0
|  |  | 181017653-2 | 2; 0 | DT | 2018-12-10 | Pseudomonas alcaligenes | 1.02 | 1 | 0
|  |  | 171011707 | 2; 0 | DT | 2017-08-31 | Pseudomonas anguilliseptica | 0.53 | 1 | 0
|  |  | 181003516-1 | 2; 0 | DT | 2018-03-12 | Pseudomonas brenneri | 0.1 | 1 | 0
|  |  | 181017934 | 2; 0 | DT | 2018-12-12 | Pseudomonas fulva | 0.64 | 1 | 0
|  |  | 181003531 | 2; 0 | DT | 2018-03-12 | Pseudomonas koreensis | 0.36 | 1 | 0
|  |  | 181012674 | 2; 0 | DT | 2018-08-30 | Pseudomonas monteilii | 1.02 | 1 | 0
|  |  | 171011722-1_1 nit | 2; 0 | DT | 2017-08-31 | Pseudomonas nitroreducens | 0.72 | 1 | 0
|  |  | 171014178 | 2; 0 | DT | 2017-10-20 | Pseudomonas oryzihabitans | 0.16 | 1 | 0
|  |  | 171011722-1_2 put | 2; 0 | DT | 2017-08-31 | Pseudomonas putida | 0.22 | 1 | 0
|  |  | 181017452 | 2; 0 | DT | 2018-12-04 | Pseudomonas taetrolens | 0.76 | 1 | 0
|  |  | CVUAS 2280,3 | 3; sequenced | eDT | 2018-03-28 | Aureimonas glaciistagni | 0.36 | 1 | 0
|  |  | A-269_08 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 02.04.2012 K8 | Ochrobactrum anthropi | 0.53 | 1 | 0
|  |  | BM373 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; D02 | Brucella abortus | 0.39 | 1 | 0
|  |  | BM369 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; E01 | Brucella abortus | 0.66 | 1 | 0
|  |  | A183-5 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; I16 | Brucella canis | 0.19 | 1 | 0
|  |  | A466-3 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 05.03.2014; E17 | Brucella canis | 0.54 | 1 | 0
|  |  | BM389 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; F04 | Brucella ceti | 0.4 | 1 | 0
|  |  | BM403 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; K04 | Brucella ceti | 0.65 | 1 | 0
|  |  | BM374 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I02 | Brucella melitensis | 0.73 | 1 | 0
|  |  | BM407 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; J05 | Brucella melitensis | 0.78 | 1 | 0
|  |  | BM408 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; H06 | Brucella melitensis | 0.37 | 1 | 0
|  |  | BM410 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I06 | Brucella melitensis | 1.04 | 1 | 0
|  |  | CCM 4915 CVUAS | 3; type strain | DT | 2013-11-26 | Brucella microti | 0.4 | 1 | 0
|  |  | CCM 4915 RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 17.03.2015 C7 | Brucella microti | 0.74 | 1 | 0
|  |  | BM421 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; F07 | Brucella neotomae | 0.75 | 1 | 0
|  |  | A148-7 (5K33) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.03.2015; C15 | Brucella neotomae | 0.72 | 1 | 0
|  |  | BM423 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; L07 | Brucella ovis | 0.35 | 1 | 0
|  |  | A138-7 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.02.2015; F16 | Brucella ovis | 0.54 | 1 | 0
|  |  | BM405 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; E05 | Brucella pinnipedialis | 0.2 | 1 | 0
|  |  | A148-8 (M1771-94-1) | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.05.2016; C03 | Brucella pinnipedialis | 0.36 | 1 | 0
|  |  | BM431 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; B10 | Brucella suis | 0.41 | 1 | 0
|  |  | BM427 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.12.2010; G08 | Brucella suis | 0.55 | 1 | 0
|  |  | CVUAS 30916 | 3; confirmed NRL Brucella FLI | DT | 2015-06-22 | Brucella suis | 0 | 1 | 0
|  |  | BM428 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 22.12.2010; I08 | Brucella suis | 0.28 | 1 | 0
|  |  | BM430 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 15.12.2010; I09 | Brucella suis | 0.62 | 1 | 0
|  |  | DSM 101715 | 3; type strain | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 15.15.2016 I7 | Brucella vulpis | 0.41 | 1 | 0
|  |  | B 1315 | 3; type strain; sequenced | DT | 2013-10-14 | Falsochrobactrum ovis | 0.33 | 1 | 0
|  |  | DSM 20150 | 3; public strain collection | 80% TFA/1:10 verdünnt/gespottet 1:1 mit Matrix (10g/l HCCA in TA2) | 03.06.2016 H12 | Ochrobactrum anthropi | 0.5 | 1 | 0
|  |  | DSM 7051 | 3; type strain | DT | 2013-09-29 | Aminobacter aganoensis | 0.42 | 1 | 0
|  |  | DSM 7048 | 3; type strain | DT | 2013-09-25 | Aminobacter aminovorans | 0.14 | 1 | 0
|  |  | RRLK05_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.04.2015 H22 | Paracoccus yeei | 0.55 | 1 | 0
|  |  | MB 2022000098 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Roseomonas mucosa | 0.47 | 1 | 0
|  |  | DSM 16000 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HC | 24.01.2007, H7, [0034] | Inquilinus limosus | 0.84 | 1 | 0
|  |  | 161013111-1 | 2; 0 | DT | 2017-11-03 | Taylorella asinigenitalis | 0.51 | 1 | 0
|  |  | DSM 30026 | 3; type strain | DT | 2013-06-25 | Achromobacter denitrificans | 0.8 | 1 | 0
|  |  | DSM 30031 | 3; public strain collection | DT | 2013-06-25 | Achromobacter sp-DSM-30031 | 0.78 | 1 | 0
|  |  | Ax 247 06 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2008 E5 | Achromobacter xylosoxidans | 1.08 | 1 | 0
|  |  | HST 115 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.02.2016 N7 | Achromobacter xylosoxidans | 0.74 | 1 | 0
|  |  | CVUAS 31213 | 3; sequenced | EFEx | 2018-09-28 | Basilea sp-CVUAS-31213 | 0.39 | 1 | 0
|  |  | CVUAS 5087 | 3; sequenced | DT | 2013-04-03 | Bordetella avium | 1.01 | 1 | 0
|  |  | CVUAS 32946 | 3; confirmed proficency test | EFEx | 2021-02-18 | Kerstersia gyiorum | 0.32 | 1 | 0
|  |  | CVUAS 3975 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.45 | 1 | 0
|  |  | CVUAS 7536 | 3; sequenced | DT | 2013-07-09 | Oligella urethralis | 0.58 | 1 | 0
|  |  | DSM 13608 | 3; type strain | DT | 2013-06-25 | Pigmentiphaga kullae | 0.29 | 1 | 0
|  |  | 040210 RC 9257 | 2; LVU 2010 | DT | 2013-05-15 | Taylorella asinigenitalis | 0.42 | 1 | 0
|  |  | CVUAS 9211 | 3; Rep-E3 / PFGE TE-A1 | eDT | 2015-11-04 | Taylorella equigenitalis | 0.67 | 1 | 0
|  |  | CVUAS 9220 | 3; Rep-E2 / PFGE TE-A1 | DT | 2013-05-15 | Taylorella equigenitalis | 0.49 | 1 | 0
|  |  | A-269_01 A | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.04.2012; K10 | Burkholderia pseudomallei | 0.56 | 1 | 0
|  |  | A333 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; H21 | Burkholderia pseudomallei | 0.41 | 1 | 0
|  |  | A-269_07 A | 3; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 2.04.2012, K6 | Burkholderia thailandensis | 0.55 | 1 | 0
|  |  | Fq 6804 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007, K7 | Burkholderia ambifaria | 0.73 | 1 | 0
|  |  | DSM 16087 LMG 19182 | 3; type starin | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K19 | Burkholderia ambifaria | 0.33 | 1 | 0
|  |  | DSM 16086 LMG 20980 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007, K21 | Burkholderia anthina | 0.74 | 1 | 0
|  |  | LMG 20983 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 06.08.2007, N10 | Burkholderia anthina | 0.74 | 1 | 0
|  |  | DSM 23435 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.11.2010, K12 | Burkholderia arboris | 0.33 | 1 | 0
|  |  | LMG 24269 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 3.4.2017; L6 | Burkholderia arboris | 0.71 | 1 | 0
|  |  | DSM 16553 LMG 16656 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K11 | Burkholderia cenocepacia | 0.54 | 1 | 0
|  |  | Fq 6403 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 02.07.2009; K21 | Burkholderia cenocepacia | 0.56 | 1 | 0
|  |  | ATCC 25608 | 2; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 4.5.2006; H15 | Burkholderia cepacia | 0.55 | 1 | 0
|  |  | DSM 7288 ts | 2; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 9.5.2007; N9 | Burkholderia cepacia | 0.57 | 1 | 0
|  |  | 103 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.11.2006; H10 | Burkholderia contaminans | 0.71 | 1 | 0
|  |  | BC 01 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E10 | Burkholderia contaminans | 0.65 | 1 | 0
|  |  | DSM 23434 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016 C13 | Burkholderia diffusa | 0.51 | 1 | 0
|  |  | LMG 24065 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L17 | Burkholderia diffusa | 0.04 | 1 | 0
|  |  | DSM 16088 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K16 | Burkholderia dolosa | 0.03 | 1 | 0
|  |  | LMG 21820 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 17.11.2006; E09 | Burkholderia dolosa | 0.42 | 1 | 0
|  |  | ATCC 10248 DSM 4285 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.01.2007; H22 | Burkholderia gladioli | 0.92 | 1 | 0
|  |  | DSM 8361 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H15 | Burkholderia gladioli | 0.69 | 1 | 0
|  |  | LMG 1277 2 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K03 | Burkholderia glumae | 1.08 | 1 | 0
|  |  | LMG 2196 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.05.2012; K22 | Burkholderia glumae | 0.68 | 1 | 0
|  |  | DSM 23089 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 31.05.2016; C10 | Burkholderia lata | 0.85 | 1 | 0
|  |  | LMG 22485 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 23.08.2016; K19 | Burkholderia lata | 0.71 | 1 | 0
|  |  | DSM 23436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.05.2016; L06 | Burkholderia latens | 0.58 | 1 | 0
|  |  | ATCC 23344 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 04.09.2009; K20 | Burkholderia mallei | 0.42 | 1 | 0
|  |  | Bogor | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 24.08.2007; H13 | Burkholderia mallei | 1.02 | 1 | 0
|  |  | DSM 23519 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.12.2010; H14 | Burkholderia metallica | 0.52 | 1 | 0
|  |  | LMG 24275 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; L21 | Burkholderia metallica | 0.15 | 1 | 0
|  |  | ATCC 17616 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.12.2006; K19 | Burkholderia multivorans | 0.54 | 1 | 0
|  |  | HST 101 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 08.11.2006; K04 | Burkholderia multivorans | 0.65 | 1 | 0
|  |  | LMG 23620 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K17 | Burkholderia oklahomensis | 0.37 | 1 | 0
|  |  | LMG 23621 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K18 | Burkholderia oklahomensis | 0.58 | 1 | 0
|  |  | LMG 10907 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K06 | Burkholderia plantarii | 0.23 | 1 | 0
|  |  | LMG 9035 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K05 | Burkholderia plantarii | 0.15 | 1 | 0
|  |  | LMG 16669 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K07 | Burkholderia pseudomultivorans | 0.35 | 1 | 0
|  |  | LMG 26883 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 21.03.2017; K08 | Burkholderia pseudomultivorans | 0.27 | 1 | 0
|  |  | CAMPA 1043 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.04.2016; K04 | Burkholderia puraquae | 0.68 | 1 | 0
|  |  | R-50214 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.03.2016; C04 | Burkholderia puraquae | 0.27 | 1 | 0
|  |  | ATCC 15958 DSM 10685 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; N06 | Burkholderia pyrrocinia | 0.55 | 1 | 0
|  |  | LMG 21822 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.04.2017; I21 | Burkholderia pyrrocinia | 0.47 | 1 | 0
|  |  | DSM 23518 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H08 | Burkholderia seminalis | 0.61 | 1 | 0
|  |  | HST 20 | 3; [0102] | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 11.01.2007; K12 | Burkholderia seminalis | 0.48 | 1 | 0
|  |  | DSM 16586 LMG 14294 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.01.2007; K13 | Burkholderia stabilis | 0.6 | 1 | 0
|  |  | M 42 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.12.2006; E08 | Burkholderia stabilis | 0.54 | 1 | 0
|  |  | LMG 28156 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; C08 | Burkholderia stagnalis | 0.4 | 1 | 0
|  |  | LMG 28157 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F14 | Burkholderia stagnalis | 0.77 | 1 | 0
|  |  | BC 2813 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007; H19 | Burkholderia taxonK | 0.73 | 1 | 0
|  |  | Fq_6584 | 2; strain collection of the RKI | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.06.2009; K13 | Burkholderia taxonK | 0.31 | 1 | 0
|  |  | LMG 28158 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; F16 | Burkholderia territorii | 0.7 | 1 | 0
|  |  | LMG 28159 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 16.03.2017; E14 | Burkholderia territorii | 0.33 | 1 | 0
|  |  | DSM 13276 n RKI | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 24.11.2010, K19 | Burkholderia thailandensis | 0.83 | 1 | 0
|  |  | DSM 17311 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 12.11.2010; H11 | Burkholderia ubonensis | 1.08 | 1 | 0
|  |  | LMG 24263 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.03.2017; I04 | Burkholderia ubonensis | 0.51 | 1 | 0
|  |  | CCUG 9631=LMG 6999 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.05.2007 N15 | Burkholderia vietnamiensis | 0.69 | 1 | 0
|  |  | LMG 10929 V | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.12.2006 N9 | Burkholderia vietnamiensis | 0.52 | 1 | 0
|  |  | LMG 14190 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 13.04.2012, N07 | Caballeronia glathei | 0.54 | 1 | 0
|  |  | LMG 14932 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H12 | Caballeronia glathei | 0.74 | 1 | 0
|  |  | DSM 16535 LMG 16407 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H15 | Pandoraea apista | 0.47 | 1 | 0
|  |  | DSM 16536 LMG 18087 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H21 | Pandoraea pnomenusa | 0.53 | 1 | 0
|  |  | DSM 16583 LMG 18106 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 24.01.2007 H16 | Pandoraea pulmonicola | 0.16 | 1 | 0
|  |  | LMG 19076 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K7 | Paraburkholderia caledonica | 1.01 | 1 | 0
|  |  | LMG 19077 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 27.04.2012; K8 | Paraburkholderia caledonica | 0.6 | 1 | 0
|  |  | LMG 18531 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N9 | Paraburkholderia caribensis | 0.77 | 1 | 0
|  |  | LMG 18532 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 13.04.2012; N10 | Paraburkholderia caribensis | 0.56 | 1 | 0
|  |  | LMG 20574 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H03 | Paraburkholderia hospita | 0.6 | 1 | 0
|  |  | LMG 20586 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 20.04.2012; H04 | Paraburkholderia hospita | 0.49 | 1 | 0
|  |  | DSM 13646 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 1.12.2010; H15 | Paraburkholderia kururiensis | 0.78 | 1 | 0
|  |  | DSM 17167 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 26.11.2010; N19 | Paraburkholderia phymatum | 0.3 | 1 | 0
|  |  | DSM 17436 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 19.11.2010; K17 | Paraburkholderia phytofirmans | 0.65 | 1 | 0
|  |  | LMG 21444 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.04.2012, H16 | Paraburkholderia tuberum | 0.7 | 1 | 0
|  |  | DSM 17367 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 26.11.2010, N21 | Paraburkholderia xenovorans | 0.86 | 1 | 0
|  |  | DSM 17714 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 E07 | Ralstonia insidiosa | 0 | 1 | 0
|  |  | DSM 17512 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 13.05.2016 C06 | Ralstonia mannitolilytica | 0.45 | 1 | 0
|  |  | ATCC 27511 DSM 6297 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 22.01.2007 H03 | Ralstonia pickettii | 0.83 | 1 | 0
|  |  | Rp 127 | 2; 0 | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 07.05.2009 N15 | Ralstonia pickettii | 0.38 | 1 | 0
|  |  | MB 2022000097 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Delftia acidovorans | 0.24 | 1 | 0
|  |  | CVUAS 5683 | 3; sequenced | Efex | 2018-07-10 | Delftia tsuruhatensis | 0.62 | 1 | 0
|  |  | DSM 5680 | 3; type strain; sequenced | eDT | 2014-04-10 | Hydrogenophaga intermedia | 0.69 | 1 | 0
|  |  | A395 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 25.07.2013 L13 | Xenophilus sp | 0.68 | 1 | 0
|  |  | DSM 24713 EFEx | 3; type strain | Extr | 2016-02-02 | Deefgea chitinilytica | 0.49 | 1 | 0
|  |  | DSM 24713 eDT | 3; type strain | eDT | 2015-09-22 | Deefgea chitinilytica | 0.7 | 1 | 0
|  |  | DSM 18356 | 3; type strain | eDT | 2015-11-18 | Deefgea rivuli | 0.71 | 1 | 0
|  |  | 08-F17,2-2 | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.46 | 1 | 0
|  |  | 09-830-5-L | 3; sequenced | eDT | 2015-03-18 | Deefgea sp-CVUAS-2703,2 | 0.19 | 1 | 0
|  |  | CVUAS 6995,3 | 3; sequenced; 16S rDNA Genbank accession no.: KX889842 | EFEx | 2015-10-28 | Deefgea sp-CVUAS-6995,3 | 0.7 | 1 | 0
|  |  | CVUAS 863 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-6995,3 | 0.67 | 1 | 0
|  |  | 6884-11 | 3; sequenced | DT | 2012-12-14 | Deefgea sp-CVUAS-9920 | 0.53 | 1 | 0
|  |  | 9447-08 | 3; sequenced | EFEx | 2015-03-18 | Deefgea sp-CVUAS-9920 | 0.2 | 1 | 0
|  |  | CVUAS 32896 | 3; sequenced | EFEx | 2021-01-22 | Iodobacter limnosediminis | 0.58 | 1 | 0
|  |  | MB 2022000095 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Kingella kingae | 0.7 | 1 | 0
|  |  | CVUAS 10067,2 | 3; sequenced | EFEx | 2018-11-16 | Neisseria animaloris | 0.72 | 1 | 0
|  |  | CVUAS 8600 | 3; sequenced | DT | 2013-07-12 | Neisseria animaloris | 0.04 | 1 | 0
|  |  | CVUAS 4931 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-4931 | 0.56 | 1 | 0
|  |  | CVUAS 8997 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-CVUAS-8997 | 0.78 | 1 | 0
|  |  | LHL 1296 | 3; sequenced | DT | 2013-07-12 | Neisseria sp-LHL-1296 | 0.17 | 1 | 0
|  |  | CVUAS 5699,3 | 3; sequenced | EFEx | 2020-03-10 | Neisseria sp-N95-16 | 0.72 | 1 | 0
|  |  | CVUAS 8636 | 3; sequenced | DT | 2013-07-22 | Neisseriaceae sp-CVUAS-8636 | 0.37 | 1 | 0
|  |  | 171015787 | 3; isolate confirmed by FLI Campylobacter | DT | 2017-11-24 | Campylobacter fetus | 1.08 | 1 | 0
|  |  | DSM 18826 DT | 3; public strain collection | DT | 2016-10-13 | Campylobacter fetus | 0.32 | 1 | 0
|  |  | NCTC 10354 | 3; public strain collection | DT | 2017-06-21 | Campylobacter fetus | 0.46 | 1 | 0
|  |  | CVFR CAMP 20 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2015-10-29 | Campylobacter hyointestinalis | 0.33 | 1 | 0
|  |  | BfR-CA-07048 | 3; confirmed by NRL-BfR | eDT | 2019-12-04 | Campylobacter peloridis | 0.26 | 1 | 0
|  |  | BfR-CA-6566 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter peloridis | 1.04 | 1 | 0
|  |  | DSM 7289 | 3; type strain | DT | 2013-04-03 | Aliarcobacter cryaerophilus | 0.35 | 1 | 0
|  |  | CVUAS 1272 | 3; specific PCR | EFEx | 2015-03-18 | Arcobacter butzleri | 0.45 | 1 | 0
|  |  | DSM 8739 | 3; type strain | EFEx | 2017-05-10 | Arcobacter butzleri | 0.42 | 1 | 0
|  |  | CVUAS 9949 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | DT | 2013-01-22 | Campylobacter coli | 0.6 | 1 | 0
|  |  | DSM 4689 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter coli | 0.58 | 1 | 0
|  |  | CIP 103751 | 3; type strain | DT | 2013-03-20 | Campylobacter jejuni | 0.24 | 1 | 0
|  |  | ATCC 49943 | 3; public strain collection | eDT | 2015-12-30 | Campylobacter jejuni | 0.69 | 1 | 0
|  |  | CVUAS 32778 | 3; confirmed NRL Campylobacter BfR | DT | 2020-11-18 | Campylobacter jejuni | 0.9 | 1 | 0
|  |  | BfR-CA-6535 | 3; confirmed by NRL-BfR | DT | 2013-07-31 | Campylobacter lari | 0.76 | 1 | 0
|  |  | BfR-CA-06556 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.81 | 1 | 0
|  |  | BfR-CA-17213 | 3; confirmed by NRL-BfR | DT | 2019-11-27 | Campylobacter lari | 0.43 | 1 | 0
|  |  | DSM 11375 CVUAS | 3; type strain | eDT | 2015-12-30 | Campylobacter lari | 0.39 | 1 | 0
|  |  | DSM 5361 CVUAS | 3; type strain | DT | 2013-09-24 | Campylobacter fetus | 1.1 | 1 | 0
|  |  | CVUAS 6673,2 | 3; sequenced | DT | 2013-08-28 | Campylobacter hyointestinalis | 0.49 | 1 | 0
|  |  | CVUAS 2868 | 3; isolate confirmed by NRL Campylobacter BfR Berlin | eDT | 2013-05-15 | Campylobacter sputorum | 0.33 | 1 | 0
|  |  | CVUAS 31484 | 3; sequenced | EFEx | 2019-01-16 | Campylobacter sp-W1 | 0.89 | 1 | 0
|  |  | CIP 103681 | 3; type strain | eDT | 2013-05-15 | Campylobacter upsaliensis | 0 | 1 | 0
|  |  | CVUAS 11185,3 | 3; sequenced | EFEx | 2019-09-06 | Helicobacter canadensis | 0 | 1 | 0
|  |  | DSM 13956 | 3; type strain | EFEx | 2015-10-14 | Aeromonas bestiarum | 0.5 | 1 | 0
|  |  | DSM 19111 | 3; type strain | eDT | 2015-07-08 | Aeromonas bivalvium | 0.55 | 1 | 0
|  |  | ATCC 15468 | 3; type strain, WGS sequenced | eDT | 2016-01-13 | Aeromonas caviae | 0.93 | 1 | 0
|  |  | DSM 17689 | 3; public strain collection, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.76 | 1 | 0
|  |  | DSM 18362 | 3; type strain, biochemistry | eDT | 2015-04-15 | Aeromonas dhakensis | 0.7 | 1 | 0
|  |  | DSM 30187 CVUAS | 4; type strain, WGS sequenced (BfR) | EFEx | 2015-08-07 | Aeromonas hydrophila | 0.83 | 1 | 0
|  |  | BSK 10 | 3; sequenced | eDT | 2015-04-15 | Aeromonas hydrophila | 0.41 | 1 | 0
|  |  | DSM 7311 | 3; type strain | eDT | 2017-07-08 | Aeromonas jandaei | 0.56 | 1 | 0
|  |  | DSM 4881 | 3; type strain | EFEx | 2015-04-22 | Aeromonas media | 0.93 | 1 | 0
|  |  | CVUAS 10488 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-03-31 | Aeromonas molluscorum | 0.4 | 1 | 0
|  |  | DSM 19604 | 3; type strain | EFEx | 2015-04-22 | Aeromonas popoffii | 0.29 | 1 | 0
|  |  | ASA 6 | 3; gyrB-sequencing, biochemistry | EFEx | 2015-02-04 | Aeromonas salmonicida | 0.98 | 1 | 0
|  |  | DSM 21281 | 3; type strain | eDT | 01.07.2015 | Aeromonas salmonicida | 0.85 | 1 | 0
|  |  | DSM 21760 | 3; reference strain collection | eDT | 2015-06-15 | Aeromonas salmonicida | 0.59 | 1 | 0
|  |  | DSM 12609 | 3; type strain | eDT | 2015-07-01 | Aeromonas salmonicida | 0.79 | 1 | 0
|  |  | CVUAS 8400,2 | 3; gyrB-sequence; biochemistry | DT | 2014-11-04 | Aeromonas salmonicida | 0.61 | 1 | 0
|  |  | DSM 19634 | 3; type strain | eDT | 2015-06-15 | Aeromonas salmonicida | 0.69 | 1 | 0
|  |  | DSM 21293 | 3; type strain | eDT | 19.07.2016 | Aeromonas salmonicida | 0.36 | 1 | 0
|  |  | UniBern 4097 | 3; intensiv untersucht; ascV+,aexT+,aopP+,aopH+. 16S / gyrB bei Einsender | eDT | 2016-03-30 | Aeromonas salmonicida | 0.69 | 1 | 0
|  |  | DSM 16559 | 3; type strain | eDT | 2017-07-08 | Aeromonas simiae | 0.61 | 1 | 0
|  |  | CVUAS 8996 | 3; 16S, gyrB; biochemistry [0045] | DT | 2013-04-25 | Aeromonas veronii | 1.17 | 1 | 0
|  |  | DSM 7386 CVUAS | 3; type strain, biochemistry | EFEx | 2015-04-29 | Aeromonas veronii | 0.61 | 1 | 0
|  |  | WA-1227 | 3; sequenced | EFEx | 2020-03-11 | Pseudaeromonas pectinilytica | 0.48 | 1 | 0
|  |  | F6-15 | 3; well defined isolat | EFEx | 2015-05-11 | Moritella viscosa | 0.44 | 1 | 0
|  |  | K58 | 3; [0093] | EFEx | 2015-06-24 | Moritella viscosa | 0.24 | 1 | 0
|  |  | CVUAS 8287 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.88 | 1 | 0
|  |  | CVUAS 9366 | 3; sequenced | DT | 2013-10-15 | Shewanella sp-CVUAS-9366 | 0.38 | 1 | 0
|  |  | CVUAS 81,2 | 3; sequenced | EFEx | 28:11.2013 | Shewanella sp-RKVR22 | 0.54 | 1 | 0
|  |  | DSM 27966 | 3; type strain | DT | 2019-03-05 | Cronobacter condimenti | 0.58 | 1 | 0
|  |  | DSM 18707 | 3; type strain | DT | 2019-03-05 | Cronobacter dublinensis | 0.13 | 1 | 0
|  |  | DSM 18702 LALLF | 3; type strain | DT | 2019-03-05 | Cronobacter malonaticus | 0.26 | 1 | 0
|  |  | DSM 18703 | 3; type strain | DT | 2019-03-05 | Cronobacter turicensis | 0.8 | 1 | 0
|  |  | E15 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 10.11.2011 E6 | Enterobacter cloacae | 0.46 | 1 | 0
|  |  | 04 08695 | 2; 0 | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, H5 | Citrobacter amalonaticus | 0.28 | 1 | 0
|  |  | ATCC 8090 | 3; type strain | DT | 2013-09-17 | Citrobacter freundii | 0.11 | 1 | 0
|  |  | CVUAS 8481 | 3; sequenced | DT | 2013-09-17 | Citrobacter freundii | 0.14 | 1 | 0
|  |  | ATCC 25408 | 3; public strain collection | 80% TFA, 30 Minuten, 1:10 verdünnt, gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 11.04.2008, N11, [0034]: Citrobacter diversus ATCC 25408\LMTG70N11_0\0_N11\1\1SLin\fid | Citrobacter koseri | 0.67 | 1 | 0
|  |  | DSM 18702 LVU 15 | 3; type strain | EFex | 2015-04-14 | Cronobacter malonaticus | 0.55 | 1 | 0
|  |  | CVUAS 32945 | 3; confirmed proficency test | EFEx | 2021-02-18 | Cronobacter sakazakii | 0.42 | 1 | 0
|  |  | DSM 4485 | 3; reference strain collection | DT | 2015-07-09 | Cronobacter sakazakii | 0.68 | 1 | 0
|  |  | DSM 30054 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E18, [0034] | Enterobacter cloacae | 0.79 | 1 | 0
|  |  | ATCC 27988 | 3; public strain collection; NCBI sequence | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K4, [0034] | Enterobacter sp-ATCC-27988 | 0.68 | 1 | 0
|  |  | CVUAS 30072 | 3; sequenced | EFex | 2017-04-25 | Enterobacteriaceae sp-CVUAS-30072 | 0.64 | 1 | 0
|  |  | CVUAS 0500,2 | 3; NRL BfR | DT | 2015-05-11 | Escherichia coli | 0.86 | 1 | 0
|  |  | CVUAS 11332 | 3; sequenced [BfR] | eDT | 2016-10-18 07:22 | Klebsiella grimontii | 0.48 | 1 | 0
|  |  | CVUAS 2486,3 | 3; sequenced [BfR] | eDT | 2017-07-19 11:27 | Klebsiella grimontii | 0.57 | 1 | 0
|  |  | CVUAS 11699 | 3; sequenced [BfR] | eDT | 2017-04-20 06:56 | Klebsiella michiganensis | 0.64 | 1 | 0
|  |  | CVUAS 2404,3 | 3; sequenced [BfR] | EFex | 2017-07-19 11:37 | Klebsiella michiganensis | 0.58 | 1 | 0
|  |  | CVUAS 10038,2 | 3; sequenced [BfR] | EFex | 2019-02-12 14:02 | Klebsiella oxytoca | 0.6 | 1 | 0
|  |  | CVUAS 11572 | 3; sequenced [BfR] | eDT | 2017-04-20 06:51 | Klebsiella oxytoca | 0.35 | 1 | 0
|  |  | CVUAS 1482,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:35 | Klebsiella pneumoniae | 0.03 | 1 | 0
|  |  | CVUAS 1616,4 | 3; sequenced [BfR] | eDT | 2017-11-22 09:55 | Klebsiella pneumoniae | 0.44 | 1 | 0
|  |  | CVUAS 1678 | 3; sequenced [BfR] | eDT | 2016-10-18 06:38 | Klebsiella pneumoniae | 0.51 | 1 | 0
|  |  | CVUAS 1701 | 3; sequenced [BfR] | eDT | 2016-10-18 06:41 | Klebsiella pneumoniae | 0.42 | 1 | 0
|  |  | CVUAS 30499 | 3; sequenced [BfR] | eDT | 2017-11-22 09:59 | Klebsiella pneumoniae | 0.3 | 1 | 0
|  |  | CVUAS 753,2 | 3; sequenced [BfR] | eDT | 2017-04-20 06:34 | Klebsiella pneumoniae | 0.18 | 1 | 0
|  |  | CVUAS 10109 | 3; sequenced [BfR] | eDT | 2016-10-18 07:00 | Klebsiella quasipneumoniae | 0.03 | 1 | 0
|  |  | CVUAS 11677 | 3; sequenced [BfR] | eDT | 2017-04-20 06:54 | Klebsiella variicola | 0.24 | 1 | 0
|  |  | CVUAS 4888,2 | 3; sequenced [BfR] | EFex | 2017-07-19 11:36 | Klebsiella variicola | 0.31 | 1 | 0
|  |  | DSM 30053 CVUAS | 3; type strain | EFEx | 2019-02-12 | Klebsiella aerogenes | 0.43 | 1 | 0
|  |  | NM 20 | 3; strain collection of RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008, E20 | Klebsiella aerogenes | 0.47 | 1 | 0
|  |  | DSM 25444 | 3; type strain | EFEx | 2017-05-16 | Klebsiella michiganensis | 0.72 | 1 | 0
|  |  | ATCC 13182 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 01.11.2013 E18 | Klebsiella oxytoca | 0.74 | 1 | 0
|  |  | ATCC 25926 | 3; public strain collection | 80% TFA/30 min Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 | 11.04.2008 N21 | Klebsiella pneumoniae | 0 | 1 | 0
|  |  | DSM 30104 | 3; type strain | EFEx | 2013-09-17 | Klebsiella pneumoniae | 0.07 | 1 | 0
|  |  | ID 28 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.07 | 1 | 0
|  |  | ID 9 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella quasipneumoniae | 0.62 | 1 | 0
|  |  | ID 49 | 3; sequenced | EFEx | 2018-09-18 | Klebsiella variicola | 0.31 | 1 | 0
|  |  | CVUAS 30071 | 3; sequenced | EFex [BRU03] | 2017-04-20 | Kosakonia sacchari | 0.54 | 1 | 0
|  |  | CVUAS 5011 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.97 | 1 | 0
|  |  | CVUAS 7882 | 3; sequenced | EFEx | 2017-01-25 | Plesiomonas shigelloides | 0.78 | 1 | 0
|  |  | ATCC 33426 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N17, [0034] (Enterobacter gergoviae) | Pluralibacter gergoviae | 1.07 | 1 | 0
|  |  | LT21 Ib 08 01796 | 3; well described strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H9 | Salmonella enterica | 0.17 | 1 | 0
|  |  | II 06 06820 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H6 | Shigella boydii | 0.7 | 1 | 0
|  |  | 3A 07 02503 | 2; strain collection of the RKI | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 09.06.2008 H7 | Shigella flexneri | 0.86 | 1 | 0
|  |  | DSM 3493 IGV | 3; type strain | DT | 2016-02-17 | Pantoea agglomerans | 0.59 | 1 | 0
|  |  | DSM 3493 CVUAS | 3; type strain | eDT | 2015-07-09 | Pantoea agglomerans | 0.33 | 1 | 0
|  |  | 11MRJ 3 6 1 | 2; sequenced | DT | 2013-03-20 | Pantoea sp | 0.67 | 1 | 0
|  |  | CVUAS 4263,2 | 3; sequenced | EFEx | 2013-05-16 | Pantoea sp-FF5 | 0.54 | 1 | 0
|  |  | CVUAS 11691 | 2; sequenced | DT | 2016-12-30 | Edwardsiella tarda | 0.45 | 1 | 0
|  |  | DSM 30052 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, N13, [0034] | Edwardsiella tarda | 0.45 | 1 | 0
|  |  | DSM 30163 | 3; type strain | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 18.04.2008, K9, [0034] | Hafnia alvei | 0.01 | 1 | 0
|  |  | LHL1270 | 3; sequenced (LHL) | DT | 2013-01-30 | Hafnia alvei | 0.93 | 1 | 0
|  |  | LHL 3574 | 3; sequenced | EFEx | 2017-12-28 | Cosenzaea myxofaciens | 0.65 | 1 | 0
|  |  | CVUAS 9555 | 3; sequenced | DT | 2013-09-11 | Morganella morganii | 0.34 | 1 | 0
|  |  | LHLL 1818b | 3; sequenced | EFEx | 2015-03-31 | Morganella morganii | 0.41 | 1 | 0
|  |  | DSM 100173 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 07:15 | Proteus cibarius | 0.38 | 1 | 0
|  |  | DSM 14437 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus hauseri | 0.07 | 1 | 0
|  |  | LHL 2003 | 3; sequenced | EFEx | 2017-12-28 | Proteus hauseri | 0.35 | 1 | 0
|  |  | CVUAS 30465 | 3; sequenced (rpoB) | eDT | 2017-11-14 | Proteus mirabilis | 0.39 | 1 | 0
|  |  | DSM 4479 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus mirabilis | 0.38 | 1 | 0
|  |  | DSM 4544 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus penneri | 0.3 | 1 | 0
|  |  | CVUAS 30222 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.78 | 1 | 0
|  |  | LHL 527 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.81 | 1 | 0
|  |  | LHL 554 | 2; sequenced | EFEx | 2017-12-28 | Proteus sp | 0.88 | 1 | 0
|  |  | DSM 29910 | 3; typestrain, sequenced (rpoB) | eDT | 2017-10-05 | Proteus terrae | 0.29 | 1 | 0
|  |  | ATCC 33420 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, E5 K16, [0034]: Proteus vulgaris ATCC 33420\LMTG72K16_0\0_K16\1\1SLin\fid | Proteus vulgaris | 0.88 | 1 | 0
|  |  | DSM 2140 | 3; strain from public collection; sequenced (rpoB) | eDT | 30.11.2017 | Proteus vulgaris | 0.8 | 1 | 0
|  |  | CVUAS 5796 | 2; sequenced | DT | 2012-11-13 | Providencia rettgeri | 0.56 | 1 | 0
|  |  | LHL 390 | 3; sequenced | EFEx | 2017-12-28 | Providencia rettgeri | 1.08 | 1 | 0
|  |  | ATCC 25827 | 3; public strain collection | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 11.04.2008, K15, [0034]: Proteus stuartii ATCC 25827\LMTG72K15_0\0_K15\1\1SLin\fid in vers 2016 | Providencia stuartii | 0.34 | 1 | 0
|  |  | 171014911 | 2; 0 | DT | 2017-08-11 | Yersinia frederiksenii | 0.54 | 1 | 0
|  |  | 161002906 | 2; 0 | DT | 2016-03-14 | Yersinia mollaretii | 0.4 | 1 | 0
|  |  | NCTC 2028 | 3; public strain collection | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H13 | Yersinia pestis | 0.35 | 1 | 0
|  |  | NCTC 5923 | 3; type strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 27.11.2009 H14 | Yersinia pestis | 0.1 | 1 | 0
|  |  | 151012363 | 2; 0 | DT | 2015-09-21 | Yersinia wautersii | 0.44 | 1 | 0
|  |  | CVUAS 3171 | 3; sequenced, S. glossinae is later synonym of S. fonticola [0050] | EFEx | 2017-01-25 | Serratia fonticola | 0.47 | 1 | 0
|  |  | DSM 30063 | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 04.11.2008, E11, [0034] | Serratia grimesii | 0.08 | 1 | 0
|  |  | 25_12_13996_1 LVU09 | 3; sequenced (LGL) | EFex | 2015-04-15 | Serratia marcescens | 0.27 | 1 | 0
|  |  | MB 2022000096 | 2; corresponds to expected result of PT | EFex [BRU03] | 2022-09-09 | Serratia marcescens | 0.57 | 1 | 0
|  |  | DSM 4597 | 3; type strain | EFEx | 2017-01-25 | Serratia quinivorans | 0.59 | 1 | 0
|  |  | CVUAS 6221 | 4; sequenced | DT | 2013-01-08 | Yersinia aldovae | 0 | 1 | 0
|  |  | DSM 14987 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 09.03.2007, H12 | Yersinia aleksiciae | 0.43 | 1 | 0
|  |  | BAG 139 K2 | 2; [0054] | DT | 2013-01-09 | Yersinia bercovieri | 0.46 | 1 | 0
|  |  | SL2011 003926 | 2; 0 | DT | 2013-01-30 | Yersinia bercovieri | 0.07 | 1 | 0
|  |  | SL2012 001153 | 3; sequenced | DT | 2013-01-30 | Yersinia bercovieri | 0.04 | 1 | 0
|  |  | DSM 4780 CVUAS | 3; type strain | DT | 2013-05-08 | Yersinia enterocolitica | 0.16 | 1 | 0
|  |  | Zp 37-1 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.33 | 1 | 0
|  |  | SZ 5108-01 | 3; [0054] | DT | 2013-05-14 | Yersinia enterocolitica | 0.35 | 1 | 0
|  |  | BAG 022 | 3; 0 | DT | 2012-12-11 | Yersinia enterocolitica | 0.17 | 1 | 0
|  |  | BAG 004 | 3; FT-IR; serotype | DT | 2012-12-05 | Yersinia enterocolitica | 0.27 | 1 | 0
|  |  | CVUAS 9295 | 3; Biotype / Serotype / FT-IR according [0054] | DT | 2012-12-11 | Yersinia enterocolitica | 0.38 | 1 | 0
|  |  | DSM 22339 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.ma_01\0_C13, [0034] | Yersinia entomophaga | 0.23 | 1 | 0
|  |  | 21 RKI | 3; well described strain | 80% TFA/ 30 min Einwirkzeit /1:10 verdünnt/1:1 HCCA:TA2 | 22.01.2007 H06 | Yersinia frederiksenii | 0.22 | 1 | 0
|  |  | 21 CVUAS | 3; well described strain | DT | 2013-01-30 | Yersinia frederiksenii | 0.12 | 1 | 0
|  |  | CVUAS 1824 | 2; sequenced | DT | 2012-12-05 | Yersinia intermedia | 0.41 | 1 | 0
|  |  | RKI 29211 | 2; API20E: 015453377 - UP Yers fr/in | DT | 2013-01-30 | Yersinia intermedia | 0.12 | 1 | 0
|  |  | CVFR 008 | 2; API20E: 115450367 - Yer kr 95,2% | DT | 2013-01-30 | Yersinia kristensenii | 0 | 1 | 0
|  |  | CVUAS 302 | 2; API20E: 1144723 - API 95,9% | DT | 2013-01-30 | Yersinia kristensenii | 0.06 | 1 | 0
|  |  | DSM 21859 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 25.11.2015, Y.ma_01\0_C11, [0034] | Yersinia massiliensis | 0.16 | 1 | 0
|  |  | VB 925699 | 3; sequenced | DT | 2013-03-19 | Yersinia massiliensis | 0.21 | 1 | 0
|  |  | DSM 18520 CVUAS | 3; type strain | DT | 16.01.2013 | Yersinia mollaretii | 0.4 | 1 | 0
|  |  | RKI 51 | 2; 0 | DT | 2013-01-16 | Yersinia mollaretii | 0.55 | 1 | 0
|  |  | DSM 22296 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.nu_04\0_E10, [0034] | Yersinia nurmii | 0.15 | 1 | 0
|  |  | DSM 22769 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.pe_04\0_E8, [0034] | Yersinia pekkanenii | 0.44 | 1 | 0
|  |  | CVUAS 0304,2 DT | 2; biochemistry | DT | 2013-01-31 | Yersinia pseudotuberculosis | 0.52 | 1 | 0
|  |  | DSM 8992 RKI | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.04.2009, E12 | Yersinia pseudotuberculosis | 0.4 | 1 | 0
|  |  | BAG 162 | 2; [0054] | DT | 2013-01-16 | Yersinia rohdei | 0.38 | 1 | 0
|  |  | DSM 18270 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 02.03.2007, E07 | Yersinia rohdei | 0.51 | 1 | 0
|  |  | CVUAS 2772 CVUAS | 3; sequenced | DT | 2013-01-30 | Yersinia ruckeri | 0.42 | 1 | 0
|  |  | EX5 | 3; well defined isolate [0053] | DT | 2013-01-30 | Yersinia ruckeri | 0.29 | 1 | 0
|  |  | DSM 18211 EFEx | 3; type strain | EFEx | 2013-02-19 | Yersinia similis | 0.63 | 1 | 0
|  |  | IMB 4354 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 13.07.2007, E22 | Yersinia similis | 0.24 | 1 | 0
|  |  | DSM 27350 | 3; type strain | 80% TFA/1:10 verdünnt/ 1:1 HCCA in TA2 [0011] | 03.12.2015, Y.wa_04\0_E5, [0034] | Yersinia wautersii | 0.02 | 1 | 0
|  |  | CVUAS 2311 | 3; sequenced | EFEx | 2014-12-15 | Legionella londiniensis | 0 | 1 | 0
|  |  | DSM 7515 | 3; type strain | Efex | 2013-08-27 | Legionella pneumophila | 0.25 | 1 | 0
|  |  | CVUAS 4122 | 3; 0 | eDT | 20.01.2016 | Legionella pneumophila | 0 | 1 | 0
|  |  | CVUAS 10250 | 3; sequenced | Efex | 2014-02-19 | Legionella taurinensis | 0.38 | 1 | 0
|  |  | CVUAS 10873,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0 | 1 | 0
|  |  | CVUAS 10874,2 | 3; sequenced | eDT | 2018-03-12 | Halomonas stevensii | 0.3 | 1 | 0
|  |  | 13URF23 | 3; sequenced | eDT | 2015-12-01 | (Pasteurella) caballi | 0.38 | 1 | 0
|  |  | DSM 22973 | 3; type strain | EFEx | 2020-10-20 | (Pasteurella) caballi | 0.8 | 1 | 0
|  |  | CVUAS 3006 | 3; sequenced | eDT | 2015-10-14 | (Pasteurella) testudinis | 0.36 | 1 | 0
|  |  | LHL 957 | 3; sequenced | eDT | 02.12.2015 | (Pasteurella) testudinis | 0.5 | 1 | 0
|  |  | 14UGF209 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus anseriformium | 0.84 | 1 | 0
|  |  | 15UGF 198 | 2; 0 | eDT | 2015-12-09 | Actinobacillus anseriformium | 0.96 | 1 | 0
|  |  | CVUAS 8124 | 3; sequenced | EFEx | 2015-10-14 | Actinobacillus arthritidis | 0.47 | 1 | 0
|  |  | 08UZF67 | 3; sequenced | eDT | 2015-12-01 | Actinobacillus capsulatus | 0.2 | 1 | 0
|  |  | CVUAS 32356 | 3; sequenced | eDT | 2020-06-05 | Actinobacillus capsulatus | 0.66 | 1 | 0
|  |  | CVUAS 352 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.86 | 1 | 0
|  |  | CVUAS 7824 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.65 | 1 | 0
|  |  | CVUAS 8470 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus equuli | 0.65 | 1 | 0
|  |  | S365 | 3; sequenced | EFEx | 2018-09-11 | Actinobacillus equuli | 0.61 | 1 | 0
|  |  | CVUAS 34202 | 3; sequenced | eDT | 2022-08-16 | Actinobacillus lignieresii | 0.63 | 1 | 0
|  |  | CVUAS 9945,2 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus lignieresii | 0.36 | 1 | 0
|  |  | CVUAS 639 | 3; sequenced | EFEx | 2018-03-01 | Actinobacillus rossii | 0.34 | 1 | 0
|  |  | DSM 17479 | 3; type strain | EFEx | 2020-10-20 | Actinobacillus rossii | 0.28 | 1 | 0
|  |  | CVUAS 32505,2 | 3; sequenced | EFEx | 2020-11-19 | Actinobacillus sp-CVUAS-32505,2 | 0.1 | 1 | 0
|  |  | 15UCF37 | 3; sequenced | eDT | 2015-12-15 | Actinobacillus sp-genomospecies-1 | 0.72 | 1 | 0
|  |  | CVUAS 8964 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-genomospecies-1 | 0.79 | 1 | 0
|  |  | 14UCF177 | 3; sequenced | EFEx | 2015-12-09 | Actinobacillus sp-MK-2012 | 0.86 | 1 | 0
|  |  | CVUAS 9304 | 3; sequenced | eDT | 2015-10-14 | Actinobacillus sp-MK-2012 | 0.92 | 1 | 0
|  |  | CVUAS 5527 | 3; sequenced | eDT | 2018-09-27 | Actinobacillus suis | 0.3 | 1 | 0
|  |  | CVUAS 5767 | 3; sequenced | eDT | 2015-04-29 | Actinobacillus suis | 0.54 | 1 | 0
|  |  | CVUAS 4184 | 3; sequenced | eDT | 2015-10-14 | Avibacterium gallinarum | 0.3 | 1 | 0
|  |  | LHL 2230 | 3; sequenced | EFEx | 2015-12-09 | Avibacterium gallinarum | 0.39 | 1 | 0
|  |  | CVUAS 537 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 0.68 | 1 | 0
|  |  | CVUAS 7850 | 3; sequenced | DT | 2013-01-30 | Bibersteinia trehalosi | 0.93 | 1 | 0
|  |  | CVUAS 3005 | 3; sequenced | EFEx | 2015-10-14 | Chelonobacter oris | 0.39 | 1 | 0
|  |  | CVUAS 8268 | 3; sequenced | DT | 2015-09-02 | Chelonobacter oris | 0.72 | 1 | 0
|  |  | CVUAS 3802,2 | 3; sequenced | eDT | 2016-03-15 | Cricetibacter osteomyelitidis | 0.85 | 1 | 0
|  |  | 15UCF12 | 3; sequenced | eDT | 2015-12-02 | Frederiksenia canicola | 0.68 | 1 | 0
|  |  | CVUAS 31989 | 3; sequenced | eDT | 2019-10-16 | Frederiksenia canicola | 0.71 | 1 | 0
|  |  | 15URF78 | 3; sequenced | eDT | 2015-12-02 | Gallibacterium anatis | 0.45 | 1 | 0
|  |  | CVUAS 6632 | 3; sequenced | eDT | 2013-07-09 | Gallibacterium anatis | 0.04 | 1 | 0
|  |  | CVUAS 11663 | 3; sequenced | EFEx | 2017-01-18 | Gallibacterium sp-F151 | 0.33 | 1 | 0
|  |  | CVUAS 11383 | 3; sequenced | eDT | 2016-11-03 | Gallibacterium trehalosifermentans | 0.26 | 1 | 0
|  |  | CVUAS 8968 | 3; sequenced | DT | 2014-04-08 | Haemophilus haemoglobinophilus | 0.25 | 1 | 0
|  |  | 2269 | 2; sequenced | eDT | 2013-07-17 | Haemophilus sp-2269 | 0.52 | 1 | 0
|  |  | 2609 | 2; sequenced | EFEx | 2013-07-17 | Haemophilus sp-2269 | 0.53 | 1 | 0
|  |  | P1066-2016 | 3; sequenced | EFEx | 2017-01-02 | Histophilus somni | 0.79 | 1 | 0
|  |  | CVUAS 10798 | 3; sequenced | eDT | 2015-11-18 | Mannheimia caviae | 0.72 | 1 | 0
|  |  | CVUAS 1064,4 | 3; sequenced | eDT | 2020-01-29 | Mannheimia glucosida | 0.19 | 1 | 0
|  |  | CVUAS 5568 | 3; sequenced; ESQ pos | DT | 2013-07-09 | Mannheimia glucosida | 0.63 | 1 | 0
|  |  | 10-12015-05612 | 3; sequenced | eDT | 2015-12-02 | Mannheimia granulomatis | 0.44 | 1 | 0
|  |  | CVUAS 6778 | 3; sequenced | EFex | 2021-05-21 | Mannheimia granulomatis | 0.55 | 1 | 0
|  |  | CVUAS 571 | 3; sequenced; ESQ neg | DT | 2013-07-09 | Mannheimia haemolytica | 0.5 | 1 | 0
|  |  | CVUAS 6197,2 | 2; ESQ neg | DT | 2013-09-13 | Mannheimia haemolytica | 0.52 | 1 | 0
|  |  | CVUAS 32053 | 3; sequenced | EFEx | 2019-11-15 | Mannheimia pernigra | 0.44 | 1 | 0
|  |  | CVUAS 32999 | 3; sequenced | EFEx | 2022-03-17 | Mannheimia pernigra | 0.4 | 1 | 0
|  |  | LLBB 14UOF72-1 | 3; sequenced | eDT | 2016-04-21 | Mannheimia ruminalis | 0.44 | 1 | 0
|  |  | 14UOF84 | 3; sequenced (LLBB) | EFEx | 2015-12-09 | Mannheimia ruminalis | 0.74 | 1 | 0
|  |  | CVUAS 10396,2 | 3; sequenced | eDT | 2020-03-24 | Mannheimia sp-10-12559-1 | 0.89 | 1 | 0
|  |  | CVUAS 10183,3 | 3; sequenced | eDT | 2020-03-13 | Mannheimia sp-14UCF195 | 0.42 | 1 | 0
|  |  | LLBB 14UCF195 | 3; sequenced | EFEx | 2015-12-09 | Mannheimia sp-14UCF195 | 0.76 | 1 | 0
|  |  | CVUAS 32092 | 3; sequenced | eDT | 2019-12-13 | Mannheimia sp-15USF15 | 0.67 | 1 | 0
|  |  | LLBB 15USF15 | 3; sequenced (LLBB) | Efex | 2017-02-07 | Mannheimia sp-15USF15 | 0.76 | 1 | 0
|  |  | CVUAS 11240 | 3; sequenced | eDT | 2016-09-16 | Mannheimia sp-CVUAS-11240 | 0.87 | 1 | 0
|  |  | CVUAS 1496 | 3; sequenced | eDT | 2013-07-09 | Mannheimia sp-CVUAS-1496 | 0.66 | 1 | 0
|  |  | CVUAS 1915 | 3; sequenced | DT | 2015-03-19 | Necropsobacter rosorum | 0.76 | 1 | 0
|  |  | CVUAS 6495 | 3; sequenced | eDT | 2015-10-14 | Necropsobacter rosorum | 0.54 | 1 | 0
|  |  | CVUAS 32124 | 3; sequenced | EFEx | 2020-01-15 | Nicoletella semolina | 0.7 | 1 | 0
|  |  | CVUAS 32185 | 3; sequenced | EFEx | 2020-01-30 | Nicoletella semolina | 0.58 | 1 | 0
|  |  | CVUAS 4060,3 | 3; sequenced | EFEx | 2019-10-01 | Pasteurella aerogenes | 0.55 | 1 | 0
|  |  | CVUAS 7933 | 3; sequenced | eDT | 2015-10-14 | Pasteurella canis | 0.43 | 1 | 0
|  |  | CVUAS 5284 | 3; sequenced | eDT | 2015-10-14 | Pasteurella mairii | 0 | 1 | 0
|  |  | ATCC 15742 | 3; public strain collection | eDT | 2018-03-22 | Pasteurella multocida | 0.49 | 1 | 0
|  |  | CVUAS 32055 | 3; specific PCR | eDT | 2019-11-14 | Pasteurella multocida | 0.28 | 1 | 0
|  |  | ATCC 12948 | 3; specific PCR | DT | 2013-04-03 | Pasteurella multocida | 0 | 1 | 0
|  |  | Bunia II | 3; REP profile id. CVUAS 7409; specific PCR (kmt), cap E, toxinA negative | eDT | 2016-03-01 | Pasteurella multocida | 0.21 | 1 | 0
|  |  | CVUAS 1556 | 3; specific PCR cap F | DT | 2016-03-01 | Pasteurella multocida | 0.12 | 1 | 0
|  |  | P 1997 | 4; specific PCR | eDT | 2018-03-23 | Pasteurella multocida | 0.17 | 1 | 0
|  |  | CVUAS 32506,2 | 3; sequenced | EFEx | 2020-11-19 | Pasteurellaceae Bisgaard-Taxon-10 | 1.02 | 1 | 0
|  |  | CVUAS 34050 | 3; sequenced | eDT | 2022-07-14 | Pasteurellaceae Bisgaard-Taxon-10 | 1.07 | 1 | 0
|  |  | LHL 1427 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae Bisgaard-Taxon-40 | 0.9 | 1 | 0
|  |  | CVUAS 9987,2 | 3; sequenced | eDT | 2015-07-15 | Pasteurellaceae canine-oral-taxon-080 | 0.33 | 1 | 0
|  |  | CVUAS 261,4 | 3; sequenced | EFEx | 2018-03-05 | Pasteurellaceae sp-CVUAS-261,4 | 0.47 | 1 | 0
|  |  | CVUAS 31988 | 2; sequenced | EFEx | 2019-10-09 | Pasteurellaceae sp-CVUAS-31988 | 0 | 1 | 0
|  |  | CVUAS 4868 | 3; sequenced | DT | 2013-06-25 | Pasteurellaceae sp-CVUAS-4868 | 0.43 | 1 | 0
|  |  | CVUAS 724,3 | 3; sequenced | eDT | 2016-07-05 | Pasteurellaceae sp-CVUAS-724,3 | 0.8 | 1 | 0
|  |  | CVUAS 9305 | 3; sequenced | DT | 2017-05-16 | Pasteurellaceae sp-CVUAS-9305 | 0.54 | 1 | 0
|  |  | CVUAS 9967,2 | 3; sequenced | EFEx | 2014-12-16 | Pasteurellaceae sp-CVUAS-9967,2 | 0.24 | 1 | 0
|  |  | CVUAS 32290 | 3; sequenced | eDT | 2020-04-28 | Pasteurellaceae sp-LHL-U-309 | 0.45 | 1 | 0
|  |  | LHL U-309 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-309 | 0.56 | 1 | 0
|  |  | LHL U-3444 | 3; sequenced | DT | 2013-10-14 | Pasteurellaceae sp-LHL-U-3444 | 0.69 | 1 | 0
|  |  | CVUAS 6779 | 3; sequenced | DT | 2013-09-11 | Pasteurellaceae sp-MCCM-2539 | 0.18 | 1 | 0
|  |  | CVUAS 308,2 | 3; sequenced | DT | 2013-11-06 | Psittacicella hinzii | 0.27 | 1 | 0
|  |  | 2031-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.64 | 1 | 0
|  |  | 2376-10 | 3; sequenced | DT | 2013-04-03 | Rodentibacter heylii | 0.62 | 1 | 0
|  |  | CVUAS 4946 | 3; sequenced | DT | 2013-06-13 | Spirabiliibacterium mucosae | 0.5 | 1 | 0
|  |  | CVUAS 5940 | 3; sequenced | eDT | 2021-03-18 | Spirabiliibacterium mucosae | 0.73 | 1 | 0
|  |  | CVUAS 8966 | 3; sequenced | DT | 23.07.2013 | Volucribacter amazonae | 0.09 | 1 | 0
|  |  | DSM 24110 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.12.2016 E22 | Acinetobacter baumannii | 0.02 | 1 | 0
|  |  | DSM 30011 | 3; public strain collection | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2014 C05 | Acinetobacter baumannii | 0.1 | 1 | 0
|  |  | CVUAS 31883 | 3; sequenced | EFEx | 2019-08-23 | Acinetobacter defluvii | 0.18 | 1 | 0
|  |  | CVUAS 7795 | 3; sequenced | DT | 2013-09-18 | Acinetobacter guillouiae | 0.44 | 1 | 0
|  |  | CVUAS 2466 | 3; sequenced | DT | 2013-09-18 | Acinetobacter johnsonii | 0.36 | 1 | 0
|  |  | CVUAS 3963 | 3; sequenced | DT | 2013-09-18 | Acinetobacter junii | 0.08 | 1 | 0
|  |  | CVUAS 8349 | 3; sequenced | DT | 2013-09-18 | Acinetobacter lwoffii | 0 | 1 | 0
|  |  | CVUAS 7379,3 | 2; sequenced | eDT | 2015-02-18 | Acinetobacter pseudolwoffii | 0.27 | 1 | 0
|  |  | CVUAS 845 | 3; sequenced | DT | 2013-09-18 | Acinetobacter pseudolwoffii | 0.83 | 1 | 0
|  |  | CVUAS 7662 | 3; sequenced | DT | 2013-09-18 | Acinetobacter variabilis | 0.6 | 1 | 0
|  |  | CVUAS 7269 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 0.76 | 1 | 0
|  |  | LHL 1171 | 3; sequenced | eDT | 2013-09-11 | Moraxella bovoculi | 0.17 | 1 | 0
|  |  | CVUAS 7014 | 3; sequenced | eDT | 2013-09-11 | Moraxella catarrhalis | 0.27 | 1 | 0
|  |  | CVUAS 11674,3 | 3; sequenced | eDT | 2022-03-22 | Moraxella cuniculi | 0.39 | 1 | 0
|  |  | CVUAS 31885 | 3; 0 | EFex | 2019-09-03 | Moraxella osloensis | 0.8 | 1 | 0
|  |  | CVUAS 7064 | 3; sequenced | DT | 2013-09-11 | Moraxella ovis | 0.46 | 1 | 0
|  |  | CVUAS 10307 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter frigidicola | 0.4 | 1 | 0
|  |  | 57z-1 | 3; sequenced | EFEx | 2016-12-28 | Psychrobacter pasteurii | 0.71 | 1 | 0
|  |  | 10-4-D-04921 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.57 | 1 | 0
|  |  | CVUAS 7715 | 3; sequenced; biochemistry | eDT | 2014-01-22 | Psychrobacter pulmonis | 0.62 | 1 | 0
|  |  | CVUAS 163 | 3; public strain collection; sequenced | EFEx | 2018-09-25 | Pseudomonas (alkylphenolica) | 0 | 1 | 0
|  |  | 11-0023135001-01 LVU 08 | 3; sequenced (LGL) | eDT | 2015-04-15 | Pseudomonas aeruginosa | 0.28 | 1 | 0
|  |  | Y 8296 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 10.01.2007 H8 | Pseudomonas aeruginosa | 0.74 | 1 | 0
|  |  | CVUAS 30029 | 2; sequenced | EFEx | 2020-01-23 | Pseudomonas baetica | 0.53 | 1 | 0
|  |  | ATCC 17809 | 3; public strain collection | in 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 1:1 HCCATA2 (A) | 20.02.2008 N11 | Pseudomonas chlororaphis | 0.52 | 1 | 0
|  |  | A-269 05 A | 2; aus Ringversuch | 80% TFA/1:10 verdünnt/ 1:1 HCCATA2 | 02.04.2012, H21 | Francisella tularensis | 0.57 | 1 | 0
|  |  | DSM 102975 | 3; type strain | 80§ TFA, 30 Minuten ,1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 12.12.2016, F19, [0096] | Allofrancisella guangzhouensis | 0.7 | 1 | 0
|  |  | DSM 22475 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 18.11.2015, L16, [0096] | Francisella hispaniensis | 0.78 | 1 | 0
|  |  | DSM 21254 | 3; type strain | 80§ TFA, 1:10 verdünnt , gespottet 1:1 mit Matrix (12g/l HCCA in TA2) | 25.11.2015, I3, [0096] | Francisella naotunensis | 0.48 | 1 | 0
|  |  | CVUAS 4680,2 | 3; FLI: specific PCR | DT | 2013-04-02 | Francisella tularensis | 0.26 | 1 | 0
|  |  | Ft32 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation/1:10 verdünnt/ 1:1 HCCA TA2 (A) | 23.04.2013, F20, [0096] | Francisella tularensis | 0.69 | 1 | 0
|  |  | Ft31 | 3; 0 | in 80% TFA/ 30 min. Einwirkzeit und Zentrifugation mit Filtrierung, 1:10 verdünnt/ 1:1 HCCA/TA2 | 26.04.2013, I20, [0096] | Francisella tularensis | 0.8 | 1 | 0
|  |  | Ft26 | 3; 0 | in 80% TFA, 30 min, Filtriert, Zentrifugiert, 1:10 Verduennt in H2O, Probe 1:1 mit Matrix (12g/l HCCA in TA2) | 01.06.2017, F3, [0096] | Francisella tularensis | 0.82 | 1 | 0
|  |  | BfR-VB-00003 LMG 7909 | 3; type strain | DT | 2019-09-01 | Vibrio aestuarianus | 0 | 1 | 0
|  |  | 17-VB00029 | 3; rpob Sequenzierung Vibrio anguillarum Speziesidentität 99% Sequenzidentität(bp) 867/871 | DT | 2017-01-18 | Vibrio anguillarum | 0.09 | 1 | 0
|  |  | 17-VB00046 | 4; WGS sequenced | DT | 2017-03-02 | Vibrio campbelli | 0.48 | 1 | 0
|  |  | 18-VB00023 | 4; WGS sequenced | DT | 2018-01-08 | Vibrio campbelli | 0.26 | 1 | 0
|  |  | 16-VB00067-0_BfR 1_eDT | 3; confirmed by NRL-BfR | eDT | 2019-06-07 | Vibrio cholerae | 0.86 | 1 | 0
|  |  | 10-VBH0202 DT-48h | 4; WGS sequenced | DT | 2021-02-02 | Vibrio cincinnatiensis | 0.94 | 1 | 0
|  |  | BfR-VB-00004 LMG 7891 | 3; type strain | DT | 2020-02-29 | Vibrio cincinnatiensis | 0.51 | 1 | 0
|  |  | 15-VB00100 | 3; rpob | DT | 2021-02-02 | Vibrio diazotrophicus | 0.81 | 1 | 0
|  |  | 20-VB00358 | 3; well defined isolate | DT | 2020-11-13 | Vibrio fluvialis | 0.67 | 1 | 0
|  |  | BfR-VB-00036 LMG 7894 | 3; type strain | DT | 2021-02-24 | Vibrio fluvialis | 0.33 | 1 | 0
|  |  | 174-4 | 2; 0 | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 16.08.2012, K12, [0034] | Vibrio harveyi | 0.31 | 1 | 0
|  |  | 18-VB00006 | 3; WGS sequenced | DT | 2018-01-08 | Vibrio harveyi | 0.24 | 1 | 0
|  |  | 20-VB00047 | 3; WGS sequenced | DT | 2020-07-07 | Vibrio jasicida | 0.69 | 1 | 0
|  |  | 20-VB00053 | 3; WGS sequenced | DT | 2020-02-07 | Vibrio jasicida | 0.96 | 1 | 0
|  |  | 12-VBH0498 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio metoecus | 0.36 | 1 | 0
|  |  | 20-VB00016 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio metoecus | 0.28 | 1 | 0
|  |  | 13-VB00021 | 4; WGS sequenced | DT | 2021-02-02 | Vibrio metschnikovii | 0.53 | 1 | 0
|  |  | 20-VB00309 | 4; WGS sequenced | DT | 2021-02-24 | Vibrio navarrensis | 0.98 | 1 | 0
|  |  | 17-VB00403 | 3; WGS sequenced | DT | 2017-12-04 | Vibrio owensii | 0.89 | 1 | 0
|  |  | 17-VB00254-0_BfR 18_DT | 3; confirmed by NRL-BfR | DT | 2019-11-22 | Vibrio parahaemolyticus | 0.25 | 1 | 0
|  |  | 11-VBH0308 | 3; well defined isolate | DT | 2021-02-02 | Vibrio vulnificus | 0.74 | 1 | 0
|  |  | CVUAS 33180 25C | 3; sequenced | DT | 2021-08-04 | Photobacterium damselae | 0.55 | 1 | 0
|  |  | CVUAS 34147 | 3; sequenced | eDT | 2022-08-11 | Photobacterium damselae | 0.97 | 1 | 0
|  |  | CVUAS 1266,2 | 3; sequenced | EFEx | 25.09.1019 | Photobacterium piscicola | 0.26 | 1 | 0
|  |  | CVUAS 842,2 | 3; sequenced | Efex | 2013-03-07 | Photobacterium swingsii | 0.51 | 1 | 0
|  |  | CVUAS 7079 sir | 3; spezific PCR (CVFR) | eDT | 2022-06-23 | Vibrio alginolyticus | 0.73 | 1 | 0
|  |  | DSM 2171 RKI | 3; type strain | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 12.07.2012 E12 | Vibrio alginolyticus | 0.47 | 1 | 0
|  |  | A8-12 sir | 3; gyrB sequence | eDT | 2022-06-23 | Vibrio anguillarum | 0.9 | 1 | 0
|  |  | CVUAS 10227 | 3; sequenced | eDT | 2013-05-22 | Vibrio cholerae | 0.83 | 1 | 0
|  |  | CVUAS 8995 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio diazotrophicus | 0.36 | 1 | 0
|  |  | DSM 14383 sir | 3; public strain collection, sequenced | eDT | 2022-06-23 | Vibrio furnissi | 0.36 | 1 | 0
|  |  | CVUAS 1754 eDT | 3; sequenced | eDT | 2013-04-16 | Vibrio furnissii | 0.2 | 1 | 0
|  |  | CVUAS 619 sir | 2; 0 | eDT | 2022-06-23 | Vibrio metschnikovii | 0.42 | 1 | 0
|  |  | CVUAS 7872 sir | 3; sequenced | eDT | 2022-06-23 | Vibrio mimicus | 0.74 | 1 | 0
|  |  | DSM 19130 RKI | 4; type strain; WGS sequenced | 80% TFA/ 30 min. Einwirkzeit/1:10 verdünnt/ 2:2 HCCA/TA2 | 24.01.2011, N4, [0034] | Vibrio mimicus | 0.81 | 1 | 0
|  |  | DSM 10027 eDT | 3; type strain | eDT | 2022-06-23 | Vibrio parahaemolyticus | 0.04 | 1 | 0
|  |  | CVUAS 871,2 | 3; sequenced | eDT | 2014-08-21 | Vibrio sp-CVUAS-871,2 | 0.35 | 1 | 0
|  |  | CVUAS 2778,2 sir | 3; sequenced | eDT | 2022-07-14 | Vibrio tubiashii | 0.22 | 1 | 0
|  |  | CVUAS 2778,2 eDT | 3; sequenced | eDT | 2019-10-24 | Vibrio tubiashii | 0.23 | 1 | 0
|  |  | LVM05 LVU05 | 3; sequenced + species specific PCR (LGL) [0069] | EFex | 2015-04-14 | Vibrio vulnificus | 0.34 | 1 | 0
|  |  | CVUAS 30333 | 3; sequenced | eDT | 2018-10-25 | Wohlfahrtiimonas sp-G9077 | 0.8 | 1 | 0
|  |  | 10-7-D-03033-2 | 3; sequenced | EFEx | 2016-12-14 | Ignatzschineria indica | 0.86 | 1 | 0
|  |  | CVUAS 9634 | 3; sequenced | DT | 2013-09-17 | Ignatzschineria ureiclastica | 0.35 | 1 | 0
|  |  | 217-U2-6 | 3; sequenced | EFex | 2017-03-15 | Stenotrophomonas chelatiphaga | 0.44 | 1 | 0
|  |  | ATCC 13637 | 3; type strain | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 03.09.2009 K17 | Stenotrophomonas maltophilia | 0.59 | 1 | 0
|  |  | Sm 207 217 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 14.05.2009 K22 | Stenotrophomonas maltophilia | 0.5 | 1 | 0
|  |  | 277-U2-2 | 3; sequenced | Efex | 2016-12-28 | Stenotrophomonas rhizophila | 0.55 | 1 | 0
|  |  | RRLK03_01 | 2; 0 | 80% TFA/1:10 verdünnt/1:1 HCCA:TA2 | 18.05.2015 C17 | Stenotrophomonas rhizophila | 0.68 | 1 | 0
|  |  | CVUAS 4239,2 | 3; sequenced, biochemistry | eDT | 2014-06-18 | Stenotrophomonas tumulicola | 0.59 | 1 | 0

## Report (Spezies)
Bestätigungsprüfung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, Stand März 2021
Parameter: Clostridium perfringens
Version:
erstellt von: | Dyk | geprüft von: | Rau
erstellt am: | 2024-08-19 | geprüft am: | 2024-08-19
Anwendung | Bakterienisolate aus Lebensmitteln & der veterinärmedizinischen Diagnostik
Umfeld | Lebensmittel-Mikrobiologie, Mastitis, Veterinärmedizinische Diagnostik
Verwendete Datenbank / Programmversion |  | BRU 11897 | BT 4.0
Validierungsset:
Validierungsisolate/-materialien (Parameter): 101 |  | Vergleichsisolate/-materialien (#Parameter): 1144
Kultivierungsbedingungen bei Mikroorganismen: SBA, 30° C, 24-48h, anaerob
Zustand der Probe bei sonstigen Materialien:
Aufreinigungsmethode: DT, eDT, EFex, 80%TFA
Es werden alle Proben für die Ermittlung der In- und Exklusivität herangezogen.
Richtig-positiv: 1.Hit = Ziel-Parameter mit Score ≥ 2 und Score 2.Hit < 1.7 oder Score 2.Hit ≥ 1.7 & Gattung gleich 1.Hit oder Score 2.Hit ≥ 2 & Art gleich 1.Hit. |  | Richtig-negativ: 1. Hit ohne Zuordnung zum Ziel-Parameter mit score ≥ 2.
Falsch-negativ: Kriterien für richtig-positiv werden nicht erfüllt. Alle nicht identifizierten Ziel-Parameter werden als falsch negativ gewertet. |  | Falsch-positiv: Identifizierung als Ziel-Parameter über 1. / 2. Hit mit score ≥ 2.
[8 empty] | Level | Spezies | Genus | Family
Validierungsergebnis | Ziel-Parameter: | Clostridium perfringens |  | SOLL | erfüllt |  | nach 16140-2:2018 = 50
Zahl aller Spektren | 101 |  |  | 50 | ja | 1 | nach 16140-6:2020 = 100 |  | 100 | 150 | 200
davon Eigenspektren | 48 |  |  | 20.2 | ja | 1
In Bezug auf alle Spektren: |  | Inklusivität
Positive Übereinstimmungen (PA) | 101 | Richtig-positiv Rate | 100% | 98% | ja | 1 |  |  | 98% |  | 98%
Negative Abweichungen (ND) | 0 | Falsch-negativ Rate | 0% | 1% | ja | 1 |  |  | 2% | 0.6667% | 2%
Nicht-Ziel-Parameter
Zahl aller Spektren | 1144 |  |  | 30 | ja | 1 |  |  | 100 | 100 | 100
davon Eigenspektren | 356 |  |  | 228.8 | ja | 1
In Bezug auf alle Spektren: |  | Exklusivität
Negative Übereinstimmungen (NA) | 1144 | Richtig-negativ Rate | 100% | 99% | ja | 1 |  |  | 98% | 98% | 98%
Positive Abweichungen (PD) | 0 | Falsch-positiv Rate | 0% | 1% | ja | 1 |  |  | 2% | 2% | 2%
Test nah verwandte Spektren | 1144 | Anteil nah verwandter Sp. | 0%
Von den 101 Proben des Parameters wurden unter Verwendung der vollständigen Datenbank 101 (=100%) richtig erkannt (Inklusivität). 0 (=0%) dieser Proben wurden fehlerhaft einer anderen Spezies zugeordnet. 0 (=0%) der Proben wurden nicht identifiziert. Damit ergibt sich für die Bestätigungsprüfung eine Falsch-negativ-Rate von 0%
In den 1144 Proben der Nicht-Ziel-Parameter wurden unter Verwendung der vollständigen Datenbank 100% richtig-negativ (Exklusivität) gezählt. Von diesen 1144 Proben wurden 0 (=0%) fehlerhaft als Parameter identifiziert.
In den 1144 Proben der Nicht-Ziel-Parameter wurden unter Verwendung einer nur den Parameter enthaltenden Datenbank bei 0 (=0%) Einträgen ein score >2 erreicht.
Die Kriterien für die | Bestätigungsprüfung | sind somit | erfüllt
Kommentar:
Die Validierung erfüllt hinsichtlich Umfang und Anforderungen auch die Vorgaben der DIN EN ISO 16140-4:2020.
Freigabe an Standort | für (Abteilung / Bereich) | am (Datum) | von (Hdz.)
CVUA S
Das Verfahren ist für den vorgesehenen Einsatzzweck geeignet und zur Verwendung freigegeben.
als Prüfmethode für Parameter in LIMS eingepflegt
Eine unterschriebene Version dieses Validierungsreports befindet sich im erstellenden Labor.
Andere Keime, die in dieser Matrix vorkommen können.
Andere Keime, die auf diesem Agar wachsen können.
Vergleichsisolate/-materialien, die ähnliche Spektren liefern können (siehe Blatt "Ähnliche zu Parameter").
ja
nein